## Overview
Metric | Value
Total Death Cases | 2277
Unique Devices | 630
Total Patient Problems | 219
Total Product Problems | 152

## Device Summary
Device Name | Death Count | Top Patient Problems | Top Product Problems
DREAMSTATION AUTO CPAP | 360 | {'Cancer': 164, 'Unspecified Respiratory Problem': 130, 'Insufficient Information': 79, 'Headache': 52, 'Dizziness': 51} | {'Degraded': 293, 'Adverse Event Without Identified Device or Use Problem': 64, 'Contamination': 2, 'Insufficient Information': 2}
DREAMSTATION CPAP | 122 | {'Cancer': 68, 'Unspecified Respiratory Problem': 43, 'Unspecified Kidney or Urinary Problem': 24, 'Insufficient Information': 21, 'Dizziness': 17} | {'Degraded': 118, 'Adverse Event Without Identified Device or Use Problem': 3}
DREAMSTATION AUTO BIPAP | 73 | {'Unspecified Respiratory Problem': 26, 'Cancer': 25, 'Insufficient Information': 18, 'No Clinical Signs, Symptoms or Conditions': 11, 'Unspecified Kidney or Urinary Problem': 10} | {'Degraded': 54, 'Adverse Event Without Identified Device or Use Problem': 19, 'Contamination': 1, 'Material Integrity Problem': 1}
REMSTAR AUTO A-FLEX | 60 | {'Cancer': 29, 'Unspecified Respiratory Problem': 15, 'Insufficient Information': 13, 'Dizziness': 11, 'Headache': 11} | {'Degraded': 56, 'Adverse Event Without Identified Device or Use Problem': 9, 'Contamination': 7}
REMSTAR PRO C-FLEX+ | 57 | {'Cancer': 32, 'Unspecified Respiratory Problem': 21, 'No Clinical Signs, Symptoms or Conditions': 10, 'Insufficient Information': 7, 'Headache': 6} | {'Degraded': 50, 'Adverse Event Without Identified Device or Use Problem': 7, 'Contamination': 2, 'Material Integrity Problem': 1, 'No Apparent Adverse Event': 1}
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | 46 | {'Insufficient Information': 20, 'Cardiac Arrest': 20, 'Dyspnea': 3, 'Myocardial Infarction': 2, 'Peritonitis': 2} | {'Adverse Event Without Identified Device or Use Problem': 45}
DAVINCI XI | 41 | {'Insufficient Information': 14, 'Hemorrhage/Blood Loss/Bleeding': 13, 'Cardiac Arrest': 6, 'Pulmonary Embolism': 4, 'Low Blood Pressure/ Hypotension': 2} | {'Adverse Event Without Identified Device or Use Problem': 28, 'Insufficient Information': 13, 'Appropriate Term/Code Not Available': 1, 'Output Problem': 1}
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | 36 | {'Unspecified Vascular Problem': 17, 'Hematoma': 9, 'Hemorrhage/Blood Loss/Bleeding': 6, 'Retroperitoneal Hemorrhage': 4} | {'Adverse Event Without Identified Device or Use Problem': 29, 'Insufficient Information': 7}
DS2ADV AUTO CPAP | 33 | {'Insufficient Information': 16, 'No Clinical Signs, Symptoms or Conditions': 15, 'Appropriate Term / Code Not Available': 1, 'Cancer': 1, 'Unspecified Respiratory Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 30, 'Insufficient Information': 2, 'No Apparent Adverse Event': 1, 'Degraded': 1}
DREAMSTATION | 31 | {'Insufficient Information': 15, 'No Clinical Signs, Symptoms or Conditions': 7, 'Cancer': 2, 'Obstruction/Occlusion': 1, 'Chronic Obstructive Pulmonary Disease (COPD)': 1} | {'Adverse Event Without Identified Device or Use Problem': 22, 'Degraded': 8, 'No Apparent Adverse Event': 1}
REMSTAR AUTOA-FLEX | 29 | {'Cancer': 14, 'Unspecified Respiratory Problem': 6, 'Insufficient Information': 6, 'Unspecified Kidney or Urinary Problem': 4, 'Eye Irritation': 4} | {'Degraded': 27, 'Adverse Event Without Identified Device or Use Problem': 2, 'Contamination': 1, 'Material Integrity Problem': 1}
RENTAL, DREAMSTATION ST30, GB | 26 | {'Insufficient Information': 22, 'No Clinical Signs, Symptoms or Conditions': 3} | {'Adverse Event Without Identified Device or Use Problem': 25, 'Degraded': 1}
PHILIPS CPAP DEVICE | 25 | {'Unspecified Respiratory Problem': 10, 'No Clinical Signs, Symptoms or Conditions': 10, 'Cancer': 8, 'Dizziness': 4, 'Headache': 4} | {'Degraded': 19, 'Adverse Event Without Identified Device or Use Problem': 7}
DREAMSTATION ST30 | 25 | {'Insufficient Information': 21, 'No Clinical Signs, Symptoms or Conditions': 4} | {'Adverse Event Without Identified Device or Use Problem': 23, 'Degraded': 2}
ALARIS SYSTEM | 24 | {'Insufficient Information': 13, 'Cardiac Arrest': 6, 'Low Blood Pressure/ Hypotension': 4, 'High Blood Pressure/ Hypertension': 1, 'Intracranial Hemorrhage': 1} | {'Contamination': 8, 'Inaccurate Delivery': 8, 'Use of Device Problem': 8, 'Device Alarm System': 4, 'Excess Flow or Over-Infusion': 4}
DREAMSTATION CPAP PRO | 21 | {'Cancer': 12, 'Unspecified Respiratory Problem': 7, 'Insufficient Information': 5, 'Asthma': 3, 'Dizziness': 3} | {'Degraded': 18, 'Adverse Event Without Identified Device or Use Problem': 5}
PATIENT INFORMATION CENTER IX | 21 | {'Insufficient Information': 13, 'Cardiac Arrest': 4, 'Low Oxygen Saturation': 1, 'No Clinical Signs, Symptoms or Conditions': 1, 'Ventricular Fibrillation': 1} | {'Defective Alarm': 8, 'No Audible Alarm': 7, 'Adverse Event Without Identified Device or Use Problem': 4, 'Device Alarm System': 2, 'Use of Device Problem': 1}
REMSTAR PLUS | 17 | {'Cancer': 11, 'Unspecified Respiratory Problem': 5, 'Dizziness': 3, 'Headache': 3, 'No Clinical Signs, Symptoms or Conditions': 3} | {'Degraded': 16, 'Adverse Event Without Identified Device or Use Problem': 1, 'Contamination': 1}
ION | 16 | {'Hemorrhage/Blood Loss/Bleeding': 7, 'Cardiac Arrest': 6, 'Pulmonary Embolism': 4, 'Insufficient Information': 4, 'Respiratory Arrest': 3} | {'Adverse Event Without Identified Device or Use Problem': 16}
REP DREAMSTATION AUTO CPAP, DOM - RECRT | 16 | {'Insufficient Information': 8, 'No Clinical Signs, Symptoms or Conditions': 4, 'Cancer': 3, 'Unspecified Kidney or Urinary Problem': 2, 'Unspecified Respiratory Problem': 2} | {'Adverse Event Without Identified Device or Use Problem': 13, 'Degraded': 3}
NXSTAGE SYSTEM ONE | 14 | {'Cardiac Arrest': 14, 'Hemorrhage/Blood Loss/Bleeding': 5, 'Chest Pain': 1, 'Dyspnea': 1, 'Low Blood Pressure/ Hypotension': 1} | {'Adverse Event Without Identified Device or Use Problem': 14}
REP DREAMSTATION AUTO CPAP | 14 | {'Insufficient Information': 7, 'No Clinical Signs, Symptoms or Conditions': 6, 'Hypersensitivity/Allergic reaction': 1, 'Cancer': 1, 'Eye Irritation': 1} | {'Adverse Event Without Identified Device or Use Problem': 10, 'Degraded': 3, 'Insufficient Information': 1}
REMSTAR SYSTEM ONE 60 SERIES | 14 | {'Insufficient Information': 12, 'Stroke/CVA': 1, 'Respiratory Failure': 1} | {'Adverse Event Without Identified Device or Use Problem': 12, 'Degraded': 2}
MICRA | 14 | {'Cardiac Perforation': 5, 'Cardiac Arrest': 5, 'Cardiac Tamponade': 4, 'Hemorrhage/Blood Loss/Bleeding': 4, 'Pulmonary Embolism': 3} | {'Adverse Event Without Identified Device or Use Problem': 9, 'Use of Device Problem': 5, 'Appropriate Term/Code Not Available': 3, 'Leak/Splash': 1}
BIPAP AUTO BIFLEX | 13 | {'Cancer': 6, 'Unspecified Respiratory Problem': 5, 'Eye Irritation': 3, 'Headache': 3, 'Inflammation': 3} | {'Degraded': 12, 'Adverse Event Without Identified Device or Use Problem': 3}
RENTAL-DREAMSTATION AUTO CPAP, GB | 13 | {'Insufficient Information': 13} | {'Adverse Event Without Identified Device or Use Problem': 12, 'Degraded': 1}
DREAMSTATION BIPAP PRO | 12 | {'Unspecified Respiratory Problem': 6, 'Cancer': 4, 'Dizziness': 3, 'Unspecified Kidney or Urinary Problem': 3, 'Inflammation': 3} | {'Degraded': 11, 'Adverse Event Without Identified Device or Use Problem': 1}
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | 12 | {'Insufficient Information': 6, 'Elevated ketones/Diabetic Ketoacidosis': 3, 'Hyperglycemia': 2, 'Hypoglycemia': 2} | {'Adverse Event Without Identified Device or Use Problem': 12}
TUBING SET | 12 | {'Insufficient Information': 4, 'Cardiac Arrest': 3, 'Low Oxygen Saturation': 2, 'Septic Shock': 2, 'Hemolysis': 2} | {'Fluid/Blood Leak': 5, 'Adverse Event Without Identified Device or Use Problem': 4, 'Air/Gas in Device': 3, 'Gas Output Problem': 2, 'Gas/Air Leak': 2}
GORE® DRYSEAL FLEX INTRODUCER SHEATH | 11 | {'Ischemia': 4, 'Vascular Dissection': 2, 'Obstruction/Occlusion': 2, 'Perforation of Vessels': 2, 'Reperfusion Injury': 1} | {'Adverse Event Without Identified Device or Use Problem': 4, 'Patient Device Interaction Problem': 3, 'Difficult to Advance': 1, 'Insufficient Information': 1, 'Device Handling Problem': 1}
TRILOGY EVO | 10 | {'Insufficient Information': 4, 'Respiratory Arrest': 3, 'Appropriate Term / Code Not Available': 1, 'Respiratory Failure': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 8, 'Insufficient Information': 2}
REP DREAMSTATION AUTO BIPAP | 10 | {'No Clinical Signs, Symptoms or Conditions': 4, 'Insufficient Information': 4, 'Heart Failure/Congestive Heart Failure': 1, 'Unspecified Respiratory Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 6, 'Degraded': 3, 'Insufficient Information': 1}
CARDIOHELP | 10 | {'Hemorrhage/Blood Loss/Bleeding': 6, 'Septic Shock': 2, 'Hemolysis': 2, 'Asystole': 2} | {'Adverse Event Without Identified Device or Use Problem': 6, 'Insufficient Information': 4, 'Noise, Audible': 2, 'Device Handling Problem': 2}
PHILIPS CPAP | 9 | {'No Clinical Signs, Symptoms or Conditions': 4, 'Cancer': 2, 'Asthma': 2, 'Inflammation': 2, 'Unspecified Respiratory Problem': 2} | {'Degraded': 8, 'Adverse Event Without Identified Device or Use Problem': 1}
REACT CATHETER | 9 | {'Intracranial Hemorrhage': 3, 'Bacterial Infection': 2, 'Hemorrhage/Blood Loss/Bleeding': 2, 'Ischemia Stroke': 2, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 8, 'Material Deformation': 2, 'Difficult to Advance': 2, 'No Apparent Adverse Event': 1, 'Physical Resistance/Sticking': 1}
VERSACROSS CONNECT LAAC ACCESS SOLUTION | 9 | {'Cardiac Arrest': 5, 'Bradycardia': 5, 'Pericardial Effusion': 4, 'Air Embolism': 4, 'Low Blood Pressure/ Hypotension': 2} | {'Adverse Event Without Identified Device or Use Problem': 9}
DAVINCI X | 9 | {'Insufficient Information': 2, 'Sepsis': 2, 'Peritonitis': 2, 'Hemodynamic instability': 2, 'Renal Failure': 1} | {'Adverse Event Without Identified Device or Use Problem': 8, 'Insufficient Information': 2}
TRILOGY EVO, JAPAN | 9 | {'Insufficient Information': 3, 'Respiratory Arrest': 2, 'Cardiac Arrest': 2, 'Respiratory Failure': 1, 'Unspecified Respiratory Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 9, 'Insufficient Information': 1}
REMSTAR AUTO | 7 | {'Cancer': 5, 'Dizziness': 2, 'Headache': 2, 'Skin Inflammation/ Irritation': 1, 'Breast Cancer': 1} | {'Degraded': 7, 'Contamination': 2}
DS2ADV AUTO CPAP W/HUMID CELL/BT | 7 | {'Insufficient Information': 5, 'No Clinical Signs, Symptoms or Conditions': 2} | {'Adverse Event Without Identified Device or Use Problem': 7}
SUREFORM | 7 | {'Unspecified Tissue Injury': 4, 'Sepsis': 3, 'Failure to Anastomose': 2, 'Impaired Healing': 1, 'Brain Injury': 1} | {'Failure to Form Staple': 3, 'Insufficient Information': 3, 'Failure to Fire': 1}
BIPAP AUTOSV ADV | 7 | {'No Clinical Signs, Symptoms or Conditions': 3, 'Cancer': 2, 'Anemia': 1, 'Respiratory Failure': 1, 'Pleural Effusion': 1} | {'Degraded': 7, 'Adverse Event Without Identified Device or Use Problem': 2, 'Contamination': 1}
MITRACLIP | 7 | {'Low Blood Pressure/ Hypotension': 5, 'Pericardial Effusion': 2, 'Cardiac Arrest': 2, 'Embolism/Embolus': 1, 'Air Embolism': 1} | {'Adverse Event Without Identified Device or Use Problem': 5, 'Leak/Splash': 1, 'Difficult to Advance': 1, 'Physical Resistance/Sticking': 1}
DREAMSTATION BIPAP AUTOSV | 7 | {'Cancer': 4, 'Unspecified Respiratory Problem': 3, 'Insufficient Information': 2, 'Eye Irritation': 1, 'Dizziness': 1} | {'Degraded': 5, 'Adverse Event Without Identified Device or Use Problem': 2}
SOLITAIRE FR4 | 6 | {'Ischemia Stroke': 3, 'Hemorrhagic Stroke': 2, 'Bacterial Infection': 1, 'Hematoma': 1, 'Aspiration Pneumonitis': 1} | {'Adverse Event Without Identified Device or Use Problem': 5, 'Break': 2, 'Physical Resistance/Sticking': 1}
ILET BIONIC PANCREAS | 6 | {'Insufficient Information': 3, 'Hypoglycemia': 2, 'Elevated ketones/Diabetic Ketoacidosis': 1} | {'Insufficient Information': 5, 'Adverse Event Without Identified Device or Use Problem': 1}
EDWARDS ESHEATH+ INTRODUCER SET | 6 | {'Unspecified Vascular Problem': 2, 'Great Vessel Perforation': 1, 'Vascular Dissection': 1, 'Stroke/CVA': 1, 'Perforation of Vessels': 1} | {'Physical Resistance/Sticking': 3, 'Material Puncture/Hole': 2, 'Material Split, Cut or Torn': 2, 'Adverse Event Without Identified Device or Use Problem': 1}
TRILOGY 100 | 6 | {'Insufficient Information': 3, 'No Clinical Signs, Symptoms or Conditions': 2, 'Asthma': 1, 'Unspecified Kidney or Urinary Problem': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 3, 'Adverse Event Without Identified Device or Use Problem': 3, 'Failure to Calibrate': 1, 'Electrical /Electronic Property Problem': 1}
HEART LUNG MACHINE | 6 | {'Cardiac Arrest': 2, 'Aneurysm': 1, 'No Clinical Signs, Symptoms or Conditions': 1, 'Adult Respiratory Distress Syndrome': 1} | {'Electrical /Electronic Property Problem': 3, 'Pumping Stopped': 2, 'Pumping Problem': 1}
CARDIOSAVE HYBRID, TYPE B PLUG | 6 | {'Insufficient Information': 3, 'Low Blood Pressure/ Hypotension': 1, 'Ischemic Heart Disease': 1, 'Mitral Valve Insufficiency/ Regurgitation': 1} | {'Fluid/Blood Leak': 2, 'Appropriate Term/Code Not Available': 1, 'Gas/Air Leak': 1, 'Overheating of Device': 1, 'Incorrect, Inadequate or Imprecise Result or Readings': 1}
FLEXCATH ADVANCE STEERABLE SHEATH | 6 | {'Low Blood Pressure/ Hypotension': 4, 'Pericardial Effusion': 2, 'Vascular Dissection': 1, 'Tachycardia': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 6}
REMSTAR PRO C-FLEX | 6 | {'Cancer': 3, 'Unspecified Respiratory Problem': 2, 'No Clinical Signs, Symptoms or Conditions': 2, 'Eye Irritation': 1, 'Dizziness': 1} | {'Degraded': 6, 'Adverse Event Without Identified Device or Use Problem': 1}
VERSACROSS ACCESS SOLUTION | 6 | {'Cardiac Perforation': 2, 'Pericardial Effusion': 2, 'Cardiac Arrest': 2, 'Ischemia Stroke': 2, 'Cardiac Tamponade': 1} | {'Adverse Event Without Identified Device or Use Problem': 4}
OMNIPOD 5 POD | 5 | {'Hypoglycemia': 3, 'Hyperglycemia': 2, 'Myocardial Infarction': 1} | {'Insufficient Information': 4, 'Adverse Event Without Identified Device or Use Problem': 1}
WATCHMAN TRUSTEER ACCESS SYSTEM | 5 | {'Cardiac Perforation': 2, 'Pericardial Effusion': 2, 'Non specific EKG/ECG Changes': 2, 'Retroperitoneal Hemorrhage': 1, 'Cardiac Tamponade': 1} | {'Adverse Event Without Identified Device or Use Problem': 4, 'Material Deformation': 1, 'Material Integrity Problem': 1}
INTELLIVUE MX450 PATIENT MONITOR | 5 | {'Insufficient Information': 3, 'No Clinical Signs, Symptoms or Conditions': 1, 'Cardiac Arrest': 1, 'Bradycardia': 1} | {'Defective Alarm': 2, 'Device Alarm System': 2, 'Unable to Obtain Readings': 1, 'No Audible Alarm': 1}
WECK VISTA | 5 | {'Bowel Perforation': 2, 'Great Vessel Perforation': 2, 'Hemorrhage/Blood Loss/Bleeding': 2, 'Sepsis': 1, 'Pulmonary Embolism': 1} | {'Insufficient Information': 4, 'Adverse Event Without Identified Device or Use Problem': 1}
CONFIDA BRECKER CURVE GUIDEWIRE | 5 | {'Perforation': 3, 'Cardiac Arrest': 2, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Obstruction/Occlusion': 1, 'Low Blood Pressure/ Hypotension': 1} | {'Patient Device Interaction Problem': 5, 'Insufficient Information': 1, 'Migration': 1}
TEMPUS LS | 5 | {'Cardiac Arrest': 4, 'No Clinical Signs, Symptoms or Conditions': 2, 'Unspecified Heart Problem': 1} | {'Failure to Deliver Shock/Stimulation': 2, 'No Pacing': 1, 'Unable to Obtain Readings': 1, 'Erratic or Intermittent Display': 1}
ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | 5 | {'Hemorrhage/Blood Loss/Bleeding': 4, 'Unspecified Tissue Injury': 1, 'Insufficient Information': 1, 'No Clinical Signs, Symptoms or Conditions': 1, 'Exsanguination': 1} | {'Adverse Event Without Identified Device or Use Problem': 4, 'Insufficient Information': 2, 'Difficult to Open or Close': 2, 'Failure to Fire': 1, 'Activation Problem': 1}
ENDO ANCHOR SYSTEM - HELI-FX AAA | 5 | {'Insufficient Information': 3, 'Multiple Organ Dysfunction Syndrome': 1, 'No Clinical Signs, Symptoms or Conditions': 1, 'Cardiac Arrest': 1, 'Stroke/CVA': 1} | {'Adverse Event Without Identified Device or Use Problem': 5, 'Leak/Splash': 1}
ALPHAVAC | 5 | {'Pulmonary Embolism': 5, 'Cardiac Arrest': 1} | {'Difficult to Advance': 3, 'Adverse Event Without Identified Device or Use Problem': 2}
EDWARDS ESHEATH PLUS INTRODUCER SHEATH | 5 | {'Unspecified Vascular Problem': 2, 'Vascular Dissection': 1, 'Great Vessel Perforation': 1, 'Insufficient Information': 1, 'Perforation of Vessels': 1} | {'Material Puncture/Hole': 4, 'Physical Resistance/Sticking': 3, 'Difficult to Remove': 1, 'Material Integrity Problem': 1}
DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | 5 | {'Insufficient Information': 3, 'Hypoglycemia': 1, 'Vomiting': 1, 'Elevated ketones/Diabetic Ketoacidosis': 1} | {'Appropriate Term/Code Not Available': 3, 'Defective Alarm': 1, 'Incorrect, Inadequate or Imprecise Result or Readings': 1, 'Insufficient Information': 1}
WATCHMAN TRUSTEER? ACCESS SYSTEM | 5 | {'Pericardial Effusion': 4, 'Low Blood Pressure/ Hypotension': 3, 'Cardiac Tamponade': 3, 'Bradycardia': 2, 'Loss of consciousness': 1} | {'Adverse Event Without Identified Device or Use Problem': 5}
WATCHMAN FXD CURVE ACCESS SYSTEM | 5 | {'Bradycardia': 2, 'Cardiac Arrest': 2, 'Low Blood Pressure/ Hypotension': 2, 'Arrhythmia': 2, 'Heart Block': 1} | {'Adverse Event Without Identified Device or Use Problem': 5}
SAFARI2¿ | 5 | {'Cardiac Perforation': 5, 'Cardiac Tamponade': 2, 'Pericardial Effusion': 2, 'Sepsis': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 4, 'Malposition of Device': 1}
BIPAP PRO BIFLEX | 5 | {'Cancer': 3, 'Unspecified Kidney or Urinary Problem': 2, 'Insufficient Information': 1} | {'Degraded': 5}
AZURION | 5 | {'Insufficient Information': 4, 'No Clinical Signs, Symptoms or Conditions': 1, 'Hemodynamic instability': 1} | {'Positioning Failure': 3, 'Radiation Output Failure': 2}
BIPAP AUTO | 4 | {'Insufficient Information': 2, 'Cancer': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 3, 'Adverse Event Without Identified Device or Use Problem': 1}
DREAMSTATION GO AUTO | 4 | {'Cancer': 2, 'Insufficient Information': 2, 'Hypersensitivity/Allergic reaction': 1, 'Vomiting': 1, 'Dizziness': 1} | {'Degraded': 3, 'Adverse Event Without Identified Device or Use Problem': 1}
FISHER & PAYKEL HEALTHCARE | 4 | {'Insufficient Information': 3, 'Respiratory Failure': 1} | {'Insufficient Information': 3, 'Adverse Event Without Identified Device or Use Problem': 1, 'Misconnection': 1, 'Misassembly by Users': 1, 'Misassembled': 1}
REMSTAR | 4 | {'Unspecified Kidney or Urinary Problem': 2, 'Cancer': 2, 'Eye Irritation': 2, 'Unspecified Respiratory Problem': 2, 'Insufficient Information': 1} | {'Degraded': 4, 'Adverse Event Without Identified Device or Use Problem': 2}
DREAM STATION CPAP | 4 | {'Cancer': 2, 'Unspecified Respiratory Problem': 2, 'Chronic Obstructive Pulmonary Disease (COPD)': 1, 'Aspiration Pneumonitis': 1, 'Vomiting': 1} | {'Degraded': 4, 'Adverse Event Without Identified Device or Use Problem': 2}
POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 4 | {'Unspecified Infection': 4, 'Insufficient Information': 2, 'Appropriate Term / Code Not Available': 2, 'Septic Shock': 1} | {'Adverse Event Without Identified Device or Use Problem': 4}
BIPAP SYSTEM ONE 60 SERIES | 4 | {'No Clinical Signs, Symptoms or Conditions': 3, 'Insufficient Information': 1} | {'Degraded': 3, 'Adverse Event Without Identified Device or Use Problem': 1}
HLS SET | 4 | {'Insufficient Information': 4} | {'Pressure Problem': 2, 'Adverse Event Without Identified Device or Use Problem': 2}
BIPAP PRO | 4 | {'Cancer': 3, 'Unspecified Respiratory Problem': 3, 'Asthma': 2, 'Hypersensitivity/Allergic reaction': 1, 'Vomiting': 1} | {'Degraded': 4}
REMSTAR PLUS C-FLEX | 4 | {'Cancer': 4, 'Insufficient Information': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 3, 'Material Integrity Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
REMSTAR PRO C-FLEX + | 4 | {'Cancer': 3, 'Unspecified Respiratory Problem': 2, 'Inflammation': 2, 'Unspecified Kidney or Urinary Problem': 1, 'Skin Inflammation/ Irritation': 1} | {'Degraded': 4, 'Adverse Event Without Identified Device or Use Problem': 2}
EVIS EXERA III DUODENOVIDEOSCOPE | 4 | {'Drug Resistant Bacterial Infection': 2, 'No Clinical Signs, Symptoms or Conditions': 1, 'Sepsis': 1, 'Encephalopathy': 1, 'Septic Shock': 1} | {'Adverse Event Without Identified Device or Use Problem': 3, 'Contamination': 1}
DREAM STATION AUTO CPAP | 4 | {'Unspecified Respiratory Problem': 3, 'Cancer': 3, 'Insufficient Information': 2, 'Heart Failure/Congestive Heart Failure': 1, 'Respiratory Failure': 1} | {'Degraded': 4}
ESSENCE RENTAL | 4 | {'No Clinical Signs, Symptoms or Conditions': 3, 'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 4, 'Degraded': 1}
OMNIPOD 5 CONTROLLER | 4 | {'Insufficient Information': 2, 'Hypoglycemia': 1, 'Elevated ketones/Diabetic Ketoacidosis': 1} | {'Insufficient Information': 4}
CONFIDENCE KIT, NO NEEDLES | 4 | {'Low Blood Pressure/ Hypotension': 4, 'Hemorrhage/Blood Loss/Bleeding': 4} | {'Adverse Event Without Identified Device or Use Problem': 4}
AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | 3 | {'Low Blood Pressure/ Hypotension': 2, 'Cardiac Arrest': 2, 'Brain Injury': 1, 'Heart Block': 1, 'Ventricular Fibrillation': 1} | {'Air/Gas in Device': 1, 'Difficult to Advance': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
REBAR | 3 | {'Intracranial Hemorrhage': 1, 'Nervous System Injury': 1, 'Ischemia Stroke': 1} | {'Adverse Event Without Identified Device or Use Problem': 3, 'Physical Resistance/Sticking': 1}
ACHIEVE MAPPING CATHETER - 20 MM | 3 | {'Low Blood Pressure/ Hypotension': 2, 'Vascular Dissection': 1, 'Tachycardia': 1, 'Cardiac Arrest': 1, 'Bradycardia': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
EMBOTRAP | 3 | {'Intracranial Hemorrhage': 2, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Perforation of Vessels': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Retraction Problem': 1, 'Difficult to Remove': 1}
PRLNE BLU 24IN 7-0 D/A BV-1 | 3 | {'Perforation of Vessels': 3} | {'Break': 3}
BIPAP AUTO SERIES ASSY | 3 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Cancer': 1, 'Vomiting': 1, 'Eye Irritation': 1, 'Asthma': 1} | {'Degraded': 3}
PHENOM CATHETER | 3 | {'Intracranial Hemorrhage': 2, 'Hemorrhage/Blood Loss/Bleeding': 2, 'Thrombosis/Thrombus': 1, 'Ruptured Aneurysm': 1, 'Rupture': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Physical Resistance/Sticking': 1}
DREAMSTATION AUTO | 3 | {'Cancer': 3, 'Insufficient Information': 1, 'Inflammation': 1} | {'Degraded': 1, 'Adverse Event Without Identified Device or Use Problem': 1, 'Material Integrity Problem': 1}
LIBERATOR | 3 | {'Hypoxia': 1, 'Appropriate Term / Code Not Available': 1, 'Insufficient Information': 1} | {'Fluid/Blood Leak': 1, 'Improper or Incorrect Procedure or Method': 1, 'Insufficient Information': 1}
RENTAL-DREAMSTATION AUTO CPAP | 3 | {'No Clinical Signs, Symptoms or Conditions': 3} | {'Adverse Event Without Identified Device or Use Problem': 3}
DREAMSTATION AUTO CPAP, JP | 3 | {'Aspiration Pneumonitis': 2, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Patient Device Interaction Problem': 1}
DUODOPA_DUOPA | 3 | {'Appropriate Term / Code Not Available': 2, 'Peritoneal Laceration(s)': 2, 'Pneumonia': 1, 'Sepsis': 1, 'Bowel Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
CENTRIMAG BLOOD PUMP | 3 | {'Air Embolism': 2, 'Insufficient Information': 1, 'Exsanguination': 1} | {'Adverse Event Without Identified Device or Use Problem': 3, 'Gas/Air Leak': 2}
VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | 3 | {'Unspecified Infection': 2, 'Post Operative Wound Infection': 1, 'Wound Dehiscence': 1, 'Respiratory Failure': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
DEVILBISS HEALTHCARE | 3 | {'Insufficient Information': 2, 'Low Oxygen Saturation': 1} | {'Mechanical Problem': 1, 'Insufficient Information': 1, 'Infusion or Flow Problem': 1}
RESPIRONICS | 3 | {'Insufficient Information': 3} | {'Tidal Volume Fluctuations': 1, 'No Audible Alarm': 1, 'Insufficient Information': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
TIGHTRAIL ROTATING DILATOR SHEATH | 3 | {'Great Vessel Perforation': 3, 'Perforation of Vessels': 1, 'Cardiac Perforation': 1, 'Pericardial Effusion': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | 3 | {'Hemorrhage/Blood Loss/Bleeding': 2, 'Insufficient Information': 1, 'Perforation of Vessels': 1, 'Cardiac Arrest': 1, 'Discomfort': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Poor Visibility': 1, 'Use of Device Problem': 1}
WATCHMAN FXD CURVE? ACCESS SYSTEM | 3 | {'Cardiac Tamponade': 2, 'Pericardial Effusion': 2, 'Air Embolism': 1, 'Bradycardia': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
REMSTAR PLUS SERIES ASSY | 3 | {'Cancer': 2, 'Hypersensitivity/Allergic reaction': 1, 'Dizziness': 1, 'Inflammation': 1, 'Skin Inflammation/ Irritation': 1} | {'Degraded': 3}
EFFICIA DFM100 | 3 | {'Unspecified Heart Problem': 2, 'Ventricular Fibrillation': 1, 'Insufficient Information': 1} | {'Failure to Deliver Shock/Stimulation': 1, 'Incorrect Measurement': 1}
FREEDOM PERIPHERAL NERVE STIMULATOR | 3 | {'No Clinical Signs, Symptoms or Conditions': 3} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Off-Label Use': 1}
SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | 3 | {'Low Blood Pressure/ Hypotension': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Perforation': 1, 'Vascular Dissection': 1, 'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
SMART PORT CT | 3 | {'Sepsis': 2, 'Infiltration into Tissue': 1, 'Fever': 1, 'Muscle Weakness/Atrophy': 1, 'Septic Shock': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Fluid/Blood Leak': 1}
NELLCOR | 3 | {'Appropriate Term / Code Not Available': 2, 'Hypoxia': 1, 'Cardiac Arrest': 1} | {'Appropriate Term/Code Not Available': 2, 'Device Alarm System': 1}
LLD EZ LEAD LOCKING DEVICE | 3 | {'Cardiac Perforation': 2, 'Pericardial Effusion': 1, 'Perforation of Vessels': 1, 'Great Vessel Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
SIGNIA | 3 | {'Hemorrhage/Blood Loss/Bleeding': 3, 'Cardiac Arrest': 3, 'Tissue Breakdown': 2, 'Low Blood Pressure/ Hypotension': 2, 'Low Oxygen Saturation': 2} | {'Misfire': 2, 'Adverse Event Without Identified Device or Use Problem': 1}
CARDIOSAVE HYBRID W/ E/F PLUG | 3 | {'Insufficient Information': 3, 'Appropriate Term / Code Not Available': 1, 'Heart Failure/Congestive Heart Failure': 1} | {'Overheating of Device': 1, 'Electrical /Electronic Property Problem': 1, 'Inability to Auto-Fill': 1, 'Electrical Power Problem': 1}
NC TREK | 3 | {'Unspecified Heart Problem': 2, 'Foreign Body In Patient': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Failure to Deflate': 1, 'Difficult to Remove': 1, 'Material Separation': 1, 'Improper or Incorrect Procedure or Method': 1}
EVERFLO | 3 | {'Insufficient Information': 2, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
VIVO 45 LS | 3 | {'Insufficient Information': 3} | {'Adverse Event Without Identified Device or Use Problem': 3}
DAVINCI 5 | 3 | {'Hemorrhage/Blood Loss/Bleeding': 2, 'Insufficient Information': 1, 'Thromboembolism': 1, 'Coagulation Disorder': 1, 'Multiple Organ Dysfunction Syndrome': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 3 | {'Insufficient Information': 2, 'Sepsis': 1, 'Pain': 1, 'Fungal Infection': 1, 'Fever': 1} | {'Adverse Event Without Identified Device or Use Problem': 3}
PROXIMATE*VASC LIN STAP RELOAD | 3 | {'Hemorrhage/Blood Loss/Bleeding': 3} | {'Adverse Event Without Identified Device or Use Problem': 3, 'Failure to Fire': 3, 'No Apparent Adverse Event': 3}
MX40 1.4 GHZ SMART HOPPING | 3 | {'Insufficient Information': 2} | {'No Audible Alarm': 1, 'Defective Alarm': 1}
AMPLATZER VASCULAR PLUG II | 3 | {'Insufficient Information': 2, 'Aortic Valve Insufficiency/ Regurgitation': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Patient Device Interaction Problem': 1}
SENSATION PLUS 8FR. 50CC IAB & ACCESSORIES (W/PRESSURE TUBES, NO STYLET) | 3 | {'Appropriate Term / Code Not Available': 2, 'Cardiac Arrest': 1} | {'Inflation Problem': 2, 'Leak/Splash': 1, 'Gas/Air Leak': 1}
DREAMSTATION HUMIDIFIER, GB-RENT | 3 | {'Insufficient Information': 2, 'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Insufficient Information': 1}
INTELLIVUE MX800 PATIENT MONITOR | 3 | {'Insufficient Information': 3} | {'No Audible Alarm': 1, 'Loss of Data': 1, 'Insufficient Information': 1}
HT70 VENTILATOR | 2 | {'Respiratory Failure': 2} | {'No Flow': 2}
CADD-SOLIS VIP AMBULATORY INFUSION PUMP | 2 | {'Insufficient Information': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
REMSTAR W/SD CARD | 2 | {'Cancer': 1, 'Insufficient Information': 1} | {'Degraded': 2}
REP DREAMSTATION AUTO BIPAP, DOM | 2 | {'Insufficient Information': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
RENTAL, DREAMSTATION ST30 | 2 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
FREESTYLE LIBRE 2 | 2 | {'Insufficient Information': 2} | {'Insufficient Information': 2}
HEARTSTART XL+ DEFIBRILLATOR/MONITOR | 2 | {'Appropriate Term / Code Not Available': 1, 'Cardiac Arrest': 1, 'Burn(s)': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1, 'Defibrillation/Stimulation Problem': 1}
CORTRAK 2 ENTERAL ACCESS SYSTEM | 2 | {'Pneumothorax': 1, 'Fistula': 1, 'Pleural Effusion': 1, 'Cardiac Arrest': 1} | {'Malposition of Device': 2}
TRILOGY | 2 | {'Insufficient Information': 2} | {'Insufficient Information': 2}
EDWARDS ESHEATH INTRODUCER SET | 2 | {'Hematoma': 1, 'Retroperitoneal Hemorrhage': 1, 'No Clinical Signs, Symptoms or Conditions': 1, 'Rupture': 1, 'Perforation of Vessels': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Deformation Due to Compressive Stress': 1, 'Physical Resistance/Sticking': 1, 'Material Puncture/Hole': 1}
FLEXCATH CONTOUR¿ STEERABLE SHEATH | 2 | {'Low Blood Pressure/ Hypotension': 1, 'Pericardial Effusion': 1, 'Cardiac Tamponade': 1, 'Cardiac Arrest': 1, 'Stroke/CVA': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
REMSTAR SE | 2 | {'Insufficient Information': 1, 'Cancer': 1} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Degraded': 2}
EPOC BLOOD ANALYSIS SYSTEM | 2 | {'Insufficient Information': 2} | {'High Test Results': 1, 'Low Test Results': 1}
DREAMSTATION 2 ADVANCED AUTO CPAP | 2 | {'Insufficient Information': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
PROGRIP | 2 | {'Burning Sensation': 2, 'Pain': 2, 'Discomfort': 2, 'Appropriate Term / Code Not Available': 2, 'Depression': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
FREESTYLE LIBRE 3 PLUS | 2 | {'Dizziness': 1, 'Hyperglycemia': 1, 'Dehydration': 1, 'Loss of consciousness': 1, 'Elevated ketones/Diabetic Ketoacidosis': 1} | {'Low Readings': 2}
OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | 2 | {'Elevated ketones/Diabetic Ketoacidosis': 2, 'Hyperglycemia': 1} | {'Insufficient Information': 2, 'Insufficient Flow or Under Infusion': 1}
PACEL BIPOLAR PACING CATHETER | 2 | {'Insufficient Information': 1, 'Cardiac Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
INTRAOSSEOUS NEEDLE KITS | 2 | {'Appropriate Term / Code Not Available': 2} | {'Defective Component': 2}
EVIS EXERA DUODENOVIDEOSCOPE | 2 | {'Cardiac Arrest': 1, 'Septic Shock': 1, 'Cholangitis': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
XENMATRIX AB | 2 | {'Obstruction/Occlusion': 2, 'Fluid Discharge': 2, 'Adhesion(s)': 2, 'Abdominal Pain': 2, 'Hernia': 2} | {'Patient-Device Incompatibility': 2}
HAKIM PROGRAMMABLE VALVE | 2 | {'Appropriate Term / Code Not Available': 2} | {'Infusion or Flow Problem': 2}
REMSTAR AUTO A-FLEX W/HUMID, SD CARD | 2 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Cancer': 1} | {'Degraded': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
FUSION HOLLOW FIBER OXYGENATOR | 2 | {'Low Blood Pressure/ Hypotension': 1, 'Low Oxygen Saturation': 1, 'Ventilator Dependent': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Cardiac Arrest': 1} | {'Gas Output Problem': 1, 'Obstruction of Flow': 1, 'Fluid/Blood Leak': 1}
SPIDER FX | 2 | {'Stroke/CVA': 1, 'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
UMBILICAL CATHETER | 2 | {'Cardiac Tamponade': 1, 'Pericardial Effusion': 1, 'Foreign Body In Patient': 1, 'Extravasation': 1} | {'Positioning Problem': 2, 'Malposition of Device': 2}
SERVO-I UNIT BASE | 2 | {'Low Oxygen Saturation': 2} | {'Output Problem': 2}
GALAXY SYSTEM | 2 | {'Bradycardia': 1, 'Low Oxygen Saturation': 1, 'Bronchial Hemorrhage': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
STEALTHSTATION¿ S8 SYSTEM | 2 | {'Appropriate Term / Code Not Available': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Incorrect, Inadequate or Imprecise Result or Readings': 1}
AGILIA SP MC WIFI GB | 2 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Appropriate Term / Code Not Available': 1} | {'Infusion or Flow Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1, 'Excess Flow or Over-Infusion': 1}
REMSTAR AUTO SERIES AFLEX ASSY | 2 | {'Insufficient Information': 1, 'Unspecified Heart Problem': 1, 'Pulmonary Dysfunction': 1, 'Unspecified Respiratory Problem': 1, 'Cancer': 1} | {'Degraded': 2}
INJECTOMAT AGILIA BR | 2 | {'Respiratory Arrest': 1, 'Cardiac Arrest': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Insufficient Information': 1, 'Excess Flow or Over-Infusion': 1}
INTELLIVUE MX40 802.11A/B/G | 2 | {'Insufficient Information': 2, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Incorrect Measurement': 1}
INFINITY ACUTE CARE SYSTEM (M540) | 2 | {'Insufficient Information': 1, 'Elevated ketones/Diabetic Ketoacidosis': 1} | {'No Audible Prompt/Feedback': 1, 'No Device Output': 1}
CLOTTRIEVER XL CATHETER | 2 | {'Pulmonary Embolism': 2} | {'Difficult to Remove': 2, 'Use of Device Problem': 2}
CONTOUR 3D | 2 | {'Insufficient Information': 2, 'Pneumonia': 1} | {'Insufficient Information': 2, 'Adverse Event Without Identified Device or Use Problem': 1}
VIVO 45LS | 2 | {'Insufficient Information': 2} | {'Insufficient Information': 2}
PARKER | 2 | {'Decreased Respiratory Rate': 1, 'Unintended Extubation': 1, 'Diminished Pulse Pressure': 1, 'Low Oxygen Saturation': 1, 'Cardiac Arrest': 1} | {'Inflation Problem': 1, 'Mechanical Problem': 1, 'Gas/Air Leak': 1}
75MM SELECTABLE NEW TLC | 2 | {'Failure to Anastomose': 1, 'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Patient Device Interaction Problem': 1}
BIPAP AUTOSV ADVANCED SYSTEM ONE 60 SERIES | 2 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
REP DREAMSTATION AUTO CPAP, DOM | 2 | {'Insufficient Information': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
DAVINCI SI | 2 | {'Unspecified Tissue Injury': 2, 'Hemorrhage/Blood Loss/Bleeding': 2} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Insufficient Information': 1}
INTELLIVUE MULTI-MEASUREMENT MODULE X3 | 2 | {'Insufficient Information': 2} | {'Wireless Communication Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
CANNULA & CATHETER | 2 | {'Air Embolism': 2, 'Hypovolemic Shock': 2} | {'Disconnection': 2}
DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | 2 | {'No Clinical Signs, Symptoms or Conditions': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
INTELLIVUE MX40 WLAN | 2 | {'Insufficient Information': 2} | {'Erratic Results': 1, 'Insufficient Information': 1}
SERVO-S BASE UNIT | 2 | {'Respiratory Arrest': 2, 'Cardiac Arrest': 2} | {'Connection Problem': 2, 'No Flow': 2}
INTELLIVUE MP70 | 2 | {'Insufficient Information': 1, 'Sepsis': 1, 'Bradycardia': 1, 'Low Oxygen Saturation': 1} | {'Device Alarm System': 1, 'No Audible Alarm': 1}
PROLENE POLYPROPYLENE SUTURE UNKNOWN PRODUCT | 2 | {'Insufficient Information': 1, 'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
FMC CASSETTE | 2 | {'Sepsis': 2, 'Low Blood Pressure/ Hypotension': 2, 'Peritonitis': 2, 'Vomiting': 1, 'Nausea': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
HEMOSPRAY ENDOSCOPIC HEMOSTAT | 2 | {'Gastrointestinal Hemorrhage': 2, 'Perforation': 1, 'Abdominal Distention': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
ERBE VIO 3 | 2 | {'Bowel Perforation': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Appropriate Term/Code Not Available': 2}
UNKNOWN_VOCERA_PRODUCT | 2 | {'Insufficient Information': 2, 'No Clinical Signs, Symptoms or Conditions': 1} | {'No Audible Alarm': 1, 'Adverse Event Without Identified Device or Use Problem': 1, 'Device Alarm System': 1}
ZIO AT | 2 | {'Insufficient Information': 1, 'Unspecified Heart Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Insufficient Information': 1}
BD ALARIS PUMP MODULE SMARTSITE INFUSION SET | 2 | {'Brain Injury': 1, 'Heart Block': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Air/Gas in Device': 2}
DREAMSTATION BIPAP ST30 | 2 | {'Insufficient Information': 2, 'Pulmonary Dysfunction': 1, 'Myocardial Infarction': 1, 'Inflammation': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
SIMPLYGO | 2 | {'Insufficient Information': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
EPIQ CVXI | 2 | {'Cardiac Arrest': 2} | {'Application Program Freezes, Becomes Nonfunctional': 2}
COLONIC FTRD SET | 2 | {'Bowel Perforation': 2} | {'Use of Device Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
BENEVISION N1 PATIENT MONITOR | 2 | {'Insufficient Information': 1, 'Asystole': 1} | {'No Audible Alarm': 2, 'Device Alarm System': 2, 'Alarm Not Visible': 1}
EDWARDS ESHEATH+ INTRODUCER | 2 | {'Unspecified Vascular Problem': 2, 'Great Vessel Perforation': 1} | {'Difficult to Advance': 1, 'Material Puncture/Hole': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
DREAMSTATION BIPAP AVAPS30 | 2 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Cancer': 1, 'Hypothyroidism': 1, 'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Degraded': 1}
CATHERA | 2 | {'Hemorrhagic Stroke': 1, 'Ruptured Aneurysm': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
SHILEY | 2 | {'Unintended Extubation': 2, 'Insufficient Information': 1} | {'Gas/Air Leak': 2}
REP DREAMSTATION AUTO BIPAP, DOM-RECRT | 2 | {'Cancer': 1, 'Insufficient Information': 1} | {'Material Integrity Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
TURNPIKE SPIRAL 135CM | 2 | {'Foreign Body In Patient': 1, 'Obstruction/Occlusion': 1, 'Arrhythmia': 1} | {'Material Separation': 2}
HIGH FLOW INSUFFLATION UNIT | 2 | {'Low Blood Pressure/ Hypotension': 2, 'Cardiac Arrest': 2, 'Pulmonary Embolism': 1, 'Air Embolism': 1} | {'Pressure Problem': 1, 'Increase in Pressure': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
AVEIR¿ INTRODUCER KIT ¿ 50 CM | 2 | {'Low Blood Pressure/ Hypotension': 2, 'Arrhythmia': 1, 'Unspecified Infection': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Cardiac Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
ROTAREX | 2 | {'Appropriate Term / Code Not Available': 1, 'Heart Failure/Congestive Heart Failure': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Break': 1, 'Detachment of Device or Device Component': 1}
TRI-STAPLE 2.0 | 2 | {'Cardiac Arrest': 1, 'Appropriate Term / Code Not Available': 1, 'Tissue Breakdown': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Failure to Cut': 1, 'Failure to Form Staple': 1}
DREAMSTATION 2 ADVANCE AUTO CPAP | 2 | {'Insufficient Information': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
ZOOM REPERFUSION CATHETER | 2 | {'Device Embedded In Tissue or Plaque': 1, 'Foreign Body In Patient': 1, 'Vascular Dissection': 1, 'Extravasation': 1} | {'Break': 1, 'Material Separation': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
DREAMSTATION BIPAP | 2 | {'Cancer': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 2}
INVACARE® PLATINUM® OXYGEN CONCENTRATOR | 2 | {'Insufficient Information': 2} | {'Insufficient Information': 2, 'Adverse Event Without Identified Device or Use Problem': 1}
AIRSEAL | 2 | {'Exsanguination': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Gas Output Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
SOLITAIRE FR | 2 | {'Unspecified Eye / Vision Problem': 1, 'Foreign Body In Patient': 1, 'No Clinical Signs, Symptoms or Conditions': 1, 'Thrombosis/Thrombus': 1} | {'Break': 2, 'Physical Resistance/Sticking': 1}
PRLNE BLU 36IN 4-0 D/A SH | 2 | {'Wound Dehiscence': 2, 'Hemorrhage/Blood Loss/Bleeding': 2} | {'Break': 2}
SIMPLEX P WITH TOBRAMYCIN 1 PACK | 2 | {'Low Blood Pressure/ Hypotension': 2, 'Bone Cement Implantation Syndrome': 2, 'Cardiac Arrest': 2} | {'Patient Device Interaction Problem': 2}
NIPRO BLOOD TUBING SET W/TP | 2 | {'Low Blood Pressure/ Hypotension': 1, 'Anaphylactic Shock': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Fluid/Blood Leak': 1, 'Detachment of Device or Device Component': 1}
FORTILINK-C IBF SYSTEM | 2 | {'Pulmonary Embolism': 2, 'Cardiac Arrest': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
EVERFLO DOM | 2 | {'No Clinical Signs, Symptoms or Conditions': 2} | {'Adverse Event Without Identified Device or Use Problem': 2, 'Degraded': 2}
TREK | 2 | {'Unspecified Heart Problem': 2} | {'Adverse Event Without Identified Device or Use Problem': 2}
FILMARRAY® PNEUMONIA PANEL PLUS | 2 | {'Insufficient Information': 2} | {'False Positive Result': 1, 'False Negative Result': 1}
AERO® TRACHEOBRONCHIAL STENT | 2 | {'Perforation': 2, 'Hemorrhage/Blood Loss/Bleeding': 2, 'Cardiac Arrest': 2} | {'Difficult to Open or Close': 2, 'Activation, Positioning or Separation Problem': 2}
SEPRAMESH IP | 2 | {'Insufficient Information': 1, 'Bacterial Infection': 1, 'Fistula': 1, 'Abscess': 1, 'Adhesion(s)': 1} | {'Patient Device Interaction Problem': 2, 'Defective Device': 2}
NITREX | 2 | {'Hemoptysis': 1, 'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
VIZIGO | 2 | {'Hemorrhage/Blood Loss/Bleeding': 2, 'Anemia': 1, 'Hematoma': 1, 'Cardiac Tamponade': 1} | {'Adverse Event Without Identified Device or Use Problem': 2}
SENTRANT INTRODUCER SHEATH | 2 | {'Vascular Dissection': 1, 'Insufficient Information': 1} | {'Difficult to Remove': 1, 'Device Damaged by Another Device': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
APEXPRO TELEMETRY SYSTEM V6 | 1 | {'Insufficient Information': 1} | {'Appropriate Term/Code Not Available': 1, 'Display or Visual Feedback Problem': 1}
CHPV RT ANG RESERVOIR UNITIZED | 1 | {'Appropriate Term / Code Not Available': 1} | {'Infusion or Flow Problem': 1}
DAVINCI SINGLE PORT | 1 | {'Embolism/Embolus': 1, 'Respiratory Failure': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
POWERED WHEELCHAIR SYSTEM | 1 | {'Heart Failure/Congestive Heart Failure': 1} | {'Insufficient Information': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
QC HEMOSTATIC DRESSING, 2X2, 10X | 1 | {'Insufficient Information': 1} | {'Improper or Incorrect Procedure or Method': 1}
EPIQ CVX | 1 | {'Stroke/CVA': 1} | {'Application Program Freezes, Becomes Nonfunctional': 1}
MESH ¿ VENTRALEX | 1 | {'Insufficient Information': 1} | {'Patient Device Interaction Problem': 1, 'Defective Device': 1}
INTELLIPAP AUTO BILEVEL W/HEATED HUMIDIFIER | 1 | {'Unspecified Respiratory Problem': 1} | {'Detachment of Device or Device Component': 1}
SPYSCOPE DS II | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PENTAX MEDICAL | 1 | {'Respiratory Arrest': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PLUM 360 (V 15.X) | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Unexpected Shutdown': 1}
TRILOGY DEVICE | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
I-NEB CF UK UNIT | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 1 | {'Appropriate Term / Code Not Available': 1, 'Unspecified Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
HOMEFILL® OXYGEN SYSTEM | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
FLEXOR RAABE GUIDING SHEATH | 1 | {'Cardiac Arrest': 1} | {'Break': 1}
LATERAL F3D 18 MM IMPLANT; COBB ELEVATOR | 1 | {'Perforation of Vessels': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
COMBISET SMARTECH BLOODLINE (NO HEPARIN) | 1 | {'Exsanguination': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DREAMSTATION BIPAP AUTOSV, DS | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
INDIGO SYSTEM FLASH ASPIRATION CATHETER | 1 | {'Pulmonary Embolism': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EMBOTRAP III 6.5 MM X 45 MM | 1 | {'Intracranial Hemorrhage': 1} | {'Retraction Problem': 1}
UL_INJECT | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ULRICH MEDICAL | 1 | {'Anaphylactoid / Anaphylaxis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
AMPLATZER TORQVUE DELIVERY SYSTEM | 1 | {'Unspecified Tissue Injury': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Pericardial Effusion': 1, 'Cardiac Arrest': 1, 'Cardiac Tamponade': 1} | {'Use of Device Problem': 1}
ENDO CLIP III | 1 | {'Laceration(s)': 1, 'Peritonitis': 1} | {'Mechanics Altered': 1}
SAPPHIRE M.T INFUSION PUMP - GERMAN | 1 | {'Appropriate Term / Code Not Available': 1, 'Multiple Organ Dysfunction Syndrome': 1} | {'Excess Flow or Over-Infusion': 1, 'Use of Incorrect Control/Treatment Settings': 1}
TANDEM MOBI INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
TRILOGY EVO, USA | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
NOVUM IQ | 1 | {'Appropriate Term / Code Not Available': 1, 'No Clinical Signs, Symptoms or Conditions': 1, 'Cardiac Arrest': 1} | {'Filling Problem': 1}
REMSTAR AUTO A-FLEXX | 1 | {'Cancer': 1} | {'Degraded': 1}
REVEAL LINQ | 1 | {'Ventricular Fibrillation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
B125 | 1 | {'Insufficient Information': 1} | {'Defective Alarm': 1}
POWERTRIALYSIS | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Device Dislodged or Dislocated': 1}
GIRAFFE OMNIBED | 1 | {'Fungal Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CIRCULAR MECH STD 29MM | 1 | {'Burn(s)': 1, 'Failure to Anastomose': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
COOLSEAL¿ GENERATOR | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
COOLSEAL¿ TRINITY | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
RHYTHM ANALYTICS - SOFTWARE (SAMD) | 1 | {'Tachycardia': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1}
HARMONIC 1100 SHEARS 36CM | 1 | {'Burn(s)': 1, 'Failure to Anastomose': 1, 'Cardiac Arrest': 1} | {'Patient Device Interaction Problem': 1}
ENSEAL X1 CURVED 25CM SHAFT | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
SYSTEM ONE HTD HUMID ASSY ,DOM | 1 | {'Skin Inflammation/ Irritation': 1, 'Dizziness': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1}
RESUSCITAIRE RADIANT WARMER | 1 | {'Insufficient Information': 1} | {'Gas Output Problem': 1, 'Infusion or Flow Problem': 1}
TI-CRON | 1 | {'Insufficient Information': 1} | {'Expiration Date Error': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
FILMARRAY® MENINGITIS/ENCEPHALITIS (ME) PANEL | 1 | {'Insufficient Information': 1} | {'False Negative Result': 1}
INTRAOSSEOUS, INTRAOSSEOUS NEEDLE KITS | 1 | {'Cardiac Arrest': 1} | {'Material Twisted/Bent': 1}
DEVILBISS 525 OXYGEN CONCENTRATOR | 1 | {'Burn(s)': 1} | {'Insufficient Information': 1}
HARMONIC ACE 5MM SHEAR | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Patient Device Interaction Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
INTRI24 INTRODUCER SHEATH | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Improper or Incorrect Procedure or Method': 1}
TRILOGY O2 VENTILATOR, JAPAN | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
FEEDING TUBE | 1 | {'Perforation': 1} | {'Material Perforation': 1}
PROWATERFLEX | 1 | {'Foreign Body In Patient': 1} | {'Material Separation': 1}
TRINAV INFUSION SYSTEM | 1 | {'Stroke/CVA': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PHOTODYNAMIC BONE STABILIZATION SYSTEM | 1 | {'Embolism/Embolus': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
REFLEX CATHETER | 1 | {'Renal Failure': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
LINEAR 7.5FR. 40CC IAB WITH ACCESSORIES, US ONLY | 1 | {'Insufficient Information': 1} | {'Fluid/Blood Leak': 1, 'Material Rupture': 1}
NOXBOXI NITRIC OXIDE DELIVERY SYSTEM | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
LEAD EXTRACTION EVOLUTION RL CONTROLLED-ROTATION DILATOR SHEATH SET | 1 | {'Great Vessel Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
SABLE | 1 | {'Appropriate Term / Code Not Available': 1} | {'Insufficient Information': 1}
HOBBS MEDICAL TOUCHSOFT COAGULATOR | 1 | {'Perforation': 1, 'Unspecified Infection': 1, 'Abdominal Pain': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PARAPAC PLUS | 1 | {'Respiratory Failure': 1} | {'Defective Alarm': 1, 'Contamination of Device Ingredient or Reagent': 1}
MEDIHONEY GEL IN TUBE, .5 FL OZ TWISTOFF - STERILE | 1 | {'Pain': 1} | {'Lack of Effect': 1}
DA VINCI ENERGY | 1 | {'Perforation of Vessels': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CAPIOX | 1 | {'Insufficient Information': 1} | {'Failure to Deliver': 1}
ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Insufficient Information': 1}
DREAMSTATION BIPAP AVAPS | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
INQUIRY¿ STEERABLE DIAGNOSTIC CATHETER | 1 | {'Arrhythmia': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP A40 | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
MINICAP WITH POVIDONE-IODINE SOLUTION | 1 | {'Unspecified Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
LIFE SCOPE PT | 1 | {'Sepsis': 1} | {'Protective Measures Problem': 1, 'No Audible Alarm': 1}
SAVVYWIRE | 1 | {'Cardiac Arrest': 1} | {'Insufficient Information': 1, 'Capturing Problem': 1}
PENDITURE LAA EXCLUSION SYSTEM | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
COMPANION 5 | 1 | {'Burn(s)': 1} | {'Use of Device Problem': 1}
CORTRAK 2 NG/NI FEED, ELECTROMAGNETIC STYLET, ENFIT, CONNECTOR | 1 | {'Foreign Body In Patient': 1} | {'Break': 1, 'Malposition of Device': 1}
INTELLISPACE EVENT MGMT PLATFORM | 1 | {'Cardiac Arrest': 1} | {'No Audible Alarm': 1}
CRAGG-MCNAMARA | 1 | {'Ischemia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DISTAL PERFUSION KIT | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CLEARVIEW BLOWER/ MISTER | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CLEARVIEW ARTERIOTOMY SHUNT | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
INJECTOR OPVSU NPS | 1 | {'Appropriate Term / Code Not Available': 1} | {'Use of Device Problem': 1}
NIO ADULT | 1 | {'Insufficient Information': 1} | {'Firing Problem': 1}
AGILIA VP MC BR | 1 | {'Appropriate Term / Code Not Available': 1} | {'No Audible Alarm': 1}
PHILIPS RESPIRONICS E30 WITH HUMIDIFIER | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
LIBERTY CYCLER SET, SINGLE CONN./EXT. DL | 1 | {'Peritonitis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ANGIOVAC | 1 | {'Pulmonary Embolism': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DREAMST ST30 H | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIOFIRE® BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL | 1 | {'Insufficient Information': 1} | {'False Negative Result': 1}
INTELLIVUE X3 | 1 | {'Bradycardia': 1} | {'No Audible Alarm': 1}
RADIOGRAPHY 7300 C | 1 | {'Fall': 1, 'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIPOLAR VENTRICULAR ENDOCARDIAL LEAD | 1 | {'Vascular Dissection': 1, 'Low Blood Pressure/ Hypotension': 1, 'Thrombosis/Thrombus': 1, 'Unspecified Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
MERGE CARDIO | 1 | {'Insufficient Information': 1, 'Unspecified Heart Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
INTRACEPT RF PROBE | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BENEVISION N17 PATIENT MONITOR | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1, 'Output Problem': 1, 'No Visual Prompts/Feedback': 1}
HEMOSTAR | 1 | {'Delirium': 1, 'Air Embolism': 1, 'Dyspnea': 1} | {'Air/Gas in Device': 1, 'Fracture': 1, 'Material Discolored': 1, 'Material Split, Cut or Torn': 1}
SENSATION PLUS 7.5FR. 40CC IAB & ACCESSORIES | 1 | {'Insufficient Information': 1} | {'Fluid/Blood Leak': 1, 'Leak/Splash': 1}
DJO SURGICAL | 1 | {'Failure of Implant': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EDISON SYSTEM | 1 | {'Renal Impairment': 1, 'Low Blood Pressure/ Hypotension': 1, 'Urinary Retention': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
HOMECHOICE APD SYSTEM | 1 | {'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
RADICAL-7 | 1 | {'Hypoxia': 1} | {'Application Program Freezes, Becomes Nonfunctional': 1}
PORT-A-CATH IMPLANTABLE ACCESS SYSTEM | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
AMPLATZER VASCULAR PLUG | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
TRANS-RAY PLUS 7.5FR 35CC IAB WITH ACCESSORIES | 1 | {'Myocardial Infarction': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1}
CARDIOSAVE HYBRID, 3.1 EDITION | 1 | {'Myocardial Infarction': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1}
AEROGEN SOLO | 1 | {'Hypoxia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
INTELLIVUE MX40 1.4GHZ | 1 | {'Insufficient Information': 1} | {'Device Alarm System': 1}
PU-681RA | 1 | {'Low Blood Pressure/ Hypotension': 1} | {'Appropriate Term/Code Not Available': 1}
ARROW CVC SET: 4-LUMEN 8.5FR X 16CM | 1 | {'Embolism/Embolus': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | 1 | {'Elevated ketones/Diabetic Ketoacidosis': 1, 'Hyperglycemia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
MEDTRONIC NAVIGATION | 1 | {'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
K-PACK II | 1 | {'Shock': 1} | {'Delivered as Unsterile Product': 1}
TRILOGY 202 VENTILATOR, INTERNATIONAL | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CHORDX UNKNOWN | 1 | {'Mitral Valve Insufficiency/ Regurgitation': 1} | {'Insufficient Information': 1}
ASTRAL 150 - AMER | 1 | {'Insufficient Information': 1} | {'Therapeutic or Diagnostic Output Failure': 1, 'No Audible Alarm': 1}
SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER, SL1¿, 63 CM LENGTH, 8.5F | 1 | {'Air Embolism': 1, 'Stroke/CVA': 1} | {'Gas/Air Leak': 1}
CAPIOX®FX | 1 | {'Low Blood Pressure/ Hypotension': 1} | {'Fluid/Blood Leak': 1}
PANORAMA CENTRAL STATION WITH TELEMETRY | 1 | {'Insufficient Information': 1} | {'Device Handling Problem': 1}
BEAD BLOCK | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
INTELLIVUE MX550 PATIENT MONITOR | 1 | {'Insufficient Information': 1} | {'Defective Alarm': 1}
LIBERTY CYCLER SET, DUAL PATIENT CONNECT | 1 | {'Swelling/ Edema': 1, 'Sepsis': 1, 'Abdominal Pain': 1, 'Dyspnea': 1, 'Peritonitis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | 1 | {'Insufficient Information': 1} | {'Fracture': 1}
INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | 1 | {'Unspecified Respiratory Problem': 1} | {'Insufficient Information': 1}
FREESTYLE LIBRE 3 | 1 | {'Vomiting': 1, 'Nausea': 1, 'Hyperglycemia': 1, 'Fatigue': 1} | {'Low Readings': 1}
CIRCULO | 1 | {'Arrhythmia': 1, 'Cardiac Perforation': 1, 'Cardiac Tamponade': 1, 'Pericardial Effusion': 1} | {'Insufficient Information': 1, 'Use of Device Problem': 1}
AVEIR¿ RETRIEVAL SYSTEM CATHETER | 1 | {'Appropriate Term / Code Not Available': 1} | {'Connection Problem': 1}
INTRODUCER KIT FOR GASTROSTOMY FEEDING TUBE | 1 | {'Aspiration Pneumonitis': 1} | {'Break': 1, 'Malposition of Device': 1}
CG FUTURE BAND | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
MESH - VENTRALEX | 1 | {'Bacterial Infection': 1, 'Fistula': 1, 'Abscess': 1} | {'Patient-Device Incompatibility': 1}
AGILIA SP MC BR | 1 | {'Respiratory Arrest': 1, 'Cardiac Arrest': 1} | {'Insufficient Information': 1}
AVALON FETAL MONITOR FM30 | 1 | {'Hypoxia in Utero': 1} | {'Incorrect Measurement': 1}
HERO® | 1 | {'Liver Damage/Dysfunction': 1, 'Thrombosis/Thrombus': 1, 'Unspecified Kidney or Urinary Problem': 1, 'Unspecified Hepatic or Biliary Problem': 1} | {'Obstruction of Flow': 1}
MAGNETOM ESPREE | 1 | {'Airway Obstruction': 1} | {'Use of Device Problem': 1}
BIPAP S/T C SERIES | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
TELESCOPE GUIDE EXTENSION CATHETER | 1 | {'Device Embedded In Tissue or Plaque': 1} | {'Difficult to Remove': 1, 'Detachment of Device or Device Component': 1}
ECHELON 3000 60MM STANDARD | 1 | {'Unspecified Tissue Injury': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
VACCESS CT LOW-PROFILE POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 8F | 1 | {'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | 1 | {'Bacterial Infection': 1, 'Sepsis': 1, 'Unspecified Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
MEGA 7.5FR. 40CC IAB | 1 | {'Myocardial Infarction': 1} | {'Migration or Expulsion of Device': 1}
QUANTUM WORKSTATION 12.1" ELITE | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
COLLECTIVE CONCENTRATE- NATURALYTE | 1 | {'Non specific EKG/ECG Changes': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EVERFLO 230V, OPI, AUS, ULTRAFILL | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PORTEX BLUSELECT SUCTIONAID CUFFED TRACHEOSTOMY | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP PRO BIFLEX, W/SD CARD, US-RENTAL | 1 | {'Insufficient Information': 1} | {'Degraded': 1}
0.14 GUIDEWIRES | 1 | {'Ischemia Stroke': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
GORE® MOLDING & OCCLUSION BALLOON CATHETER | 1 | {'Endoleaks': 1, 'Rupture': 1} | {'Use of Device Problem': 1}
CS100 INTRA-AORTIC BALLOON PUMP, ENGLISH, 220V | 1 | {'Insufficient Information': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1}
TRILOGY EVO, O2, JAPAN | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DRAGONFLY OPSTAR¿ | 1 | {'Vascular Dissection': 1, 'Cardiac Arrest': 1} | {'Difficult to Advance': 1}
TANDEM MOBI INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | 1 | {'Elevated ketones/Diabetic Ketoacidosis': 1, 'Myocardial Infarction': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CATH F5.2+ PIG 145° 110CM 6SH | 1 | {'Insufficient Information': 1} | {'Leak/Splash': 1}
BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PREMICATH | 1 | {'Insufficient Information': 1} | {'Fluid/Blood Leak': 1}
NIPRO SURDIAL DX HEMODIALYSIS SYSTEM | 1 | {'Low Blood Pressure/ Hypotension': 1, 'Anaphylactic Shock': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
NAVIGATOR HD | 1 | {'Sepsis': 1, 'Septic Shock': 1, 'Low Blood Pressure/ Hypotension': 1, 'Ischemia': 1, 'Abdominal Pain': 1} | {'Peeled/Delaminated': 1, 'Material Split, Cut or Torn': 1}
REMSTAR PRO SERIES C-FLEX + ASSY | 1 | {'Cancer': 1, 'Unspecified Hepatic or Biliary Problem': 1} | {'Degraded': 1}
DREAMSTATION BIPAP S/T 30 W/HEATED HUMIDIFIER, GB | 1 | {'Unspecified Respiratory Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
OPTICROSS HD | 1 | {'Foreign Body In Patient': 1} | {'Physical Resistance/Sticking': 1, 'Detachment of Device or Device Component': 1}
STEERABLE GUIDE CATHETER | 1 | {'Perforation': 1, 'Pericardial Effusion': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
OMNIPOD 5 AUTOMATED INSULIN DELIVERY SYSTEM | 1 | {'Elevated ketones/Diabetic Ketoacidosis': 1} | {'Insufficient Information': 1}
OPTICROSS? HD | 1 | {'Non specific EKG/ECG Changes': 1, 'Air Embolism': 1, 'Reduced Blood Flow': 1, 'Cardiac Arrest': 1} | {'Poor Quality Image': 1}
MEDICALYTE | 1 | {'Bacteremia': 1} | {'Device Contaminated During Manufacture or Shipping': 1}
EVERFLO INTL OPI | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
LIFEPAK 15 DEFIBRILLATOR/MONITOR | 1 | {'Cardiac Arrest': 1} | {'Device Sensing Problem': 1}
INNOSIGHT | 1 | {'Insufficient Information': 1} | {'Poor Quality Image': 1}
DREAM STATION AUTO BIPAP | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
SIMUFORM | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 1 | {'Pain': 1, 'Insufficient Information': 1, 'Thrombosis/Thrombus': 1, 'Pulmonary Embolism': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CLOTTRIEVER BOLD CATHETER | 1 | {'Pulmonary Embolism': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EUPHORA RX | 1 | {'Insufficient Information': 1, 'Device Embedded In Tissue or Plaque': 1} | {'Entrapment of Device': 1, 'Difficult to Remove': 1, 'Detachment of Device or Device Component': 1, 'Deflation Problem': 1}
PORTEX ENDOTRACHEAL TUBE, SILICONIZED | 1 | {'Hydrocephalus': 1, 'Intracranial Hemorrhage': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CS300 INTRA-AORTIC BALLOON PUMP, SPANISH, 110V | 1 | {'Insufficient Information': 1} | {'Unexpected Shutdown': 1}
MRIDIUM | 1 | {'Insufficient Information': 1} | {'Improper or Incorrect Procedure or Method': 1}
REMSTAR PRO C-FLEX+ SYSTEM ONE 60 SERIES | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
TOUCHSOFT COAGULATOR | 1 | {'Bowel Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP AUTOSV ADVANCED | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIOFLO VORTEX | 1 | {'Bacterial Infection': 1, 'Endocarditis': 1, 'Muscle Weakness/Atrophy': 1, 'Pericardial Effusion': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EXTRACTOR PRO RX | 1 | {'Perforation': 1, 'Peritonitis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
COBLATION HALO WAND | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
HEMOSIL HIT- AB(PF4-H) | 1 | {'Thrombosis/Thrombus': 1, 'Stroke/CVA': 1} | {'False Negative Result': 1}
BIPAP PRO SERIES ASSY | 1 | {'Vomiting': 1, 'Eye Irritation': 1, 'Asthma': 1, 'Dizziness': 1, 'Nausea': 1} | {'Degraded': 1}
POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS | 1 | {'Hypoxia': 1} | {'Fluid/Blood Leak': 1}
VISISHEATH DILATOR SHEATH | 1 | {'Great Vessel Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
AVALON FM50 FETAL MONITOR | 1 | {'Encephalopathy': 1, 'Hypoxia in Utero': 1} | {'Unable to Obtain Readings': 1}
MYOSURE TISSUE REMOVAL DEVICE | 1 | {'Arrhythmia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PROTEK DUO VENOUS CANNULA | 1 | {'Hypovolemic Shock': 1} | {'Patient Device Interaction Problem': 1}
REMSTAR PLUS W/C-FLEX | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PERCUFLEX URINARY DIVERSION STENT SET | 1 | {'Sepsis': 1, 'Foreign Body In Patient': 1} | {'Break': 1, 'Difficult to Remove': 1}
BIA-GEN | 1 | {'Swelling/ Edema': 1, 'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CNS-2101 CENTRAL MONITOR | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Protective Measures Problem': 1, 'No Audible Alarm': 1}
NIPRO ELISIO-H DIALYZER | 1 | {'Low Blood Pressure/ Hypotension': 1, 'Anaphylactic Shock': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
MAIN ASSEMBLY, VENTILATOR, COMMERCIAL, ZVENT, BASI | 1 | {'Insufficient Information': 1} | {'Failure of Device to Self-Test': 1}
INTELLIVUE PATIENT MONITOR MX750 | 1 | {'Insufficient Information': 1, 'Respiratory Failure': 1} | {'Defective Alarm': 1}
GORE® TRI LUMEN CATHETER | 1 | {'Ischemia Stroke': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
VALLEYLAB | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1, 'Hypovolemic Shock': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
OP TELESCOPE, 7°, 11 FR. X 220 MM, 7.5 FR. CHANNEL, ANGLED OCULAR | 1 | {'Sepsis': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Low Blood Pressure/ Hypotension': 1, 'Ventilator Dependent': 1, 'Septic Shock': 1} | {'Material Twisted/Bent': 1}
ANGIOJET? SOLENT? OMNI | 1 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Ischemia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
WALLFLEX ESOPHAGEAL | 1 | {'Aspiration/Inhalation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 1 | {'Unspecified Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
UNKNOWN_NEUROVASCULAR_PRODUCT | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DURAMATRIX-ONLAY COLLAGEN DURA SUBSTITUTE MEMBRANE | 1 | {'Septic Shock': 1, 'Low Blood Pressure/ Hypotension': 1} | {'Insufficient Information': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP A40 PRO | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
KANGAROO | 1 | {'Dehydration': 1} | {'Insufficient Flow or Under Infusion': 1}
PULSE | 1 | {'Sepsis': 1, 'Multiple Organ Dysfunction Syndrome': 1, 'Pneumonia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | 1 | {'Perforation of Vessels': 1} | {'Difficult to Remove': 1, 'Physical Resistance/Sticking': 1, 'Material Puncture/Hole': 1, 'Difficult to Insert': 1}
CS300 INTRA-AORTIC BALLOON PUMP, PORTUGUESE, 220V | 1 | {'Insufficient Information': 1} | {'No Device Output': 1}
SOLITAIRE X | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1, 'Myocardial Infarction': 1, 'Cerebral Edema': 1, 'Ischemia Stroke': 1, 'Intracranial Hemorrhage': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PRLNE BLU 36IN 2-0 D/A V-7 MP | 1 | {'Insufficient Information': 1} | {'Break': 1}
ECHELON FLEX 60 POWERED PLUS ENDOPATH STAPLER ARTICULATING ENDOSCOPIC LINEAR CUT | 1 | {'Insufficient Information': 1} | {'Patient Device Interaction Problem': 1}
CAPIOX FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Use of Device Problem': 1}
CAPTIVATOR EMR | 1 | {'Perforation': 1, 'Pain': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
REMSTAR PRO C-FLEX+ W/HUMID, SD CARD | 1 | {'Unspecified Kidney or Urinary Problem': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1}
COCHLEAR OSIA OSI200 IMPLANT | 1 | {'Appropriate Term / Code Not Available': 1} | {'Appropriate Term/Code Not Available': 1}
OBSIDIO | 1 | {'Embolism/Embolus': 1, 'Necrosis': 1, 'Multiple Organ Dysfunction Syndrome': 1, 'Unspecified Kidney or Urinary Problem': 1, 'Abdominal Pain': 1} | {'Material Separation': 1}
LUCAS¿ CPR CHEST COMPRESSION SYSTEM | 1 | {'Cardiac Arrest': 1} | {'Detachment of Device or Device Component': 1}
UNKNOWN RENASYS TOUCH DEVICE | 1 | {'Exsanguination': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PRESSIO ICP AND ICT MONITORING KIT, PARENCHYMAL WITH BOLT | 1 | {'Intracranial Pressure Increased': 1} | {'Unstable': 1}
PRLNE BLU 24IN 6-0 D/A BV-1 | 1 | {'Insufficient Information': 1, 'Intraoperative Cardiac Injury': 1} | {'Break': 1}
ALLURA XPER FD | 1 | {'Insufficient Information': 1} | {'Radiation Output Failure': 1}
TRICLIP G4 SYSTEM | 1 | {'Pain': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Fatigue': 1, 'Anemia': 1, 'Hematoma': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
LINEAR 7.5FR. 40CC IAB | 1 | {'Cardiac Arrest': 1} | {'Inability to Auto-Fill': 1}
FLOWTRIEVER CATHETER, XL | 1 | {'Pulmonary Embolism': 1} | {'Difficult to Remove': 1, 'Use of Device Problem': 1}
INTELLIVUE MX500 PATIENT MONITOR | 1 | {'Insufficient Information': 1} | {'No Audible Alarm': 1}
INTELLIVUE MX40 2.4GHZ | 1 | {'Asystole': 1} | {'Complete Loss of Power': 1}
TEMPUS LS-MANUAL | 1 | {'No Clinical Signs, Symptoms or Conditions': 1, 'Unspecified Heart Problem': 1} | {'No Pacing': 1}
INFINITY CENTRAL STATION | 1 | {'Unspecified Heart Problem': 1} | {'Device Handling Problem': 1, 'No Audible Prompt/Feedback': 1}
REMSTAR PRO SERIES | 1 | {'Cancer': 1, 'Unspecified Hepatic or Biliary Problem': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1}
OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | 1 | {'Hypoglycemia': 1} | {'Insufficient Information': 1}
REM STAR AUTO A-FLEX | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
CVAC ASPIRATION SYSTEM | 1 | {'Pain': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Low Blood Pressure/ Hypotension': 1, 'Disseminated Intravascular Coagulation (DIC)': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
STAR¿ DRIVE MA | 1 | {'Post Operative Wound Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PACIFIC PLUS | 1 | {'Insufficient Information': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Detachment of Device or Device Component': 1, 'Burst Container or Vessel': 1}
DREAMSTATION BIPAP AVAP | 1 | {'Unspecified Gastrointestinal Problem': 1, 'Unspecified Kidney or Urinary Problem': 1, 'Unspecified Respiratory Problem': 1, 'Cancer': 1, 'Unspecified Hepatic or Biliary Problem': 1} | {'Degraded': 1}
DREAMSTATION HUMIDIFIER | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP AUTO BI-FLEX | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
PERFIX PLUG LIGHT | 1 | {'Obstruction/Occlusion': 1, 'Urinary Retention': 1} | {'Patient Device Interaction Problem': 1, 'Defective Device': 1}
IPV-2C | 1 | {'Pneumothorax': 1, 'Cardiac Arrest': 1} | {'Use of Device Problem': 1}
AVALON FETAL MONITOR FM50 | 1 | {'Insufficient Information': 1, 'Hypoxia': 1} | {'Insufficient Information': 1, 'Incorrect, Inadequate or Imprecise Result or Readings': 1}
EPSTAR FIXED ELECTROPHYSIOLOGY CATHETER | 1 | {'Cardiac Perforation': 1, 'Low Blood Pressure/ Hypotension': 1, 'Cardiac Tamponade': 1, 'Pericardial Effusion': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ADMIRAL XTREME OTW | 1 | {'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EVERCROSS 035 | 1 | {'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
NANOCROSS 014 | 1 | {'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BTK TURBOHAWK | 1 | {'Sepsis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
STREAMLINE TEMPORARY PACING LEADS | 1 | {'Vomiting': 1, 'Heart Block': 1, 'Multiple Organ Dysfunction Syndrome': 1, 'Cardiac Arrest': 1, 'Septic Shock': 1} | {'Patient Device Interaction Problem': 1}
VISUALASE SYSTEM | 1 | {'Swelling/ Edema': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Loss of Range of Motion': 1, 'Cardiac Arrest': 1, 'Cancer': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ESHEATH PLUS INTRODUCER SHEATH | 1 | {'Rupture': 1} | {'Failure to Advance': 1, 'Material Puncture/Hole': 1}
POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | 1 | {'Septic Shock': 1, 'Sepsis': 1, 'Bacteremia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
HEATER/COOLER SYSTEM | 1 | {'Bacterial Infection': 1} | {'Microbial Contamination of Device': 1}
MO.MA ULTRA | 1 | {'Obstruction/Occlusion': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Stroke/CVA': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER SL 1 CURVE, 63 CM LENGTH, 8F | 1 | {'Cardiac Perforation': 1, 'Great Vessel Perforation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EMBOSHIELD NAV6 | 1 | {'Unspecified Vascular Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
STARTER GUIDEWIRE | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Insufficient Information': 1}
BENTSON TYPE PTFE COATED GUIDEWIRE | 1 | {'Appropriate Term / Code Not Available': 1} | {'Insufficient Information': 1}
BIPAP AUTO BIFLEX, | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
LEVEEN SUPERSLIM | 1 | {'Bradycardia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CERELINK ICP PROBE 1L M BOLT | 1 | {'Visual Disturbances': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1}
LEAD EXTRACTION BYRD POLYPROPYLENE DILATOR SHEATH SET | 1 | {'Cardiac Perforation': 1, 'Low Blood Pressure/ Hypotension': 1, 'Pericardial Effusion': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
IMPEDE-FX EMBOLIZATION PLUG | 1 | {'Ruptured Aneurysm': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
REP DREAMSTATION AUTO CPAP, | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CERELINK ICP MONITOR | 1 | {'Visual Disturbances': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1}
CRESCENT RA DUAL LUMEN CATHETER 13FR | 1 | {'Air Embolism': 1} | {'Fracture': 1}
MC3 NAUTILUS ECMO OXYGENATOR | 1 | {'Appropriate Term / Code Not Available': 1} | {'Infusion or Flow Problem': 1}
SIMPLYGO MINI | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
HAMILTON-T1 | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
BIPAP PRO BI-FLEX | 1 | {'Unspecified Kidney or Urinary Problem': 1} | {'Degraded': 1}
LTV 1150 VENTILATOR | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Patient Device Interaction Problem': 1}
SENSATION 7FR. 40CC IAB | 1 | {'Appropriate Term / Code Not Available': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Leak/Splash': 1, 'Gas/Air Leak': 1}
HEMOCHRON ACTIVATED CLOTTING TIME PLUS (ACT+) TEST | 1 | {'Ischemic Heart Disease': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PROXIMATE*75MM LINEAR CUTTER | 1 | {'Ischemia': 1, 'Failure to Anastomose': 1, 'Sepsis': 1} | {'Patient Device Interaction Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP AUTOSV ADV SYS ONE | 1 | {'Cancer': 1} | {'Degraded': 1}
RING PROFILE 3D 30MM | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
INTELLIVUE MP30 | 1 | {'Insufficient Information': 1} | {'Poor Quality Image': 1}
EDWARDS TRANSFEMORAL BALLOON CATHETER | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
INTELLAMAP ORION HIGH RESOLUTION MAPPING CATHETER | 1 | {'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BRILLIANCE | 1 | {'Respiratory Insufficiency': 1} | {'Intermittent Communication Failure': 1}
POWERPORT VUE M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | 1 | {'Erythema': 1, 'Insufficient Information': 1, 'Pain': 1} | {'Fluid/Blood Leak': 1, 'Fracture': 1}
STRATTICE LAPAROSCOPIC 10CM X 16CM | 1 | {'Appropriate Term / Code Not Available': 1, 'Fistula': 1, 'Abscess': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIVONA TRACHEOSTOMY TUBE, NEONATAL/PEDIATRIC | 1 | {'Respiratory Insufficiency': 1} | {'Defective Device': 1}
VERSYS FEMORAL STEM CEMENTED REVISION/CALCAR 12/14 NECK TAPER SIZE 13 250MM | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ESSENCE RENTAL WITH SSD | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
ESSENCE WITH SSD | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
LINEAR 7.5FR. 40CC IAB (CHINA) | 1 | {'Appropriate Term / Code Not Available': 1} | {'Pressure Problem': 1}
2008T HEMODIALYSIS SYS., WITH CDX | 1 | {'Tachycardia': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
NEOTRACT UROLIFT 2 SYSTEM | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP AUTOSV ADVANCED SYSTEM ONE | 1 | {'Insufficient Information': 1} | {'Degraded': 1}
2008T MACHINE | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ENDO GIA | 1 | {'Tissue Breakdown': 1, 'Fistula': 1, 'Bowel Perforation': 1, 'Failure to Anastomose': 1, 'Unspecified Infection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
REMSTAR PRO C FLEX | 1 | {'Insufficient Information': 1} | {'Degraded': 1}
TABLO HEMODIALYSIS SYSTEM | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
REMSTAR LITE | 1 | {'Vomiting': 1, 'Eye Irritation': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1}
REM AUTO A-FLEX | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
SJM SÃ©GUIN SEMI-RIGID ANNULOPLASTY RING | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
SJM RIGID SADDLE RING | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DREAMSTATION AUTO ,CPAP | 1 | {'Eye Irritation': 1, 'Asthma': 1, 'Dizziness': 1, 'Inflammation': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1}
REMSTAR AUTO SERIES AFLEX | 1 | {'Myocardial Infarction': 1, 'Pneumonia': 1} | {'Degraded': 1}
DREAMSTATION HUMID | 1 | {'Cancer': 1} | {'Degraded': 1}
LIFEPAK® 15 DEFIBRILLATOR/MONITOR | 1 | {'Cardiac Arrest': 1} | {'No Audible Prompt/Feedback': 1}
VUE PACS | 1 | {'Appropriate Term / Code Not Available': 1, 'Gastrointestinal Hemorrhage': 1} | {'Insufficient Information': 1, 'Application Program Problem': 1}
XENOSURE BIOLOGIC PATCH | 1 | {'Ventricular Fibrillation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
REMSTAR AUTO A FLEX | 1 | {'Cancer': 1, 'Unspecified Hepatic or Biliary Problem': 1} | {'Degraded': 1}
HAMILTON-C1 | 1 | {'Low Oxygen Saturation': 1} | {'Gas Output Problem': 1}
MX40 PATIENT WEARABLE MONITOR | 1 | {'Insufficient Information': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Failure to Transmit Record': 1, 'No Device Output': 1}
MARK 5 NUVO LITE | 1 | {'Burn(s)': 1} | {'Improper or Incorrect Procedure or Method': 1}
BIPAP AUTO SV ADV SYSTEM ONE | 1 | {'Cancer': 1, 'Headache': 1, 'Dizziness': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1}
BIPAP AUTO BI FLEX | 1 | {'Eye Irritation': 1, 'Cardiovascular Dysfunction/Insufficiency': 1, 'Dizziness': 1, 'Skin Inflammation/ Irritation': 1, 'Unspecified Respiratory Problem': 1} | {'Degraded': 1}
REMSTARAUTO | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
RECERT, DREAMSTATION AUTO, FR | 1 | {'Insufficient Information': 1} | {'Degraded': 1}
AVALON FM30 FETAL MONITOR | 1 | {'Insufficient Information': 1} | {'Device Alarm System': 1}
2008T HEMODIALYSIS SYSTEM W/BIBAG | 1 | {'Loss of consciousness': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER | 1 | {'Vascular Dissection': 1, 'Thrombosis/Thrombus': 1, 'Stroke/CVA': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
SENDIT DELIVERY DEVICE | 1 | {'Vascular Dissection': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
FLAT CUT TALAR DOME SZ 2 INFINITY ADAPTIS | 1 | {'Skin Tears': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
POWERPORT ISP MRI 8CF INT. W/SP, ATT, SL | 1 | {'Insufficient Information': 1, 'Low Blood Pressure/ Hypotension': 1, 'Dyspnea': 1, 'Pleural Effusion': 1} | {'Migration': 1, 'Disconnection': 1}
ENDOWRIST | 1 | {'Perforation of Vessels': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Failure to Deliver Energy': 1}
POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | 1 | {'Caustic/Chemical Burns': 1} | {'Fluid/Blood Leak': 1, 'Disconnection': 1}
6VT-D | 1 | {'Perforation of Esophagus': 1} | {'Overheating of Device': 1, 'Positioning Problem': 1}
MONARCH ENDOSCOPY PLATFORM | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EDWARDS ESHEATH PLUS | 1 | {'Vascular Dissection': 1} | {'Difficult to Insert': 1}
TD60 TELEMETRY TRANSMITTER | 1 | {'Insufficient Information': 1} | {'Device Alarm System': 1}
PLUM 360 | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CAREEVENT | 1 | {'Insufficient Information': 1} | {'No Audible Alarm': 1}
ARISTOTLE 24 | 1 | {} | {}
EVITA 4 | 1 | {'Insufficient Information': 1} | {'Unexpected Therapeutic Results': 1, 'Human-Device Interface Problem': 1}
MILLENNIUM M600 WITHOUT OPI DOMESTIC | 1 | {'Unspecified Heart Problem': 1, 'Dizziness': 1, 'Unspecified Kidney or Urinary Problem': 1, 'Unspecified Respiratory Problem': 1, 'Headache': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EMERGE? | 1 | {'Low Blood Pressure/ Hypotension': 1, 'Bradycardia': 1, 'Foreign Body In Patient': 1, 'Chest Pain': 1} | {'Detachment of Device or Device Component': 1}
S5 HEART-LUNG-MACHINE | 1 | {'Air Embolism': 1} | {'No Apparent Adverse Event': 1}
INTELLIVUE MP5 | 1 | {'Insufficient Information': 1} | {'Defective Alarm': 1}
STELLANT DUAL CT INJECTOR | 1 | {} | {}
OPTIMA IGS 320 | 1 | {} | {}
MICROTARGETING¿ ELECTRODE | 1 | {} | {}
TELEHEALTH CENTER PRODUCT LICENSE | 1 | {'Cardiac Arrest': 1} | {'Intermittent Communication Failure': 1}
SECHRIST INDUSTRIES INC. | 1 | {'Insufficient Information': 1, 'Burn(s)': 1} | {'Insufficient Information': 1, 'Use of Device Problem': 1}
SPECTRUM IQ | 1 | {'Appropriate Term / Code Not Available': 1} | {'Complete Loss of Power': 1}
AERO TRACHEOBRONCHIAL STENT | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
2008T HD SYS. CDX W/BIBAG BLUE STAR | 1 | {'Loss of consciousness': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
GORE® TRI-LOBE BALLOON CATHETER | 1 | {'Vascular Dissection': 1} | {'Patient Device Interaction Problem': 1}
INFINITY DELTA | 1 | {'Asystole': 1} | {'No Audible Prompt/Feedback': 1}
GORE-TEX® STRETCH VASCULAR GRAFT | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ACIST | 1 | {'Air Embolism': 1} | {'Insufficient Information': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
MINI TREK¿ | 1 | {'Unspecified Heart Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ESSENCE RENTAL WITH SSD 230V50HZ UK | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
OXYLOG 3000PLUS | 1 | {'Hypoxia': 1} | {'Failure to Deliver': 1, 'Failure to Run on Battery': 1, 'Complete Loss of Power': 1}
AMNISURE ROM TEST | 1 | {'Chorioamnionitis': 1} | {'Human-Device Interface Problem': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE | 1 | {'Arrhythmia': 1, 'Bradycardia': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
RING ANNULOPLSTY TRICUSP CONTOUR 3D 28MM | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
HAMILTON G5 | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
CARESCAPE B450 | 1 | {'Respiratory Failure': 1} | {'No Audible Alarm': 1}
EXETER V40(TM). FEM. STEM 44MM | 1 | {'Embolism/Embolus': 1, 'Pulmonary Embolism': 1, 'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EMERGE | 1 | {'Low Blood Pressure/ Hypotension': 1, 'Pericardial Effusion': 1, 'Perforation': 1, 'Cardiac Arrest': 1, 'Extravasation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
NOVALUNG CONSOLE (LINUX BOARD) | 1 | {'Cardiac Arrest': 1} | {'Insufficient Information': 1, 'Material Integrity Problem': 1, 'Crack': 1}
INOGEN ONE G5 PORTABLE OXYGEN CONCENTRATOR | 1 | {'Dyspnea': 1} | {'Failure to Charge': 1}
PIEZOLITH 3000 TRIPLE FOCUS LITHOTRIPTER | 1 | {'Insufficient Information': 1} | {'Appropriate Term/Code Not Available': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
CAPIOX® FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Use of Device Problem': 1}
MEDIHONEY GEL IN TUBE, 1.5 FL OZ FLIPCAP - STERILE | 1 | {'Unspecified Infection': 1} | {'Insufficient Information': 1}
PROVOX2 | 1 | {'Osteomyelitis': 1, 'Sepsis': 1, 'Abscess': 1} | {'Inadequacy of Device Shape and/or Size': 1}
SINGLE USE GUIDE SHEATH KIT | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
ULTRASONIC PROBE | 1 | {'Insufficient Information': 1} | {'Detachment of Device or Device Component': 1, 'Peeled/Delaminated': 1}
GIA | 1 | {'Ascites': 1, 'Unspecified Tissue Injury': 1, 'Emotional Changes': 1, 'Pain': 1, 'Appropriate Term / Code Not Available': 1} | {'Mechanics Altered': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
TA | 1 | {'Ascites': 1, 'Unspecified Tissue Injury': 1, 'Emotional Changes': 1, 'Pain': 1, 'Appropriate Term / Code Not Available': 1} | {'Insufficient Information': 1, 'Mechanics Altered': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
POWERED 60 ECHELON +, 340MM SHAFT | 1 | {'Insufficient Information': 1} | {'Patient Device Interaction Problem': 1}
SPACELABS ULTRAVIEW SL COMMAND MODULE | 1 | {'Arrhythmia': 1} | {'Capturing Problem': 1}
SWAN-GANZ CCOMBO V | 1 | {'Laceration(s)': 1, 'Low Blood Pressure/ Hypotension': 1, 'Hemorrhage/Blood Loss/Bleeding': 1} | {'Difficult to Remove': 1, 'Physical Resistance/Sticking': 1, 'Use of Device Problem': 1}
KUGEL PATCH | 1 | {'Cellulitis': 1, 'Abscess': 1, 'Fistula': 1, 'Adhesion(s)': 1, 'Bacterial Infection': 1} | {'Patient Device Interaction Problem': 1, 'Defective Device': 1}
GLIDEPATH | 1 | {'Exsanguination': 1, 'Cardiac Arrest': 1} | {'Fluid/Blood Leak': 1}
EQUISTREAM ALPHACURVE DL CATHETER 14.5FR, 28CM | 1 | {'Exsanguination': 1} | {'Fluid/Blood Leak': 1, 'Material Split, Cut or Torn': 1}
DXM 1040 MICROSCAN WALKAWAY | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Incorrect, Inadequate or Imprecise Result or Readings': 1}
GELSOFT | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
MESH ¿ COMPOSIX KUGEL | 1 | {'Insufficient Information': 1} | {'Patient Device Interaction Problem': 1, 'Defective Device': 1}
COMPOSIX MESH E/X | 1 | {'Insufficient Information': 1} | {'Patient Device Interaction Problem': 1, 'Defective Device': 1}
DURAN ANCORE | 1 | {'Cardiac Arrest': 1} | {'Insufficient Information': 1, 'Patient Device Interaction Problem': 1}
NC TRAVELER CORONARY DILATATION CATHETER | 1 | {'Unspecified Heart Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CARDIOROOT WOVEN | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
WAVELINQ 4FR P | 1 | {'Appropriate Term / Code Not Available': 1, 'Stenosis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
V-18 CONTROL WIRE | 1 | {'Vascular Dissection': 1, 'Low Blood Pressure/ Hypotension': 1, 'Hemorrhage/Blood Loss/Bleeding': 1, 'Cardiac Arrest': 1} | {'Material Deformation': 1}
REP DREAMSTATION GO AUTO W/BT, DOM | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
LIFELINE | 1 | {'Insufficient Information': 1} | {'Insufficient Information': 1}
MSERIES HEATED HUMIDIFIER | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
EVERFLO 120V OPI | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
AVALON US TRANSDUCER | 1 | {'Insufficient Information': 1} | {'False Positive Result': 1, 'High Readings': 1}
GEM PREMIER 5000 PAK | 1 | {'Hypoglycemia': 1, 'Pericardial Effusion': 1} | {'Missing Test Results': 1, 'Adverse Event Without Identified Device or Use Problem': 1}
BIPAP AUTOSV | 1 | {'Insufficient Information': 1} | {'Degraded': 1}
ATRICLIP® PRO-V® LAA EXCLUSION SYSTEM ( | 1 | {'Ventricular Fibrillation': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
KYPHON® HV-R® BONE CEMENT | 1 | {'Cardiac Arrest': 1} | {'Leak/Splash': 1}
ASTRAL 150 DBL KIT FRA | 1 | {'Insufficient Information': 1} | {'Device Alarm System': 1}
2008T HEMODIALYSIS SYSTEM WITHOUT CDX | 1 | {'Insufficient Information': 1, 'Low Blood Pressure/ Hypotension': 1, 'Loss of consciousness': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
MULTI-FUNCTIONAL RESPIRATORY P | 1 | {'Insufficient Information': 1} | {'Delayed Program or Algorithm Execution': 1}
SURPASS EVOLVE 4.0MM X 17MM - PMAS | 1 | {'Perforation of Vessels': 1, 'Stroke/CVA': 1} | {'Malposition of Device': 1, 'Activation Failure': 1}
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | 1 | {'Convulsion/Seizure': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
LUCAS CPR CHEST COMPRESSION SYSTEM | 1 | {'Insufficient Information': 1, 'No Clinical Signs, Symptoms or Conditions': 1} | {'Mechanical Jam': 1}
DUAL CHAMBER TEMPORARY PACEMAKER | 1 | {'Heart Block': 1, 'Multiple Organ Dysfunction Syndrome': 1} | {'No Device Output': 1}
VOLUMAT MC AGILIA US | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Insufficient Flow or Under Infusion': 1}
BIPAP AVAPS C SERIES | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Degraded': 1}
REP DREAMSTATION AVAPS30 | 1 | {'Insufficient Information': 1} | {'Degraded': 1}
AVALON CABLELESS ECG/IUP TRANSDUCER | 1 | {'Insufficient Information': 1, 'Uterine Perforation': 1} | {'Incorrect Measurement': 1}
EVIS EXERA II ULTRASOUND GASTROVIDEOSCOPE | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
TFNA FEM NAIL Ø12 R 130° L400 TIMO15 | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
HEMODIALYSIS BLOODLINES | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1, 'Cardiac Arrest': 1} | {'Leak/Splash': 1, 'Detachment of Device or Device Component': 1}
PILLCAM | 1 | {'Foreign Body In Patient': 1} | {'Entrapment of Device': 1}
NOVOPEN 6 | 1 | {'Elevated ketones/Diabetic Ketoacidosis': 1} | {'Break': 1}
TM80 | 1 | {'Insufficient Information': 1} | {'Device Alarm System': 1}
BNX | 1 | {'Hemorrhage/Blood Loss/Bleeding': 1, 'Pancreatitis': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CADD DISPOSABLE | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
FLUENT PRO FLUID MANAGEMENT SYSTEM | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
MYOSURE REACH TISSUE REMOVAL DEVICE | 1 | {'Cardiac Arrest': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DELUXE 230V 50HZ IN | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
INNOSPIRE DELUXE RENTAL UNIT | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DELUXE 230V 50HZ INT. UK PLUG | 1 | {'No Clinical Signs, Symptoms or Conditions': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
TRAVELER RX CORONARY DILATATION CATHETER | 1 | {'Unspecified Heart Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
DREAMSTATION CPAP | 1 | {'Cancer': 1} | {'Degraded': 1}
GST BLUE RELOAD 45MM 6 ROW | 1 | {'Failure to Anastomose': 1} | {'Patient Device Interaction Problem': 1}
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH 4.5F | 1 | {'Heart Failure/Congestive Heart Failure': 1} | {'Material Frayed': 1}
CUSTOM COMBI SET | 1 | {'Insufficient Information': 1, 'Low Blood Pressure/ Hypotension': 1, 'Loss of consciousness': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
OPTIFLUX F180NRE HF 12/CS 1.7SA EBEAM | 1 | {'Insufficient Information': 1, 'Low Blood Pressure/ Hypotension': 1, 'Loss of consciousness': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
RECERT SIMPLYGO MINI | 1 | {'Insufficient Information': 1} | {'Degraded': 1}
MINI TREK | 1 | {'Unspecified Heart Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CARESCAPE | 1 | {'Insufficient Information': 1} | {'Insufficient Flow or Under Infusion': 1}
NC TREK NEO¿ | 1 | {'Unspecified Heart Problem': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
PHOENIX/INNOVA MACHINES | 1 | {'Appropriate Term / Code Not Available': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
CEREPAK FREEFORM | 1 | {'Intracranial Hemorrhage': 1, 'Ruptured Aneurysm': 1} | {'Separation Failure': 1, 'Positioning Failure': 1}
BIPAP S/T C-SERIES | 1 | {'Insufficient Information': 1} | {'Adverse Event Without Identified Device or Use Problem': 1}
STEALTH | 1 | {'Vascular Dissection': 1} | {'Activation Problem': 1, 'Unintended System Motion': 1, 'Material Separation': 1}

## Patient Problems
Problem | Count
Insufficient Information | 621
Cancer | 445
Unspecified Respiratory Problem | 320
No Clinical Signs, Symptoms or Conditions | 237
Cardiac Arrest | 176
Dizziness | 137
Headache | 131
Unspecified Kidney or Urinary Problem | 131
Hemorrhage/Blood Loss/Bleeding | 121
Appropriate Term / Code Not Available | 85
Low Blood Pressure/ Hypotension | 76
Skin Inflammation/ Irritation | 72
Eye Irritation | 71
Inflammation | 70
Vomiting | 69
Nausea | 67
Unspecified Hepatic or Biliary Problem | 61
Asthma | 59
Unspecified Heart Problem | 51
Pericardial Effusion | 39
Sepsis | 38
Pulmonary Embolism | 33
Pulmonary Dysfunction | 31
Respiratory Failure | 31
Hypersensitivity/Allergic reaction | 30
Cardiac Perforation | 28
Cardiac Tamponade | 26
Bradycardia | 24
Air Embolism | 24
Unspecified Vascular Problem | 24
Stroke/CVA | 23
Perforation of Vessels | 20
Myocardial Infarction | 20
Hematoma | 20
Heart Failure/Congestive Heart Failure | 19
Vascular Dissection | 18
Respiratory Arrest | 17
Septic Shock | 17
Dyspnea | 16
Elevated ketones/Diabetic Ketoacidosis | 16
Unspecified Infection | 16
Pneumonia | 16
Perforation | 15
Bowel Perforation | 14
Unspecified Tissue Injury | 14
Arrhythmia | 14
Intracranial Hemorrhage | 14
Ventricular Fibrillation | 13
Ischemia Stroke | 13
Pain | 13
Low Oxygen Saturation | 13
Asystole | 13
Great Vessel Perforation | 13
Cardiovascular Dysfunction/Insufficiency | 13
Bacterial Infection | 12
Hypoxia | 12
Failure to Anastomose | 12
Thrombosis/Thrombus | 12
Ischemia | 12
Chronic Obstructive Pulmonary Disease (COPD) | 12
Hypoglycemia | 11
Peritonitis | 11
Foreign Body In Patient | 11
Loss of consciousness | 10
Non specific EKG/ECG Changes | 9
Lymphoma | 9
Obstruction/Occlusion | 9
Liver Damage/Dysfunction | 9
Abdominal Pain | 9
Multiple Organ Dysfunction Syndrome | 8
Respiratory Tract Infection | 8
Hyperglycemia | 8
Exsanguination | 8
Tachycardia | 8
Fistula | 8
Embolism/Embolus | 8
Heart Block | 7
Burn(s) | 7
Fibrosis | 7
Aspiration Pneumonitis | 6
Shock | 6
Retroperitoneal Hemorrhage | 6
Pleural Effusion | 6
Ruptured Aneurysm | 5
Fever | 5
Rupture | 5
Abscess | 5
Adhesion(s) | 5
Bronchitis | 4
Breast Cancer | 4
Respiratory Insufficiency | 4
Gastrointestinal Hemorrhage | 4
Pneumothorax | 4
Brain Injury | 4
Hemorrhagic Stroke | 4
Hypovolemic Shock | 4
Pancreatitis | 4
Swelling/ Edema | 4
Hemolysis | 4
Unspecified Gastrointestinal Problem | 4
Chest Pain | 4
Tissue Breakdown | 4
Hemodynamic instability | 4
Ventilator Dependent | 3
Device Embedded In Tissue or Plaque | 3
Hernia | 3
Aspiration/Inhalation | 3
High Blood Pressure/ Hypertension | 3
Adult Respiratory Distress Syndrome | 3
Anemia | 3
Anaphylactic Shock | 3
Dehydration | 3
Fungal Infection | 3
Unintended Extubation | 3
Wound Dehiscence | 3
Extravasation | 3
Discomfort | 3
Ischemic Heart Disease | 3
Urinary Retention | 3
Cough | 3
Sore Throat | 3
Thromboembolism | 3
Laceration(s) | 3
Renal Failure | 3
Coagulation Disorder | 2
Diaphoresis | 2
Fatigue | 2
Pericarditis | 2
Endocarditis | 2
Mitral Valve Insufficiency/ Regurgitation | 2
Ascites | 2
Nervous System Injury | 2
Fluid Discharge | 2
Depression | 2
Emotional Changes | 2
Burning Sensation | 2
Anxiety | 2
Stenosis | 2
Hypoxia in Utero | 2
Peritoneal Laceration(s) | 2
Hypothyroidism | 2
Cholangitis | 2
Cerebral Edema | 2
Unspecified Immune System Problem | 2
Visual Disturbances | 2
Bacteremia | 2
Bone Cement Implantation Syndrome | 2
Encephalopathy | 2
Airway Obstruction | 2
Drug Resistant Bacterial Infection | 2
Bronchopneumonia | 2
Muscle Weakness/Atrophy | 2
Convulsion/Seizure | 2
Post Operative Wound Infection | 2
Perforation of Esophagus | 1
Hydrocephalus | 1
Respiratory Acidosis | 1
Sleepiness/Drowsiness/Somnolence | 1
Aortic Valve Insufficiency/ Regurgitation | 1
Failure of Implant | 1
Movement Disorder | 1
Anaphylactoid / Anaphylaxis | 1
Paresis | 1
Laceration(s) of Esophagus | 1
Renal Impairment | 1
Bronchial Hemorrhage | 1
Thrombocytopenia | 1
Medical device site infection | 1
Endoleaks | 1
Pulmonary Edema | 1
Aortic dissection | 1
Atrial Fibrillation | 1
Constipation | 1
Cardiogenic Shock | 1
Altered State of Consciousness | 1
Reduced Blood Flow | 1
Paralysis | 1
Breathing-related sleep disorder | 1
Unspecified Eye / Vision Problem | 1
Rash | 1
Oversedation | 1
Chorioamnionitis | 1
Decreased Level of Consciousness | 1
Hematuria | 1
Malaise | 1
Decreased Respiratory Rate | 1
Unspecified Mental, Emotional or Behavioural Problem | 1
Necrosis | 1
Aneurysm | 1
Uterine Perforation | 1
Cardiomyopathy | 1
Transient Ischemic Attack | 1
Hematemesis | 1
Sleep Dysfunction | 1
Infiltration into Tissue | 1
Erythema | 1
Skin Tears | 1
Intracranial Pressure Increased | 1
Intraoperative Cardiac Injury | 1
Syncope/Fainting | 1
Unspecified Nervous System Problem | 1
Impaired Healing | 1
Reperfusion Injury | 1
Loss of Range of Motion | 1
Fall | 1
Disseminated Intravascular Coagulation (DIC) | 1
Caustic/Chemical Burns | 1
Osteomyelitis | 1
Diminished Pulse Pressure | 1
Hemoptysis | 1
Seroma | 1
Upper Respiratory Tract Irritation | 1
Urinary Tract Infection | 1
Cellulitis | 1
Abdominal Distention | 1
Liver Failure | 1
Epistaxis | 1
Sinusitis | 1
Delirium | 1

## Product Problems
Problem | Count
Adverse Event Without Identified Device or Use Problem | 912
Degraded | 824
Insufficient Information | 118
Use of Device Problem | 31
Patient Device Interaction Problem | 31
Contamination | 26
No Audible Alarm | 24
Break | 21
Fluid/Blood Leak | 20
Device Alarm System | 18
Defective Alarm | 17
Appropriate Term/Code Not Available | 16
Physical Resistance/Sticking | 16
Difficult to Remove | 13
Detachment of Device or Device Component | 12
Incorrect, Inadequate or Imprecise Result or Readings | 12
Gas/Air Leak | 11
No Apparent Adverse Event | 11
Material Puncture/Hole | 10
Difficult to Advance | 10
Material Integrity Problem | 9
Leak/Splash | 9
Defective Device | 8
Malposition of Device | 8
Inaccurate Delivery | 8
Material Separation | 7
Infusion or Flow Problem | 7
Excess Flow or Over-Infusion | 7
Air/Gas in Device | 7
Improper or Incorrect Procedure or Method | 6
Gas Output Problem | 6
Pumping Stopped | 5
Pressure Problem | 5
Device Handling Problem | 5
Fracture | 5
Material Split, Cut or Torn | 5
Failure to Fire | 5
Electrical /Electronic Property Problem | 5
Positioning Failure | 4
Failure to Form Staple | 4
Insufficient Flow or Under Infusion | 4
Incorrect Measurement | 4
False Negative Result | 4
No Flow | 4
Difficult to Open or Close | 4
Application Program Freezes, Becomes Nonfunctional | 4
Disconnection | 4
No Device Output | 4
Output Problem | 4
Material Deformation | 4
No Audible Prompt/Feedback | 4
Inflation Problem | 3
Failure to Deliver Shock/Stimulation | 3
Poor Quality Image | 3
Crack | 3
Complete Loss of Power | 3
Positioning Problem | 3
Unable to Obtain Readings | 3
Radiation Output Failure | 3
Patient-Device Incompatibility | 3
Mechanics Altered | 3
Low Readings | 3
Connection Problem | 3
Overheating of Device | 3
Unexpected Shutdown | 2
Capturing Problem | 2
Activation Problem | 2
Misassembly During Maintenance/Repair | 2
Difficult to Insert | 2
Peeled/Delaminated | 2
Protective Measures Problem | 2
Obstruction of Flow | 2
Entrapment of Device | 2
Corroded | 2
Failure to Deliver | 2
Mechanical Problem | 2
Inability to Auto-Fill | 2
Migration | 2
Defective Component | 2
Retraction Problem | 2
Misfire | 2
Intermittent Communication Failure | 2
False Positive Result | 2
Activation Failure | 2
Human-Device Interface Problem | 2
No Pacing | 2
Activation, Positioning or Separation Problem | 2
Material Twisted/Bent | 2
Noise, Audible | 2
Erratic Results | 1
Device Sensing Problem | 1
Deflation Problem | 1
Alarm Not Visible | 1
Material Discolored | 1
Failure to Prime | 1
Incomplete or Inadequate Priming | 1
Filling Problem | 1
Pacing Problem | 1
Migration or Expulsion of Device | 1
Device Dislodged or Dislocated | 1
No Visual Prompts/Feedback | 1
Device Contaminated During Manufacture or Shipping | 1
Expiration Date Error | 1
Wireless Communication Problem | 1
Low Test Results | 1
Failure to Deflate | 1
Display or Visual Feedback Problem | 1
High Test Results | 1
Use of Incorrect Control/Treatment Settings | 1
Delivered as Unsterile Product | 1
Electrical Power Problem | 1
Erratic or Intermittent Display | 1
Failure to Calibrate | 1
Failure of Device to Self-Test | 1
Firing Problem | 1
Unstable | 1
Material Perforation | 1
Material Rupture | 1
Loss of Data | 1
Failure to Charge | 1
Free or Unrestricted Flow | 1
Tidal Volume Fluctuations | 1
Failure to Cut | 1
Misassembled | 1
Misassembly by Users | 1
Misconnection | 1
Pumping Problem | 1
Burst Container or Vessel | 1
Failure to Deliver Energy | 1
Increase in Pressure | 1
Poor Visibility | 1
Device Damaged by Another Device | 1
Therapeutic or Diagnostic Output Failure | 1
Failure to Transmit Record | 1
Application Program Problem | 1
Off-Label Use | 1
Material Frayed | 1
Contamination of Device Ingredient or Reagent | 1
Lack of Effect | 1
Inadequacy of Device Shape and/or Size | 1
Failure to Run on Battery | 1
Unexpected Therapeutic Results | 1
Defibrillation/Stimulation Problem | 1
Separation Failure | 1
Microbial Contamination of Device | 1
High Readings | 1
Missing Test Results | 1
Delayed Program or Algorithm Execution | 1
Deformation Due to Compressive Stress | 1
Mechanical Jam | 1
Failure to Advance | 1
Unintended System Motion | 1

## Device-Patient Matrix
| REMSTAR PRO C-FLEX+ | INJECTOMAT AGILIA BR | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | AGILIA SP MC BR | REP DREAMSTATION AUTO CPAP | DREAMSTATION BIPAP AVAPS30 | REMSTAR PLUS | EDWARDS ESHEATH PLUS INTRODUCER SHEATH | MITRACLIP | CATH F5.2+ PIG 145° 110CM 6SH | REMSTAR AUTO A-FLEX | ASTRAL 150 - AMER | NIPRO SURDIAL DX HEMODIALYSIS SYSTEM | POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS | DREAMSTATION AUTO CPAP | NIPRO BLOOD TUBING SET W/TP | NIPRO ELISIO-H DIALYZER | DAVINCI XI | BIPAP AUTO BIFLEX | REACT CATHETER | NXSTAGE SYSTEM ONE | BIPAP A40 PRO | TEMPUS LS | VISISHEATH DILATOR SHEATH | DREAMSTATION AUTO BIPAP | SUREFORM | ION | REMSTAR AUTOA-FLEX | BIPAP AUTO | COBLATION HALO WAND | EDWARDS ESHEATH+ INTRODUCER | INTELLIVUE MX450 PATIENT MONITOR | SENSATION PLUS 8FR. 50CC IAB & ACCESSORIES (W/PRESSURE TUBES, NO STYLET) | EMBOTRAP III 6.5 MM X 45 MM | VERSACROSS CONNECT LAAC ACCESS SOLUTION | DREAMSTATION CPAP PRO | REMSTAR AUTO A-FLEXX | GIRAFFE OMNIBED | INVACARE® PLATINUM® OXYGEN CONCENTRATOR | HOMEFILL® OXYGEN SYSTEM | SHILEY | PRLNE BLU 36IN 4-0 D/A SH | SOLITAIRE FR | EVERFLO | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | PHILIPS CPAP | SAFARI2¿ | DREAMSTATION GO AUTO | PHILIPS CPAP DEVICE | GORE® DRYSEAL FLEX INTRODUCER SHEATH | MESH ¿ VENTRALEX | SEPRAMESH IP | TRILOGY 100 | OMNIPOD 5 POD | DREAMSTATION BIPAP AUTOSV | DREAMSTATION ST30 | SYSTEM ONE HTD HUMID ASSY ,DOM | DS2ADV AUTO CPAP | AZURION | FORTILINK-C IBF SYSTEM | COMPANION 5 | DJO SURGICAL | DREAMSTATION BIPAP PRO | AGILIA VP MC BR | OMNIPOD 5 CONTROLLER | DREAMST ST30 H | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | NELLCOR | TRILOGY O2 VENTILATOR, JAPAN | VIVO 45 LS | BIPAP AUTOSV ADV | SENTRANT INTRODUCER SHEATH | TUBING SET | RENTAL, DREAMSTATION ST30, GB | SIGNIA | BIPAP A40 | SOLITAIRE X | 6VT-D | ZOOM REPERFUSION CATHETER | MX40 1.4 GHZ SMART HOPPING | HEART LUNG MACHINE | CARDIOSAVE HYBRID, TYPE B PLUG | MILLENNIUM M600 WITHOUT OPI DOMESTIC | REMSTAR PRO C-FLEX | S5 HEART-LUNG-MACHINE | DREAMSTATION CPAP | HIGH FLOW INSUFFLATION UNIT | TELEHEALTH CENTER PRODUCT LICENSE | AERO TRACHEOBRONCHIAL STENT | TREK | MINI TREK¿ | NC TRAVELER CORONARY DILATATION CATHETER | TRAVELER RX CORONARY DILATATION CATHETER | MINI TREK | NC TREK | ILET BIONIC PANCREAS | NC TREK NEO¿ | HEMODIALYSIS BLOODLINES | DREAMSTATION 2 ADVANCE AUTO CPAP | AIRSEAL | V-18 CONTROL WIRE | FILMARRAY® PNEUMONIA PANEL PLUS | ASTRAL 150 DBL KIT FRA | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | DREAMSTATION HUMIDIFIER, GB-RENT | PATIENT INFORMATION CENTER IX | SURPASS EVOLVE 4.0MM X 17MM - PMAS | INTELLIVUE MX800 PATIENT MONITOR | RENTAL-DREAMSTATION AUTO CPAP, GB | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | AMNISURE ROM TEST | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | MICRA | PERFIX PLUG LIGHT | ENDOWRIST | NITREX | SIMPLEX P WITH TOBRAMYCIN 1 PACK | VISUALASE SYSTEM | LIBERATOR | RENTAL-DREAMSTATION AUTO CPAP | POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | MO.MA ULTRA | SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER SL 1 CURVE, 63 CM LENGTH, 8F | EMBOSHIELD NAV6 | STARTER GUIDEWIRE | BENTSON TYPE PTFE COATED GUIDEWIRE | BIPAP AUTO BIFLEX, | DREAMSTATION | CERELINK ICP PROBE 1L M BOLT | FISHER & PAYKEL HEALTHCARE | CERELINK ICP MONITOR | SOLITAIRE FR4 | CATHERA | CAPTIVATOR EMR | REMSTAR PRO C-FLEX+ W/HUMID, SD CARD | COCHLEAR OSIA OSI200 IMPLANT | OBSIDIO | UNKNOWN RENASYS TOUCH DEVICE | TRILOGY EVO | TEMPUS LS-MANUAL | EVERFLO DOM | AERO® TRACHEOBRONCHIAL STENT | PACIFIC PLUS | HAMILTON-T1 | VIZIGO | DREAMSTATION AUTO ,CPAP | ROTAREX | ALARIS SYSTEM | TURNPIKE SPIRAL 135CM | REP DREAMSTATION AUTO BIPAP, DOM-RECRT | HAMILTON-C1 | 2008T HEMODIALYSIS SYSTEM W/BIBAG | AVEIR¿ INTRODUCER KIT ¿ 50 CM | SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | DREAMSTATION AUTO CPAP, JP | REBAR | FREEDOM PERIPHERAL NERVE STIMULATOR | POWERPORT ISP MRI 8CF INT. W/SP, ATT, SL | TRI-STAPLE 2.0 | DREAMSTATION BIPAP | LTV 1150 VENTILATOR | SENSATION 7FR. 40CC IAB | HEMOCHRON ACTIVATED CLOTTING TIME PLUS (ACT+) TEST | RESPIRONICS | EDWARDS TRANSFEMORAL BALLOON CATHETER | EVIS EXERA III DUODENOVIDEOSCOPE | INTELLAMAP ORION HIGH RESOLUTION MAPPING CATHETER | DREAM STATION AUTO CPAP | VERSYS FEMORAL STEM CEMENTED REVISION/CALCAR 12/14 NECK TAPER SIZE 13 250MM | WATCHMAN TRUSTEER ACCESS SYSTEM | ESSENCE RENTAL | ESSENCE RENTAL WITH SSD | CENTRIMAG BLOOD PUMP | ESSENCE WITH SSD | LINEAR 7.5FR. 40CC IAB (CHINA) | REMSTAR PLUS C-FLEX | DEVILBISS HEALTHCARE | 2008T HEMODIALYSIS SYS., WITH CDX | BIPAP AUTO SERIES ASSY | NEOTRACT UROLIFT 2 SYSTEM | SMART PORT CT | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | RESUSCITAIRE RADIANT WARMER | INOGEN ONE G5 PORTABLE OXYGEN CONCENTRATOR | ENDO GIA | 2008T MACHINE | BIPAP AUTOSV ADVANCED SYSTEM ONE | BIPAP PRO BIFLEX | REMSTAR | REMSTAR AUTO | REMSTAR PRO C-FLEX + | ACHIEVE MAPPING CATHETER - 20 MM | FLEXCATH ADVANCE STEERABLE SHEATH | BIVONA TRACHEOSTOMY TUBE, NEONATAL/PEDIATRIC | STRATTICE LAPAROSCOPIC 10CM X 16CM | POWERPORT VUE M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | BRILLIANCE | PACEL BIPOLAR PACING CATHETER | REMSTAR SE | INTELLIVUE MP30 | RING PROFILE 3D 30MM | BIPAP AUTOSV ADV SYS ONE | PROXIMATE*75MM LINEAR CUTTER | TIGHTRAIL ROTATING DILATOR SHEATH | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | EDWARDS ESHEATH+ INTRODUCER SET | BIPAP PRO BI-FLEX | REMSTAR PRO C FLEX | CADD-SOLIS VIP AMBULATORY INFUSION PUMP | TABLO HEMODIALYSIS SYSTEM | REMSTAR LITE | TRILOGY | MARK 5 NUVO LITE | HEARTSTART XL+ DEFIBRILLATOR/MONITOR | FLAT CUT TALAR DOME SZ 2 INFINITY ADAPTIS | REP DREAMSTATION AUTO BIPAP | SENDIT DELIVERY DEVICE | PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER | REMSTAR W/SD CARD | AVALON FM30 FETAL MONITOR | RECERT, DREAMSTATION AUTO, FR | REMSTARAUTO | BIPAP AUTO BI FLEX | BIPAP AUTO SV ADV SYSTEM ONE | MX40 PATIENT WEARABLE MONITOR | REM AUTO A-FLEX | CORTRAK 2 ENTERAL ACCESS SYSTEM | EDWARDS ESHEATH INTRODUCER SET | SPIDER FX | REMSTAR AUTO A FLEX | XENOSURE BIOLOGIC PATCH | SERVO-I UNIT BASE | VUE PACS | LIFEPAK® 15 DEFIBRILLATOR/MONITOR | DREAMSTATION HUMID | REMSTAR AUTO SERIES AFLEX | ENDO ANCHOR SYSTEM - HELI-FX AAA | SJM RIGID SADDLE RING | SJM SÃ©GUIN SEMI-RIGID ANNULOPLASTY RING | EMBOTRAP | SIMPLYGO | SIMPLYGO MINI | MC3 NAUTILUS ECMO OXYGENATOR | CRESCENT RA DUAL LUMEN CATHETER 13FR | DREAMSTATION BIPAP AVAP | DREAMSTATION BIPAP ST30 | BIPAP PRO | ZIO AT | STAR¿ DRIVE MA | CVAC ASPIRATION SYSTEM | REM STAR AUTO A-FLEX | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | REMSTAR PRO SERIES | INFINITY CENTRAL STATION | UNKNOWN_VOCERA_PRODUCT | INTELLIVUE MX40 2.4GHZ | INTELLIVUE MX500 PATIENT MONITOR | FLOWTRIEVER CATHETER, XL | LINEAR 7.5FR. 40CC IAB | TRICLIP G4 SYSTEM | ALLURA XPER FD | PRLNE BLU 24IN 6-0 D/A BV-1 | PRESSIO ICP AND ICT MONITORING KIT, PARENCHYMAL WITH BOLT | AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | LUCAS¿ CPR CHEST COMPRESSION SYSTEM | FMC CASSETTE | BIPAP SYSTEM ONE 60 SERIES | EFFICIA DFM100 | DAVINCI X | DREAMSTATION HUMIDIFIER | CAPIOX FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | REMSTAR PLUS SERIES ASSY | BIPAP AUTO BI-FLEX | IPV-2C | REP DREAMSTATION AUTO CPAP, DOM - RECRT | REP DREAMSTATION AUTO CPAP, | IMPEDE-FX EMBOLIZATION PLUG | LEAD EXTRACTION BYRD POLYPROPYLENE DILATOR SHEATH SET | REMSTAR SYSTEM ONE 60 SERIES | LEVEEN SUPERSLIM | DUODOPA_DUOPA | DS2ADV AUTO CPAP W/HUMID CELL/BT | REP DREAMSTATION AUTO CPAP, DOM | BIPAP AUTOSV ADVANCED SYSTEM ONE 60 SERIES | 75MM SELECTABLE NEW TLC | PARKER | HEATER/COOLER SYSTEM | CONFIDA BRECKER CURVE GUIDEWIRE | ESHEATH PLUS INTRODUCER SHEATH | DREAMSTATION AUTO | COLONIC FTRD SET | STREAMLINE TEMPORARY PACING LEADS | BTK TURBOHAWK | NANOCROSS 014 | EVERCROSS 035 | ADMIRAL XTREME OTW | WECK VISTA | EPSTAR FIXED ELECTROPHYSIOLOGY CATHETER | AVALON FETAL MONITOR FM50 | PRLNE BLU 24IN 7-0 D/A BV-1 | TRILOGY EVO, JAPAN | WATCHMAN FXD CURVE ACCESS SYSTEM | STEALTHSTATION¿ S8 SYSTEM | AVALON CABLELESS ECG/IUP TRANSDUCER | REP DREAMSTATION AVAPS30 | BIPAP AVAPS C SERIES | VOLUMAT MC AGILIA US | DUAL CHAMBER TEMPORARY PACEMAKER | LUCAS CPR CHEST COMPRESSION SYSTEM | ALPHAVAC | AGILIA SP MC WIFI GB | MULTI-FUNCTIONAL RESPIRATORY P | 2008T HEMODIALYSIS SYSTEM WITHOUT CDX | BENEVISION N1 PATIENT MONITOR | KYPHON® HV-R® BONE CEMENT | ATRICLIP® PRO-V® LAA EXCLUSION SYSTEM ( | INTELLIVUE MX40 802.11A/B/G | BIPAP AUTOSV | GEM PREMIER 5000 PAK | AVALON US TRANSDUCER | EVERFLO 120V OPI | MSERIES HEATED HUMIDIFIER | LIFELINE | REP DREAMSTATION GO AUTO W/BT, DOM | INFINITY ACUTE CARE SYSTEM (M540) | CLOTTRIEVER XL CATHETER | WAVELINQ 4FR P | CONTOUR 3D | CARDIOROOT WOVEN | EVIS EXERA II ULTRASOUND GASTROVIDEOSCOPE | TFNA FEM NAIL Ø12 R 130° L400 TIMO15 | VIVO 45LS | GST BLUE RELOAD 45MM 6 ROW | BIPAP S/T C-SERIES | CEREPAK FREEFORM | PHOENIX/INNOVA MACHINES | DAVINCI SI | CARESCAPE | DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | RECERT SIMPLYGO MINI | OPTIFLUX F180NRE HF 12/CS 1.7SA EBEAM | CUSTOM COMBI SET | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH 4.5F | DREAMSTATION CPAP | PILLCAM | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | DELUXE 230V 50HZ INT. UK PLUG | INNOSPIRE DELUXE RENTAL UNIT | DELUXE 230V 50HZ IN | MYOSURE REACH TISSUE REMOVAL DEVICE | FLUENT PRO FLUID MANAGEMENT SYSTEM | CADD DISPOSABLE | BNX | TM80 | NOVOPEN 6 | DREAM STATION CPAP | DURAN ANCORE | COMPOSIX MESH E/X | MESH ¿ COMPOSIX KUGEL | VERSACROSS ACCESS SOLUTION | ACIST | GORE-TEX® STRETCH VASCULAR GRAFT | INFINITY DELTA | GORE® TRI-LOBE BALLOON CATHETER | SERVO-S BASE UNIT | 2008T HD SYS. CDX W/BIBAG BLUE STAR | INTELLIVUE MP70 | SPECTRUM IQ | SECHRIST INDUSTRIES INC. | CARDIOHELP | CARDIOSAVE HYBRID W/ E/F PLUG | INTELLIVUE MP5 | HEMOSPRAY ENDOSCOPIC HEMOSTAT | EMERGE? | BD ALARIS PUMP MODULE SMARTSITE INFUSION SET | EVITA 4 | CANNULA & CATHETER | CAREEVENT | PLUM 360 | TD60 TELEMETRY TRANSMITTER | EDWARDS ESHEATH PLUS | ESSENCE RENTAL WITH SSD 230V50HZ UK | MONARCH ENDOSCOPY PLATFORM | OXYLOG 3000PLUS | VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE | GELSOFT | HT70 VENTILATOR | DXM 1040 MICROSCAN WALKAWAY | EQUISTREAM ALPHACURVE DL CATHETER 14.5FR, 28CM | GLIDEPATH | KUGEL PATCH | SWAN-GANZ CCOMBO V | SPACELABS ULTRAVIEW SL COMMAND MODULE | POWERED 60 ECHELON +, 340MM SHAFT | TA | GIA | ULTRASONIC PROBE | SINGLE USE GUIDE SHEATH KIT | PROVOX2 | MEDIHONEY GEL IN TUBE, 1.5 FL OZ FLIPCAP - STERILE | CAPIOX® FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | PIEZOLITH 3000 TRIPLE FOCUS LITHOTRIPTER | PRLNE BLU 36IN 2-0 D/A V-7 MP | NOVALUNG CONSOLE (LINUX BOARD) | EMERGE | EXETER V40(TM). FEM. STEM 44MM | CARESCAPE B450 | HAMILTON G5 | RING ANNULOPLSTY TRICUSP CONTOUR 3D 28MM | REP DREAMSTATION AUTO BIPAP, DOM | ECHELON FLEX 60 POWERED PLUS ENDOPATH STAPLER ARTICULATING ENDOSCOPIC LINEAR CUT | TOUCHSOFT COAGULATOR | RENTAL, DREAMSTATION ST30 | FREESTYLE LIBRE 2 | INQUIRY¿ STEERABLE DIAGNOSTIC CATHETER | DREAMSTATION BIPAP AVAPS | ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING | CAPIOX | DA VINCI ENERGY | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | MEDIHONEY GEL IN TUBE, .5 FL OZ TWISTOFF - STERILE | PARAPAC PLUS | HOBBS MEDICAL TOUCHSOFT COAGULATOR | SABLE | LEAD EXTRACTION EVOLUTION RL CONTROLLED-ROTATION DILATOR SHEATH SET | NOXBOXI NITRIC OXIDE DELIVERY SYSTEM | LINEAR 7.5FR. 40CC IAB WITH ACCESSORIES, US ONLY | REFLEX CATHETER | PHOTODYNAMIC BONE STABILIZATION SYSTEM | TRINAV INFUSION SYSTEM | PROWATERFLEX | FEEDING TUBE | DAVINCI 5 | INTRI24 INTRODUCER SHEATH | HARMONIC ACE 5MM SHEAR | DEVILBISS 525 OXYGEN CONCENTRATOR | INTRAOSSEOUS, INTRAOSSEOUS NEEDLE KITS | MINICAP WITH POVIDONE-IODINE SOLUTION | FILMARRAY® MENINGITIS/ENCEPHALITIS (ME) PANEL | LIFE SCOPE PT | PENDITURE LAA EXCLUSION SYSTEM | SENSATION PLUS 7.5FR. 40CC IAB & ACCESSORIES | HEMOSTAR | BENEVISION N17 PATIENT MONITOR | INTRACEPT RF PROBE | MERGE CARDIO | BIPOLAR VENTRICULAR ENDOCARDIAL LEAD | RADIOGRAPHY 7300 C | INTELLIVUE X3 | BIOFIRE® BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL | LLD EZ LEAD LOCKING DEVICE | ANGIOVAC | LIBERTY CYCLER SET, SINGLE CONN./EXT. DL | PHILIPS RESPIRONICS E30 WITH HUMIDIFIER | EVIS EXERA DUODENOVIDEOSCOPE | NIO ADULT | INJECTOR OPVSU NPS | CLEARVIEW ARTERIOTOMY SHUNT | CLEARVIEW BLOWER/ MISTER | DISTAL PERFUSION KIT | CRAGG-MCNAMARA | XENMATRIX AB | CORTRAK 2 NG/NI FEED, ELECTROMAGNETIC STYLET, ENFIT, CONNECTOR | AMPLATZER VASCULAR PLUG II | HAKIM PROGRAMMABLE VALVE | SAVVYWIRE | TI-CRON | RHYTHM ANALYTICS - SOFTWARE (SAMD) | REMSTAR AUTO A-FLEX W/HUMID, SD CARD | PHENOM CATHETER | POWERED WHEELCHAIR SYSTEM | QC HEMOSTATIC DRESSING, 2X2, 10X | FUSION HOLLOW FIBER OXYGENATOR | INTELLIPAP AUTO BILEVEL W/HEATED HUMIDIFIER | PROGRIP | I-NEB CF UK UNIT | TRILOGY DEVICE | POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PENTAX MEDICAL | SPYSCOPE DS II | EPIQ CVX | APEXPRO TELEMETRY SYSTEM V6 | DAVINCI SINGLE PORT | CHPV RT ANG RESERVOIR UNITIZED | PLUM 360 (V 15.X) | DREAMSTATION 2 ADVANCED AUTO CPAP | FLEXOR RAABE GUIDING SHEATH | LATERAL F3D 18 MM IMPLANT; COBB ELEVATOR | WATCHMAN TRUSTEER? ACCESS SYSTEM | ENSEAL X1 CURVED 25CM SHAFT | HARMONIC 1100 SHEARS 36CM | UMBILICAL CATHETER | COOLSEAL¿ TRINITY | COOLSEAL¿ GENERATOR | CIRCULAR MECH STD 29MM | HLS SET | CONFIDENCE KIT, NO NEEDLES | POWERTRIALYSIS | B125 | REVEAL LINQ | PROXIMATE*VASC LIN STAP RELOAD | NOVUM IQ | TRILOGY EVO, USA | TANDEM MOBI INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | SAPPHIRE M.T INFUSION PUMP - GERMAN | ENDO CLIP III | AMPLATZER TORQVUE DELIVERY SYSTEM | ULRICH MEDICAL | UL_INJECT | GALAXY SYSTEM | INDIGO SYSTEM FLASH ASPIRATION CATHETER | DREAMSTATION BIPAP AUTOSV, DS | COMBISET SMARTECH BLOODLINE (NO HEPARIN) | FLEXCATH CONTOUR¿ STEERABLE SHEATH | INTELLISPACE EVENT MGMT PLATFORM | HOMECHOICE APD SYSTEM | CHORDX UNKNOWN | EPIQ CVXI | EXTRACTOR PRO RX | BIOFLO VORTEX | BIPAP AUTOSV ADVANCED | EDISON SYSTEM | REMSTAR PRO C-FLEX+ SYSTEM ONE 60 SERIES | MRIDIUM | CS300 INTRA-AORTIC BALLOON PUMP, SPANISH, 110V | PORTEX ENDOTRACHEAL TUBE, SILICONIZED | EUPHORA RX | CLOTTRIEVER BOLD CATHETER | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | SIMUFORM | DREAM STATION AUTO BIPAP | INNOSIGHT | LIFEPAK 15 DEFIBRILLATOR/MONITOR | EVERFLO INTL OPI | MEDICALYTE | OPTICROSS? HD | OMNIPOD 5 AUTOMATED INSULIN DELIVERY SYSTEM | STEERABLE GUIDE CATHETER | OPTICROSS HD | DREAMSTATION BIPAP S/T 30 W/HEATED HUMIDIFIER, GB | HEMOSIL HIT- AB(PF4-H) | REMSTAR PRO SERIES C-FLEX + ASSY | BIPAP PRO SERIES ASSY | REMSTAR AUTO SERIES AFLEX ASSY | CS300 INTRA-AORTIC BALLOON PUMP, PORTUGUESE, 220V | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | PULSE | KANGAROO | INTELLIVUE MULTI-MEASUREMENT MODULE X3 | DURAMATRIX-ONLAY COLLAGEN DURA SUBSTITUTE MEMBRANE | UNKNOWN_NEUROVASCULAR_PRODUCT | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | WALLFLEX ESOPHAGEAL | ANGIOJET? SOLENT? OMNI | OP TELESCOPE, 7°, 11 FR. X 220 MM, 7.5 FR. CHANNEL, ANGLED OCULAR | VALLEYLAB | GORE® TRI LUMEN CATHETER | INTELLIVUE PATIENT MONITOR MX750 | MAIN ASSEMBLY, VENTILATOR, COMMERCIAL, ZVENT, BASI | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | CNS-2101 CENTRAL MONITOR | BIA-GEN | PERCUFLEX URINARY DIVERSION STENT SET | INTELLIVUE MX40 WLAN | REMSTAR PLUS W/C-FLEX | PROTEK DUO VENOUS CANNULA | MYOSURE TISSUE REMOVAL DEVICE | AVALON FM50 FETAL MONITOR | WATCHMAN FXD CURVE? ACCESS SYSTEM | NAVIGATOR HD | PROLENE POLYPROPYLENE SUTURE UNKNOWN PRODUCT | PREMICATH | CIRCULO | FREESTYLE LIBRE 3 | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | LIBERTY CYCLER SET, DUAL PATIENT CONNECT | INTELLIVUE MX550 PATIENT MONITOR | BEAD BLOCK | PANORAMA CENTRAL STATION WITH TELEMETRY | CAPIOX®FX | SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER, SL1¿, 63 CM LENGTH, 8.5F | ERBE VIO 3 | TRILOGY 202 VENTILATOR, INTERNATIONAL | RADICAL-7 | K-PACK II | MEDTRONIC NAVIGATION | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | ARROW CVC SET: 4-LUMEN 8.5FR X 16CM | PU-681RA | INTELLIVUE MX40 1.4GHZ | AEROGEN SOLO | CARDIOSAVE HYBRID, 3.1 EDITION | TRANS-RAY PLUS 7.5FR 35CC IAB WITH ACCESSORIES | AMPLATZER VASCULAR PLUG | PORT-A-CATH IMPLANTABLE ACCESS SYSTEM | AVEIR¿ RETRIEVAL SYSTEM CATHETER | INTRODUCER KIT FOR GASTROSTOMY FEEDING TUBE | CG FUTURE BAND | MESH - VENTRALEX | BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | FREESTYLE LIBRE 3 PLUS | TANDEM MOBI INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | DRAGONFLY OPSTAR¿ | TRILOGY EVO, O2, JAPAN | CS100 INTRA-AORTIC BALLOON PUMP, ENGLISH, 220V | GORE® MOLDING & OCCLUSION BALLOON CATHETER | 0.14 GUIDEWIRES | EPOC BLOOD ANALYSIS SYSTEM | BIPAP PRO BIFLEX, W/SD CARD, US-RENTAL | PORTEX BLUSELECT SUCTIONAID CUFFED TRACHEOSTOMY | EVERFLO 230V, OPI, AUS, ULTRAFILL | COLLECTIVE CONCENTRATE- NATURALYTE | QUANTUM WORKSTATION 12.1" ELITE | MEGA 7.5FR. 40CC IAB | M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | VACCESS CT LOW-PROFILE POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 8F | ECHELON 3000 60MM STANDARD | TELESCOPE GUIDE EXTENSION CATHETER | BIPAP S/T C SERIES | MAGNETOM ESPREE | HERO® | AVALON FETAL MONITOR FM30 | INTRAOSSEOUS NEEDLE KITS | STEALTH
Abdominal Distention | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Abdominal Pain | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Abscess | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Adhesion(s) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Adult Respiratory Distress Syndrome | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Air Embolism | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Airway Obstruction | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0
Altered State of Consciousness | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anaphylactic Shock | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anaphylactoid / Anaphylaxis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anemia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Aneurysm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Anxiety | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Aortic Valve Insufficiency/ Regurgitation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Aortic dissection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Appropriate Term / Code Not Available | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 18 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0
Arrhythmia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ascites | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Aspiration Pneumonitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Aspiration/Inhalation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asthma | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 29 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Asystole | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Atrial Fibrillation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bacteremia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bacterial Infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bone Cement Implantation Syndrome | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bowel Perforation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bradycardia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Brain Injury | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Breast Cancer | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Breathing-related sleep disorder | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bronchial Hemorrhage | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bronchitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Bronchopneumonia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Burn(s) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Burning Sensation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cancer | 32 | 0 | 0 | 0 | 1 | 1 | 11 | 0 | 0 | 0 | 29 | 0 | 0 | 0 | 164 | 0 | 0 | 0 | 6 | 1 | 0 | 0 | 0 | 0 | 25 | 0 | 0 | 14 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 2 | 8 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 68 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 2 | 5 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cardiac Arrest | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 2 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 6 | 0 | 1 | 14 | 0 | 4 | 0 | 0 | 1 | 6 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 20 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 3 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cardiac Perforation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cardiac Tamponade | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cardiogenic Shock | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cardiomyopathy | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cardiovascular Dysfunction/Insufficiency | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Caustic/Chemical Burns | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cellulitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cerebral Edema | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Chest Pain | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cholangitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Chorioamnionitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Chronic Obstructive Pulmonary Disease (COPD) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Coagulation Disorder | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Constipation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Convulsion/Seizure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Cough | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Decreased Level of Consciousness | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Decreased Respiratory Rate | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dehydration | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Delirium | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Depression | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Device Embedded In Tissue or Plaque | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Diaphoresis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Diminished Pulse Pressure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Discomfort | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Disseminated Intravascular Coagulation (DIC) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dizziness | 6 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 11 | 0 | 0 | 0 | 51 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 9 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 4 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Drug Resistant Bacterial Infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Dyspnea | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Elevated ketones/Diabetic Ketoacidosis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Embolism/Embolus | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Emotional Changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Encephalopathy | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Endocarditis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Endoleaks | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Epistaxis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Erythema | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Exsanguination | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Extravasation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Eye Irritation | 5 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 22 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure of Implant | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Anastomose | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fall | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fatigue | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fever | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fibrosis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fistula | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fluid Discharge | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Foreign Body In Patient | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fungal Infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Gastrointestinal Hemorrhage | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Great Vessel Perforation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Headache | 6 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 11 | 0 | 0 | 0 | 52 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 4 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Heart Block | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Heart Failure/Congestive Heart Failure | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hematemesis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hematoma | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 9 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hematuria | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hemodynamic instability | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hemolysis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hemoptysis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hemorrhage/Blood Loss/Bleeding | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 13 | 0 | 2 | 5 | 0 | 0 | 0 | 0 | 1 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 3 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 4 | 1 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hemorrhagic Stroke | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hernia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
High Blood Pressure/ Hypertension | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hydrocephalus | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hyperglycemia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hypersensitivity/Allergic reaction | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hypoglycemia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hypothyroidism | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hypovolemic Shock | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hypoxia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Hypoxia in Utero | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
Impaired Healing | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Infiltration into Tissue | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Inflammation | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 23 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 12 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Insufficient Information | 7 | 0 | 1 | 0 | 7 | 0 | 1 | 1 | 0 | 1 | 13 | 1 | 0 | 0 | 79 | 0 | 0 | 14 | 1 | 1 | 0 | 1 | 0 | 0 | 18 | 1 | 4 | 6 | 2 | 1 | 0 | 3 | 0 | 0 | 0 | 5 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 2 | 20 | 0 | 0 | 2 | 4 | 0 | 1 | 1 | 3 | 0 | 2 | 21 | 0 | 16 | 4 | 0 | 0 | 0 | 2 | 0 | 2 | 0 | 2 | 0 | 0 | 3 | 1 | 1 | 4 | 22 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 3 | 0 | 1 | 0 | 21 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 2 | 0 | 0 | 2 | 1 | 0 | 2 | 13 | 0 | 3 | 13 | 6 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 15 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 13 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 3 | 0 | 0 | 1 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 1 | 1 | 0 | 2 | 1 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 12 | 0 | 0 | 5 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 3 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 2 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 3 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Intracranial Hemorrhage | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Intracranial Pressure Increased | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Intraoperative Cardiac Injury | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ischemia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ischemia Stroke | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ischemic Heart Disease | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Laceration(s) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Laceration(s) of Esophagus | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Liver Damage/Dysfunction | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
Liver Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Loss of Range of Motion | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Loss of consciousness | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Low Blood Pressure/ Hypotension | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Low Oxygen Saturation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lymphoma | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Malaise | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Medical device site infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mitral Valve Insufficiency/ Regurgitation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Movement Disorder | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Multiple Organ Dysfunction Syndrome | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Muscle Weakness/Atrophy | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Myocardial Infarction | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Nausea | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 24 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Necrosis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Nervous System Injury | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Clinical Signs, Symptoms or Conditions | 10 | 1 | 0 | 0 | 6 | 1 | 3 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 43 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 2 | 0 | 11 | 0 | 0 | 4 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 4 | 0 | 0 | 10 | 0 | 0 | 0 | 2 | 0 | 0 | 4 | 0 | 15 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 3 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 3 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 4 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0
Non specific EKG/ECG Changes | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Obstruction/Occlusion | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Osteomyelitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Oversedation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pain | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pancreatitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Paralysis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Paresis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Perforation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Perforation of Esophagus | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Perforation of Vessels | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pericardial Effusion | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pericarditis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Peritoneal Laceration(s) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Peritonitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pleural Effusion | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pneumonia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pneumothorax | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Post Operative Wound Infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pulmonary Dysfunction | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 9 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pulmonary Edema | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pulmonary Embolism | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rash | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Reduced Blood Flow | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Renal Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Renal Impairment | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Reperfusion Injury | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Respiratory Acidosis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Respiratory Arrest | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Respiratory Failure | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Respiratory Insufficiency | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Respiratory Tract Infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Retroperitoneal Hemorrhage | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Rupture | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ruptured Aneurysm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sepsis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Septic Shock | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Seroma | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Shock | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sinusitis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Skin Inflammation/ Irritation | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 27 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Skin Tears | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sleep Dysfunction | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sleepiness/Drowsiness/Somnolence | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Sore Throat | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Stenosis | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Stroke/CVA | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Swelling/ Edema | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Syncope/Fainting | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Tachycardia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Thrombocytopenia | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Thromboembolism | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Thrombosis/Thrombus | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
Tissue Breakdown | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Transient Ischemic Attack | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unintended Extubation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Eye / Vision Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Gastrointestinal Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Heart Problem | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 23 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 2 | 1 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Hepatic or Biliary Problem | 6 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 18 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 10 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
Unspecified Immune System Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Kidney or Urinary Problem | 6 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 47 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 10 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 3 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 24 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
Unspecified Mental, Emotional or Behavioural Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Nervous System Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Respiratory Problem | 21 | 0 | 0 | 0 | 1 | 0 | 5 | 0 | 0 | 0 | 15 | 0 | 0 | 0 | 130 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 26 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 1 | 10 | 0 | 0 | 0 | 1 | 0 | 3 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 43 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 3 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Tissue Injury | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unspecified Vascular Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 17 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Upper Respiratory Tract Irritation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Urinary Retention | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Urinary Tract Infection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Uterine Perforation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Vascular Dissection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
Ventilator Dependent | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Ventricular Fibrillation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Visual Disturbances | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Vomiting | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 23 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Wound Dehiscence | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Device-Product Matrix
| REMSTAR PRO C-FLEX+ | INJECTOMAT AGILIA BR | AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | AGILIA SP MC BR | REP DREAMSTATION AUTO CPAP | DREAMSTATION BIPAP AVAPS30 | REMSTAR PLUS | EDWARDS ESHEATH PLUS INTRODUCER SHEATH | MITRACLIP | CATH F5.2+ PIG 145° 110CM 6SH | REMSTAR AUTO A-FLEX | ASTRAL 150 - AMER | NIPRO SURDIAL DX HEMODIALYSIS SYSTEM | POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS | DREAMSTATION AUTO CPAP | NIPRO BLOOD TUBING SET W/TP | NIPRO ELISIO-H DIALYZER | DAVINCI XI | BIPAP AUTO BIFLEX | REACT CATHETER | NXSTAGE SYSTEM ONE | BIPAP A40 PRO | TEMPUS LS | VISISHEATH DILATOR SHEATH | DREAMSTATION AUTO BIPAP | SUREFORM | ION | REMSTAR AUTOA-FLEX | BIPAP AUTO | COBLATION HALO WAND | EDWARDS ESHEATH+ INTRODUCER | INTELLIVUE MX450 PATIENT MONITOR | SENSATION PLUS 8FR. 50CC IAB & ACCESSORIES (W/PRESSURE TUBES, NO STYLET) | EMBOTRAP III 6.5 MM X 45 MM | VERSACROSS CONNECT LAAC ACCESS SOLUTION | DREAMSTATION CPAP PRO | REMSTAR AUTO A-FLEXX | GIRAFFE OMNIBED | INVACARE® PLATINUM® OXYGEN CONCENTRATOR | HOMEFILL® OXYGEN SYSTEM | SHILEY | PRLNE BLU 36IN 4-0 D/A SH | SOLITAIRE FR | EVERFLO | LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | PHILIPS CPAP | SAFARI2¿ | DREAMSTATION GO AUTO | PHILIPS CPAP DEVICE | GORE® DRYSEAL FLEX INTRODUCER SHEATH | MESH ¿ VENTRALEX | SEPRAMESH IP | TRILOGY 100 | OMNIPOD 5 POD | DREAMSTATION BIPAP AUTOSV | DREAMSTATION ST30 | SYSTEM ONE HTD HUMID ASSY ,DOM | DS2ADV AUTO CPAP | AZURION | FORTILINK-C IBF SYSTEM | COMPANION 5 | DJO SURGICAL | DREAMSTATION BIPAP PRO | AGILIA VP MC BR | OMNIPOD 5 CONTROLLER | DREAMST ST30 H | POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | NELLCOR | TRILOGY O2 VENTILATOR, JAPAN | VIVO 45 LS | BIPAP AUTOSV ADV | SENTRANT INTRODUCER SHEATH | TUBING SET | RENTAL, DREAMSTATION ST30, GB | SIGNIA | BIPAP A40 | SOLITAIRE X | 6VT-D | ZOOM REPERFUSION CATHETER | MX40 1.4 GHZ SMART HOPPING | HEART LUNG MACHINE | CARDIOSAVE HYBRID, TYPE B PLUG | MILLENNIUM M600 WITHOUT OPI DOMESTIC | REMSTAR PRO C-FLEX | S5 HEART-LUNG-MACHINE | DREAMSTATION CPAP | HIGH FLOW INSUFFLATION UNIT | TELEHEALTH CENTER PRODUCT LICENSE | AERO TRACHEOBRONCHIAL STENT | TREK | MINI TREK¿ | NC TRAVELER CORONARY DILATATION CATHETER | TRAVELER RX CORONARY DILATATION CATHETER | MINI TREK | NC TREK | ILET BIONIC PANCREAS | NC TREK NEO¿ | HEMODIALYSIS BLOODLINES | DREAMSTATION 2 ADVANCE AUTO CPAP | AIRSEAL | V-18 CONTROL WIRE | FILMARRAY® PNEUMONIA PANEL PLUS | ASTRAL 150 DBL KIT FRA | EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | DREAMSTATION HUMIDIFIER, GB-RENT | PATIENT INFORMATION CENTER IX | SURPASS EVOLVE 4.0MM X 17MM - PMAS | INTELLIVUE MX800 PATIENT MONITOR | RENTAL-DREAMSTATION AUTO CPAP, GB | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | AMNISURE ROM TEST | POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | MICRA | PERFIX PLUG LIGHT | ENDOWRIST | NITREX | SIMPLEX P WITH TOBRAMYCIN 1 PACK | VISUALASE SYSTEM | LIBERATOR | RENTAL-DREAMSTATION AUTO CPAP | POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | MO.MA ULTRA | SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER SL 1 CURVE, 63 CM LENGTH, 8F | EMBOSHIELD NAV6 | STARTER GUIDEWIRE | BENTSON TYPE PTFE COATED GUIDEWIRE | BIPAP AUTO BIFLEX, | DREAMSTATION | CERELINK ICP PROBE 1L M BOLT | FISHER & PAYKEL HEALTHCARE | CERELINK ICP MONITOR | SOLITAIRE FR4 | CATHERA | CAPTIVATOR EMR | REMSTAR PRO C-FLEX+ W/HUMID, SD CARD | COCHLEAR OSIA OSI200 IMPLANT | OBSIDIO | UNKNOWN RENASYS TOUCH DEVICE | TRILOGY EVO | TEMPUS LS-MANUAL | EVERFLO DOM | AERO® TRACHEOBRONCHIAL STENT | PACIFIC PLUS | HAMILTON-T1 | VIZIGO | DREAMSTATION AUTO ,CPAP | ROTAREX | ALARIS SYSTEM | TURNPIKE SPIRAL 135CM | REP DREAMSTATION AUTO BIPAP, DOM-RECRT | HAMILTON-C1 | 2008T HEMODIALYSIS SYSTEM W/BIBAG | AVEIR¿ INTRODUCER KIT ¿ 50 CM | SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | DREAMSTATION AUTO CPAP, JP | REBAR | FREEDOM PERIPHERAL NERVE STIMULATOR | POWERPORT ISP MRI 8CF INT. W/SP, ATT, SL | TRI-STAPLE 2.0 | DREAMSTATION BIPAP | LTV 1150 VENTILATOR | SENSATION 7FR. 40CC IAB | HEMOCHRON ACTIVATED CLOTTING TIME PLUS (ACT+) TEST | RESPIRONICS | EDWARDS TRANSFEMORAL BALLOON CATHETER | EVIS EXERA III DUODENOVIDEOSCOPE | INTELLAMAP ORION HIGH RESOLUTION MAPPING CATHETER | DREAM STATION AUTO CPAP | VERSYS FEMORAL STEM CEMENTED REVISION/CALCAR 12/14 NECK TAPER SIZE 13 250MM | WATCHMAN TRUSTEER ACCESS SYSTEM | ESSENCE RENTAL | ESSENCE RENTAL WITH SSD | CENTRIMAG BLOOD PUMP | ESSENCE WITH SSD | LINEAR 7.5FR. 40CC IAB (CHINA) | REMSTAR PLUS C-FLEX | DEVILBISS HEALTHCARE | 2008T HEMODIALYSIS SYS., WITH CDX | BIPAP AUTO SERIES ASSY | NEOTRACT UROLIFT 2 SYSTEM | SMART PORT CT | POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | RESUSCITAIRE RADIANT WARMER | INOGEN ONE G5 PORTABLE OXYGEN CONCENTRATOR | ENDO GIA | 2008T MACHINE | BIPAP AUTOSV ADVANCED SYSTEM ONE | BIPAP PRO BIFLEX | REMSTAR | REMSTAR AUTO | REMSTAR PRO C-FLEX + | ACHIEVE MAPPING CATHETER - 20 MM | FLEXCATH ADVANCE STEERABLE SHEATH | BIVONA TRACHEOSTOMY TUBE, NEONATAL/PEDIATRIC | STRATTICE LAPAROSCOPIC 10CM X 16CM | POWERPORT VUE M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | BRILLIANCE | PACEL BIPOLAR PACING CATHETER | REMSTAR SE | INTELLIVUE MP30 | RING PROFILE 3D 30MM | BIPAP AUTOSV ADV SYS ONE | PROXIMATE*75MM LINEAR CUTTER | TIGHTRAIL ROTATING DILATOR SHEATH | DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | EDWARDS ESHEATH+ INTRODUCER SET | BIPAP PRO BI-FLEX | REMSTAR PRO C FLEX | CADD-SOLIS VIP AMBULATORY INFUSION PUMP | TABLO HEMODIALYSIS SYSTEM | REMSTAR LITE | TRILOGY | MARK 5 NUVO LITE | HEARTSTART XL+ DEFIBRILLATOR/MONITOR | FLAT CUT TALAR DOME SZ 2 INFINITY ADAPTIS | REP DREAMSTATION AUTO BIPAP | SENDIT DELIVERY DEVICE | PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER | REMSTAR W/SD CARD | AVALON FM30 FETAL MONITOR | RECERT, DREAMSTATION AUTO, FR | REMSTARAUTO | BIPAP AUTO BI FLEX | BIPAP AUTO SV ADV SYSTEM ONE | MX40 PATIENT WEARABLE MONITOR | REM AUTO A-FLEX | CORTRAK 2 ENTERAL ACCESS SYSTEM | EDWARDS ESHEATH INTRODUCER SET | SPIDER FX | REMSTAR AUTO A FLEX | XENOSURE BIOLOGIC PATCH | SERVO-I UNIT BASE | VUE PACS | LIFEPAK® 15 DEFIBRILLATOR/MONITOR | DREAMSTATION HUMID | REMSTAR AUTO SERIES AFLEX | ENDO ANCHOR SYSTEM - HELI-FX AAA | SJM RIGID SADDLE RING | SJM SÃ©GUIN SEMI-RIGID ANNULOPLASTY RING | EMBOTRAP | SIMPLYGO | SIMPLYGO MINI | MC3 NAUTILUS ECMO OXYGENATOR | CRESCENT RA DUAL LUMEN CATHETER 13FR | DREAMSTATION BIPAP AVAP | DREAMSTATION BIPAP ST30 | BIPAP PRO | ZIO AT | STAR¿ DRIVE MA | CVAC ASPIRATION SYSTEM | REM STAR AUTO A-FLEX | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | REMSTAR PRO SERIES | INFINITY CENTRAL STATION | UNKNOWN_VOCERA_PRODUCT | INTELLIVUE MX40 2.4GHZ | INTELLIVUE MX500 PATIENT MONITOR | FLOWTRIEVER CATHETER, XL | LINEAR 7.5FR. 40CC IAB | TRICLIP G4 SYSTEM | ALLURA XPER FD | PRLNE BLU 24IN 6-0 D/A BV-1 | PRESSIO ICP AND ICT MONITORING KIT, PARENCHYMAL WITH BOLT | AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | LUCAS¿ CPR CHEST COMPRESSION SYSTEM | FMC CASSETTE | BIPAP SYSTEM ONE 60 SERIES | EFFICIA DFM100 | DAVINCI X | DREAMSTATION HUMIDIFIER | CAPIOX FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | REMSTAR PLUS SERIES ASSY | BIPAP AUTO BI-FLEX | IPV-2C | REP DREAMSTATION AUTO CPAP, DOM - RECRT | REP DREAMSTATION AUTO CPAP, | IMPEDE-FX EMBOLIZATION PLUG | LEAD EXTRACTION BYRD POLYPROPYLENE DILATOR SHEATH SET | REMSTAR SYSTEM ONE 60 SERIES | LEVEEN SUPERSLIM | DUODOPA_DUOPA | DS2ADV AUTO CPAP W/HUMID CELL/BT | REP DREAMSTATION AUTO CPAP, DOM | BIPAP AUTOSV ADVANCED SYSTEM ONE 60 SERIES | 75MM SELECTABLE NEW TLC | PARKER | HEATER/COOLER SYSTEM | CONFIDA BRECKER CURVE GUIDEWIRE | ESHEATH PLUS INTRODUCER SHEATH | DREAMSTATION AUTO | COLONIC FTRD SET | STREAMLINE TEMPORARY PACING LEADS | BTK TURBOHAWK | NANOCROSS 014 | EVERCROSS 035 | ADMIRAL XTREME OTW | WECK VISTA | EPSTAR FIXED ELECTROPHYSIOLOGY CATHETER | AVALON FETAL MONITOR FM50 | PRLNE BLU 24IN 7-0 D/A BV-1 | TRILOGY EVO, JAPAN | WATCHMAN FXD CURVE ACCESS SYSTEM | STEALTHSTATION¿ S8 SYSTEM | AVALON CABLELESS ECG/IUP TRANSDUCER | REP DREAMSTATION AVAPS30 | BIPAP AVAPS C SERIES | VOLUMAT MC AGILIA US | DUAL CHAMBER TEMPORARY PACEMAKER | LUCAS CPR CHEST COMPRESSION SYSTEM | ALPHAVAC | AGILIA SP MC WIFI GB | MULTI-FUNCTIONAL RESPIRATORY P | 2008T HEMODIALYSIS SYSTEM WITHOUT CDX | BENEVISION N1 PATIENT MONITOR | KYPHON® HV-R® BONE CEMENT | ATRICLIP® PRO-V® LAA EXCLUSION SYSTEM ( | INTELLIVUE MX40 802.11A/B/G | BIPAP AUTOSV | GEM PREMIER 5000 PAK | AVALON US TRANSDUCER | EVERFLO 120V OPI | MSERIES HEATED HUMIDIFIER | LIFELINE | REP DREAMSTATION GO AUTO W/BT, DOM | INFINITY ACUTE CARE SYSTEM (M540) | CLOTTRIEVER XL CATHETER | WAVELINQ 4FR P | CONTOUR 3D | CARDIOROOT WOVEN | EVIS EXERA II ULTRASOUND GASTROVIDEOSCOPE | TFNA FEM NAIL Ø12 R 130° L400 TIMO15 | VIVO 45LS | GST BLUE RELOAD 45MM 6 ROW | BIPAP S/T C-SERIES | CEREPAK FREEFORM | PHOENIX/INNOVA MACHINES | DAVINCI SI | CARESCAPE | DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | RECERT SIMPLYGO MINI | OPTIFLUX F180NRE HF 12/CS 1.7SA EBEAM | CUSTOM COMBI SET | UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH 4.5F | DREAMSTATION CPAP | PILLCAM | ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | DELUXE 230V 50HZ INT. UK PLUG | INNOSPIRE DELUXE RENTAL UNIT | DELUXE 230V 50HZ IN | MYOSURE REACH TISSUE REMOVAL DEVICE | FLUENT PRO FLUID MANAGEMENT SYSTEM | CADD DISPOSABLE | BNX | TM80 | NOVOPEN 6 | DREAM STATION CPAP | DURAN ANCORE | COMPOSIX MESH E/X | MESH ¿ COMPOSIX KUGEL | VERSACROSS ACCESS SOLUTION | ACIST | GORE-TEX® STRETCH VASCULAR GRAFT | INFINITY DELTA | GORE® TRI-LOBE BALLOON CATHETER | SERVO-S BASE UNIT | 2008T HD SYS. CDX W/BIBAG BLUE STAR | INTELLIVUE MP70 | SPECTRUM IQ | SECHRIST INDUSTRIES INC. | CARDIOHELP | CARDIOSAVE HYBRID W/ E/F PLUG | INTELLIVUE MP5 | HEMOSPRAY ENDOSCOPIC HEMOSTAT | EMERGE? | BD ALARIS PUMP MODULE SMARTSITE INFUSION SET | EVITA 4 | CANNULA & CATHETER | CAREEVENT | PLUM 360 | TD60 TELEMETRY TRANSMITTER | EDWARDS ESHEATH PLUS | ESSENCE RENTAL WITH SSD 230V50HZ UK | MONARCH ENDOSCOPY PLATFORM | OXYLOG 3000PLUS | VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE | GELSOFT | HT70 VENTILATOR | DXM 1040 MICROSCAN WALKAWAY | EQUISTREAM ALPHACURVE DL CATHETER 14.5FR, 28CM | GLIDEPATH | KUGEL PATCH | SWAN-GANZ CCOMBO V | SPACELABS ULTRAVIEW SL COMMAND MODULE | POWERED 60 ECHELON +, 340MM SHAFT | TA | GIA | ULTRASONIC PROBE | SINGLE USE GUIDE SHEATH KIT | PROVOX2 | MEDIHONEY GEL IN TUBE, 1.5 FL OZ FLIPCAP - STERILE | CAPIOX® FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | PIEZOLITH 3000 TRIPLE FOCUS LITHOTRIPTER | PRLNE BLU 36IN 2-0 D/A V-7 MP | NOVALUNG CONSOLE (LINUX BOARD) | EMERGE | EXETER V40(TM). FEM. STEM 44MM | CARESCAPE B450 | HAMILTON G5 | RING ANNULOPLSTY TRICUSP CONTOUR 3D 28MM | REP DREAMSTATION AUTO BIPAP, DOM | ECHELON FLEX 60 POWERED PLUS ENDOPATH STAPLER ARTICULATING ENDOSCOPIC LINEAR CUT | TOUCHSOFT COAGULATOR | RENTAL, DREAMSTATION ST30 | FREESTYLE LIBRE 2 | INQUIRY¿ STEERABLE DIAGNOSTIC CATHETER | DREAMSTATION BIPAP AVAPS | ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING | CAPIOX | DA VINCI ENERGY | OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | MEDIHONEY GEL IN TUBE, .5 FL OZ TWISTOFF - STERILE | PARAPAC PLUS | HOBBS MEDICAL TOUCHSOFT COAGULATOR | SABLE | LEAD EXTRACTION EVOLUTION RL CONTROLLED-ROTATION DILATOR SHEATH SET | NOXBOXI NITRIC OXIDE DELIVERY SYSTEM | LINEAR 7.5FR. 40CC IAB WITH ACCESSORIES, US ONLY | REFLEX CATHETER | PHOTODYNAMIC BONE STABILIZATION SYSTEM | TRINAV INFUSION SYSTEM | PROWATERFLEX | FEEDING TUBE | DAVINCI 5 | INTRI24 INTRODUCER SHEATH | HARMONIC ACE 5MM SHEAR | DEVILBISS 525 OXYGEN CONCENTRATOR | INTRAOSSEOUS, INTRAOSSEOUS NEEDLE KITS | MINICAP WITH POVIDONE-IODINE SOLUTION | FILMARRAY® MENINGITIS/ENCEPHALITIS (ME) PANEL | LIFE SCOPE PT | PENDITURE LAA EXCLUSION SYSTEM | SENSATION PLUS 7.5FR. 40CC IAB & ACCESSORIES | HEMOSTAR | BENEVISION N17 PATIENT MONITOR | INTRACEPT RF PROBE | MERGE CARDIO | BIPOLAR VENTRICULAR ENDOCARDIAL LEAD | RADIOGRAPHY 7300 C | INTELLIVUE X3 | BIOFIRE® BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL | LLD EZ LEAD LOCKING DEVICE | ANGIOVAC | LIBERTY CYCLER SET, SINGLE CONN./EXT. DL | PHILIPS RESPIRONICS E30 WITH HUMIDIFIER | EVIS EXERA DUODENOVIDEOSCOPE | NIO ADULT | INJECTOR OPVSU NPS | CLEARVIEW ARTERIOTOMY SHUNT | CLEARVIEW BLOWER/ MISTER | DISTAL PERFUSION KIT | CRAGG-MCNAMARA | XENMATRIX AB | CORTRAK 2 NG/NI FEED, ELECTROMAGNETIC STYLET, ENFIT, CONNECTOR | AMPLATZER VASCULAR PLUG II | HAKIM PROGRAMMABLE VALVE | SAVVYWIRE | TI-CRON | RHYTHM ANALYTICS - SOFTWARE (SAMD) | REMSTAR AUTO A-FLEX W/HUMID, SD CARD | PHENOM CATHETER | POWERED WHEELCHAIR SYSTEM | QC HEMOSTATIC DRESSING, 2X2, 10X | FUSION HOLLOW FIBER OXYGENATOR | INTELLIPAP AUTO BILEVEL W/HEATED HUMIDIFIER | PROGRIP | I-NEB CF UK UNIT | TRILOGY DEVICE | POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | PENTAX MEDICAL | SPYSCOPE DS II | EPIQ CVX | APEXPRO TELEMETRY SYSTEM V6 | DAVINCI SINGLE PORT | CHPV RT ANG RESERVOIR UNITIZED | PLUM 360 (V 15.X) | DREAMSTATION 2 ADVANCED AUTO CPAP | FLEXOR RAABE GUIDING SHEATH | LATERAL F3D 18 MM IMPLANT; COBB ELEVATOR | WATCHMAN TRUSTEER? ACCESS SYSTEM | ENSEAL X1 CURVED 25CM SHAFT | HARMONIC 1100 SHEARS 36CM | UMBILICAL CATHETER | COOLSEAL¿ TRINITY | COOLSEAL¿ GENERATOR | CIRCULAR MECH STD 29MM | HLS SET | CONFIDENCE KIT, NO NEEDLES | POWERTRIALYSIS | B125 | REVEAL LINQ | PROXIMATE*VASC LIN STAP RELOAD | NOVUM IQ | TRILOGY EVO, USA | TANDEM MOBI INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | SAPPHIRE M.T INFUSION PUMP - GERMAN | ENDO CLIP III | AMPLATZER TORQVUE DELIVERY SYSTEM | ULRICH MEDICAL | UL_INJECT | GALAXY SYSTEM | INDIGO SYSTEM FLASH ASPIRATION CATHETER | DREAMSTATION BIPAP AUTOSV, DS | COMBISET SMARTECH BLOODLINE (NO HEPARIN) | FLEXCATH CONTOUR¿ STEERABLE SHEATH | INTELLISPACE EVENT MGMT PLATFORM | HOMECHOICE APD SYSTEM | CHORDX UNKNOWN | EPIQ CVXI | EXTRACTOR PRO RX | BIOFLO VORTEX | BIPAP AUTOSV ADVANCED | EDISON SYSTEM | REMSTAR PRO C-FLEX+ SYSTEM ONE 60 SERIES | MRIDIUM | CS300 INTRA-AORTIC BALLOON PUMP, SPANISH, 110V | PORTEX ENDOTRACHEAL TUBE, SILICONIZED | EUPHORA RX | CLOTTRIEVER BOLD CATHETER | POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | SIMUFORM | DREAM STATION AUTO BIPAP | INNOSIGHT | LIFEPAK 15 DEFIBRILLATOR/MONITOR | EVERFLO INTL OPI | MEDICALYTE | OPTICROSS? HD | OMNIPOD 5 AUTOMATED INSULIN DELIVERY SYSTEM | STEERABLE GUIDE CATHETER | OPTICROSS HD | DREAMSTATION BIPAP S/T 30 W/HEATED HUMIDIFIER, GB | HEMOSIL HIT- AB(PF4-H) | REMSTAR PRO SERIES C-FLEX + ASSY | BIPAP PRO SERIES ASSY | REMSTAR AUTO SERIES AFLEX ASSY | CS300 INTRA-AORTIC BALLOON PUMP, PORTUGUESE, 220V | EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | PULSE | KANGAROO | INTELLIVUE MULTI-MEASUREMENT MODULE X3 | DURAMATRIX-ONLAY COLLAGEN DURA SUBSTITUTE MEMBRANE | UNKNOWN_NEUROVASCULAR_PRODUCT | X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | WALLFLEX ESOPHAGEAL | ANGIOJET? SOLENT? OMNI | OP TELESCOPE, 7°, 11 FR. X 220 MM, 7.5 FR. CHANNEL, ANGLED OCULAR | VALLEYLAB | GORE® TRI LUMEN CATHETER | INTELLIVUE PATIENT MONITOR MX750 | MAIN ASSEMBLY, VENTILATOR, COMMERCIAL, ZVENT, BASI | VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | CNS-2101 CENTRAL MONITOR | BIA-GEN | PERCUFLEX URINARY DIVERSION STENT SET | INTELLIVUE MX40 WLAN | REMSTAR PLUS W/C-FLEX | PROTEK DUO VENOUS CANNULA | MYOSURE TISSUE REMOVAL DEVICE | AVALON FM50 FETAL MONITOR | WATCHMAN FXD CURVE? ACCESS SYSTEM | NAVIGATOR HD | PROLENE POLYPROPYLENE SUTURE UNKNOWN PRODUCT | PREMICATH | CIRCULO | FREESTYLE LIBRE 3 | INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | LIBERTY CYCLER SET, DUAL PATIENT CONNECT | INTELLIVUE MX550 PATIENT MONITOR | BEAD BLOCK | PANORAMA CENTRAL STATION WITH TELEMETRY | CAPIOX®FX | SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER, SL1¿, 63 CM LENGTH, 8.5F | ERBE VIO 3 | TRILOGY 202 VENTILATOR, INTERNATIONAL | RADICAL-7 | K-PACK II | MEDTRONIC NAVIGATION | T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | ARROW CVC SET: 4-LUMEN 8.5FR X 16CM | PU-681RA | INTELLIVUE MX40 1.4GHZ | AEROGEN SOLO | CARDIOSAVE HYBRID, 3.1 EDITION | TRANS-RAY PLUS 7.5FR 35CC IAB WITH ACCESSORIES | AMPLATZER VASCULAR PLUG | PORT-A-CATH IMPLANTABLE ACCESS SYSTEM | AVEIR¿ RETRIEVAL SYSTEM CATHETER | INTRODUCER KIT FOR GASTROSTOMY FEEDING TUBE | CG FUTURE BAND | MESH - VENTRALEX | BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | FREESTYLE LIBRE 3 PLUS | TANDEM MOBI INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | DRAGONFLY OPSTAR¿ | TRILOGY EVO, O2, JAPAN | CS100 INTRA-AORTIC BALLOON PUMP, ENGLISH, 220V | GORE® MOLDING & OCCLUSION BALLOON CATHETER | 0.14 GUIDEWIRES | EPOC BLOOD ANALYSIS SYSTEM | BIPAP PRO BIFLEX, W/SD CARD, US-RENTAL | PORTEX BLUSELECT SUCTIONAID CUFFED TRACHEOSTOMY | EVERFLO 230V, OPI, AUS, ULTRAFILL | COLLECTIVE CONCENTRATE- NATURALYTE | QUANTUM WORKSTATION 12.1" ELITE | MEGA 7.5FR. 40CC IAB | M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | VACCESS CT LOW-PROFILE POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 8F | ECHELON 3000 60MM STANDARD | TELESCOPE GUIDE EXTENSION CATHETER | BIPAP S/T C SERIES | MAGNETOM ESPREE | HERO® | AVALON FETAL MONITOR FM30 | INTRAOSSEOUS NEEDLE KITS | STEALTH
Activation Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Activation Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
Activation, Positioning or Separation Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Adverse Event Without Identified Device or Use Problem | 7 | 0 | 2 | 0 | 10 | 1 | 1 | 0 | 5 | 0 | 9 | 0 | 1 | 0 | 64 | 1 | 1 | 28 | 3 | 8 | 14 | 1 | 0 | 1 | 19 | 0 | 16 | 2 | 1 | 1 | 1 | 0 | 0 | 0 | 9 | 5 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 3 | 45 | 1 | 4 | 1 | 7 | 4 | 0 | 0 | 3 | 1 | 2 | 23 | 0 | 30 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 1 | 3 | 0 | 1 | 3 | 2 | 1 | 4 | 25 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 3 | 1 | 0 | 1 | 2 | 1 | 1 | 1 | 1 | 2 | 1 | 1 | 0 | 2 | 1 | 0 | 0 | 0 | 29 | 2 | 4 | 0 | 0 | 12 | 12 | 1 | 1 | 0 | 9 | 0 | 0 | 2 | 0 | 1 | 0 | 3 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 22 | 0 | 1 | 0 | 5 | 2 | 1 | 0 | 0 | 0 | 1 | 8 | 0 | 2 | 0 | 0 | 0 | 2 | 0 | 1 | 2 | 0 | 1 | 0 | 1 | 2 | 3 | 2 | 3 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 3 | 1 | 0 | 1 | 4 | 4 | 0 | 3 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 2 | 4 | 0 | 0 | 1 | 1 | 0 | 0 | 2 | 0 | 2 | 3 | 6 | 0 | 1 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 1 | 3 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 6 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 5 | 1 | 1 | 2 | 2 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 2 | 1 | 0 | 8 | 1 | 0 | 0 | 0 | 0 | 13 | 1 | 1 | 1 | 12 | 1 | 3 | 7 | 2 | 2 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 9 | 5 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 2 | 0 | 1 | 1 | 0 | 0 | 0 | 4 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 4 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 6 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 2 | 0 | 1 | 2 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 0 | 3 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 3 | 1 | 1 | 1 | 2 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 1 | 2 | 1 | 0 | 0 | 0 | 2 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 1 | 5 | 1 | 0 | 0 | 1 | 1 | 1 | 2 | 4 | 0 | 0 | 1 | 3 | 0 | 1 | 1 | 0 | 0 | 0 | 1 | 1 | 2 | 1 | 1 | 1 | 2 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 3 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 3 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0
Air/Gas in Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Alarm Not Visible | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Application Program Freezes, Becomes Nonfunctional | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Application Program Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Appropriate Term/Code Not Available | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Break | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Burst Container or Vessel | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Capturing Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Complete Loss of Power | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Connection Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Contamination | 2 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 7 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Contamination of Device Ingredient or Reagent | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Corroded | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Crack | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Defective Alarm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Defective Component | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0
Defective Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Defibrillation/Stimulation Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Deflation Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Deformation Due to Compressive Stress | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Degraded | 50 | 0 | 0 | 0 | 3 | 1 | 16 | 0 | 0 | 0 | 56 | 0 | 0 | 0 | 293 | 0 | 0 | 0 | 12 | 0 | 0 | 0 | 0 | 0 | 54 | 0 | 0 | 27 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 18 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 3 | 19 | 0 | 0 | 0 | 3 | 0 | 5 | 2 | 1 | 1 | 0 | 0 | 0 | 0 | 11 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 6 | 0 | 118 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 5 | 4 | 7 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 2 | 0 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 4 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 3 | 1 | 0 | 3 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Delayed Program or Algorithm Execution | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Delivered as Unsterile Product | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Detachment of Device or Device Component | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Device Alarm System | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Device Contaminated During Manufacture or Shipping | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Device Damaged by Another Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Device Dislodged or Dislocated | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Device Handling Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Device Sensing Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Difficult to Advance | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Difficult to Insert | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Difficult to Open or Close | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Difficult to Remove | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0
Disconnection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Display or Visual Feedback Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Electrical /Electronic Property Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Electrical Power Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Entrapment of Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Erratic Results | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Erratic or Intermittent Display | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Excess Flow or Over-Infusion | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Expiration Date Error | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure of Device to Self-Test | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Advance | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Calibrate | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Charge | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Cut | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Deflate | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Deliver | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Deliver Energy | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Deliver Shock/Stimulation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Fire | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Form Staple | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Prime | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Run on Battery | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Failure to Transmit Record | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
False Negative Result | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
False Positive Result | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Filling Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Firing Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fluid/Blood Leak | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Fracture | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Free or Unrestricted Flow | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Gas Output Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Gas/Air Leak | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
High Readings | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
High Test Results | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Human-Device Interface Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Improper or Incorrect Procedure or Method | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Inability to Auto-Fill | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Inaccurate Delivery | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Inadequacy of Device Shape and/or Size | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Incomplete or Inadequate Priming | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Incorrect Measurement | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0
Incorrect, Inadequate or Imprecise Result or Readings | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Increase in Pressure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Inflation Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Infusion or Flow Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Insufficient Flow or Under Infusion | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Insufficient Information | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 13 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 2 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Intermittent Communication Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Lack of Effect | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Leak/Splash | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Loss of Data | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Low Readings | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Low Test Results | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Malposition of Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Deformation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Discolored | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Frayed | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Integrity Problem | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Perforation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Puncture/Hole | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 4 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Rupture | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Separation | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
Material Split, Cut or Torn | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Material Twisted/Bent | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mechanical Jam | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mechanical Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Mechanics Altered | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Microbial Contamination of Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Migration | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Migration or Expulsion of Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Misassembled | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Misassembly During Maintenance/Repair | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Misassembly by Users | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Misconnection | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Misfire | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Missing Test Results | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Apparent Adverse Event | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Audible Alarm | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 7 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Audible Prompt/Feedback | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Device Output | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Flow | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Pacing | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
No Visual Prompts/Feedback | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Noise, Audible | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Obstruction of Flow | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0
Off-Label Use | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Output Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Overheating of Device | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Patient Device Interaction Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 1 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 5 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Patient-Device Incompatibility | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Peeled/Delaminated | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Physical Resistance/Sticking | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Poor Quality Image | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Poor Visibility | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Positioning Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Positioning Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pressure Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Protective Measures Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pumping Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Pumping Stopped | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 3 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Radiation Output Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Retraction Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Separation Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Therapeutic or Diagnostic Output Failure | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Tidal Volume Fluctuations | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unable to Obtain Readings | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unexpected Shutdown | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unexpected Therapeutic Results | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Unintended System Motion | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1
Unstable | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Use of Device Problem | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 5 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 8 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 2 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0
Use of Incorrect Control/Treatment Settings | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
Wireless Communication Problem | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0

## Top 20 Devices Detail
Brand Name | K Number | Patient Problems | Product Problems | Event Narrative | Manufacturer Narrative
REMSTAR PRO C-FLEX+ | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A REMSTAR PRO C-FLEX+ DEVICE. THERE WAS AN ALLEGATION OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR PRO C-FLEX+ DEVICE'S SOUND ABATEMENT FOAM. THERE WAS AN ALLEGATION OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. ADVERSE EVENT/PRODUCT PROBLEM (UPDATED FROM ADVERSE EVENT TO BOTH), OUTCOMES ATTRIBUTED TO AE, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID WERE UPDATED.
REMSTAR PLUS | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND GASTROESOPHAGEAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH DUE TO CANCER FOR USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED, AND MANUFACTURER COULD NOT CONFIRM PARTICLES IN AIR PATH DURING DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. H6 UPDATED/ CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Insufficient Information | Insufficient Information | IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED SURGICAL PROCEDURE, THE PATIENT EXPIRED. THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE WAS NOTIFIED OF THE PATIENT DEATH BY A SURGEON; NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS TO OBTAIN FURTHER INFORMATION WERE MADE BUT HAVE NOT BEEN SUCCESSFUL. | THE FOLLOWING INVESTIGATIONS COULD NOT BE PERFORMED DUE TO INSUFFICIENT INFORMATION PROVIDED (I.E. PROCEDURE/EVENT DATE, UNCONFIRMED SURGEON NAME, PROCEDURE TYPE): EVENT VERIFICATION, SYSTEM/INSTRUMENT LOG REVIEW FOR THE EVENT, PROCEDURE INFORMATION. THE USER FACILITY HAS TWO DAVINCI ROBOTIC SYSTEMS. THE INFORMATION IN SECTION D, SUSPECT MEDICAL DEVICE, IS POTENTIALLY THE DEVICE USED IN THE EVENT. THEIR OTHER DAVINCI SYSTEM IS MODEL# 380652-55, SERIAL (B)(6), UDI (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED ON BOTH OF THE DAVINCI SYSTEMS CURRENTLY IN USE AT THIS HOSPITAL WITH NO NON-CONFORMANCE'S IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SITE'S COMPLAINT HISTORY DOES NOT SHOW ANY REPORTED EVENTS RELATED TO THIS EVENT. A REVIEW OF THE EVENT WAS PERFORMED BY AN INTUITIVE SURGICAL, INC. (ISI) MEDICAL SAFETY OFFICER (MSO) WHO CONCLUDED THAT THE TYPE OF PROCEDURE AND THE EVENTS AROUND THIS PATIENT¿S DEATH WERE NOT REPORTED. NO ALLEGATIONS AGAINST ANY INTUITIVE PRODUCTS WERE MADE. INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL SYSTEM OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | Insufficient Information | Insufficient Information | IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED SURGICAL PROCEDURE, THE PATIENT EXPIRED. THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE WAS NOTIFIED OF THE PATIENT DEATH BY A SURGEON; NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS TO OBTAIN FURTHER INFORMATION WERE MADE BUT HAVE NOT BEEN SUCCESSFUL. | THE FOLLOWING INVESTIGATIONS COULD NOT BE PERFORMED DUE TO INSUFFICIENT INFORMATION PROVIDED (I.E. PROCEDURE/EVENT DATE, UNCONFIRMED SURGEON NAME, PROCEDURE TYPE): EVENT VERIFICATION, SYSTEM/INSTRUMENT LOG REVIEW FOR THE EVENT, PROCEDURE INFORMATION. THE USER FACILITY HAS TWO DAVINCI ROBOTIC SYSTEMS. THE INFORMATION IN SECTION D, SUSPECT MEDICAL DEVICE, IS POTENTIALLY THE DEVICE USED IN THE EVENT. THEIR OTHER DAVINCI SYSTEM IS MODEL# 380652-55, SERIAL# (B)(6), UDI (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED ON BOTH OF THE DAVINCI SYSTEMS CURRENTLY IN USE AT THIS HOSPITAL WITH NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SITE'S COMPLAINT HISTORY DOES NOT SHOW ANY REPORTED EVENTS RELATED TO THIS EVENT. A REVIEW OF THE EVENT WAS PERFORMED BY AN INTUITIVE SURGICAL, INC. (ISI) MEDICAL SAFETY OFFICER (MSO) WHO CONCLUDED THAT THE TYPE OF PROCEDURE AND THE EVENTS AROUND THIS PATIENT¿S DEATH WERE NOT REPORTED. NO ALLEGATIONS AGAINST ANY INTUITIVE PRODUCTS WERE MADE. INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL SYSTEM OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS ZERO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT CORRECTED. SECTION H6 UPDATED IN THIS REPORT.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Vomiting; Nausea; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NAUSEA / VOMITING, LIVER DISEASE/TOXICITY, CANCER, DEATH AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ION | K182188 | Respiratory Arrest; Airway Obstruction; Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING AND SUBSEQUENTLY EXPIRED. WHILE THE PHYSICIAN WAS PERFORMING A BIOPSY OF A LESION IN THE RIGHT MIDDLE LOBE, BLUSHING WAS VISIBLE ON 2-D FLUOROSCOPY IMAGING. THE ION VISION PROBE WAS REINSERTED INTO THE ION CATHETER AND A SMALL AMOUNT OF BLOOD WAS NOTED. BRONCHOALVEOLAR LAVAGE (BAL) WAS PERFORMED TO RETRIEVE THE BIOPSY SPECIMEN, AFTER WHICH APPROXIMATELY 20ML OF BLOOD WAS NOTED. THE ION SYSTEM WAS UNDOCKED AND A MANUAL THERAPEUTIC BRONCHOSCOPE WAS PLACED. BLOOD WAS OBSERVED IN THE ENDOTRACHEAL TUBE AND ¿A LOT¿ OF BLOOD WAS SUCTIONED OUT. THE PATIENT STOPPED BREATHING AND THE LUNGS COULD NOT BE OXYGENATED; ADDITIONAL BLOOD WAS SUCTIONED ALLOWING THE PATIENT TO SUCCESSFULLY BE PLACED ON VENTILATOR SUPPORT. AFTER STABILIZING FOR APPROXIMATELY 60 MINUTES, THE PATIENT CODED, REPORTEDLY DUE TO MANY BLOOD CLOTS THAT CREATED PRESSURE IN THE PATIENT'S LUNGS. RESUSCITATION WAS PERFORMED SUCCESSFULLY. THE PATIENT WAS TRANSFERRED TO THE ICU AND EXPIRED APPROXIMATELY 90 MINUTES LATER. | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING AND SUBSEQUENTLY EXPIRED. WHILE THE PHYSICIAN WAS PERFORMING A BIOPSY OF A LESION IN THE RIGHT MIDDLE LOBE, BLUSHING WAS VISIBLE ON 2-D FLUOROSCOPY IMAGING. THE ION VISION PROBE WAS REINSERTED INTO THE ION CATHETER AND A SMALL AMOUNT OF BLOOD WAS NOTED. BRONCHOALVEOLAR LAVAGE (BAL) WAS PERFORMED TO RETRIEVE THE BIOPSY SPECIMEN, AFTER WHICH APPROXIMATELY 20ML OF BLOOD WAS NOTED. THE ION SYSTEM WAS UNDOCKED AND A MANUAL THERAPEUTIC BRONCHOSCOPE WAS PLACED. BLOOD WAS OBSERVED IN THE ENDOTRACHEAL TUBE AND ¿A LOT¿ OF BLOOD WAS SUCTIONED OUT. THE PATIENT STOPPED BREATHING AND THE LUNGS COULD NOT BE OXYGENATED; ADDITIONAL BLOOD WAS SUCTIONED ALLOWING THE PATIENT TO SUCCESSFULLY BE PLACED ON VENTILATOR SUPPORT. AFTER STABILIZING FOR APPROXIMATELY 60 MINUTES, THE PATIENT CODED, REPORTEDLY DUE TO MANY BLOOD CLOTS THAT CREATED PRESSURE IN THE PATIENT'S LUNGS. RESUSCITATION WAS PERFORMED SUCCESSFULLY. THE PATIENT WAS TRANSFERRED TO THE ICU AND EXPIRED APPROXIMATELY 90 MINUTES LATER. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
REMSTAR AUTOA-FLEX | K131982 | Cancer; Liver Damage/Dysfunction; Cardiovascular Dysfunction/Insufficiency; Unspecified Respiratory Problem | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH . MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS TWENTY-FOUR ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Heart Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information; Respiratory Arrest; Unspecified Respiratory Problem; Pneumonia | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PNEUMONIA, HYPERCARBIA, RESPIRATORY ARREST, RESPIRATORY TRACT IRRITATION AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Material Integrity Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED. ON THE PREVIOUSLY SUBMITTED REPORTED, THE WRONG TYPE OF REPORTED COMPLAINT WAS SELECTED IN SECTION H. IT IS CORRECTED ON THIS REPORT.
REMSTAR AUTO A-FLEX | K091319 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON A REPLACEMENT CPAP DEVICE ISSUED TO A PATIENT THAT IS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REP DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | UPON REVIEW OF THIS COMPLAINT, THERE WAS NO SERIOUS INJURY ALLEGED. IN ADDITION, THE MALFUNCTION WILL NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY IF THE EVENT WERE TO REOCCUR. THIS COMPLAINT DOES NOT MEET THE DEFINITION OF A REPORTABLE EVENT.
DREAMSTATION AUTO BIPAP | K131982 | Respiratory Failure; Unspecified Respiratory Problem | Adverse Event Without Identified Device or Use Problem; Degraded; Contamination | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DREAMSTATION AUTO BIPAP THAT THE PATIENT ALLEGING RESPIRATORY DISTRESS, RESPIRATORY FAILURE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGES RESPIRATORY DISTRESS, RESPIRATORY FAILURE AND DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, THE MANUFACTURER OBSERVED THE CONTROL KNOB WAS MISSING. UNKNOWN BLACK CONTAMINATION, INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, IN THE AIR INLET OF THE BLOWER BOX AND ON THE UI (USER INTERFACE) PANEL. UNKNOWN BROWN RESIDUE ON THE SD (SECURE DIGITAL) CARD FLIP DOOR. PHILIPS INVESTIGATION LABORATORY WAS NOT ABLE TO GET CARE ORCHESTRATOR INFORMATION FROM THE DEVICE. POTENTIAL DRIED LIQUID SPOTS OR MINERAL DEPOSITS ON THE OUTSIDE AND INSIDE OF THE BLOWER MOTOR, INDICATING POTENTIAL LIQUID INGRESS. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET AND RUBBER GROMMET RETAINMENT SLOT OF THE BLOWER BOX. HAIR-LIKE FIBER ON THE BOTTOM ENCLOSURE. DUST-LIKE CONTAMINATION THROUGHOUT THE DEVICE. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. 0 ERRORS WERE LOGGED. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW PHILIPS INVESTIGATION LABORATORY USED A FITT (FOAM INTEGRITY TEST TOOL) AND THE ASSOCIATED PROCEDURE AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGESTS THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE AND WAS UNABLE TO ADDRESS THE PATIENT'S SYMPTOMS. IN BOX D: DEVICE AVAILABLE FOR EVAL? AND DATE RETURNED TO MFG HAS BEEN UPDATED. IN BOX H: DEVICE EVAL. BY MFG?, (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER AND LUNG DISEASE. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO LUNG DISEASE AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO LUNG DISEASE AND CANCER. PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED WITH WRONG TYPE OF REPORTED COMPLAINT AS SERIOUS INJURY. IN THIS REPORT, THE TYPE OF REPORTED COMPLAINT HAS BEEN UPDATED CORRECTLY AS DEATH. SECTION TYPE OF REPORTED COMPLAINT IN BOX H HAS BEEN UPDATED/ CORRECTED.
REMSTAR PLUS | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH AND GASTROESOPHAGEAL CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | ON 30/JUL/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2024 WHILE UNDERGOING CCPD THERAPY. DURING INTAKE THE PATIENT¿S PD REGISTERED NURSE (PDRN) REPORTED THE PATIENT¿S EXPIRATION WAS UNRELATED TO PD THERAPY. NO ADDITIONAL INFORMATION PROVIDED DURING INTAKE. ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. | PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | Cancer | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY DUE TO CANCER, SO HIS WIFE WILL NEED A RETURN LABEL TO RETURN THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR PRO C-FLEX+ | K091319 | Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 09/11/2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID WERE UPDATED IN THIS REPORT.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Material Integrity Problem | THE MANUFACTURER RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY, BUT THE PATIENT DID NOT PASS AWAY DUE TO CPAP DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY, BUT THE PATIENT DID NOT PASS AWAY DUE TO CPAP DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
PHILIPS CPAP DEVICE | K131982 | Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY DUE TO NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY DUE TO NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY AND CANCER. PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED WITH WRONG TYPE OF REPORTED COMPLAINT AS SERIOUS INJURY. IN THIS REPORT, THE TYPE OF REPORTED COMPLAINT HAS BEEN UPDATED CORRECTLY AS DEATH. SECTION TYPE OF REPORTED COMPLAINT IN BOX H HAS BEEN UPDATED/ CORRECTED.
REMSTAR PRO C-FLEX+ | K131982 | Vomiting; Dizziness; Nausea; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Chronic Obstructive Pulmonary Disease (COPD) | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG, COPD. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTOA-FLEX DEVICE'S SOUND ABATEMENT FOAM THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH. THE PATIENT ALSO ALLEGED KIDNEY DISEASE/TOXICITY, CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. | REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Skin Inflammation/ Irritation; Eye Irritation; Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Skin Inflammation/ Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF ASTHMA, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, REDUCED CARDIOPULMONARY RESERVE. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Eye Irritation; Asthma; Dizziness; Nausea; Cardiovascular Dysfunction/Insufficiency; Unspecified Immune System Problem; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA AND REDUCED CARDIOPULMONARY RESERVE, INFLAMMATORY RESPONSE, HYPERSENSITIVITY, DIZZINESS/HEADACHE, NAUSEA/VOMITING AND EYE, SKIN, NOSE AND RESPIRATORY TRACT IRRITATION. THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE HARM REPORTED. BASED ON AVAILABLE INFORMATION, AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE HAS BEEN NO COMMUNICATION FROM THE CUSTOMER ON THE RETURN OF THE DEVICE. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Skin Inflammation/ Irritation; Eye Irritation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, ASTHMA (NEW OR WORSENING). THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | Arrhythmia; Low Blood Pressure/ Hypotension; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING A PARA-ESOPHAGEAL HERNIA REPAIR PROCEDURE, THE PATIENT'S VITAL SIGNS STARTED TO DROP REQUIRING CARDIOPULMONARY RESUSCITATION; THE PATIENT EXPIRED. THE PROCEDURE WAS ALMOST COMPLETED WHEN IT WAS OBSERVED THAT THE PATIENT WAS HYPOTENSIVE AND HAD AN ERRATIC HEART RATE WITH AN UNSPECIFIED CARDIAC ARRHYTHMIA. WHILE UNDOCKING THE DA VINCI ROBOT TO CONVERT TO AN EMERGENT OPEN PROCEDURE, THE PATIENT "FLAT LINED" AND CPR WAS PERFORMED FOR APPROXIMATELY 30 MINUTES WITHOUT SUCCESS. THE PATIENT'S FAMILY DECLINED AN AUTOPSY AND INFORMED THE SURGEON THAT THE PATIENT'S UNCLE HAD DIED DURING A SURGERY IN THE SAME WAY YEARS PRIOR. ALTHOUGH NO TESTING WAS DONE, THE SURGEON STATED THAT THE PATIENT DEATH WAS POTENTIALLY EITHER A MYOCARDIAL INFARCTION/CARDIAC EVENT OR A MASSIVE PULMONARY EMBOLISM. THE SURGEON STATED NO DA VINCI PRODUCTS CONTRIBUTED TO THE EVENT. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, TIP-UP FENESTRATED GRASPER, LARGE NEEDLE DRIVER, CADIERE FORCEPS, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED DURING A PARA-ESOPHAGEAL HERNIA PROCEDURE. THERE WERE NO INTRAOPERATIVE ISSUES UP UNTIL THE TIME OF THE EVENT AND NO INJURIES WERE REPORTED. THERE WERE NO ALLEGATIONS OF ANY ISSUES OR PROBLEMS WITH ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. THE CAUSE OF THE DEATH IS UNKNOWN. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DS2ADV AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION THAT AN END USER DEVELOPED DEATH WHILE USING AN DS2ADV AUTO CPAP. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY BECAME AWARE OF AN ALLEGATION THAT AN END USER DEVELOPED DEATH WHILE USING AN DS2ADV AUTO CPAP. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IN PREVIOUS REPORT, MANUFACTURER MENTIONED BOX B (ADVERSE EVENT/PRODUCT PROBLEM) INCORRECTLY, WHICH HAS BEEN CORRECTED IN THIS REPORT. THE DEVICE IS NOT A PART OF RECALL. IN THIS REPORT, BOX B, H HAS BEEN UPDATED OR CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | Skin Inflammation/ Irritation; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE DEVICE LOGS WERE DOWNLOADED AND NO ERRORS CODE WERE FOUND. THERE WAS NO EVIDENCE OF FOAM PARTICLES DURING THE TESTING OF THE DEVICE. NO OTHER DEVICE PROBLEMS WERE FOUND. THE DEVICE WAS SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Dizziness; Nausea; Cardiovascular Dysfunction/Insufficiency; Unspecified Immune System Problem; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING IRRITATION (EYE, NOSE, SKIN, RESPIRATORY TRACT), DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA/VOMITING, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, LUNG DISEASE (UNSPECIFIED). THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Unspecified Heart Problem; Dizziness; Nausea; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA / VOMITINGINFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Eye Irritation; Unspecified Heart Problem; Dizziness; Nausea; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED OF EYE IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA/VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, RESPIRATORY TRACT IRRITATION AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Fungal Infection | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED FUNGAL INFECTION IN LUNGS. THE PATIENT DIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Inflammation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING INFLAMMATORY RESPONSE, CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Inflammation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY DUE TO INFLAMMATORY RESPONSE, CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Heart Problem; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE PATIENT ALSO ALLEGED EYE IRRITATION, RESPIRATORY TRACT IRRITATION, REDUCED CARDIOPULMONARY RESERVE, CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED CANCER, STAGE 4 ESOPHAGEAL CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Vomiting; Eye Irritation; Asthma; Dizziness; Nausea; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE AND CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Unspecified Gastrointestinal Problem; Insufficient Information; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN PRODUCT UDI AND PATIENT OUTCOME CODE GRID WHICH WERE CORRECTED IN THIS REPORT. DESCRIPTION EVENT OR PROBLEM HAS BEEN CORRECTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER, STOMACH, SPLEEN, INTESTINE, CANCER (UNSPECIFIED) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Skin Inflammation/ Irritation; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Appropriate Term / Code Not Available; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, LUNG, KIDNEY, LIVER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Eye Irritation; Asthma; Appropriate Term / Code Not Available; Liver Damage/Dysfunction; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, DEATH NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING AND NOT NEEDING IT ANYMORE. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL-DREAMSTATION AUTO CPAP PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Eye Irritation; Asthma; Appropriate Term / Code Not Available; Dizziness; Nausea; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING REDUCED CARDIOPULMONARY RESERVE. IN ADDITION, THE PATIENT ALSO ALLEGED EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), LUNG DISEASE. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER, WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE, DUE TO A PATIENT PASSING. AND NOT NEEDING IT ANYMORE. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION, THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED, FOR AN ALLEGATION OF A PATIENT DEATH ONLY, DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED, WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION, NO FAULTS WERE FOUND. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS RETURNED TO STOCK.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT DECEASED DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 |  |  | THE MANUFACTURER BECAME AWARE OF A RENTAL DREAMSTATION AUTO CPAP DEVICE. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 |  |  | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH, LIVER CANCER. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 |  |  | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2024 WHILE UNDERGOING CCPD THERAPY. DURING INTAKE IT WAS REPORTED THE PATIENT¿S EXPIRATION WAS UNRELATED TO PD THERAPY. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME ON (B)(6) 2024 WHILE UNDERGOING CCPD THERAPY. THE PATIENTS¿ SPOUSE REPORTEDLY DISCOVERED HER HUSBAND HAD PASSED IN HIS SLEEP AFTER SHE AROSE ¿DURING THE MIDDLE OF THE NIGHT¿ ON (B)(6) 2024. DESPITE THE PRIMARY AND SECONDARY CAUSE OF DEATH BEING UNKNOWN, THE PDRN REPORTED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO PD THERAPY. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT¿S EXPIRATION OCCURRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS CURRENTLY UNKNOWN; THEREFORE, CAUSALITY CANNOT BE FIRMLY ESTABLISHED. HOWEVER, DURING FOLLOW-UP THE PDRN STATED THE PATIENT¿S DEATH WAS UNRELATED TO PD THERAPY, NOR WAS THERE ANY REPORT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY OCCURRED. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | Cancer; Unspecified Respiratory Problem; Respiratory Failure; Pneumonia | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY DUE TO RESPIRATORY TRACT IRRITATION, CANCER, NASAL CANCER, COVID/SEVERE SEPSIS, PNEUMONIA AND RESPIRATORY FAILURE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Pneumonia | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY AND PNEUMONIA. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K091319 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY; CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Pneumonia; Aspiration/Inhalation; Stroke/CVA; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | THE PROGRAM MANAGER FOR THE ACUTE CLINIC MADE NOTIFICATION THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE ACUTE CLINIC¿S LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PROGRAM MANAGER. THE PATIENT WAS INITIALLY IN TREATMENT ON HIS HOME LIBERTY SELECT CYCLER ON (B)(6) 2024 WHEN HE EXPERIENCED APHASIA AND STROKE. THE PATIENT WAS ADMITTED TO THE HOSPITAL. NO ADDITIONAL INFORMATION PERTAINING TO THAT EVENT WAS AVAILABLE. DURING THE HOSPITALIZATION, ON (B)(6) 2024. THE ON-CALL DIALYSIS NURSE FROM THE ACUTE CLINIC WAS CALLED BY HOSPITAL STAFF TO RESPOND TO THE PATIENT¿S ROOM. THE PATIENT NEEDED TO BE DISCONNECTED FROM TREATMENT ON THE LIBERTY SELECT CYCLER TO GO FOR A COMPUTED TOMOGRAPHY (CT) SCAN. WHEN THE NURSE ARRIVED AT THE PATIENT¿S ROOM, THE PATIENT WAS CODING. THE HOSPITAL STAFF WAS ACTIVELY RUNNING THE CODE. THEY WERE UNSUCCESSFUL IN RESUSCITATING THE PATIENT AND THE PATIENT WAS PRONOUNCED DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST LIKELY SECONDARY TO ASPIRATION PNEUMONIA VERSUS WORSENING STROKE AND QUESTIONABLE PERITONITIS. IT IS UNKNOWN WHAT CYCLE OR PHASE OF TREATMENT THE PATIENT WAS IN AT THE TIME OF THE EVENT. THE NURSE DID NOT MAKE NOTE PRIOR TO BREAKING DOWN THE CYCLER FROM THE TREATMENT. IT IS UNKNOWN IF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY WERE RELATED TO THE PATIENT DEATH, BUT IT IS NOT LIKELY PER THE PROGRAM MANAGER. THIS PATIENT HAD COMORBID CONDITIONS OF THIS PATIENT HAD COMORBID CONDITIONS OF CORONARY ARTERY DISEASE (CAD), CONGESTIVE HEART FAILURE (CHF), HYPERTENSION (HTN), DEPRESSION, ALCOHOL ABUSE, STROKE, LEFT FRONTAL INFARCT, AND HOSPITAL ACQUIRED ASPIRATION PNEUMONIA. A FOLLOW-UP WITH THE PATIENT¿S CLINIC YIELDED NO ADDITIONAL INFORMATION AS THEY WERE NOT INFORMED OF THE PATIENT¿S ADMITTING DIAGNOSIS OR THE CIRCUMSTANCES RELATED TO THE PATIENT DEATH. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS AND UTILIZATION OF THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF APHASIA AND STOKE WITH HOSPITALIZATION AND SUBSEQUENT CARDIAC ARREST AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE INITIAL HOSPITALIZATION FOR STROKE AND THE SUBSEQUENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE PATIENT WAS INITIALLY HOSPITALIZED FOR A STROKE AND SUBSEQUENTLY EXPIRED AFTER EXPERIENCING CARDIAC ARREST. THE PATIENT HAD COMORBID CONDITIONS OF HYPERTENSION, CORONARY ARTERY DISEASE (CAD), AND CONGESTIVE HEART FAILURE (CHF). CARDIOVASCULAR DISEASE IS THE LEADING CAUSE OF DEATH IN PATIENTS WITH KIDNEY FAILURE WITH THE RISK OF CARDIOVASCULAR DISEASES, INCLUDING STROKE IN PATIENTS WITH END STAGE RENAL DISEASE (ESRD) BEING 5-30 TIMES GREATER THAN THE GENERAL PUBLIC. HYPERTENSION CONTINUES TO BE THE MAJOR MODIFIABLE RISK FACTOR FOR BOTH ISCHEMIC AND HEMORRHAGIC STROKE. THE INFORMATION MADE MENTION OF POSSIBLE PERITONITIS; HOWEVER, THERE WAS NO DOCUMENTATION TO CONFIRM THAT THE PATIENT WAS EVER DIAGNOSED WITH PERITONITIS DURING THE HOSPITALIZATION. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S INITIAL HOSPITALIZATION FOR APHASIA AND STROKE AND SUBSEQUENT CARDIAC ARREST AND DEATH.
PHILIPS CPAP DEVICE | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO KIDNEY DISEASE/TOXICITY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Liver Damage/Dysfunction; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, LUNG DISEASE AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER AND KIDNEY DISEASE/TOXICITY. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION KIDNEY DISEASE, CANCER AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Skin Inflammation/ Irritation; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF KIDNEY DISEASE/TOXICITY. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | PREVIOUSLY REPORTED BY THE MANUFACTURER: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. NEW INFORMATION: CORRECTION TO BOX B, DATE OF DEATH.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. A GOOD FAITH EFFORT ATTEMPT WAS MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT THE CUSTOMER DECLINED TO RESPOND TO THE QUESTIONS AND TERMINATED THE CALL. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K091319 | Cancer; Liver Damage/Dysfunction | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY; CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING KIDNEY DISEASE/TOXICITY. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT EXPIRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT EXPIRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, AND THE PATIENT EXPIRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE TECHNICIAN FOUND NO VISUAL EVIDENCE OF FOAM PARTICLES. THERE WAS NO MALFUNCTION NOTED. THERE WAS EVIDENCE OF ENVIRONMENTAL CONTAMINATION. NO COMPONENTS WERE REPLACED. THE DEVICE WAS SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED. BOX H WAS UPDATED IN THIS REPORTED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Liver Damage/Dysfunction; Vomiting; Nausea | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NAUSEA/VOMITING, LIVER DISEASE/TOXICITY, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION NOSE IRRITATION, DEATH, CANCER. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, REDUCED CARDIOPULMONARY RESERVE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Skin Inflammation/ Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED NOSE IRRITATION, SKIN IRRITATION, ASTHMA, LUNG DISEASE. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K091319 | Insufficient Information; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION INFLAMMATORY RESPONSE DEATH OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Asthma; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), LUNG DISEASE, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), LUNG DISEASE, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. SECONDARY FINDINGS WERE OBSERVED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, CANCER. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE FLIP DOOR, PCA, TOP ENCLOSURE, REAR PANEL, UI PANEL, BOTTOM ENCLOSURE, INSIDE ISO PORT, BLOWER, AND BLOWER SEAL. MANUFACTURE OBSERVED BLACK HAIRLIKE CONTAMINANT ON THE FLIP DOOR, TOP ENCLOSURE, UI PANEL, REAR PANEL, BOTTOM ENCLOSURE, AND BLOWER. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE. THE MANUFACTURER OBSERVED DUST CONTAMINATION IN THE DEVICE AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID ARE UPDATED
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Retroperitoneal Hemorrhage | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 68-91 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 4 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 5.67. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 2.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 4 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 78-86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE AND 3 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 1.25. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(6). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(6). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
REMSTAR AUTO A-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, CONCLUSION CODE GRID HAS BEEN UPDATED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 18 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 61-97 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALES AND 15 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 18 MAJOR VASCULAR COMPLICATION DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 11.94. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER (B)(4) AND 1 MAJOR VASCULAR COMPLICATION DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
REMSTAR AUTO A-FLEX | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | CORRECTED DATA: LEGACY COMPLAINT ID FROM "BLANK" TO (B)(4).
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Insufficient Information | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 75-86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 7 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.44. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Retroperitoneal Hemorrhage | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 56-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 5 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 3.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 6 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 71-82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALES AND 3 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 1.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENT FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 7 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 72-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALES AND 3 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 7 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.86. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 20 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 56-(B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 10 FEMALES AND 10 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 20 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 2.55. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 21 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 56-89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 11 FEMALES AND 10 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 21 MAJOR VASCULAR COMPLICATION DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 66.62. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Retroperitoneal Hemorrhage | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENT FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - RETROPERITONEAL DEATH EVENT FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 77 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - RETROPERITONEAL DEATH EVENT FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4).THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Insufficient Information | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENT FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 77-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 77-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 3 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 64-84 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 3 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 1.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(6). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(6). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - RETROPERITONEAL AT ACCESS SITER DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 71 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - RETROPERITONEAL AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 55-74 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 2 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 20 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 61-95 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 9 FEMALES AND 11 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 20 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 7.8. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4).ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Insufficient Information | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 61-84 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALES AND 5 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.89. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 67 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 3.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 6.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Insufficient Information | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 67 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 3.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4).ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 67 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 2.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST BLEEDING REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED LIVER DISEASE/TOXICITY; CANCER; DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF NOSE IRRITATION, RESPIRATORY TRACK IRRITATION, CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PATIENT INFORMATION CENTER IX | K211900 | Insufficient Information | Defective Alarm | THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS FIELD SERVICE PERSONNEL AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON A PATIENT INFORMATION CENTER IX SYSTEM INDICATING THAT THE DEVICE FAILED TO ALARM. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. THE PATIENT WAS NOTED TO HAVE DIED. HOWEVER, IT WAS CONFIRMED BY A CUSTOMER NURSE THAT THE DEVICE DID NOT CONTRIBUTE TO THE DEATH. | A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT TO THE CUSTOMER SITE, AND TESTED THE DEVICE. THE DEVICE PASSED TESTING. ALARM LOGS WERE PROVIDED AND WERE REVIEWED BY A PHILIPS PRODUCT SUPPORT ENGINEER (PSE). THE PSE REVIEWED THE LOGS FOR A REQUESTED TIME FROM (B)(6) 2024 16:23:47 TO (B)(6) 2024 16:23:13 FOR BED LABEL TELB142. THE PSE FOUND THAT THE DEVICE HAD ALARMED FOR MULTIPLE ALARM TYPES STARTING AT 16:42:02 ON (B)(6) 2024 THROUGH 14:40:19 ON (B)(6) 2024; ALARMS INCLUDE INOPERATIVE ALARMS, SUCH AS "CANNOT ANALYZE ECG", YELLOW LEVEL ALARMS, SUCH AS "IRREGULAR HEART RATE", AND RED LEVEL ALARMS, SUCH AS "VENTRICULAR TACHYCARDIA (VTACH)", "VENTRICULAR FIBRILLATION (VENT FIB)", AND "EXTREME TACHYCARDIA (XTACHY)". BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, PHILIPS WAS UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
DREAMSTATION AUTO BIPAP | K131982 | Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED NOSE IRRITATION, SKIN IRRITATION, KIDNEY DISEASE/TOXICITY, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NON-HODGKIN'S LYMPHOMA AND CANCER. THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NON-HODGKIN'S LYMPHOMA AND CANCER. THERE WAS A REPORT OF PATIENT DEATH. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING SERVICING OF THE DEVICE, THE UNIT PASSED FINAL TESTS. THE DEVICE HAS BEEN REPAIRED. SECTION G2 AND H6 HAVE BEEN UPDATED ACCORDINGLY.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE ARE NO ALLEGATIONS THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN EVALUATED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP | K200480 | Appropriate Term / Code Not Available | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE ARE NO ALLEGATIONS THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE EVALUATION HAS NOT BEGUN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Skin Inflammation/ Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. SECTION-H HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Dizziness; Inflammation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Unspecified Heart Problem; Dizziness; Inflammation; Respiratory Tract Infection; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, NOSE IRRITATION, DIZZINESS, HEADACHE, ASTHMA, LUNG DISEASE, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Dizziness; Nausea; Respiratory Tract Infection; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, DIZZINESS, HEADACHE, CANCER, NAUSEA, VOMITING, AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, LUNG DISEASE AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ION | K182188 | Respiratory Arrest; Hemorrhage/Blood Loss/Bleeding; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER AN UNEVENTFUL ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING IN THE RIGHT LUNG AND WENT INTO CARDIAC ARREST. THE INTUITIVE ENDOLUMINAL CLINICAL SALES REPRESENTATIVE WAS INFORMED THAT THE ION PROCEDURE WAS COMPLETED AS PLANNED AND THE ION SYSTEM WAS UNDOCKED. THE PHYSICIAN INSERTED A THERAPEUTIC BRONCHOSCOPE PER NORMAL PRACTICE AND, PER A COLLEAGUE PHYSICIAN THAT WAS NOT AT THE PROCEDURE, A LARGE CLOT WAS DISCOVERED AT THE END OF THE ENDOTRACHEAL TUBE, WHICH RESULTED IN DECREASED VISUALIZATION. THE THERAPEUTIC SCOPE WAS ADVANCED INTO THE RIGHT LUNG AND IT APPEARED TO BE FILLED WITH BLOOD. THE PHYSICIAN PLACED AN ENDOBRONCHIAL BLOCKER IN THE RIGHT LUNG AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS REVIVED AND TRANSFERRED TO THE INTENSIVE CARE UNIT INTUBATED, WHERE PARALYTIC AGENTS WERE ADMINISTERED UNTIL A NEUROLOGICAL TEST COULD BE CONDUCTED. THE PATIENT EXPIRED AN UNSPECIFIED TIME LATER. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN THAT PERFORMED THE PROCEDURE HAVE BEEN SUCCESSFUL. | THE ION BIOPSY PROCEDURE WAS REPORTED TO HAVE BEEN COMPLETED AND THE SYSTEM WAS UNDOCKED WITH NO INDICATION OF ANY PRODUCT ISSUES. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED THAT WOULD BE RELEVANT TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE ION PROCEDURE FOUND THERE WERE NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THE EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT OF UNKNOWN CO-MORBIDITIES UNDERWENT AN ION ENDOLUMINAL BIOPSY OF AN UNKNOWN TARGET. THE BIOPSY WAS COMPLETED WITHOUT ISSUE AND THE ION CATHETER WAS REMOVED, THE SYSTEM WAS UNDOCKED AND THE PROCEDURE WAS TRANSITIONED TO MANUAL BRONCHOSCOPY. A CLOT WAS NOTED AT THE END OF THE ENDOTRACHEAL TUBE WITH MANUAL BRONCHOSCOPY. BLOOD WAS ALSO NOTED IN THE RIGHT LUNG FOR WHICH AN ENDOBRONCHIAL BLOCKER WAS PLACED. THE PATIENT THEN SUFFERED A CARDIAC ARREST, WAS SUCCESSFULLY RESUSCITATED AND TRANSFERRED TO THE ICU FOR ONGOING MANAGEMENT AND ULTIMATELY DIED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. NO FURTHER DATA IS AVAILABLE DESPITE MULTIPLE ATTEMPTS. BASED ON THE AVAILABLE DATA THE EVENTS ARE PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. SECTION B2 DATE OF DEATH IS ESTIMATED.
ION | K182188 | Respiratory Arrest; Bradycardia; Pulmonary Embolism | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT FOLLOWING AN UNEVENTFUL ION ASSISTED TRANSBRONCHIAL LUNG BIOPSY PROCEDURE THE PATIENT EXPIRED. THE ION BIOPSY PROCEDURE AND SUBSEQUENT EBUS PROCEDURES WERE COMPLETED WITHOUT ISSUE. A POST-PROCEDURAL SCAN DID NOT IDENTIFY A PNEUMOTHORAX AND THE PATIENT WAS EXTUBATED AND MOVED TO RECOVERY IN STABLE CONDITION. THE PHYSICIAN REPORTED THAT AFTER 16 MINUTES THE PATIENT BEGAN EXPERIENCING DIFFICULTY BREATHING AND WENT INTO RESPIRATORY ARREST, LEADING TO A SUSPECTED BRADYCARDIC EVENT (HEART RATE 40 BPM) AND SUBSEQUENTLY CODED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR APPROXIMATELY 20 MINUTES WITHOUT REVIVAL. THE PHYSICIAN SUSPECTED A MASSIVE PULMONARY EMBOLISM BECAUSE OF THE SEQUENCE OF RESPIRATORY DISTRESS/ARREST LEADING TO BRADYCARDIA AND THEN CARDIAC ARREST. PER THE PHYSICIAN, THE ION SYSTEM DID NOT CAUSE OR CONTRIBUTE THE DEATH. THE PATIENT HAD A MEDICAL HISTORY OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND EMPHYSEMA. | A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE ION SYSTEM FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH COPD AND EMPHYSEMA UNDERWENT AN ION ENDOLUMINAL BIOPSY AND EBUS. THE PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS AND IMAGING REVEALED NO PNEUMOTHORAX. THE PATIENT WAS EXTUBATED AND TRANSPORTED TO RECOVERY. IN RECOVERY THE PATIENT REPORTED DIFFICULTY BREATHING FOLLOWED BY CARDIOPULMONARY ARREST. EFFORTS AT RESUSCITATION FOR 20 MINUTES WERE UNSUCCESSFUL AND THE PATIENT DIED. BASED ON THE AVAILABLE DATA A CARDIOVASCULAR EVENT INCLUDING PULMONARY EMBOLISM IS THE MOST LIKELY TRIGGERING EVENT. THERE WERE NO MALFUNCTIONS OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. THE EVENT WAS LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF (B)(4) WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). --OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DREAMSTATION AUTO CPAP | K131982 | Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Dizziness; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING) LUNG DISEASE, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Unspecified Heart Problem; Appropriate Term / Code Not Available; Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Unspecified Heart Problem; Dizziness; Respiratory Tract Infection; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA, LUNG DISEASE, AND REDUCED CARDIOPULMONARY RESERVE. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Appropriate Term / Code Not Available; Dizziness; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), CANCER, DEATH NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED TO FRESENIUS CUSTOMER SERVICE THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS FOUND UNRESPONSIVE BY A RELATIVE WHEN THEY WENT TO DISCONNECT THE PATIENT FROM THE LIBERTY SELECT CYCLER. THE PATIENT HAD REPORTEDLY PASSED AWAY. ADDITIONAL INFORMATION WAS PROVIDED UPON FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). IT WAS CONFIRMED THE PATIENT WAS CONNECTED TO THE CYCLER WHEN THEY WERE FOUND UNRESPONSIVE BY THE RELATIVE. THE PATIENT¿S TREATMENT HAD ALREADY BEEN COMPLETED. 911 WAS CALLED AND PARAMEDICS RESPONDED TO THE PATIENT¿S HOME. THE PATIENT WAS PRONOUNCED DECEASED UPON ARRIVAL OF EMERGENCY MEDICAL SERVICES (EMS). THERE WERE NO KNOWN ISSUES WITH THE LIBERTY SELECT CYCLER PRIOR TO THE TREATMENT. THE CAUSE OF DEATH IS UNKNOWN; HOWEVER, THE PATIENT HAD SIGNIFICANT CARDIAC HISTORY INCLUDING ATRIAL FIBRILLATION (A FIB), ISCHEMIC CARDIOMYOPATHY, CORONARY ARTERY DISEASE WITH STENT, AND HEART FAILURE WITH REDUCED LEFT VENTRICULAR EJECTION FRACTION (LVEF) = 30% WITH A HISTORY OF PLEURAL FLUID. THE PATIENT RECENTLY HAD A THORACENTESIS PROCEDURE DUE TO THIS REASON. THE PATIENT HAD BEEN INSTRUCTED TO USE 2.5% DIALYSATE SOLUTION FOR THREE DAYS AFTER BEING SEEN BY THEIR PRIMARY CARE PHYSICIAN (PCP). IT WAS THOUGHT THERE WAS EXTRA FLUID BUILDING UP IN THE LUNGS AGAIN DUE TO THE SAME ISSUE. IT WAS REPORTED THAT THE PATIENT¿S DEATH WAS NOT RELATED TO PD THERAPY OR THE USE OF ANY FRESENIUS PRODUCTS OR DEVICES. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT ON THE PUMP, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED A CASSETTE MEMBRANE. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST, A VALVE ACTUATION TEST, AND A TEACH PUMPS TEST. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE CYCLER¿S MUSHROOM HEADS. AN INTERNAL VISUAL INSPECTION IDENTIFIED EVIDENCE OF DRIED FLUID UNDER THE PUMP ASSEMBLY AND THE BOTTOM COVER. THE CAUSE OF THE DRIED FLUID COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
REMSTAR PRO C-FLEX+ | K091319 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE; CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem; Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS THIRTY-TWO ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT CORRECTED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Dizziness; Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA, KIDNEY DISEASE /TOXICITY, LUNG DISEASE, INFLAMMATORY RESPONSE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED AND BOX G: REPORT SOURCE - OTHER & 510K (RFB),BOX E : REPORTING INSTITUTION NAME UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Cardiovascular Dysfunction/Insufficiency; Dizziness; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING) LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DS2ADV AUTO CPAP DEVICE. THERE WAS AN ALLEGATION OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PLUS | K091319 | Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF NOSE IRRITATION, SKIN IRRITATION, DIZZINESS AND/OR HEADACHE, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Dizziness; Inflammation; Respiratory Tract Infection; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, INFLAMMATORY RESPONSE, CANCER AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, INFLAMMATORY RESPONSE, CANCER AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH DUE TO RESPIRATORY TRACT IRRITATION, LUNG DISEASE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED
DREAMSTATION AUTO BIPAP | K131982 | Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K163584 | Insufficient Information | No Audible Alarm | IT WAS REPORTED THE ALARMS WERE SILENCED AT THE CENTRAL STATION WITHOUT CLINICIAN INPUT. NURSES INDICATED NO ONE AT THE CENTRAL STATION SILENCED THE ALARMS OR TOUCHED THE MOUSE, YET THE ALARMS AUTO-SILENCED, PER THE CUSTOMER. THE CUSTOMER'S ALARM CONFIGURATION WAS UNKNOWN AT THE TIME OF THIS ASSESSMENT. GOOD FAITH EFFORTS ARE ONGOING. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. PATIENT CALLED REGARDING THE PATIENT LISTED AS CATHERINE. WANTS ALL COMMUNICATION STOPPED BECAUSE THE PATIENT PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Vomiting; Dizziness; Nausea; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA/VOMITING. THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Asthma; Unspecified Heart Problem; Dizziness; Inflammation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA, VOMITING, ASTHMA, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Dizziness; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA, LUNG DISEASE, DEATH AND CANCER. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Asthma; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, ASTHMA AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION REGARDING LUNG DISEASE, CANCER AND "DEATH". IN ADDITION, THE PATIENT ALSO ALLEGED RESPIRATORY TRACT IRRITATION, DIZZINESS AND HEADACHE. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE CUSTOMER REPORTED THAT THE DEVICE WILL BE RETURNED DUE TO THE USER PASSING AWAY CURRENTLY THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. MORE INFORMATION REGARDING THE EVENT IS BEING REQUESTED. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | BASED ON THE AVAILABLE INFORMATION, THIS COMPLAINT IS A DUPLICATE OF MFR #2518422-2024-100070. REFER TO MFR2518422-2024-100070 FOR ADDITIONAL INFORMATION.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem; Respiratory Tract Infection | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED RESPIRATORY TRACT IRRITATION, REDUCED CARDIOPULMONARY. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. | THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTRE FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO PARTICLES IN AIR PATH. THERE WAS NO ERROR CODE FOUND, UNIT WAS CORRECTED. EVALUATION METHOD CODE, RESULT CODE AND CONCLUSION CODE HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Appropriate Term / Code Not Available; Unspecified Heart Problem; Liver Damage/Dysfunction; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Vomiting; Unspecified Heart Problem; Nausea; Pulmonary Dysfunction; Unspecified Respiratory Problem; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. BOX H: PATIENT OUTCOME CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED. IN PREVIOUS REPORT REDUCED CARDIOPULMONARY RESERVE DIDN'T CAPTURED IN PATIENT OUTCOME CODE GRID. IN THIS HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Dizziness; Nausea; Unspecified Respiratory Problem; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Heart Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Dizziness; Nausea; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER,DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Skin Inflammation/ Irritation; Eye Irritation; Chronic Obstructive Pulmonary Disease (COPD); Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT TRRITATION AND COPD. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Lymphoma | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF CANCER, NON- HODGKIN'S LYMPHOMA, PROSTATE CANCER AND PATIENT PASSING AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Unspecified Heart Problem; Dizziness; Nausea; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Dizziness; Nausea; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, HEART ATTACK. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | Vomiting; Eye Irritation; Dizziness; Nausea; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LIVER DISEASE/TOXICITY, CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Cough; Appropriate Term / Code Not Available; Inflammation; Unspecified Kidney or Urinary Problem; Respiratory Failure; Unspecified Respiratory Problem; Pneumonia | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, LUNG DISEASE, KIDNEY DISEASE, TOXICITY, COUGH, SINUS INFECTIONS, PNEUMONIA, ACUTE RESPIRATORY FAILURE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Respiratory Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY FAILURE. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Sore Throat; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, THROAT, LUNG DISEASE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Asthma; Dizziness; Nausea; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, ASTHMA, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND THE SECTION H11 SHOULD BE UPDATED AS, THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, ASTHMA, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Unspecified Heart Problem; Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Insufficient Information; Hemorrhage/Blood Loss/Bleeding | Insufficient Information | REFER TO H11 FOR FOLLOW-UP INFORMATION. | THE VIDEO RECEIVED FOR THIS EVENT WAS REVIEWED BY A CLINICAL DEVELOPMENT ENGINEER (CDE). THE SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT FIRES A WHITE RELOAD ON A VESSEL, THEN EXPERIENCES A PAUSE FOR COMPRESSION WHEN THE STAPLER REACHES 20MM. A MESSAGE INDICATES THE TARGET TISSUE IS TOO THICK. WHEN THE STAPLER INSTRUMENT UNCLAMPS FROM THE VESSEL, MALFORMED STAPLES AND BLEEDING FROM THE VESSEL IS OBSERVED. THE SURGEON PLACES GAUZE OVER THE BLEED. UPON SLOWLY RELEASING THE GAUZE, THE BLEEDING CONTINUES. LATER IN THE VIDEO, THE STAPLER INSTRUMENT IS REINSTALLED WITH A WHITE RELOAD. THE STAPLER INSTRUMENT SUCCESSFULLY COMPLETES THE FIRE ON THE VESSEL. LATER IN THE VIDEO, THE STAPLER INSTRUMENT IS INSTALLED AGAIN WITH A GREEN RELOAD AND COMPLETES THE RELOAD FIRE ON ANOTHER LOCATION OF THE LUNG. PROFUSE BLEEDING IS OBSERVED IN THE BACKGROUND. THE ORIGIN OF THE BLEEDING IS NOT SHOWN IN THE VIDEO. SUCTIONING OF THE BLOOD WITH A LAPAROSCOPIC SUCTION IRRIGATOR IS PERFORMED. THE VIDEO ENDS. BASED ON THIS VIDEO REVIEW, THE TISSUE WAS TOO THICK TO BE TRANSECTED WITH THE WHITE RELOAD. THE DA VINCI SUREFORM USER MANUAL WARNS AGAINST IMPROPER SELECTION OF RELOAD COLOR AND INCLUDES A TABLE DETAILING THE APPROPRIATE TISSUE THICKNESS RANGE FOR EACH RELOAD COLOR.
ION | K182188 | Hemorrhage/Blood Loss/Bleeding; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE THE PATIENT BEGAN BLEEDING, WENT INTO CARDIAC ARREST AND EXPIRED. THE PATIENT HAD A HISTORY OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE WAS INFORMED THAT AT THE TIME OF THE BIOPSY PROCEDURE, EXCESSIVE BLEEDING ORIGINATED FROM THE RIGHT MIDDLE LOBE AND THE PROCEDURE WAS ABORTED. IT WAS REPORTED THAT EPINEPHRINE MAY HAVE BEEN INSTILLED TO HELP SLOW DOWN THE BLEEDING. THE PATIENT SUBSEQUENTLY WENT INTO CARDIAC ARREST AND CPR WAS PERFORMED WITHOUT SUCCESS. THERE WERE NO REPORTED ERRORS OR MALFUNCTIONS OF THE SYSTEM OR ANY INSTRUMENTS IN USE DURING THE PROCEDURE. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN HAVE BEEN UNSUCCESSFUL. | A REVIEW OF THE ION SYSTEM LOGS FOR THE EVENT FOUND THERE WERE NO OBSERVED EVENTS THAT WOULD SUGGEST A PRODUCT ISSUE, AND LOGGED EVENTS WERE IN LINE WITH NORMAL SYSTEM FUNCTIONALITY. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL BIOPSY, DURING THE PROCEDURE BLEEDING WAS NOTED FROM THE RIGHT MIDDLE LOBE AND THE PROCEDURE WAS ABORTED. IT IS UNCLEAR WHAT INTERVENTIONS WERE TAKEN FOR THE BLEEDING. THE PATIENT SUFFERED A CARDIAC ARREST. RESUSCITATIVE EFFORTS WERE UNSUCCESSFUL AND THE PATIENT DIED. THERE WERE NO REPORTED MALFUNCTIONS OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. GIVEN THE AVAILABLE DATA THE REPORTED EVENT WAS PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. ---FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. ---FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. ---KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. ---BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. ---BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Heart Failure/Congestive Heart Failure; Chronic Obstructive Pulmonary Disease (COPD) | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING COPD, CAD, SUDDEN CARDIAC ARREST, DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE; CANCER; DEATH, LARGE-CELL LYMPH ARM PRIMARY VITREORETINAL LYMPH ARM PLEURAL FLUID. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPOR TED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL, DREAMSTATION ST30, GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Respiratory Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING IRRITATION (SKIN), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, ASTHMA (NEW OR WORSENING), LUNG DISEASE (UNSPECIFIED). THERE WAS REPORT OF PATIENT DEATH. . DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | Asthma | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Skin Inflammation/ Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DIED DUE TO CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING LUNG CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PATIENT INFORMATION CENTER IX | K183387 | Insufficient Information | Device Alarm System | IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE OBTAINING CLINICAL AUDIT LOGS AFTER A PATIENT DEATH. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE NURSE MANAGER REQUESTED VERIFICATION IF LEADS OFF ALARMS OCCURRED FOR THIS PATIENT BEING MONITORED ON TELEMETRY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | A PHILPS REMOTE SERVICE ENGINEER (RSE) AND A PHILIPS REMOTE CLINICAL SUPPORT (RCS) REPRESENTATIVE SPOKE WITH THE CUSTOMER. THE RSE RETRIEVED THE CLINICAL AUDIT LOGS AND PROVIDED THESE TO THE CUSTOMER. THE RSE AND RCS ASSISTED THE CUSTOMER IN REVIEWING LOG INFORMATION AND IDENTIFIED YELLOW LEVEL ECG ALARMS AND SPO2 ALARMS. THE RCS PROVIDED INFORMATION TO THE CUSTOMER ON HOW TO SEARCH THE PATIENT DATA BY THE PATIENT MEDICAL RECORD NUMBER. THE CUSTOMER INDICATED THERE WAS NO PRODUCT MALFUNCTION AND DECLINED ONSITE SERVICE. THE LOGS WERE PROVIDED TO A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) FOR TECHNICAL INVESTIGATION. THE PSE VERIFIED THAT ON (B)(6) 2024, ¿ECG LEADS OFF¿ (CYAN INOP) STARTED AT 05:06:32 AND THEN IT ESCALATED TO ¿!!ECG LEADS OFF¿ (YELLOW INOP) AT 05:06:36. THIS ESCALATED ¿!!ECG LEADS OFF¿ (YELLOW INOP) ENDED AT 06:16:50. THE REPORTED PROBLEM WAS NOT CONFIRMED. INFORMATION WAS PROVIDED TO THE CUSTOMER TO RESOLVE THE ISSUE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION CANCER, DEATH, LUNG CANCER. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF LIVER DISEASE/TOXICITY, CANCER AND DEATH "SQUAMOUS ESOPHAGEAL CANCER, STAGE 4B". THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION REGARDING THYROID CANCER AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Nausea; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING SKIN IRRITATION, NAUSEA / VOMITING, CANCER, LIVER & KIDNEY RELATED ISSUE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Headache; Appropriate Term / Code Not Available; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO CANCER, HEADACHE, DIFFICULTY BREATHING. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Respiratory Tract Infection; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED RESPIRATORY TRACT IRRITATION, LUNG DISEASE, LUNG INFECTION, EXACERBATION LEADING TO DEATH. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SHORTNESS OF BREATH, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Respiratory Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO RESPIRATORY FAILURE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF CANCER, DEATH AND ESOPHAGEAL CANCER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Headache; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS AND/OR HEADACHE; CANCER; DEATH, BRAIN CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Breast Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, METASTATIC BREAST CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION REGARDING RESPIRATORY FAILURE AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Appropriate Term / Code Not Available; Dyspnea; Stroke/CVA | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BREATHING ISSUES, BLOOD CLOTS, AND STROKE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. PATIENT CALLED STATING THAT HE WAS CALLING ABOUT THE PATIENT NAMED CATHERINE. WANTS ALL COMMUNICATION STOPPED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF KIDNEY CANCER AND PATIENT PASSING AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING MULTIPLE MYELOMA. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION CPAP PRO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING BLADDER CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS COMPLETED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, CANCER, LIVER CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Dyspnea | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED METASTATIC BREATHING ISSUES. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Vomiting; Dizziness; Nausea; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, THROAT AND BLADDER CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER SERVICE CENTER FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED BY THE SERVICE TECHNICIAN. EVALUATION RESULT STATED THAT THE COMPLAINT WAS NOT CONFIRMED AND FOUND NO EVIDENCE OF VISIBLE FOAM DEGRADATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER SERVICE CENTER FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED BY THE SERVICE TECHNICIAN. EVALUATION RESULT STATED THAT THE COMPLAINT WAS NOT CONFIRMED AND FOUND NO EVIDENCE OF VISIBLE FOAM DEGRADATION. THE CORRECT EVENT DESCRIPTION SHOULD BE - THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED BY THE SERVICE TECHNICIAN. EVALUATION RESULT STATED THAT THE DEVICE PASSED ALL TESTS. THE COMPLAINT WAS ORIGINALLY REPORTED UNDER THE RECALL, BUT AFTER FURTHER REVIEW IT WAS DETERMINED THIS COMPLAINT SHOULD NOT HAVE BEEN FILED UNDER THE RECALL. THEREFORE, BOX D, G AND H HAS BEEN UPDATED OR CORRECTED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER AND PROSTATE CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Unspecified Heart Problem; Nausea; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, KIDNEY DISEASE/TOXICITY AND REDUCED CARDIOPULMONARY RESERVE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALSO ALLEGED OF HAVING CANCER, LUNG ADENOCARCINOMA. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Dizziness; Inflammation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND METASTATIC SQUAMOUS CELL CARCINOMA. THE PATIENT HAS BEEN PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, BRAIN CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF LUNG DISEASE, CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO CANCER STAGE IV LUNG DISEASE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Hypersensitivity/Allergic reaction; Appropriate Term / Code Not Available; Unspecified Heart Problem; Dizziness; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Respiratory Failure; Unspecified Respiratory Problem; Cancer; Headache; Heart Failure/Congestive Heart Failure; Chronic Obstructive Pulmonary Disease (COPD) | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS ,HEADACHE,HYPERSENSITIVITY,KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER DEATH, INJURIES INCLUDE LUNG DAMAGE, COPD, HEART DISEASE, RESPIRATORY FAILURE, AND CONGESTIVE HEART FAILURE. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND CANCER SARCOMA (SOFT TISSUE CANCER), SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING AND INFLAMMATORY RESPONSE. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Respiratory Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND RESPIRATORY FAILURE. THE PATIENT HAS BEEN PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND UNSPECIFIED LUNG ISSUES. THE PATIENT HAS BEEN PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, SMALL CELL LUNG CANCER, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Liver Damage/Dysfunction; Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, METASTATIC LIVER. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Vomiting; Eye Irritation; Dizziness; Nausea; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Dizziness; Nausea; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA AND VOMITING, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Pneumonia; Asthma; Unspecified Heart Problem; Appropriate Term / Code Not Available; Inflammation; Renal Failure; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Heart Failure/Congestive Heart Failure; Fibrosis | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, SKIN CANCER, DEATH, LUNG, INFLAMMATION, PRE-DIABETIC, PULMONARY FIBROSIS, PNEUMONIA, STROKE, HEART FAILURE, QUADRUPLE BYPASS, CHRONIC KIDNEY FAILURE, STAGE 4; ACUTE RENAL FAILURE NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO KIDNEY DISEASE/TOXICITY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Chronic Obstructive Pulmonary Disease (COPD); Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, LUNG CANCER, KIDNEY CANCER, COPD. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER,DEATH AND LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Headache; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, CANCER, AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, CANCER, AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT GOT CORRECTED. SECTION H6 UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ANAPLASTIC THYROID CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE AND CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Transient Ischemic Attack; Pulmonary Dysfunction; Pleural Effusion | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DAMAGE, HEART ATTACK, PLEURAL EFFUSION AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Skin Inflammation/ Irritation; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER, MULTIPLE MYELOMA AND PLASMA CELL LEUKEMIA. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem; Chronic Obstructive Pulmonary Disease (COPD) | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF KIDNEY DISEASE/TOXICITY AND COPD. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Fibrosis | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PULMONARY FIBROSIS, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Appropriate Term / Code Not Available; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE, LIVER DISEASE, TOXICITY, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, AND MOUTH/LYMPH NODE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ALARIS SYSTEM | K133532 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS INITIALLY REPORTED A REQUEST FOR LOG REVIEW ON THE PCA MODULE TO DETERMINE THE DOSES ADMINISTERED TO THE PATIENT FOR PAIN CONTROL. PER REPORT, THE PATIENT WAS IN POST-OP FOR KNEE SURGERY. THE CUSTOMER STATES THERE WAS NO DEVICE MALFUNCTION SUSPECTED. IT WAS REPORTED A "DEATH" OCCURRED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT THE PCA MODULE WAS USED FOR MORPHINE INFUSION. THE MEDICATION WAS CONTAINED IN TERUMO 60ML SYRINGE. THE INFUSION WAS PROGRAMMED AS PCA DOSE WITH CONTINUOUS INFUSION. PER FACILITY'S BIOMEDICAL ENGINEER, IT IS NOT BELIEVED THAT THE DEVICE CAUSED OR CONTRIBUTED TO PATIENT'S DEATH. PER REPORT, THE PRIMARY CAUSE OF DEATH WAS "CARDIAC ARREST POST KNEE SURGERY." THE PATIENT DIED ON (B)(6) 2024. ALTHOUGH ASKED, THE CUSTOMER DID NOT PROVIDE THE SPECIFIC REASON FOR WHY THE FACILITY WISHED TO REVIEW THE PCA DEVICE EVENT LOGS OTHER THAN THE CUSTOMER REQUESTED TO HAVE THE INFUSION EVENT DETAILS TO REVIEW INTERNALLY. THE CUSTOMER INDICATED THAT THEY ARE REQUIRED (BY THEIR FACILITY) TO CONDUCT EVENT LOG REVIEW IMMEDIATELY "DUE TO THE PATIENT BEING HIGH PROFILE." THE EVENT OCCURRED IN UNITED ARAB EMIRATES. MULTIPLE COMMUNICATIONS HAVE BEEN MADE BY BD REQUESTING FOR ADDITIONAL INFORMATION. THE CUSTOMER REFUSED TO PROVIDE FURTHER INFORMATION CITING CONFIDENTIALITY. | OMIT : C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. CORRECTION: MEDICAL DEVICE TYPE, MEDICAL DEVICE CATALOG #, UNIQUE IDENTIFIER (UDI) #, MEDICAL DEVICE SERIAL #, MEDICAL DEVICE MODEL #, CONCOMITANT MED PROD DATA, PMA / 510(K)#, DEVICE MANUFACTURE DATE. ADDITIONAL INFORMATION: IMDRF ANNEX B, C, D CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: DURING REVIEW OF THE DEVICE LOG FOR THIS INVESTIGATION NO ERRORS OR MALFUNCTIONS WERE IDENTIFIED, AND NO ISSUES OR ANOMALIES WERE OBSERVED DURING THE KEY PRESS REPLICATION PERFORMED ON LABORATORY EQUIPMENT. ¿ NO DEVICES WERE RETURNED FOR LABORATORY TESTING; THEREFORE, LOG REVIEW WAS PERFORMED USING THE PCA AND PCU LOG EVENTS PROVIDED BY THE CUSTOMER. ¿ A LOG REVIEW WAS PERFORMED FOR THE REQUESTED TIME FRAME FROM OF (B)(6) 2024 TO (B)(6) 2024. NO ERRORS OR MALFUNCTIONS WERE IDENTIFIED, AND NO ISSUES OR ANOMALIES WERE OBSERVED DURING THE KEY PRESS REPLICATION PERFORMED ON LABORATORY EQUIPMENT ¿ THERE WERE TWO CONFIRMED MORPHINE PROGRAMMING ACTIVITIES, 1ST ON (B)(6) 2024 AT 4:34:49 PM AND 2ND (B)(6) 2024 AT 2:45:26 PM, BOTH PROGRAMS HAD PARAMETERS SYRINGE SELECTION: TERUMO 50; INFUSION MODE PCA AND CONTINUOUS; CONTINUOUS RATE:1ML/H; DOSE:2MG/H, DOSE CONCENTRATION 2MG PROGRAMMED; LOCKOUT INTERVAL: 10 MINUTES. ¿ FROM THE EVENT LOG, THE INFUSION PROGRAMMED WAS MORPHINE WITH A CONCENTRATION OF 100MG/50ML AND A PCA DOSE OF 2 MG WITH A LOCKOUT PERIOD OF 10 MINUTES WAS CONFIRMED, NO MAX LIMIT WAS PROGRAMMED. ¿ THE EVENT LOG SHOWS A TOTAL ESTIMATED INFUSED VOLUME OF 96.64 ML FROM (B)(6) 2024 TO (B)(6) 2024, ¿ THERE WERE 293 UNSUCCESSFUL ATTEMPTS TO REQUEST A DOSE DURING LOCKOUT PERIOD WITHIN THE TIME FRAME REVIEWED. ROOT CAUSE: NO MALFUNCTION WAS ALLEGED ON THE DEVICE INVOLVED IN THE COMPLAINT; THE CUSTOMER REQUESTED AN EVENT LOG REVIEW.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Appropriate Term / Code Not Available; Nausea; Respiratory Tract Infection; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NAUSEA, VOMITING, CANCER, ESOPHAGEAL CANCER, CHEST PRESSURE, AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Eye Irritation; Unspecified Heart Problem; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF EYE IRRITATION, SKIN IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING METASTATIC PANCREATIC CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PATIENT INFORMATION CENTER IX | K183387 | Insufficient Information | No Audible Alarm | IT WAS REPORTED THE PATIENT ON THE MEDICAL TELEMETRY UNIT WAS FOUND UNRESPONSIVE WHILE THE TELEMETRY DEVICE WAS IN STANDBY MODE. THE PATIENT HAD BEEN OFF THE MONITOR WITH MX40 IN STANDBY MODE, WHICH WAS INITIATED BY THE TELEMETRY TECHNICIAN, FOR SEVERAL HOURS. THE SECOND TIME THE MX40 WAS PLACED IN STANDBY MODE, IT REMAINS UNKNOWN WHAT USER INITIATED THIS AS THE TECHNICIAN DID NOT REMEMBER PLACING THE MX40 IN STANDBY MODE THE SECOND TIME. THE PATIENT WAS FOUND UNRESPONSIVE, TRANSFERRED TO ICU, CODED, AND LATER EXPIRED. THERE WAS NO DEVICE MALFUNCTION FOUND UPON CUSTOMER INSPECTION. BASED ON THIS INFORMATION, THE LACK OF MONITORING DUE TO THE DEVICE BEING PLACED IN STANDBY MODE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT; HOWEVER, IT CANNOT BE DETERMINED HOW THE MX40 WAS PLACED IN STANDBY AND IT REMAINS UNKNOWN WHETHER THERE WAS A DEVICE MALFUNCTION, CLINICAL WORKFLOW OR USE ERROR THAT CONTRIBUTED. IF ADDITIONAL INFORMATION IS OBTAINED, PLEASE REQUEST RE-ASSESSMENT OF THE RECORD BY THE PMS CLINICAL EXPERT. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED
DREAMSTATION AUTO CPAP | K131982 | Cancer; Cardiac Arrest | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SMALL LUNG CANCER, HEART ATTACKS. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | Heart Failure/Congestive Heart Failure; Unspecified Kidney or Urinary Problem; Respiratory Failure; Pleural Effusion | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, HEART FAILURE, PLEURAL EFFUSION, AND RESPIRATORY FAILURE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, HEART FAILURE, PLEURAL EFFUSION, AND RESPIRATORY FAILURE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE. PIL INITIATED THERAPY WITH THE DEVICE AND VERIFIED AIRFLOW, USING A KNOWN GOOD PIL SUPPLIED POWER SUPPLY AND AC POWER CORD. PIL ALSO OBSERVED CUSTOMERS HUMIDIFIER HEATER PLATE DID HEAT. PIL OBSERVED THE FOLLOWING SOUND ABATEMENT DEGRADED FOAM INDICATIONS: BLACK SUBSTANCE, CONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION, WAS OBSERVED ON THE OUTSIDE BOTTOM OF THE BLOWER BOX AND WAS STUCK TO IT. TINY BLACK PIECES OF A BLACK SUBSTANCE, CONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION, WERE FOUND ON THE BOTTOM OF THE ENCLOSURE. TINY BLACK PIECES OF A BLACK SUBSTANCE, CONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION, WERE FOUND IN THE HUMIDIFIER WATER TANK. DURING THE EVALUATION SECONDARY FINDINGS (32 ERRORS WERE LOGGED. DUST-LIKE CONTAMINATION AT THE AIR INLET, ON THE PCA, IN THE BLOWER BOX AND THROUGHOUT THE INSIDE ENCLOSURE. DUST-LIKE CONTAMINATION INSIDE HUMIDIFIER ENCLOSURE, ON AND UNDER THE HEATER PLATE, AND ON THE DRY BOX. POTENTIAL FLUID SPOTS IN BLOWER BOX AND BLOWER MOTOR INDICATING POSSIBLE WATER INGRESS. PIL CAN CONFIRM THE PRESENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM FROM THE SECONDARY FINDINGS. PIL WAS UNABLE TO ADDRESS THE SYMPTOMS SPECIFIED) WERE OBSERVED. THE ISSUE OF DEGRADED SOUND ABATEMENT FOAM IS ADDRESSED BY CAPA 7211. IN BOX D: DEVICE AVAILABLE FOR EVAL? AND DATE RETURNED TO MFG HAS BEEN UPDATED. IN BOX H: DEVICE EVAL. BY MFG? (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
DAVINCI XI | K131861 | Great Vessel Perforation; Hemorrhage/Blood Loss/Bleeding | Appropriate Term/Code Not Available | IT WAS REPORTED THAT DURING A DA VINCI ASSISTED APICO-DORSAL LEFT LUNG SEGMENTECTOMY, A BRANCH OF THE PULMONARY ARTERY (PA) WAS INJURED; MASSIVE BLEEDING OCCURRED RESULTING IN AN EMERGENT CONVERSION TO AN OPEN PROCEDURE AND THE PATIENT ULTIMATELY EXPIRED. THE INTUITIVE CLINICAL TERRITORY ASSOCIATE WAS PRESENT AT THE PROCEDURE AND REPORTED THAT DURING DISSECTION AROUND THE PA THE VESSEL RUPTURED AND BLOOD IMMEDIATELY WAS PROJECTED ONTO THE ENDOSCOPE, OBSTRUCTING VISION. THE ENDOSCOPE WAS REMOVED FOR CLEANING OF THE TIP TO RESTORE VISION AND ATTEMPTS TO REINSTALL THE ENDOSCOPE WERE UNSUCCESSFUL, LEADING TO AN EMERGENT CONVERSION TO OPEN THORACOTOMY. THE ESTIMATED BLOOD LOSS WAS MORE THAN 3L AND THE PATIENT WAS IN HEMORRHAGIC SHOCK. THE PATIENT WENT INTO CARDIAC ARREST, INTERNAL CARDIAC MASSAGE WAS PERFORMED, AND THE PA WAS CLAMPED. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND THE SEGMENTECTOMY PROCEDURE WAS COMPLETED OPEN. THE PATIENT EXPIRED ON POSTOPERATIVE DAY 10. | A VIDEO OF THE EVENT WAS PROVIDED BY THE CUSTOMER AND REVIEW SHOWED THE SURGEON WAS DISSECTING AND TRYING TO GO AROUND A BRANCH OF THE PA WHEN THE INJURED AREA STARTED TO OOZE. THE SURGICAL INSTRUMENT WAS MOVED AWAY FROM THE PA AND LATER REPOSITIONED IN SUCH A WAY THAT IT APPEARS TO PUSH DOWN ON THE MAIN BRANCH OF THE ARTERY, CAUSING THE MAJOR BLEEDING EVENT. THE ENDOSCOPE WAS REMOVED, CLEANED, AND REINSTALLED. THE ROBOT UNIVERSAL SURGICAL MANIPULATOR (USM 2), WHICH WAS HOUSING THE ENDOSCOPE, DISPLAYED A USER FACING MESSAGE SAYING, ¿ADVANCE ENDOSCOPE TO EARLIER POSITION¿. HOWEVER, THE BLEEDING FROM THE PA WAS OBSCURING THE VIEW OF THE SURGICAL WORKSPACE. THE ENDOSCOPE WAS REMOVED AGAIN AND REINSTALLED BACK ONTO USM 2 WHEN ¿ADVANCE ENDOSCOPE TO EARLIER POSITION¿ AND ¿DISENGAGEMENT OF THE ARM AT THE PSM¿ MESSAGING APPEARS. A SUCTION INSTRUMENT IS SEEN TRYING TO SUCTION THE BLOOD FROM THE SITE, FOLLOWED BY REMOVAL OF THE OTHER INSTRUMENTS AND ACCESSORIES, THEN REMOVAL OF THE SUCTION AND THE ENDOSCOPE. THE ACTIONS NOTED ABOVE OCCURRED OVER THE COURSE OF APPROXIMATELY TWO MINUTES. REVIEW OF THE DVSTREAM LOG AND KINEMATIC DATA SHOWED ENDOSCOPE REMOVAL, FOLLOWED BY RE-INSTALLATION. APPROXIMATELY 8 SECONDS LATER THE ENDOSCOPE WAS REMOVED AND REINSTALLED AGAIN, FOLLOWED BY A SHORT PERIOD OF ENDOSCOPE CONTROL ACTIVATION, THEN ENDOSCOPE REMOVAL WITHIN APPROXIMATELY 10 SECONDS. REVIEW OF THE SYSTEM LOGS FOUND THERE WERE NO ERRORS RELATED TO THE REPORTED PULMONARY ARTERY INJURY EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED FOUND THERE WERE NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THIS EVENT. VIDEO REVIEW COINCIDES WITH AN ENDOSCOPE THAT LOST ITS CONNECTION TO THE STERILE ADAPTER, PROMPTING THE USER WITH THE ¿ARM NOT READY. REMOVE THE ENDOSCOPE TO CONTINUE¿ MESSAGE. THIS PROMPT WOULD NOT COINCIDE WITH AN ERROR. THE FOLLOWING CONCOMITANT DEVICES WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, TIP-UP FENESTRATED GRASPER, MARYLAND BIPOLAR FORCEPS, FENESTRATED BIPOLAR FORCEPS, SUREFORM 45 CURVED TIP STAPLER. A SITE HISTORY REVIEW SHOWED THAT WITH THE EXCEPTION OF THE SINGLE-USE STAPLER, THE OTHER INSTRUMENTS AND THE ENDOSCOPE WERE ALL USED IN SUBSEQUENT PROCEDURES WITH NO REPORTED COMPLAINTS. REVIEW OF THE EVENT BY AN INTUITIVE MEDICAL SAFETY OFFICER CONCLUDED THAT THE PATIENT WAS UNDERGOING A PULMONARY SEGMENTECTOMY WHEN THERE WAS AN INJURY TO THE PULMONARY ARTERY. THE MOMENT THE PULMONARY ARTERY RUPTURED WAS CAPTURED ON VIDEO WHICH OCCURRED DURING DISSECTION OF THIS VESSEL AND THE SEGMENTAL BRANCHES. THIS RESULTED IN SPLASHING THE CAMERA WHICH REQUIRED IMMEDIATE CLEANING OF THE LENS. THE SURGICAL TEAM SUBSEQUENTLY HAD DIFFICULTY RE-INSTALLING THE CAMERA AFTER CLEANING THE LENS AS EVIDENCED BY THE LOG INFORMATION. THE PATIENT WAS CONVERTED TO AN OPEN THORACOTOMY AND LATER REQUIRED ECMO. THE PATIENT EVENTUALLY DIED 10 DAYS LATER. THE ROOT CAUSE OF THE DEATH IS RELATED TO THE BLEEDING FROM THE PULMONARY ARTERY. ALTHOUGH THIS OCCURRED DURING DISSECTION OF THE VESSEL, HOW THE VESSEL WAS INITIALLY WEAKENED AND SUBSEQUENTLY RUPTURED IS UNKNOWN. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION A2 AGE: THE PATIENT AGE WAS REPORTED AS "SIXTIES".
DAVINCI XI | K131861 | Ischemic Heart Disease; Insufficient Information; Bronchopneumonia; Chronic Obstructive Pulmonary Disease (COPD) | Adverse Event Without Identified Device or Use Problem | A PATIENT ENROLLED IN A POST-MARKET STUDY UNDERWENT A DA VINCI ASSISTED PULMONARY SEGMENTECTOMY FOR LEFT LOWER LOBE LUNG CANCER. THE MEDICAL RECORD DOCUMENTED THAT ADHESIONS WERE ENCOUNTERED DURING SURGERY. A CHEST TUBE DRAIN WAS PLACED POST-SURGERY AS STANDARD PROCEDURE AND WAS REMOVED ON POST-OPERATIVE DAY (POD) 4. THE PATIENT WAS DISCHARGED HOME THE SAME DAY WITHOUT COMPLICATIONS. THERE WAS NO REPORT OF ANY DA VINCI DEVICE MALFUNCTION DURING THE PROCEDURE. DURING THE 30-DAY POST-OPERATIVE STUDY VISIT, IT WAS FOUND THAT THE PATIENT HAD PASSED AWAY ON POD 5, THE DAY AFTER DISCHARGE FROM THE HOSPITAL. THE DETAILS OF THE PATIENT¿S COURSE AFTER DISCHARGE AND THE CAUSE OF DEATH WERE NOT PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED BUT HAS NOT BEEN MADE AVAILABLE. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD LIKELY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS RELATED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, FENESTRATED BIPOLAR FORCEPS, PERMANENT CAUTERY HOOK, TWO CADIERE FORCEPS, SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT. ALL MULTI-USE INSTRUMENTS HAVE BEEN USED IN SUBSEQUENT PROCEDURES. THE SUREFORM 45 CURVED-TIP STAPLER LOGS SHOWED THE INSTRUMENT FIRED 7 RELOADS. ALL FIRINGS WERE COMPLETED SUCCESSFULLY WITHOUT ISSUES. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. THE PATIENT IN THIS REPORT HAD MULTIPLE COMORBIDITIES AND UNDERWENT A PULMONARY SEGMENTECTOMY. THE PATIENT DIED 5 DAYS AFTER SURGERY, 1 DAY AFTER BEING DISCHARGED FROM THE HOSPITAL. NO COMPLICATIONS NOR ISSUES WITH ANY INTUITIVE SURGICAL PRODUCTS WERE REPORTED AND NO CAUSE OF DEATH IS PROVIDED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BRAIN CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS ZERO ERROR CODE FOUND AND DURING THE EVALUATION NO SECONDARY FINDINGS WAS OBSERVED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED. BOX G: MANUFACTURE DATE (RFB) UPDATED. BOX E WAS UPDATED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | ON 09/SEP/0224 IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6) 2024. IT IS UNKNOWN IF THE PATIENT HAD COMPLETED TREATMENT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE UNSUCCESSFUL. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN USE OF THE LIBERTY SELECT CYCLER AND THE PATIENT DEATH. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THE EVENT. IT IS UNKNOWN IF THE PATIENT WAS FOUND DECEASED OR UNRESPONSIVE OR IF THE TREATMENT WAS COMPLETED AT THE TIME OF DEATH. THE MORTALITY RATES FOR PATIENTS WITH END-STAGE RENAL DISEASE ARE SIGNIFICANTLY HIGHER THAN THOSE WITHOUT THE DISEASE. CARDIOVASCULAR DISEASE ACCOUNTS FOR APPROXIMATELY 50% OF THOSE DEATHS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S), PLEASE RE-SUBMIT FOR A CLINICAL REVIEW AND THE NEED FOR A CLINICAL INVESTIGATION WILL BE REASSESSED ACCORDINGLY.
DAVINCI XI | K131861 | Insufficient Information; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING A DA VINCI- CHOLECYSTECTOMY, THE PATIENT EXPERIENCED CARDIAC ARREST AND EXPIRED. THE INTUITIVE CLINICAL SALES REPRESENTATIVE (CSR) WAS INFORMED OF THE PATIENT EVENT AND TOLD THAT THE DETAILS ARE CONFIDENTIAL; THEREFORE, WAS NOT PROVIDED ANY SPECIFIC INFORMATION. THE CSR SPOKE WITH THE HEAD OF THE DEPARTMENT AND THE MEDICAL DIRECTOR OF THE HOSPITAL WHO STATED THAT THE PATIENT PROBABLY SUFFERED A HEART ATTACK AND DIED AFTER ATTEMPTED RESUSCITATION. THEY CONDUCTED THEIR OWN INTERNAL INVESTIGATION, AN AUTOPSY WAS CONDUCTED, AND STATED THAT NO FURTHER DETAILS WOULD BE DISCLOSED. THEY REPORTED THAT THE DA VINCI ROBOT HAD NO INFLUENCE ON THE INCIDENT, THERE IS NO CONNECTION BETWEEN DA VINCI AND THE INCIDENT, AND IT IS NOT INTUITIVE¿S CONCERN. | NO SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE WAS PROVIDED; THEREFORE, NO DAVINCI SYSTEM OR INSTRUMENT LOG FILES ARE AVAILABLE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DA VINCI SYSTEM USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED DURING A CHOLECYSTECTOMY REPORTEDLY FROM A MYOCARDIAL INFARCTION. HOW THAT DETERMINATION WAS MADE IS NOT PROVIDED. NO DEVICE RELATED ISSUES DURING THE CASE WERE REPORTED. THERE IS NO ALLEGATION OF ANY ISSUE WITH ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B3 DATE OF EVENT: EVENT DATE IS ESTIMATED AS THE INFORMATION PROVIDED ON 20-AUG-2024 WAS THAT THE EVENT OCCURRED "ABOUT 2 MONTHS AGO IN JUNE".
DREAMSTATION AUTO CPAP | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE TYPE OF REPORTED COMPLAINT WAS PREVIOUSLY MARKED INCORRECT AS SERIOUS INJURY. IN THIS REPORT IT HAS BEEN UPDATED AND CORRECTED AS DEATH.
REMSTAR PRO C-FLEX+ | K091319 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE REPORTER ALLEGED OF HAVING CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT PASSED AWAY AND THE DEVICE WAS GOING TO BE RETURNED AS IT WAS NO LONGER IN USE BY THE CUSTOMER. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT OUTCOME, HOWEVER LIMITED INFORMATION HAS BEEN PROVIDED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A PATIENT PASSED AWAY AND THE DEVICE WAS GOING TO BE RETURNED AS IT WAS NO LONGER IN USE BY THE CUSTOMER. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT OUTCOME, HOWEVER LIMITED INFORMATION HAS BEEN PROVIDED AT THIS TIME. THE DEVICE WAS RETURNED AND EVALUATED BY A THIRD-PARTY SERVICE CENTER. THE DEVICE WAS FOUND TO PASS ALL DEVICE TESTING. NO ABNORMALITIES WERE FOUND. THE DEVICE PERFORMED AS DESIGNED AND AS INTENDED. THE FOLLOWING CODES HAVE BEEN UPDATED AS FOLLOWS: EVALUATION METHOD CODE - TESTING OF ACTUAL/SUSPECTED DEVICE. EVALUATION RESULTS CODE - ANALYSIS OF INFORMATION PROVIDED BY USER/THIRD PARTY. COMPONENT GRID - NO INFORMATION. CONCLUSION CODE - NO PROBLEM DETECTED.
DREAMSTATION AUTO CPAP | K131982 | Pleural Effusion | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND PLEURAL EFFUSION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Skin Inflammation/ Irritation; Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION RESPIRATORY TRACT IRRITATION DIZZINESS AND/OR HEADACHE KIDNEY DISEASE/TOXICITY LUNG DISEASE DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | ON 18/SEP/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2024. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2024. THE PATIENT¿S DAUGHTER DISCOVERED THE PATIENT HAD EXPIRED SEVERAL HOURS PRIOR (BASED ON RIGAMORTUS) WHEN SHE ARRIVED TO DISCONNECT THE PATIENT FROM THE LIBERTY SELECT CYCLER (AS SHE DID EVERY MORNING). THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES, AND THE PATIENT WAS TAKEN DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS ¿NATURAL CAUSES¿ BUT WOULD NOT ELABORATE ANY FURTHER. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE EVENTS OCCURRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS DUE TO NATURAL CAUSES AND WAS UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Fibrosis | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PULMONARY FIBROSIS, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED ABOUT THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ABOUT AN ALLEGATION OF CANCER AND DEATH. THERE WAS NO REPORT OF SPECIFIC MEDICAL INTERVENTIONS. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE CARRIED OUT. IF ANY ADDITIONAL INFORMATION IS RECEIVED A FOLLOW-UP REPORT WILL BE FILED. | CORRECTION TO PREVIOUS INITIAL SUBMISSION DUE DATE: ACTUAL DUE DATE IS 11NOV2023
DREAMSTATION ST30 | K102465 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR DEVICE EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION AND NO OTHERS CONTAMINATION FOUND. 0 ERROR LOG. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REP DREAMSTATION AUTO BIPAP. THE REPORTER ALLEGED THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K183387 | Cardiac Arrest | Defective Alarm | IT WAS REPORTED THE PATIENT WENT INTO CARDIAC ARREST AND PASSED AWAY; HOWEVER, THE DEVICE DID NOT GENERATE AN ALARM. THE PATIENT WAS BEING MONITORED WITH TELEMETRY MX40 WHEN THE PATIENT DEVELOPED ASYSTOLE, FOR WHICH NO ALARM WAS GENERATED. THE NURSING STAFF WAS ALERTED TO THE CHANGE IN CONDITION BY THE PATIENT'S ROOMMATE. RESUSCITATION WAS ATTEMPTED BUT WAS UNSUCCESSFUL AND THE PATIENT SUBSEQUENTLY PASSED AWAY. GOOD FAITH EFFORTS ARE ONGOING. | A REVIEW OF THE DEVICE LOGS REVEALED THE DEVICE WAS WORKING AS DESIGNED. THE MX40 WORKED CORRECTLY AND SENT ALL INFORMATION TO THE CENTRAL. YELLOW AND RED ALARMS WERE GENERATED AT THE CENTRAL AND AUDIBLE TONES WERE GENERATED AS WELL. IN ADDITION, IT WAS NOTED THE ALARMS WERE ALSO SILENCED AT THE CENTRAL STATION. IT IS REITERATED THE NURSING STAFF DID INDICATE ALARMS WERE COMING THROUGH TO THEIR PHONES BUT THEY COULD NOT CONFIRM WHICH ALARMS. BASED ON THIS INFORMATION, IT REMAINS UNKNOWN WHETHER DEVICE BEHAVIOR, USE ERROR, ALARM MANAGEMENT, OR CLINICAL WORKFLOW MAY HAVE BEEN FACTORS IN THE EVENT. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND CHRONIC MYELOID LEUKEMIA. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT EXPIRED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND CHRONIC MYELOID LEUKEMIA. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT EXPIRED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX D AND H HAS BEEN UPDATED.
REMSTAR PRO C-FLEX+ | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Appropriate Term / Code Not Available; Dizziness; Nausea; Lymphoma; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER, DEATH, LARGE B-CELL LYMPHOMA WITH HIGH GRADE FEATURES INVOLVING THE STOMACH, PROSTATE CANCER, AND NASAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON THE CYCLER. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT (MALE) EXPIRED PEACEFULLY AT HOME ON (B)(6) 2024 WHILE SLEEPING. THE PATIENT¿S SPOUSE DISCOVERED HER HUSBAND HAD EXPIRED SOMETIME DURING THE NIGHT. THE PATIENT¿S SPOUSE DECLINED HAVING AN AUTOPSY PERFORMED, AND THE PATIENT WAS TAKEN DIRECTLY TO THE FUNERAL HOME. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 SHOWED THE TREATMENT WAS COMPLETED DESPITE THE PATIENT¿S PASSING, AND NO ALARMS WERE NOTED. THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS A CARDIAC ARREST, SECONDARY TO MULTIPLE CARDIAC COMORBID CONDITIONS. THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER HAS NOT BEEN SCHEDULED FOR PICK-UP. THE PDRN CONFIRMED THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN HE EXPIRED; HOWEVER, THE PDRN REPORTED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). THE PDRN WAS UNABLE TO PROVIDE ANY PATIENT DEMOGRAPHICS, MEDICAL HISTORY, OR CCPD ORDERS DUE TO THE PATIENT¿S ELECTRONIC MEDICAL RECORD BEING CLOSED. ADDITIONALLY, THE PATIENT TRANSFERRED FROM AN UNKNOWN OUTPATIENT HOME DIALYSIS TO CLINIC TO (B)(6) ONLY A FEW DAYS PRIOR TO HIS EXPIRATION. | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE PATIENT¿S SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. PER THE PATIENT¿S PDRN, THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT¿S MULTIPLE CARDIAC COMORBID CONDITIONS. THE PDRN CONFIRMED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ION | K182188 | Pulmonary Embolism; Respiratory Failure | Adverse Event Without Identified Device or Use Problem | A PATIENT ENROLLED IN A REGISTRY STUDY UNDERWENT AN UNEVENTFUL ION BIOPSY PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS OR DEVICE MALFUNCTIONS. APPROXIMATELY TWO WEEKS AFTER THE PROCEDURE, THE PATIENT EXPERIENCED RESPIRATORY FAILURE THAT REQUIRED HOSPITALIZATION IN THE INTENSIVE CARE UNIT, AND PASSED AWAY 5 DAYS LATER (21 DAYS POST-PROCEDURE). THE CAUSE OF DEATH WAS DETERMINED BY THE PRINCIPLE INVESTIGATOR TO BE RESPIRATORY FAILURE FROM ACUTE PULMONARY EMBOLISM RELATED TO THE COMORBID CONDITION OF NON-SMALL CELL LUNG CANCER. THE EVENT WAS ASSESSED AS DEFINITELY RELATED TO THE PATIENT'S PRE-EXISTING CONDITION, BUT NOT RELATED TO THE ION DEVICES OR THE PROCEDURE. | REVIEW OF THE ION SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE AVAILABLE DATA IT IS POSSIBLE THE PROCEDURE MAY HAVE BEEN A CONTRIBUTING FACTOR TO THE VENOUS THROMBOEMBOLISM (VTE) EVENT AND THEREFORE THE PULMONARY EMBOLISM 2 WEEKS POST BIOPSY MAY HAVE BEEN PROCEDURE RELATED. THERE IS NO COMPELLING EVIDENCE THE EVENT WAS DEVICE RELATED. THE CAUSE OF DEATH WAS REPORTED TO BE DUE TO ACUTE RESPIRATORY FAILURE DUE TO PULMONARY EMBOLISM IN THE SETTING OF LUNG CANCER. THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES ASSOCIATED WITH THE REPORTED COMPLICATION. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 CASES REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. THERE WERE NO REPORTED EVENTS OF VENOUS THROMBOEMBOLISM (VTE) INCLUDING DEEP VENOUS THROMBOSIS OR PULMONARY EMBOLISM. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS AND NO VTE. HOWEVER, IT IS POSSIBLE THAT IMMOBILITY DURING GENERAL ANESTHESIA AND THE SMALL AMOUNT OF BLEEDING AND SUBSEQUENT CLOTTING ASSOCIATED WITH LUNG BIOPSIES MAY CONFER SOME DEGREE OF INCREASED RISK OF VTE EVENTS. LIKELY OF MORE RELEVANCE, IT IS ESTIMATED THAT CANCER CONFERS A 4-7X HIGHER RISK FOR VTE. THE HIGHER RATE OF VTE HAS BEEN REPORTED TO RANGE BETWEEN 3-13.9% IN PATIENTS WITH LUNG CANCER. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLOM JW, DOGGEN CJ, OSANTO S, ROSENDAAL FR. MALIGNANCIES, PROTHROMBOTIC MUTATIONS, AND THE RISK OF VENOUS THROMBOSIS. JAMA. 2005. BAGCHI A, KHAN MS, SARASWAT A, ANSARI A, NAI Q, IYER V, HAMOUDA D, KHUDER S, VERGHESE C. INCREASED INCIDENCE OF THROMBOTIC COMPLICATIONS WITH NON-SMALL CELL LUNG CANCER NECESSITATES CONSIDERATION OF PROPHYLACTIC ANTICOAGULATION IN YOUNG INDIVIDUALS. CUREUS. 2021. CHEW HK, DAVIES AM, WUN T, HARVEY D, ZHOU H, WHITE RH. THE INCIDENCE OF VENOUS THROMBOEMBOLISM AMONG PATIENTS WITH PRIMARY LUNG CANCER. J THROMB HAEMOST. 2008. CONNOLLY GC, DALAL M, LIN J, KHORANA AA. INCIDENCE AND PREDICTORS OF VENOUS THROMBOEMBOLISM (VTE) AMONG AMBULATORY PATIENTS WITH LUNG CANCER. LUNG CANCER. 2012.
ALARIS SYSTEM | K133532 | High Blood Pressure/ Hypertension; Intracranial Hemorrhage; Hematemesis | Free or Unrestricted Flow | IT WAS INITIALLY REPORTED A FREE FLOW EVENT. NO FURTHER INFORMATION WAS PROVIDED. RECEIVED A COPY OF THE CUSTOMER'S MAUDE DATABASE REPORT FROM FDA WHICH STATES: "THE PATIENT WAS ADMITTED ON (B)(6) 2024 WITH ACUTE HYPOXIC RESPIRATORY FAILURE SECONDARY TO A CONGESTIVE HEART FAILURE EXACERBATION, NEW ONSET ATRIAL FIBRILLATION, AND CHEST PAIN. THE PATIENT WAS STARTED ON A HEPARIN DRIP IN THE EMERGENCY ROOM AND THE PATIENT RECEIVED 3 BAGS OF HEPARIN OVER 12 HOURS INSTEAD OF 26 HOURS. THE INTENDED INFUSION RATE WAS 9.3ML/HOUR. THE TOTAL INFUSION WAS 3 BAGS OF 250ML FOR A TOTAL OF 750ML. THE CONCENTRATION FOR THE INFUSION WAS 25,000 UNITS IN 5% DEXTROSE PER 250ML BAG. NO OTHER INFUSIONS WERE RUNNING. THE PUMP WAS PROGRAMMED VIA BAR CODE SCANNING. ON THE MORNING OF (B)(6) 2024, THE PATIENT WAS FOUND TO BE DIAPHORETIC, AGITATED, VOMITING, COMPLAINING OF HEADACHE AND "NOT FEELING GOOD". BLOOD PRESSURE AND HEART RATE WERE ELEVATED. LABORATORY RESULTS REVEALED HIGH ANTIXA LEVELS. THE PATIENT PROCEEDED TO HAVE EPISODES OF HEMATEMESIS. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) AND A CT (COMPUTED TOMOGRAPHY) OF THE HEAD REVEALED SIGNIFICANT BLEEDING IN MULTIPLE LOCATIONS WITH ACTIVE BLEEDING. NEUROSURGERY WAS CONSULTED AND THERE WAS NO INDICATION FOR SURGERY OR DECOMPRESSION, AND NO TREATMENT COULD BE OFFERED. THE PATIENT EXPERIENCED DECLINING MENTAL STATUS WITH AIRWAY COMPROMISE AND WAS INTUBATED. PHYSICIANS SPOKE WITH FAMILY AND UPDATED ON THE PATIENT'S DECLINE. A REPEAT HEAD CT WAS OBTAINED THAT REVEALED WORSENING BLEEDING IN THE BRAIN. PALLIATIVE CARE WAS CONSULTED, AND THE DECISION WAS MADE BY FAMILY TO WITHDRAW CARE. THE PATIENT WAS PRONOUNCED DEAD ON (B)(6) 2024 1725. POST-EVENT, A REVIEW OF THE PUMP DATA BY THE (B)(6) SYSTEM DIRECTOR DRUG POLICY AND MEDICATION SAFETY, FOUND THE CORRECT ADULT PROFILE WAS UTILIZED, THE CORRECT DRUG LIBRARY LINE ITEM WAS UTILIZED, THE CORRECT THERAPY NAME WAS USED, CORRECT UNITS WERE USED TO PROGRAM THE PUMP AND DOSES WERE PROGRAMMED WITHIN THE GUARDRAILS. IT WAS FURTHER NOTED THAT THE PUMP WAS PROGRAMMED WITHIN NORMAL RANGES AND IT WAS NOTED THAT THE RAPID ADMINISTRATION DOES NOT APPEAR TO BE THE RESULT OF A PUMP PROGRAMMING ERROR. NURSING STAFF REPORT NO PUMP ALARMS SOUNDED DURING INFUSIONS TO INDICATE PUMP WAS MALFUNCTIONING. ANNUAL PREVENTATIVE MAINTENANCE ASSESSMENT OF THE INVOLVED INFUSION PUMP MODULE WAS PERFORMED BY A GE BIOMED TECHNICIAN AT (B)(6) ON (B)(6) 2024. NO MAINTENANCE ISSUES WERE IDENTIFIED. ANNUAL PREVENTATIVE MAINTENANCE ASSESSMENT OF THE INVOLVED INFUSION PUMP BRAIN WAS PERFORMED BY A GE BIOMED TECHNICIAN AT (B)(6) ON (B)(6) 2024. NO MAINTENANCE ISSUES WERE IDENTIFIED. THE INFUSION PUMP TUBING (BD/ALARIS, LOT# 24055720) AND THREE EMPTY BAGS OF HEPARIN WERE NOT SEQUESTERED AND WERE DISCARDED. THE TEST WAS PERFORMED AT 1546 ON (B)(6) 2024 2243 ON (B)(6) 2024 0731 ON (B)(6) 2024 0952 ON (B)(6) 2024." MULTIPLE ATTEMPTS HAVE BEEN MADE REQUESTING THE CUSTOMER TO PROVIDE ADDITIONAL INFORMATION AND DEVICE RETURN FOR INVESTIGATION. | OMIT : MEDICAL DEVICE OTHER OPERATOR. ADDITIONAL INFORMATION : MANUFACTURER NARRATIVE. ROOT CAUSE : THE ROOT CAUSE FOR THE REPORTED ISSUE OF AN DEATH - FREE FLOW EVENT - HEPARIN WAS NOT DETERMINED BECAUSE NO PRODUCTS OR DEVICE LOGS WERE RETURNED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. THE PATIENT CONTACT REQUESTED THAT THE SUPPLIES BE PICKED UP. IN ADDITIONAL FOLLOW-UP, IT WAS CONFIRMED WITH THE PATIENT¿S PD NURSE THAT ON (B)(6) 2024 THE PATIENT EXPIRED WHILE CONNECTED TO THE FRESENIUS CYCLER DURING AN UNKNOWN PHASE OF PD TREATMENT. THE NURSE REPORTED THAT THERE WERE NO ISSUES WITH THE FRESENIUS CYCLER IN RELATION TO THE DEATH. THE NURSE PROVIDED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS THAT INCLUDED END STAGE RENAL DISEASE (ESRD), CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), OBESITY, AND HYPERTENSION (HTN). THE CAUSE OF THE DEATH WAS REPORTED AS UNKNOWN AND THE ESRD DEATH CERTIFICATE WAS NOT AVAILABLE. THE PATIENT HAS BEEN DISCHARGED FROM THE CLINIC SYSTEM AND NO OTHER INFORMATION WAS AVAILABLE. THE FRESENIUS CYCLER HAS BEEN RETURNED TO MANUFACTURER (VIA ASSOCIATED CLINIC) AT THE TIME FOLLOW-UP WAS COMPLETED. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A (AS RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED ON THE CYCLER AND PASSED. THE CYCLER UNDERWENT AND PASSED A TOUCHSCREEN TEST, VOLTAGE VERIFICATION CHECK, TEACH PUMPS, SYSTEM AIR LEAK TEST, VALVE ACTUATION TEST, AND PATIENT SENSOR CHECK. AN INVESTIGATION OF THE CYCLER MUSHROOM HEADS VERIFIED THAT THE SURFACE CONDITIONS AND ALIGNMENTS WERE WITHIN SPECIFICATION. AN INTERNAL VISUAL INSPECTION IDENTIFIED EVIDENCE OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE DRIED FLUID COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
PATIENT INFORMATION CENTER IX | K183387 | Low Oxygen Saturation | Defective Alarm | IT WAS CONFIRMED THE PATIENT HAD DESATURATED AND THE NURSES WERE NOT AWARE OF THE CHANGE IN CONDITION. THE PATIENT WAS EMERGENTLY INTUBATED AND TRANSFERRED TO CVICU WHERE SHE LATER PASSED AWAY. IT WAS INDICATED THE ALARM WAS NOT AUDIBLE AT THE BEDSIDE OR CENTRAL STATION. THE CUSTOMER WAS ALSO NOT SURE WHETHER THE ALARMS WERE SET AND FAILED OR NOT PROPERLY SET UP. WHEN STAFF REVIEWED THE WAVEFORMS, IT WAS NOTED THE SPO2 HAD BEEN TRENDING DOWNWARD FOR APPROXIMATELY 10 MINUTES PRIOR TO THE BRADYCARDIA ALARM. A PRELIMINARY REVIEW OF THE CLINICAL AUDIT LOGS CONFIRMED THE LOW HEART RATE AND BRADYCARDIA ALARM WAS GENERATED AT THE TIME OF THE EVENT BUT NO SPO2 ALARMS WERE NOTED AT THAT TIME. | INITIAL REVIEW OF THE LOGS DETERMINED THAT THERE WERE TECHNICAL INOPS FROM 11:05 UNTIL 11:55 BUT THE LOGS ARE STILL UNDER FURTHER REVIEW. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | Respiratory Arrest; Cardiac Arrest | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Swelling/ Edema; Vomiting; Eye Irritation; Asthma; Dizziness; Nausea; Inflammation; Convulsion/Seizure; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, DEATH, OTHER, SWELLING IN ABDOMEN, INSOMNIA AND SEIZURES. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | Eye Irritation; Cardiovascular Dysfunction/Insufficiency; Dizziness; Inflammation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. WE ARE SENDING A FINAL REPORT SINCE NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Breast Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, KIDNEY DISEASE/TOXICITY, CANCER, LUNG AND BREAST CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Respiratory Problem; Fibrosis | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, DEATH, OTHER ACUTE PULMONARY FIBROSIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN SECTION H- TYPE OF REPORTED COMPLAINT AS ONLY ADVERSE EVENT. AFTER REVIEW, IT WAS DETERMINED THAT SECTION H- TYPE OF REPORTED COMPLAINT SHOULD BE FILED AS DEATH. SECTION H- TYPE OF REPORTED COMPLAINT CORRECTED IN THIS REPORT.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | A REGISTERED NURSE (RN) REPORTED TO FRESENIUS THAT THIS UNKNOWN PERITONEAL DIALYSIS (PD) PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY EXPIRED ON THE PRIOR NIGHT WHILE CONNECTED TO A HOSPITAL-PROVIDED LIBERTY SELECT CYCLER. IT WAS STATED THE PATIENT¿S DEATH WAS NOT CYCLER-RELATED AND THE CALL TO FRESENIUS WAS INTENDED FOR CYCLER PROCESSING FOLLOWING THE PATIENT DEATH. NO FURTHER INFORMATION WAS PROVIDED. THERE WAS NO SPECIFIC ALLEGATION THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. MULTIPLE ATTEMPTS WERE MADE TO ACQUIRE THE PATIENT¿S IDENTITY AND DETAILS CONCERNING THIS REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS OBTAINED. AS A RESULT, PATIENT DEMOGRAPHICS AND CAUSE OF DEATH REMAIN UNKNOWN. | PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | UDI NUMBER IS UPDATED IN THIS REPORT.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR PRO C-FLEX+ | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | Cancer | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | Cancer | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX D: COMMON DEVICE NAME WAS CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI XI | K131861 | Pulmonary Embolism; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A REVIEW OF A LITERATURE ARTICLE DESCRIBING A RETROSPECTIVE ANALYSIS OF ALL PATIENTS UNDERGOING DA VINCI ASSISTED ROBOTIC DISTAL PANCREATECTOMY (DP) AND PANCREATICODUODENECTOMY (PD) PROCEDURES PERFORMED BY A SINGLE SURGEON, WAS COMPLETED. THE STUDY EVALUATED THE SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREATECTOMY PROGRAM. THE STUDY INCLUDED SURGERIES FROM MAY 2014 TO DECEMBER 2020, DURING WHICH 62 PATIENTS UNDERWENT ROBOTIC PANCREATECTOMY, 34 PATIENTS WERE IN THE PD GROUP AND 28 PATIENTS WERE IN DP GROUP. ONE PATIENT WHO UNDERWENT PD EXPERIENCED A CARDIAC ARREST FOLLOWING A PULMONARY EMBOLISM ON POST-OPERATIVE DAY 5 AND COULD NOT BE REVIVED. THERE WERE NO DA VINCI DEVICE ISSUES REPORTED IN THE ARTICLE. ADDITIONAL INFORMATION WAS REQUESTED FROM THE DESIGNATED AUTHOR; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | THIS MEDWATCH REPORT REFLECTS ONE OF TWO DEATHS THAT OCCURRED, BOTH IN THE PD GROUP. SEE RELATED REPORT NUMBER 1 FOR THE OTHER MORTALITY IDENTIFIED IN THE ARTICLE. NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR THE DA VINCI SYSTEM IDENTIFIERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. A REVIEW OF THE ARTICLE PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF A PULMONARY EMBOLUS. NO ALLEGATION OF ANY ISSUE WITH ANY INTUITIVE SURGICAL PRODUCTS WAS MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. CITATION: MCKAY B, BROUGH D, KILBURN D, CAVALLUCCI D. SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREAS PROGRAM IN THE AUSTRALIAN SETTING: A CASE SERIES AND NARRATIVE REVIEW. ANZ J SURG. 2024 JUL-AUG;94(7-8):1247-1253. DOI: 10.1111/ANS.18998. EPUB 2024 MAR 26. PMID: 38529778. SECTION A - PATIENT INFORMATION - NO SPECIFIC PATIENT DEMOGRAPHICS WERE PROVIDED FOR THIS PATIENT. STUDY DEMOGRAPHICS IN THE PD GROUP INCLUDED 34 PATIENTS WITH A MEDIAN AGE OF 68 YEARS (RANGE 42-84); THE GENDER IS CITED AS 13 FEMALE, PRESUME 21 MALE. SECTION D - SUSPECT MEDICAL DEVICE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE DA VINCI SYSTEM ASSOCIATED WITH THE ADVERSE EVENT, A GENERIC SYSTEM MATERIAL NUMBER WAS USED. BLANK MDR FIELDS: THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELDS G5 AND G7 ARE NOT APPLICABLE.
DAVINCI XI | K131861 | Gastrointestinal Hemorrhage; Fistula | Adverse Event Without Identified Device or Use Problem | A STUDY TO EVALUATE THE SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREATECTOMY PROGRAM, WAS SUMMARIZED IN A LITERATURE ARTICLE DESCRIBING A RETROSPECTIVE ANALYSIS OF ALL PATIENTS UNDERGOING DA VINCI ASSISTED ROBOTIC DISTAL PANCREATECTOMY (DP) AND PANCREATICODUODENECTOMY (PD) PROCEDURES PERFORMED BY A SINGLE SURGEON. THE STUDY INCLUDED SURGERIES FROM MAY 2014 TO DECEMBER 2020, DURING WHICH 62 PATIENTS UNDERWENT ROBOTIC PANCREATECTOMY, 34 PATIENTS WERE IN THE PD GROUP AND 28 PATIENTS WERE IN DP GROUP. ONE PATIENT WHO UNDERWENT PD EXPERIENCED GRADE C POST-OPERATIVE PANCREATIC FISTULA AND GASTRODUODENAL ARTERY HEMORRHAGE. THE PATIENT DIED FROM MULTIORGAN FAILURE ON POST-OPERATIVE DAY 18. THERE WERE NO DA VINCI DEVICE ISSUES REPORTED IN THE ARTICLE. ADDITIONAL INFORMATION WAS REQUESTED FROM THE DESIGNATED AUTHOR; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | THIS MEDWATCH REPORT REFLECTS ONE OF TWO DEATHS THAT OCCURRED, BOTH IN THE PD GROUP. SEE RELATED REPORT NUMBER 1 FOR THE OTHER MORTALITY IDENTIFIED IN THE ARTICLE. NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR THE DA VINCI SYSTEM IDENTIFIERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. A REVIEW OF THE ARTICLE PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD DIED AS A RESULT OF PANCREATIC FISTULA. THE DETAILS OF HOW THIS MAY HAVE OCCURRED ARE NOT PROVIDED. NO ALLEGATION OF ANY ISSUE WITH ANY INTUITIVE SURGICAL PRODUCTS WAS MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. CITATION: MCKAY B, BROUGH D, KILBURN D, CAVALLUCCI D. SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREAS PROGRAM IN THE AUSTRALIAN SETTING: A CASE SERIES AND NARRATIVE REVIEW. ANZ J SURG. 2024 JUL-AUG;94(7-8):1247-1253. DOI: 10.1111/ANS.18998. EPUB 2024 MAR 26. PMID: 38529778. SECTION A - PATIENT INFORMATION - NO SPECIFIC PATIENT DEMOGRAPHICS WERE PROVIDED FOR THIS PATIENT. STUDY DEMOGRAPHICS IN THE PD GROUP INCLUDED 34 PATIENTS WITH A MEDIAN AGE OF 68 YEARS (RANGE 42-84); THE GENDER IS CITED AS 13 FEMALE, PRESUME 21 MALE. SECTION D - SUSPECT MEDICAL DEVICE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE DA VINCI SYSTEM ASSOCIATED WITH THE ADVERSE EVENT, A GENERIC SYSTEM MATERIAL NUMBER WAS USED.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY AND THE EQUIPMENT WAS BEING RETURNED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO THE SERVICE CENTER FOR INVESTIGATION. THE DEVICE WAS VISUALLY INSPECTED. NO FAULTS WERE REPORTED. THERE WERE 632 BLOWER HOURS WITH NO LOGGED ERRORS. THE USER INTERFACE PANEL MAIN AND ALARM RIBBON CABLE HAD DAMAGE. THE USER INTERFACE PANEL WAS REPLACED. THE FINE AND POLLEN FILTERS WERE REPLACED. THE DEVICE WAS CLEANED AND PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A FEW DAYS AFTER A DA VINCI-ASSISTED SIGMOID COLECTOMY PROCEDURE, THE PATIENT EXPIRED. THE DISTRIBUTOR CLINICAL SUPPORT SPECIALIST WAS INFORMED THAT THE PATIENT HAD AN UNSPECIFIED SERIOUS COMPLICATION AFTER THE PROCEDURE AND WAS TOLD THAT THE SURGEON DOES NOT THINK THAT A DA VINCI SYSTEM, INSTRUMENT, AND/OR ACCESSORY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. NO SPECIFIC INFORMATION WAS AVAILABLE. ADDITIONAL INFORMATION WAS REQUESTED BUT HAS NOT BEEN PROVIDED. | NO INTUITIVE PRODUCT WILL BE RETURNED FOR INVESTIGATION, AND NO PHOTOS OR VIDEO OF THE PROCEDURE ARE AVAILABLE. THE DA VINCI SYSTEM SERIAL NUMBER WAS PROVIDED AND REVIEW OF THE DEVICE HISTORY RECORD FOUND NO NON-CONFORMANCES WERE IDENTIFIED THAT WOULD BE RELATED TO AN ADVERSE EVENT. THE PROCEDURE DATE WAS REQUESTED BUT WAS NOT PROVIDED. REVIEW OF THE AVAILABLE SURGEON¿S PROCEDURE HISTORY FOUND THAT MULTIPLE SIGMOID COLECTOMY PROCEDURES HAVE BEEN PERFORMED BY THIS SURGEON. THUS, THE SPECIFIC PROCEDURE FOR THE REPORTED EVENT COULD NOT BE DETERMINED, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED A FEW DAYS AFTER A SIGMOID COLECTOMY. NO INFORMATION IS PROVIDED REGARDING THE DETAILS OF THE PROCEDURE, POST OPERATIVE COURSE, OR THE CIRCUMSTANCES LEADING TO THE DEATH. NO CAUSE OF DEATH IS PROVIDED. BASED ON THE LIMITED DETAILS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
REMSTAR AUTOA-FLEX | K131982 | Vomiting; Headache; Nausea; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING. ADDITIONALLY, PATIENT HAS PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ABOUT AN ALLEGATION OF A PATIENT DEATH. IN ADDITION, THERE WAS ALSO ALLEGATIONS OF LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Eye Irritation; Asthma; Dizziness; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Unspecified Kidney or Urinary Problem | Degraded | "THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED."
REMSTAR AUTOA-FLEX | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE MANUFACTURER RECEIVED INFORMATION REGARDING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. ADDITIONALLY, FOAM PARTICLES WERE NOT OBSERVED DURING THE EVALUATION OF THE DEVICE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, INFLAMMATORY RESPONSE AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THERE WAS 1 ERROR CODE FOUND. THE SECONDARY FINDINGS WERE OBSERVED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. IN BOX G: DATE RECEIVED BY MFG (RFB) HAS BEEN UPDATED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR AUTOA-FLEX | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PREVIOUSLY ALLEGED CANCER AND DEATH. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS, AND DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS THREE ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTRE CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS CORRECTED. IN BOX H: (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Dyspnea; Unspecified Respiratory Problem | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH AND SHORTNESS OF BREATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO PARTICLES IN AIR PATH. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS ZERO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTRE CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT CORRECTED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION CPAP PRO | K131982 | Asthma; Insufficient Information; Cancer; Unspecified Kidney or Urinary Problem | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, KIDNEY, BLADDER AND PROSTATE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Bronchitis; Pneumonia | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RECURRING PNEUMONIA, BRONCHITIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Dizziness; Nausea; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, NAUSEA / VOMITING, LIVER DISEASE/TOXICITY, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer; Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING COLON CANCER, DEATH, OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | No Clinical Signs, Symptoms or Conditions | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED AND DUST LIKE CONTAMINATION WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS THIRTY-TWO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. DEVICE PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID WAS UPDATED.
REMSTAR PRO C-FLEX+ | K131982 | Appropriate Term / Code Not Available | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS AND DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS 32 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. BOX H: PROBLEM CODE GRID, EVALUATION METHOD CODE, EVALUATION RESULTS CODE, AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 07/19/2024, AND SECTION H 11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT IS ALLEGING CANCER. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
ION | K182188 | Insufficient Information; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNEVENTFUL ION ENDOLUMINAL PROCEDURE. AFTER THE ION PROCEDURE, ENDOBRONCHIAL ULTRASOUND (EBUS) WAS PERFORMED WITH A BRONCHOSCOPE. THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE WAS INFORMED THAT THE PATIENT CODED AND EXPIRED WHILE IN POST-PROCEDURE RECOVERY. THERE WERE NO REPORTED INTRAPROCEDURAL ION SYSTEM MALFUNCTIONS/ISSUES. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. | A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS PERFORMED FOR THIS PROCEDURE AND NO RELATED NON-CONFORMANCES WERE IDENTIFIED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE UNDERWENT AN ION ENDOLUMINAL BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. IN RECOVERY, THE PATIENT DIED AFTER SUFFERING A CODE. NO FURTHER INFORMATION IS AVAILABLE BUT IT IS LIKELY PROCEDURE RELATED GIVEN TEMPORAL ASSOCIATION. THERE IS INSUFFICIENT INFORMATION TO DETERMINE A RELATIONSHIP TO THE DEVICES, BUT THERE WAS NO REPORTED MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. ATTEMPTS AT OBTAINING FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). --OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Cardiovascular Dysfunction/Insufficiency; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI XI | K131861 | Unspecified Tissue Injury | Insufficient Information | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNEVENTFUL DA VINCI-ASSISTED SACROCOLPOPEXY WITH HYSTERECTOMY AND SALPINGO-OOPHORECTOMY PROCEDURE. THE PROCEDURE WAS COMPLETED WITH A GYNECOLOGY SURGEON AND A SECONDARY GENERAL SURGEON WAS PRESENT TO PERFORM LYSIS OF ADHESIONS. THE PATIENT REQUIRED A SECOND PROCEDURE FOR OPEN SURGERY DUE TO AN UNSPECIFIED BOWEL INJURY. THE SECOND PROCEDURE WAS COMPLETED AND THE PATIENT EXPIRED ON POD 2. THERE WAS NO ALLEGATION OF A DA VINCI PRODUCT ISSUE RELATED TO THIS EVENT. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, CADIERE FORCEPS, MONOPOLAR CURVED SCISSORS, TIP-UP FENESTRATED GRASPER, MARYLAND BIPOLAR FORCEPS, LARGE NEEDLE DRIVER, LARGE SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A ROBOTIC SACROCOLPOPEXY AND WAS RE-EXPLORED ON POD 1 FOR A BOWEL INJURY. THE SECTION OF BOWEL AND HOW THE BOWEL INJURY MAY HAVE OCCURRED ARE NOT PROVIDED. THE INTRAOPERATIVE FINDINGS AT THE TIME OF THE RE-EXPLORATION WERE NOT PROVIDED. NEITHER THE DETAILS OF THE HOSPITAL COURSE AFTER THE INITIAL OPERATION NOR AFTER THE RE-EXPLORATION ARE PROVIDED. THE PATIENT DIED ON POD 2 AND NO DETAILS AS TO HOW THEY DIED ARE PROVIDED. BASED ON THE INFORMATION PROVIDED, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
REMSTAR PLUS | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR PRO C-FLEX+ DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF KIDNEY DISEASE/TOXICITY. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON THE CYCLER. UPON FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN), IT WAS REPORTED THIS PATIENT EXPIRED AT HOME DUE TO AN UNREPORTED CAUSE ON (B)(6) 2024. IT WAS AFFIRMED THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF DEATH. THE PATIENT WAS FOUND PULSELESS AND UNRESPONSIVE ON THE MORNING OF (B)(6) 2024 WHILE CONNECTED TO HER CYCLER. EMERGENCY SERVICES WERE NOT ACTIVATED, AND THE PATIENT WAS PRONOUNCED DECEASED IN HER HOME. THOUGH THE EXACT CAUSE OF DEATH WAS NOT REPORTED, IT WAS CONFIRMED THE PATIENT¿S DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF NAY FRESENIUS PRODUCT(S) OR DEVICE(S). | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CANNOT BE DETERMINED; HOWEVER, THERE WAS NO INDICATION THIS PATIENT¿S DEATH WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ABOUT AN ALLEGATION OF A PATIENT DEATH. IN ADDITION, THERE WAS ALSO ALLEGATIONS OF LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | Hypersensitivity/Allergic reaction; Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache; Unspecified Hepatic or Biliary Problem | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER INSPECTED EXTERNALLY AND INTERNALLY, OBSERVED INITIATED THERAPY TO THE DEVICE, THERE WAS A SLIGHT WHINY NOISE, POTENTIALLY DUE TO THE DUST LIKE CONTAMINATE ON THE INSIDE OF THE BLOWER. DUST LIKE CONTAMINATE ON THE PHILIPS WHITE ULTRA FINE FILTER, EXTERIOR AND INTERIOR OF THE TOP ENCLOSURE, PCA (PRINTED CIRCUIT ASSEMBLY), FLOW MANIFOLD, EXTERIOR AND INTERIOR OF THE BLOWER HOUSING, BLOWER GROMMET, BLOWER, RIGHT SIDE ASSEMBLY, HUMIDIFIER CABLE, AIR INLET SEAL, EXTERIOR AND INTERIOR OF THE BOTTOM ENCLOSURE, ISO PORT (INTERNATIONAL ORGANIZATION OF STANDARDIZATION), FLIP LID SEAL, DRY BOX, DRY BOX SEAL, DRY BOX INLET SEAL, TOP HOUSING, EXTERIOR AND INTERIOR OF THE HUMIDIFIER BOTTOM ENCLOSURE, HEATER PLATE, HEATER PLATE SPRING, EXTERIOR AND INTERIOR OF THE HUMIDIFIER LOWER BASE, AND THE SOUND ABATEMENT FOAM. MISSING THE UI KNOB (USER INTERFACE), SD CARD, AND THE SD CARD COVER. HAIRLIKE FIBERS ON THE FLIP LID SEAL AND THE HUMIDIFIER LOWER BASE. AN UNKNOWN CONTAMINATE ON THE EXTERIOR OF THE TOP ENCLOSURE WHERE THE UI KNOB WOULD SIT, EXTERIOR OF THE SIDE PANEL, AND THE EXTERIOR OF THE HUMIDIFIER SIDE PANEL. A YELLOWISH TINT ON THE FLOW MANIFOLD, BLOWER GROMMET, BLOWER BELLOW, AIR INLET SEAL, FLIP LID SEAL, DRY BOX SEAL, AND THE DRY BOX INLET SEAL. A PAPER-LIKE SUBSTANCE ON THE INTERIOR OF THE WATER TANK. POTENTIAL BIOCONTAMINATION ON THE BLOWER GROMMET, BLOWER BELLOW, AND THE AIR INLET SEAL. POTENTIAL LIQUID INGRESS WITH A WHITE POWDER CONSISTENT WITH MINERAL DEPOSITS ON THE EXTERIOR AND INTERIOR OF THE HUMIDIFIER FLIP LID, FLIP LID SEAL, WATER TANK, INTERIOR OF THE TOP ENCLOSURE, PCA, BLOWER HOUSING, BOTTOM ENCLOSURE, ISO PORT, DRY BOX SEAL, DRY BOX, AND DRY BOX INLET SEAL. THE SCREW BOSS WAS STRIPPED ON THE RIGHT-SIDE ASSEMBLY. THERMAL DAMAGE TO THE J9 AND ON THE HUMIDIFIER CABLE. THERMAL DAMAGE TO THE J9 HAS BEEN PREVIOUSLY INVESTIGATED AS DOCUMENTED. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE AND THEY WERE UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS.
REMSTAR PLUS | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR PRO C-FLEX+ DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Heart Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH DUE TO CANCER, REDUCED CARDIOPULMONARY RESERVE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Insufficient Information | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE DEVICE RETURNED FOR EVALUATION AND THERE WAS NO EVIDENCE OF FOAM PARTICLE OBSERVED. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Unspecified Heart Problem; Unspecified Respiratory Problem | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS 32 ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. SECTION H6 WAS UPDATED AND REPAIR EVALUATION CAPTURED.
REMSTAR AUTO A-FLEX | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE SERVICE CENTER FOR EVALUATION. DURING SERVICING OF THE DEVICE, NO EVIDENCE OF FOAM PARTICLES WAS OBSERVED. THE DEVICE HAS BEEN REPAIRED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER,DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | BOX D IS UPDATED IN THIS REPORT.
REMSTAR PRO C-FLEX+ | K091319 | Vomiting; Eye Irritation; Nausea; Unspecified Respiratory Problem; Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED LUNG DISEASE, LIVER DISEASE, EYE IRRITATION, NOSE IRRITATION, CANCER. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED LUNG DISEASE, LIVER DISEASE/TOXICITY, EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING AND CANCER. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: PATIENT OUTCOME CODE GRID EVALUATION METHOD CODE GRID EVALUATION RESULTS CODE GRID CONCLUSION CODE GRID HAS BEEN UPDATED IN PREVIOUS REPORT PRODUCT BRAND NAME CAPTURED WRONGLY, IN THIS REPORT HAS BEEN UPDATED CORRECTLY.
REMSTAR PRO C-FLEX+ | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. DUE DATE HAS BEEN UPDATED. IN BOX G: DATE RECEIVED BY MFG HAS BEEN UPDATED. IN BOX H6: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAVE BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Vomiting; Eye Irritation; Dizziness; Nausea; Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Insufficient Information; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, DEATH AND OTHER. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE CUSTOMER (DISTRIBUTOR) DECLINED TO PROVIDE ADDITIONAL INFORMATION DURING A GOOD FAITH EFFORT ATTEMPT. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Hypersensitivity/Allergic reaction; Insufficient Information; Dizziness; Inflammation; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, NOSE IRRITATION AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED TO FRESENIUS THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., DISCHARGE SUMMARY, DEATH CERTIFICATE, TREATMENT RECORDS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. ADDITIONALLY, NO ON-SITE EVALUATION WAS PERFORMED BY A FRESENIUS FIELD SERVICE TECHNICIAN (FST). A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Myocardial Infarction | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER DUE TO HEART FAILURE. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PDRN STATED THAT THE PATIENT LIVED WITH THE GRANDDAUGHTER. THE GRANDDAUGHTER CONTACTED THE PDRN AND REPORTED THAT ON THE MORNING OF (B)(6) 2024 911 WAS CALLED. THE REASON FOR CALLING 911 WAS NOT PROVIDED TO THE PDRN. THE PATIENT WAS IN TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER AT THE TIME OF THE EVENT (UNKNOWN PHASE AND CYCLE). PARAMEDICS ARRIVED AND DISCONNECTED THE PATIENT FROM TREATMENT. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL AND WAS LATER PRONOUNCED DECEASED. THE GRANDDAUGHTER STATED THE PATIENT DID NOT HAVE A DO NOT RESUSCITATE (DNR) ORDER AT THE TIME OF THE EVENT; HOWEVER, THE PATIENT¿S CARDIOLOGIST ADVISED HER HEART WAS NOT WORKING ANYMORE AND IT WAS JUST A MATTER OF TIME BEFORE SHE PASSED. THE PATIENT¿S FAMILY SAID HEART FAILURE WAS THE CAUSE OF DEATH. THE PATIENT REPORTEDLY WAS NOT STRONG ENOUGH TO SURVIVE ANY SURGERY. THE PATIENT HAD TWO PRIOR HEART VALVES OPERATED ON WHEN SHE WAS YOUNGER. THERE WERE NO REPORTED ISSUES WITH THE LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS AVAILABLE. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. THERE WAS NO EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT AND THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A SIMULATED TREATMENT WITH REDUCED DWELL WAS PERFORMED ON THE CYCLER AND PASSED. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST AND VALVE ACTUATION TEST. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DAVINCI XI | K131861 | Perforation of Vessels; Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED MEDIASTINAL MASS RESECTION (THYMECTOMY), ABOUT 2 HOURS INTO THE PROCEDURE, WHILE THE SURGEON WAS DISSECTING TISSUE A MAJOR VESSEL WAS INJURED AND BLEEDING OCCURRED; THE DIRECT CAUSE OF THE INJURY IS UNKNOWN. THE INTUITIVE CLINICAL TERRITORY ASSOCIATE WAS INFORMED BY THE ROBOTICS COORDINATOR THAT BLOOD SQUIRTED AND COVERED THE ENDOSCOPE LENS, OBSTRUCTING VISION. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE TO ADDRESS THE BLEEDING. AT SOME POINT IN TIME DURING THE PROCEDURE THE PATIENT CODED. THE PATIENT EXPIRED ON THE OPERATING ROOM TABLE ABOUT ONE HOUR INTO THE UNSPECIFIED INTERVENTIONS. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS (RECOVERABLE OR NON-RECOVERABLE FAULTS) OCCURRED DURING THOSE PROCEDURES. A REVIEW OF THE INSTRUMENT LOG FOUND THAT THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, CADIERE FORCEPS, LONG BIPOLAR GRASPER, TIP-UP FENESTRATED GRASPER, SUREFORM STAPLER 60, SYNCHROSEAL, SUCTION IRRIGATOR. ALL MULTI-USE INSTRUMENTS USED IN THE CASE HAVE BEEN USED IN SUBSEQUENT PROCEDURES. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THESE INSTRUMENTS. REVIEW OF THE DEVICE HISTORY RECORDS (DHR) FOUND THERE WERE NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THE DEVICES USED DURING THE EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT WAS UNDERGOING A ROBOTIC THYMECTOMY WHEN THEY GOT INTO BLEEDING. THERE ARE SEVERAL MAJOR VESSELS IN CLOSE PROXIMITY TO THE THYMUS. WHICH VESSEL WAS INVOLVED AND HOW THE SURGEON GOT INTO THE BLEEDING WAS NOT PROVIDED. THE CASE WAS CONVERTED TO OPEN BUT THE PATIENT DIED DURING THE PROCEDURE. NO OTHER DETAILS AS TO HOW THE INJURY OCCURRED NOR THE CAUSE OF DEATH WERE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
REMSTAR PRO C-FLEX+ | K091319 | Cancer; Eye Irritation; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE CUSTOMER (DISTRIBUTOR) DECLINED TO PROVIDE ADDITIONAL INFORMATION DURING A GOOD FAITH EFFORT ATTEMPT. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Vomiting; Nausea; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Pulmonary Dysfunction; Eye Irritation; Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REP DREAMSTATION AUTO BIPAP. THE PATIENT PASSED AWAY 2 YEARS AGO. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A REP DREAMSTATION AUTO BIPAP. THE PATIENT PASSED AWAY 2 YEARS AGO. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IN THE PREVIOUS REPORT, SECTION D: MODEL NUMBER, CATALOG ITEM IDENTIFIER, SERIAL NUMBER AND PRODUCT UDI WAS UPDATED/CORRECTED IN THIS REPORT.
REMSTAR PRO C-FLEX+ | K091319 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REMSTAR PRO C-FLEX+ DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REMSTAR PRO C-FLEX+ DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED THAT THE DEVICE WILL BE AVAILABLE FOR EVALUATION, BUT IT HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. AFTER DEVICE RETURN AND COMPLETION OF EVALUATION, A FOLLOW-UP REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Dizziness; Nausea; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Insufficient Information; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, DEATH, KIDNEY DAMAGE, LUNG DAMAGE AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
REMSTAR AUTOA-FLEX | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, SARCOMA CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Hypersensitivity/Allergic reaction; Asthma; Insufficient Information; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, HYPERSENSITIVITY, ASTHMA (NEW OR WORSENING), CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Asthma; Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING) LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS ELDERLY PATIENT ON PERITONEAL DIALYSIS (PD) EXPIRED ON (B)(6) 2024 WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN FACT, IT WAS REPORTED THAT THE DEATH WAS NOT PD RELATED. IN ADDITIONAL FOLLOW-UP, IT WAS CONFIRMED WITH THE PATIENT¿S PD NURSE THAT ON (B)(6) 2024 THE PATIENT EXPIRED WHILE CONNECTED TO THE FRESENIUS CYCLER DURING AN UNKNOWN PHASE OF PD TREATMENT. THE NURSE REPORTED THAT THERE WERE NO MALFUNCTIONS OR ISSUES WITH THE PD TREATMENT IN RELATION TO THE DEATH. ADDITIONALLY, THE NURSE STATED THAT THE PATIENT WAS COMPLETING PD TREATMENTS WITHOUT ANY ADVERSE EFFECTS PRIOR TO DEATH. THE NURSE PROVIDED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS THAT INCLUDED END STAGE RENAL DISEASE (ESRD), CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), DIABETES MELLITUS, HEART FAILURE AND EMPHYSEMA. THE CAUSE OF THE DEATH WAS REPORTED AS UNKNOWN AND THE ESRD DEATH CERTIFICATE WAS NOT AVAILABLE. HOWEVER, THE PD NURSE INDICATED THE DEATH IS SUSPECTED TO BE ASSOCIATED WITH THE PATIENT¿S PRE-EXISTING MEDICAL CONDITIONS. THE FRESENIUS CYCLER WAS PENDING RETURN TO MANUFACTURER AT THE TIME FOLLOW-UP WAS COMPLETED. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP MAY EXIST BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. BASED ON THE REPORTED INFORMATION, THE PATIENT¿S CAUSE OF DEATH CANNOT BE DETERMINED AND THE ESRD DEATH CERTIFICATE WAS NOT AVAILABLE. CURRENTLY, THERE IS NO ALLEGATION THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED PRE-EXISTING COPD, DIABETES MELLITUS, HEART FAILURE AND EMPHYSEMA WHICH ALL INCREASE MORTALITY RISK. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE IS CURRENTLY AT THE SERVICE CENTER AWAITING EVALUATION. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER HAS RECEIVED THE DEVICE EVALUATION FROM THE SERVICE CENTER.
REMSTAR AUTO A-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Heart Problem; Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | Cancer; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO LUNG AND LIVER CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR AUTO A-FLEX | K091319 | Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND LIVER DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR AUTO A-FLEX | K131982 | Vomiting; Dizziness; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER, DEATH THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Vomiting; Asthma; Dizziness; Nausea; Inflammation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Cancer; Unspecified Respiratory Problem | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A DS2ADV AUTO CPAP WITH AN ALLEGATION OF LUNG DISEASE; CANCER; DEATH THIS DEVICE IS NOT PART OF THE RECALL. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Eye Irritation; Asthma; Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA AND LUNG DISEASE, DIZZINESS/HEADACHE AND EYE, SKIN, NOSE AND RESPIRATORY TRACT IRRITATION. THERE WAS REPORT OF PATIENT DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer | Adverse Event Without Identified Device or Use Problem | DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. COSMETIC DEFECT FOUND DEBRIS IN LCD SCREEN. THE DEVICE WAS SCRAPPED. THE MANUFACTURER RECEIVED INFORMATION CANCER; DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, LUNG DISEASE, CANCER, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | Myocardial Infarction; Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED RADICAL PROSTATECTOMY (WITHOUT LYMPHADENECTOMY), THE PATIENT REQUIRED A POST-OPERATIVE OPEN PROCEDURE AT AN UNSPECIFIED TIME DUE TO BLEEDING AND LATER EXPIRED DUE TO A MYOCARDIAL INFARCTION. THE SURGEON REPORTED THAT WHILE THERE WAS SOME INTRA-OPERATIVE BLEEDING DURING THE DA VINCI-ASSISTED PROCEDURE, THE BLEEDING WAS CONTROLLED AND THE PROCEDURE WAS COMPLETED. THE SURGEON STATED THAT THERE WAS NO CONCERN FOR ANY DA VINCI ROBOTIC ISSUES. POST-OPERATIVELY, THE PATIENT'S HEMOGLOBIN WAS STABLE AND CHECKED TWICE AT 8.7 AND 8.8G/DL FOUR HOURS APART AND THE PATIENT HAD NORMAL VITAL SIGNS AND ABDOMINAL EXAM. HOWEVER, THE PATIENT EXPERIENCED A VENTRICULAR FIBRILLATION ARREST AND WAS CODED FOR 25 MINUTES AFTER WHICH THE PATIENT¿S ABDOMEN WAS NOTED TO BE DISTENDED. THE SURGEON STATED THAT THE DISTENTION WAS LIKELY DUE TO THE BAGGING EFFORTS DURING CPR. DUE TO CONCERN FOR POSSIBLE BLEEDING AN EMERGENT EXPLORATORY LAPAROSCOPY WAS PERFORMED WHERE AN UNSPECIFIED AMOUNT OF BLOOD WAS FOUND IN THE ABDOMEN AND EVACUATED; THE BLOOD WAS REPORTED TO LIKELY BE OLD. THE SURGEON STATED THAT THE PATIENT EXPIRED DUE TO A MYOCARDIAL INFARCTION. NO AUTOPSY WAS PERFORMED. | A REVIEW OF THE DA VINCI SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, 30 DEGREE ENDOSCOPE PLUS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, TWO MONOPOLAR CURVED SCISSORS, TWO LARGE NEEDLE DRIVERS. A SITE HISTORY REVIEW FOUND NO SUBSEQUENT COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT A ROBOTIC PROSTATECTOMY. BLEEDING WAS NOTED TO HAVE OCCURRED INTRAOPERATIVELY BUT THE SURGEON STATED THERE WAS NO CONCERN FOR A ROBOTIC ISSUE. THE DETAILS OF WHAT BLED OR HOW THE BLEEDING MAY HAVE OCCURRED WERE NOT PROVIDED. THE HEMOGLOBIN WAS STABLE ON TWO OCCASIONS IN THE POST OPERATIVE PERIOD BUT THE TIMELINE AS TO WHEN THOSE LEVELS WERE DRAWN WAS NOT PROVIDED. VITALS WERE ALSO REPORTEDLY NORMAL. HOWEVER, THE PATIENT DEVELOPED A DYSRHYTHMIA, CODED, AND WAS NOTED TO BE DISTENDED. THE PATIENT WAS EVENTUALLY RE-EXPLORED AT WHICH TIME BLOOD WAS FOUND. NEITHER THE VOLUME OF BLOOD DISCOVERED AT RE-EXPLORATION, THE SOURCE OF THE BLOOD LOSS, NOR ANY ADDITIONAL FINDINGS WERE PROVIDED. ALTHOUGH THERE WAS NO AUTOPSY, THE SURGEON REPORTED THE PATIENT DIED AS A RESULT OF A MYOCARDIAL INFARCTION. NO ADDITIONAL SUPPORTING EVIDENCE FOR THIS CONCLUSION WAS PROVIDED. BASED ON THE INFORMATION PROVIDED, INSUFFICIENT EVIDENCE IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION B2 - DATE OF DEATH IS ESTIMATED AS SAME DATE OF PROCEDURE AS COMPLICATION OCCURRED IN THE PACU POST-PROCEDURE. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
ION | K182188 | Respiratory Insufficiency; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A PATIENT ENROLLED IN A REGISTRY STUDY UNDERWENT AN UNEVENTFUL ION BIOPSY PROCEDURE AND WAS DISCHARGED HOME 3.5 HOURS POST-PROCEDURE. NO ION SYSTEM MALFUNCTIONS OR OTHER ISSUES OCCURRED DURING THE BIOPSY PROCEDURE. IT WAS REPORTED THAT 8 DAYS POST-DISCHARGE THE PATIENT BEGAN TO EXPERIENCE RESPIRATORY DISTRESS IN THE EVENING AT HOME. THE PATIENT WENT INTO CARDIAC ARREST AND WAS UNABLE TO BE REVIVED. THE STUDY PRINCIPAL INVESTIGATOR CLASSIFIED THE EVENT AS NOT RELATED TO THE ION DEVICES, POSSIBLY RELATED TO THE ION PROCEDURE, AND PROBABLY RELATED TO THE PATIENT¿S PRE-EXISTING COMORBIDITIES. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. CONCOMITANT PRODUCTS: TWO LESIONS IN THE LEFT UPPER LOBE WERE BIOPSIED: ONE LOCATED 0 MM FROM THE PLEURA (18 X 24 X 20 MM) ¿ PRODUCTS USED WERE A 19G FLEXISION NEEDLE, 1.1 MM - CRYOPROBE, CYTOLOGY BRUSH; AND ONE LOCATED 13 MM FROM THE PLEURA (17 X 15 X 18 MM) ¿ 23G FLEXISION NEEDLE; 19G FLEXISION NEEDLE, 1.1 - MM CRYOPROBE, CYTOLOGY BRUSH. A REVIEW OF THE EVENT WAS CONDUCTED BY AN INTUITIVE SURGICAL MEDICAL OFFICER WHO CONCLUDED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL BIOPSY AS PART OF A REGISTRY CLINICAL STUDY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS DISCHARGED HOME A FEW HOURS AFTER THE PROCEDURE. THE PATIENT DEVELOPED RESPIRATORY DISTRESS AT HOME ABOUT 1 WEEK LATER PROMPTING EMERGENCY SERVICES TO BE CONTACTED. THE PATIENT WENT INTO CARDIAC ARREST. ATTEMPTS AT RESUSCITATION WERE UNSUCCESSFUL AND THE PATIENT DIED. THE PATIENT HAD SIGNIFICANT MEDICAL CO-MORBIDITIES INCLUDING SEVERE PULMONARY FIBROSIS/INTERSTITIAL PNEUMONITIS, CHRONIC BRONCHITIS, PULMONARY HYPERTENSION, RHEUMATOID ARTHRITIS, DIABETES, HYPERTENSION AND A 90 PACK-YEAR SMOKING HISTORY. BASED ON THE AVAILABLE DATA THE EVENT WAS POSSIBLY PROCEDURE RELATED IN THE SETTING OF MULTIPLE SIGNIFICANT CO-MORBIDITIES AND WAS NOT DEVICE RELATED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF (B)(4) BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE (B)(4) WITH (B)(4) TOTAL DEATHS (B)(4). A MULTICENTER STUDY REPORTED (B)(4) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH (B)(4) CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF (B)(4) REPORTED (B)(4). ANOTHER SURVEY-BASED STUDY OF (B)(4) BRONCHOSCOPIES DESCRIBED (B)(4) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN (B)(4) PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF (B)(4). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING (B)(4) CASES REPORTED NO DEATHS. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF (B)(4) BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). --OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, CANCER, DEATH, PANCREATIC AND METASTISIZED TO LIVER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem; Fibrosis | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE, CANCER AND IDIOPATHIC PULMONARY FIBROSIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Pulmonary Dysfunction; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CHRONIC OBSTRUCTIVE PULMONARY DISEASE AND ACUTE RESPIRATORY DISTRESS SYNDROME. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Heart Failure/Congestive Heart Failure; Respiratory Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY FAILURE AND HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Gastrointestinal Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH AND INTESTINAL METAPLASIA THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. USING A KNOWN GOOD POWER SUPPLY AND POWER CORD,THE MANUFACTURER WAS ABLE TO CONFIRM THE DEVICE POWERS ON AND PROVIDED AIRFLOW. THE MANUFACTURER CONFIRMED THE OPERATION OF THE HEATER PLATE. DURING REVIEW OF THE ERROR LOGS, THE MANUFACTURER DETERMINED THAT THE DEVICE HAD 13923.6 MACHINE HOURS, 13923.6 BLOWER HOURS AND 13899.1 THERAPY HOURS. THE ERROR LOG SHOWED 1 ERROR LOGGED. 1 INSTANCE OF E-53 (ERR_COMP_LOG_SEM_TIMEOUT) A CONTINUE ERROR. DURING A CHANGE TO THE COMPLIANCE LOGGER, THE SEMAPHORE TIMER EXPIRES, AND THE SEMAPHORE IS RELEASED. THIS CAN OCCUR OF THE MODEM IS REMOVED WHILE THE COMPLIANCE LOGGER HAS CONTROL OF THE SEMAPHORE. CARE ORCHESTRATOR DATA SHOWS THE PATIENT HAD A COMPLIANCE RATE OF (B)(4) AND USED A SYSTEM ONE HUMIDIFICATION MODE, WITH A HUMIDIFIER SETTING OF 3. DURING THE INVESTIGATION THE MANUFACTURER OBSERVED DRIED LIQUID SPOTS ON THE TOP ENCLOSURE. AN UNKNOWN PIECE OF DEBRIS WAS OBSERVED ON THE BACKSIDE OF THE UI PANEL. POTENTIAL DRIED LIQUID SPOTS WERE OBSERVED ON THE BLOWER HOUSING, ON THE BOTTOM OF AND INSIDE THE BLOWER MOTOR. THE MANUFACTURER OBSERVED A WHITISH POWDERY DUST-LIKE CONTAMINATION IN THE AIR INLET, INSIDE THE BLOWER MOTOR, ON THE BOTTOM OF THE BLOWER HOUSING, ON THE SOUND ABATEMENT FOAM AND THROUGHOUT THE BOTTOM ENCLOSURE. DUST-LIKE CONTAMINATION WAS OBSERVED ON THE RIGHT-SIDE PANEL, IN THE ISO PORT, ON THE INTERIOR SIDE OF THE TOP ENCLOSURE ON THE PCA AND ON THE BLOWER CAP. THE MANUFACTURER OBSERVED A SMALL AMOUNT OF POTENTIAL DEGRADED SOUND ABATEMENT FOAM ON THE BOTTOM OF THE BLOWER HOUSING. THE MANUFACTURER CONFIRMED THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. (DS6HFLG) THE MANUFACTURER OBSERVED POTENTIAL DRIED LIQUID SPOTS ON THE WATER TANK, INTERIOR SIDE OF THE FLIP LID ASSEMBLY AND IN THE LOWER BASE. DUST-LIKE CONTAMINATION WAS OBSERVED THROUGHOUT THE WATER TANK, ON THE INTERIOR SIDE OF THE FLIP LID SEAL, IN THE BOTTOM ENCLOSURE, ON THE DRY BOX, AND IN THE LOWER BASE. A BROWNISH CONTAMINATION WAS OBSERVED ON THE FLIP LID SEAL. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE AND DUST CONTAMINATION WAS ABLE TO CONFIRM THE CUSTOMER'S COMPLAINT. IN BOX G: DATE RECEIVED BY MFG. HAS BEEN UPDATED. IN BOX H: (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER AND LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | Cancer | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING DEATH AND CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 02/01/2025, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING DEATH AND CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING EVALUATION SECONDARY FINDINGS AND DUST CONTAMINATION WERE OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 11 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. IN BOX H: DEVICE PROBLEM CODE, EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, DEATH, CANCER AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | IN PREVIOUS REPORT, AE OUTCOME WAS CAPTURED INCORRECTLY, AND IT IS CAPTURED CORRECTLY IN THIS REPORT. DEATH.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information; Dizziness; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Insufficient Information; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, PANCREATIC. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Asthma; Dizziness; Nausea; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), HYPERSENSITIVITY, NAUSEA, VOMITING, KIDNEY DISEASE/TOXICITY, LUNG, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Pulmonary Embolism | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY CANCER, BILATERAL PULMONARY EMBOLISM AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K091319 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED UNKNOWN INFORMATION. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER AN ELECTIVE DA VINCI-ASSISTED INGUINAL HERNIA REPAIR, THE PATIENT ¿WOULD NOT WAKE UP FROM ANESTHESIA¿ AND PASSED AWAY. THE PATIENT HAD REPORTEDLY EXPERIENCED CHEST PAIN PRIOR TO THE PROCEDURE. THE CAUSE OF THE CHEST PAIN AND THE PATIENT¿S MEDICAL HISTORY WERE NOT PROVIDED. THE SITE ROBOTICS COORDINATOR INFORMED THE INTUITIVE CLINICAL SALES ASSOCIATE THAT THE PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS AND THE DA VINCI ROBOT WAS UNDOCKED. THE HOSPITAL OR STAFF ATTEMPTED TO WAKE THE PATIENT BUT WERE UNSUCCESSFUL. CARDIOCEREBRAL RESUSCITATION (CCR) WAS ADMINISTERED AND AN INTERVENTION TEAM WAS CALLED TO ASSIST; HOWEVER, THE PATIENT EXPIRED. THE ROBOTICS COORDINATOR REPORTED THAT THE EVENT HAD NOTHING TO DO WITH THE DA VINCI PRODUCTS. ADDITIONAL INFORMATION WAS REQUESTED BUT HAS NOT BEEN PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, FENESTRATED BIPOLAR FORCEPS, MONOPOLAR CURVED SCISSORS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED SHORTLY AFTER AN INGUINAL HERNIA REPAIR. THE PATIENT HAD CHEST PAINS PRIOR TO SURGERY BUT THE PROCEDURE WAS CARRIED OUT ANYWAY. NO ISSUES OR OTHER COMPLICATIONS OCCURRED DURING THE PROCEDURE. THE CAUSE OF DEATH IS NOT GIVEN BUT THE PREOPERATIVE CHEST PAIN IS CONCERNING FOR A CARDIAC EVENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Asthma; Nausea; Pulmonary Dysfunction; Unspecified Respiratory Problem; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND CHRONIC LYMPHOCYTIC LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, NEUROENDOCRINE (LIVER AND SPINE). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, ANAPLASTIC CANCER AND MULTIPLE MYELOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Dizziness; Nausea; Lymphoma; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING AND CANCER CUTANEOUS T-CELL LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Headache; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, CANCER, DEATH, LUNG CANCER AND BRAIN CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Asthma; Dizziness; Nausea; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING) DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Heart Failure/Congestive Heart Failure; Insufficient Information; Pulmonary Dysfunction; Ventricular Fibrillation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, REDUCED CARDIOPULMONARY RESERVE, HEART FAILURE/SUDDEN EMERGENCY SITUATION, RESULTING IN VFIB AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, LUNG CANCER PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RENAL CELL CARCINOMA, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER (DIFFUSE LARGE B-CELL LYMPHOMA) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | IN THE PERVIOUS REPORT DATE OF DEATH (RFB) WAS CAPTURE INCORRECTLY, IT IS CORRECTED IN THIS REPORT DATE OF DEATH (RFB) - NO INFORMATION.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Asthma; Dizziness; Nausea; Pulmonary Dysfunction; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND SURGERY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Dizziness; Nausea; Cardiovascular Dysfunction/Insufficiency; Inflammation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAD DIED WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE DEVICE IS NOT PART OF RECALL.
REMSTAR AUTO A-FLEX | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LIVER DISEASE/TOXICITY AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. IN BOX H: TYPE OF REPORTED COMPLAINT IS UPDATED IN THIS FOLLOW-UP.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Myocardial Infarction | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. UPON FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN), IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6)2024 WHILE CONNECTED TO HER LIBERTY SELECT CYCLER. THE PATIENT¿S CAUSE OF DEATH WAS FROM A MYOCARDIAL INFARCTION (MI) THAT OCCURRED DURING A PD TREATMENT ON THE EVENING OF (B)(6)2024. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS PRONOUNCED DECEASED IN HER HOME WITH NO TRANSPORT TO HOSPITAL. THE PATIENT DID NOT HAVE ANY KNOWN HISTORY OF CARDIAC DISEASE AND DID NOT EXPERIENCE ANY PRIOR MIS. THE PDRN STATED THE PATIENT¿S PHYSICIAN RULED OUT ANY INVOLVEMENT WITH PD THERAPY IN RELATION (OTHER THAN TEMPORAL) TO THE PATIENT¿S MI AND SUBSEQUENT DEATH. IT WAS CONFIRMED, THOUGH THE EXACT CAUSE OF THE PATIENT¿S MI WAS NOT REPORTED, THERE WAS NO INDICATION THE PATIENT¿S MI LEADING TO HER DEATH WAS RELATED TO PD THERAPY OR DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CAN BE ATTRIBUTED TO AN MI WITH NO INDICATION THE MI AND SUBSEQUENT DEATH WERE DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO BIPAP | K131982 | Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Lymphoma | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER AND NON HODGKINS LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer; Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, KIDNEY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX E: HEALTH PROFESSIONAL (RFB) WAS CORRECTED. BOX H: MANUFACTURE DATE (RFB) WAS CORRECTED.
REMSTAR AUTO A-FLEX | K131982 | Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Pulmonary Dysfunction; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Hypersensitivity/Allergic reaction; Headache; Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Bronchitis; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND CHRONIC BRONCHITIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND THROAT CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PATIENT INFORMATION CENTER IX | K211900 | Insufficient Information | No Audible Alarm | THE CUSTOMER REPORTED THAT THE STAFF DID NOT HEAR THE ALARM FROM THE EMERGENCY ROOM SURVEILLANCE AND THE PATIENT PASSED AWAY. THE ENGINEER TESTED THE SYSTEM AND FOUND IT TO BE IN WORKING ORDER. THE UNIT MANAGER REQUESTED THAT THE VOLUME OF ALARMS BE INCREASED, THE ENGINEER INCREASED THE VOLUME FROM 4 TO 6. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | A PHILIPS FIELD SERVICE ENGINEER WAS DISPATCHED TO THE CUSTOMER SITE. THE MONITOR AND THE PATIENT INFORMATION CENTER WERE BOTH SUCCESSFULLY TESTED FOR ALARMING AND FOR SOUND. AUDIT LOGS WERE RETRIEVED AND SHARED WITH THE HOSPITAL ADMINISTRATION. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH STAGE 4 METASTATIC CANCER OF THE LIVER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A NURSE REPORTED A PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2024. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING OF (B)(6) 2024. THE PATIENT¿S DAUGHTER DISCOVERED THE PATIENT HAD EXPIRED WHEN ARRIVING TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE PATIENT WAS IN THE FINAL DWELL CYCLE WHEN IT STOPPED, AND NO PRIOR ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO PATIENT'S ADVANCED AGE AND SIGNIFICANT CARDIAC HISTORY. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, DUE TO HER POOR CARDIAC HEALTH AND ADVANCED AGE. THE SERIOUS ADVERSE EVENTS WERE REPORTEDLY UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION | K131982 | Obstruction/Occlusion; Chronic Obstructive Pulmonary Disease (COPD); Syncope/Fainting | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED DUE TO COPD. THE FAMILY STATED THAT HER CONDITION GOT WORSE WHILE USING THE DEVICE. THERE ARE OTHER ALLEGATIONS THAT THE PATIENT PASSED OUT DUE TO LACK OF OXYGEN AND ALSO HAD A PULMONARY OBSTRUCTION. THE PATIENT SPENT TIME IN THE HOSPITAL AND UNDERWENT TEST. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND STOMACH CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Inflammation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN B5 VERBIAGE WHICH WAS CORRECTED IN THIS REPORT.
DREAMSTATION CPAP | K131982 | Cancer; Lymphoma | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, LARGE B CELL NON-HODGKIN'S LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION FROM CUSTOMER IN REFERENCE TO A DS2ADV AUTO CPAP WITH AN ALLEGATION PATIENT PASSED AWAY DUE TO THE DEVICE. THERE WAS NO REPORT OF SERIOUS INJURY OR HARM. THERE IS NO MEDICAL INTERVENTION WAS SPECIFIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, AND CONCLUSION CODE GRID WAS UPDATED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | Appropriate Term / Code Not Available; Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE, TOXICITY, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PHILIPS CPAP DEVICE | K131982 | Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PHILIPS CPAP DEVICE | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR AUTO A-FLEX | K091319 | Vomiting; Appropriate Term / Code Not Available; Dizziness; Nausea; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, NAUSEA, VOMITING, LIVER DISEASE, TOXICITY, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | Pneumothorax; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | REFER TO H11 FOR FOLLOW-UP INFORMATION. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS THAT COULD HAVE CONTRIBUTED TO THE REPORTED EVENT OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Asthma; Insufficient Information; Dizziness; Nausea; Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, LUNG CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY IN 2023 AND THE PATIENT'S WIFE IS REQUESTING TO BE REMOVED FROM FURTHER CORRESPONDENCE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | Insufficient Information | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED AND DUST LIKE CONTAMINATION WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS SEVEN ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION, BUT DUST-LIKE CONTAMINATION WAS OBSERVED AND UNIT GOT CORRECTED. DEVICE PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID WAS UPDATED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, CONSTANT RESPIRATORY INFECTIONS, SINUS INFECTIONS, AGGRAVATED ASTHMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION CPAP DEVICE WAS THROWN AWAY A LONG TIME AGO AND THE USER HAD PASSED AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DECEASED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ALARIS SYSTEM | K133532 | Cardiac Arrest | Device Alarm System; No Apparent Adverse Event | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6) 2024 AT 0255. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 1119." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | Cardiac Arrest | Misassembly During Maintenance/Repair; Device Alarm System; Contamination; Break; Crack | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6)2024 AT 0255. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 1119." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | Cardiac Arrest | Misassembly During Maintenance/Repair; Device Alarm System; Contamination; Break; Crack | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6) 2024 AT 02:55. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 11"19." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | Cardiac Arrest | Device Alarm System; No Apparent Adverse Event | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6) 2024 AT 0255. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 1119." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | A DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. DURING THE EVALUATION, SECONDARY FINDINGS WERE OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Insufficient Information; No Apparent Adverse Event | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION RECEIVED ON 02/05/2025. H11 UPDATED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H6 HAS BEEN UPDATED.
DREAMSTATION CPAP PRO | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION CPAP PRO DEVICE WAS RETURNED TO THE PATIENT'S DME, AND THE USER HAD PASSED AWAY A YEAR AND A HALF AGO. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT'S WIFE REQUESTED THAT NO FURTHER CORRESPONDENCE BE SENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Hematoma; Sepsis | Adverse Event Without Identified Device or Use Problem | A LITERATURE ARTICLE DESCRIBED A PROSPECTIVE SINGLE-CENTER STUDY OF PATIENTS WITH AORTOILIAC OCCLUSIVE DISEASE (AIOD) THAT UNDERWENT ROBOTIC-ASSISTED LAPAROSCOPIC (RAL) REVASCULARIZATION PROCEDURES. THE STUDY WAS CONDUCTED FROM FEBRUARY 2014 TO FEBRUARY 2019 AND INCLUDED 70 PATIENTS. THE AIM OF THE STUDY WAS TO ASSESS THE SURGEON LEARNING CURVE AND THE FEASIBILITY, SAFETY, AND MIDTERM OUTCOMES. THE ARTICLE DESCRIBED ONE PATIENT DEATH IN THE ROBOTIC COHORT. THE PATIENT HAD MULTIPLE COMORBIDITIES INCLUDING CHRONIC ISCHEMIC HEART DISEASE, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE WITH CENTRO-LOBULAR EMPHYSEMA, A HISTORY OF ACUTE PANCREATITIS, ACTIVE SMOKING, AND TASC D AORTOILIAC LESIONS (REQUIRES SURGICAL INTERVENTION). THE PATIENT PRESENTED WITH A RETROPERITONEAL HEMATOMA ON POSTOPERATIVE DAY (POD) 1, WHICH REQUIRED SURGICAL DRAINAGE; ON POD 2, ILEAL AND COLONIC ISCHEMIA DEVELOPED, LEADING TO AN ILEAL AND RIGHT COLONIC RESECTION, FOLLOWED BY A LEFT COLECTOMY ON POD 13, ULTIMATELY SUCCUMBING TO SEPTIC SHOCK IN INTENSIVE CARE ON POD 51. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS OR ISSUES STATED, NOR DID THE AUTHORS SUGGEST THAT ANY INTUITIVE DA VINCI PRODUCTS CAUSED OR CONTRIBUTED TO THE ADVERSE EVENTS. THE DESIGNATED CONTACT AUTHOR HAS NOT RESPONDED TO REQUESTS FOR ADDITIONAL INFORMATION. THE ARTICLE DOCUMENTS THAT THE STUDY RESULTS ARE COMPARABLE TO STANDARD SURGICAL CONVENTIONAL OPEN OR LAPAROSCOPIC SURGERY. | NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM OR INSTRUMENT IDENTIFIERS. NO SPECIFIC INFORMATION REGARDING THE PROCEDURE SITE OR DATE WAS PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE LITERATURE ARTICLE DESCRIBES THE DEATH OF A PATIENT WITH A LONG LIST OF SIGNIFICANT COMORBIDITIES. THE DEATH OCCURRED A PROLONGED PERIOD AFTER AORTOILIAC ROBOTIC SURGERY AND NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS ARE MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SUTTER W, ALSAC JM, BEN ABDALLAH I, MICHEL C, JULIA P, EMPANA JP, EL BATTI S. TREATMENT OF AORTOILIAC OCCLUSIVE LESIONS BY AORTIC ROBOTIC SURGERY: LEARNING CURVE AND MIDTERM OUTCOME. ANN VASC SURG. 2024 JUL; 104:258-267. DOI: 10.1016/J.AVSG.2024.02.018. EPUB 2024 APR 7. PMID: 38593921. SECTION B2 DATE OF DEATH: THE PATIENT DATE OF DEATH WAS NOT PROVIDED; THE STUDY WAS CONDUCTED BETWEEN FEBRUARY 2014 AND FEBRUARY 2019. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
DREAMSTATION | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT PASSED AWAY, AND HIS WIFE WANTED TO RETURN IT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION AUTO CPAP DEVICE. THE SPOUSE OF A PATIENT REPORTED THE PATIENT HAS PASSED. THE REPORTER STATED THE PATIENT NEVER RECEIVED A RETURN LABEL AND DID NOT RECEIVE A REPLACEMENT. VETERANS AFFAIRS WOULD NOT TAKE THE DEVICE, THEREFORE THE REPORTER STATED THE DEVICE WAS "THROWN OUT". THE PATIENT'S CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION AUTO CPAP DEVICE. THE CLINIC HAS REPORTED THAT THE PATIENT HAS PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO A DREAMSTATION AUTO CPAP DEVICE. THE CLINIC HAS REPORTED THAT THE PATIENT HAS PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS NOT REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. UPDATED: EVALUATION METHOD CODE UPDATED. EVALUATION RESULTS CODE UPDATED. CONCLUSION CODE UPDATED.
DAVINCI XI | K131861 | Pulmonary Embolism; Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE COHORT STUDY TO EVALUATE THE SAFETY AND EFFECTIVENESS OF THE USE OF THE DA VINCI SYNCHROSEAL INSTRUMENT COMPARED WITH THE DA VINCI VESSEL SEALER EXTEND INSTRUMENT FOR ROBOTIC PARENCHYMAL TRANSECTION OF THE LIVER. THE STUDY INCLUDED A TOTAL OF 155 ROBOTIC LIVER RESECTION (RLR) PATIENTS USING THE SYNCHROSEAL FROM JUNE 2021 TO MARCH 2023, AND 161 RLR PATIENTS USING THE VESSEL SEALER FROM FEBRUARY 2020 TO MARCH 2023 PERFORMED AT TWO INSTITUTIONS. DURING THE STUDY, ONE PATIENT IN THE VESSEL SEALER COHORT DIED AS A RESULT OF A COMPLICATED POSTOPERATIVE COURSE WITH A LUNG EMBOLUS AND BLEEDING FROM THE INFERIOR EPIGASTRIC ARTERY THAT REQUIRED A REOPERATION DURING WHICH IATROGENIC DAMAGE OF THE BLADDER OCCURRED. LATER, THE PATIENT DIED OF HEMORRHAGIC SHOCK STEMMING FROM BLEEDING AT THE SITE OF THE NEPHROSTOMY CATHETER. THERE WERE NO DEVICE ISSUES MENTIONED IN EITHER OF THE TWO GROUPS IN THE ARTICLE. THE ARTICLE CONCLUSION STATED THAT THE SYNCHROSEAL IS A SAFE AND EFFECTIVE DEVICE FOR ROBOTIC LIVER PARENCHYMAL TRANSECTION. | NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM IDENTIFIERS OR INSTRUMENT LOT NUMBERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. THERE IS NO INDICATION THAT ANY DA VINCI PRODUCTS CONTRIBUTED TO THE REPORTED PATIENT COMPLICATIONS. THIS ARTICLE REPRESENTS THE PATIENT DEATH REPORT. SEE SECTION H10, RELATED REPORT NUMBER 1 FOR ADDITIONAL EVENTS FROM THIS ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THIS REPORT IS FROM A LITERATURE ARTICLE IN WHICH ONE PATIENT DIED AFTER A SERIES OF POST-OPERATIVE COMPLICATIONS INCLUDING A PULMONARY EMBOLISM, BLEEDING FROM THE INFERIOR EPIGASTRIC ARTERY, AND AN IATROGENIC BLADDER INJURY ON A RE-EXPLORATION. NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. CITATION: PILZ DA CUNHA G, DE MEYERE C, D¿HONDT M, ET AL. ROBOTIC LIVER PARENCHYMAL. TRANSECTION USING THE SYNCHROSEAL. SURG ENDOSC 38, 4947¿4955 (2024). HTTPS://DOI.ORG/10.1007/S00464-024-11005-4. B2: PATIENT DATE OF DEATH - THE DATE OF DEATH IS UNKNOWN; THE STUDY WAS CONDUCTED BETWEEN FEBRUARY 2020 TO MARCH 2023. SECTION B3: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE EVENT DATE ASSOCIATED WITH THE ADVERSE EVENT, AN ALTERNATE DATE, PUBLISHED DATE HAS BEEN USED. SECTION D: SUSPECT MEDICAL DEVICE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE DA VINCI SYSTEM ASSOCIATED WITH THE ADVERSE EVENT, A GENERIC SYSTEM MATERIAL NUMBER WAS USED. SECTION E: THE PROCEDURES WERE PERFORMED AT TWO INSTITUTIONS: (B)(6) HOSPITAL.
REMSTAR AUTOA-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A NURSE REPORTED A PERITONEAL DIALYSIS (PD) PATIENT SAT UP TO DISCONNECT FROM TREATMENT ON THE LIBERTY SELECT CYCLER AND SLUMPED OVER. THE PATIENT PASSED AWAY AND CARDIAC ARREST WAS SUSPECTED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). ON (B)(6) 2024, THE PATIENT WAS NOT IN ACTIVE TREATMENT AND FELL WHILE OUTSIDE THE HOME. THE PATIENT WAS EXPERIENCING CHEST AND SHOULDER PAIN. THE PATIENT WAS BROUGHT TO THE EMERGENCY ROOM (ER). AN X-RAY REVEALED THE PATIENT HAD BROKEN HIS COLLAR BONE. NO ADDITIONAL INTERVENTIONS OR TESTING WERE PERFORMED. NO LABS WERE DRAWN. THE PATIENT WAS PLACED IN A SLING AND PRESCRIBED PAIN MEDICATION (UNSPECIFIED). THE PATIENT WAS DISCHARGED. LATER IN THE EVENING, THE PATIENT CONNECTED TO THE LIBERTY SELECT CYCLER FOR TREATMENT. IN THE MORNING ON (B)(6) 2024 THE PD TREATMENT WAS COMPLETE. THE PATIENT WENT TO STAND UP OUT OF THE BED TO DISCONNECT FROM THE LIBERTY SELECT CYCLER AND BEGAN SHAKING. THE PATIENT SUBSEQUENTLY FELL TO THE FLOOR. THE PATIENT COULD NOT BE REVIVED AND WAS PRONOUNCED DECEASED. THE SUSPECTED CAUSE OF DEATH IS CARDIAC ARREST; HOWEVER, THE CORONER HAS NOT COMPLETED THE REPORT. THE PATIENT WAS STILL CONNECTED TO THE CYCLER AT THE TIME OF DEATH. THE PDRN STATED THERE WERE NO ISSUES WITH THE PATIENT¿S PD TREATMENT AND THE PATIENT¿S SPOUSE AND DAUGHTER STATED THE PATIENT DID NOT HAVE ANY ISSUES THROUGHOUT THE NIGHT WITH THE TREATMENT. NO FURTHER INFORMATION WAS PROVIDED. | ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR INVESTIGATION A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. THE TEACH PUMPS TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. THERE WERE VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY AND THE BOTTOM COVER. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED. THERE IS NO ALLEGATION OF DEVICE CONTRIBUTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | Hypersensitivity/Allergic reaction; Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Eye Irritation; Asthma; Dizziness; Pulmonary Dysfunction; Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING),INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS ONE ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | Eye Irritation; Dizziness; Unspecified Respiratory Problem; Headache; Adult Respiratory Distress Syndrome | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, DEATH AND ACUTE RESPIRATORY DISTRESS SYNDROME. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WERE OBSERVED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WERE 22 ERRORS FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX D: PRODUCT UDI WAS CORRECTED IN THIS REPORT. BOX E: REPORTER COUNTRY WAS CORRECTED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, AND CHOLANGIOCARCINOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information; Skin Inflammation/ Irritation; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WERE OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY A YEAR AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAS PASSED AWAY COUPLE OF YEARS AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAS PASSED AWAY COUPLE OF YEARS AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER PREVIOUSLY INCORRECTLY REPORTED HEALTH IMPACT GRID CODE AS C50675. IN THIS FOLLOW UP REPORT HEALTH IMPACT GRID IS UPDATED AS C28554. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | Vomiting; Headache; Nausea; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, NAUSEA, VOMITING AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER BECAME AWARE OF A DREAMSTATION AUTO BIPAP STATING THAT HUSBAND PASSED AWAY AND THAT THERE IS NO ALLEGATION THAT THE DEVICE HAS CONTRIBUTED TO THE DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED. THERE IS NO ALLEGATION OF DEVICE CONTRIBUTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER BECAME AWARE OF A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Eye Irritation; Dizziness; Pulmonary Dysfunction; Inflammation; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem; Eye Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ACUTE MYELOID LEUKEMIA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER, KIDNEY CANCER AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Dizziness; Lymphoma; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER, LYMPHOMA, LUNG DISEASE, DIZZINESS AND HEADACHE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, DEATH, RESPIRATORY ISSUES AND PULMONARY ISSUES. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information; Heart Failure/Congestive Heart Failure; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DEATH, HEART FAILURE AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER BECAME AWARE OF A DREAMSTATION AUTO CPAP STATING THAT HUSBAND PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION AUTO BIPAP DEVICE WAS NO LONGER AVAILABLE AND THE USER PASSED AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION | K131982 | Insufficient Information | Degraded | THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S WIFE STATING THAT HER HUSBAND PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | Eye Irritation; Pulmonary Dysfunction; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | REPAIR EVALUATION INFORMATION WAS RECEIVED ON 12/4/2024 AND H11 IS UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS ONE ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G AND H IS UPDATED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | Headache; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER BECAME AWARE OF A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Respiratory Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER AND RESPIRATORY FAILURE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED WITH INFORMATION THAT THE PATIENT HAD PASSED AWAY. THE PATIENT'S SON CONTACTED THE MANUFACTURER ABOUT A CHECK HE RECEIVED THAT IS IN HIS FATHER'S NAME, SEEKING TO CHANGE THE NAME ON THE CHECK. THAT IS HOW THE MANUFACTURER WAS INFORMED OF THE PATIENT'S DEATH. PER THE PMS CLINICAL EXPERT, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. ADDITIONAL INFORMATION HAS BEEN REQUESTED TO VERIFY IF AN EVENT RELATED TO DEVICE USAGE MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE REP DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT'S WIFE REQUESTED THAT NO FURTHER CORRESPONDENCE BE SENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED WITH INFORMATION THAT THE PATIENT HAD PASSED AWAY - NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED TO VERIFY IF AN EVENT RELATED TO DEVICE USAGE MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED REGARDING BEING CONTACTED WITH INFORMATION THAT THE PATIENT HAD PASSED AWAY - NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED TO VERIFY IF AN EVENT RELATED TO DEVICE USAGE MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE ATTEMPTS WERE UNSUCCESSFUL. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER, WHICH ASSESSED IT AND FOUND NO ERRORS. ADDITIONALLY, NO FAULTS WERE IDENTIFIED, AND THE DEVICE SUCCESSFULLY PASSED ALL TESTS BEFORE BEING RETURNED TO LOAN. NO FURTHER INVESTIGATION WILL BE CONDUCTED CURRENTLY. HOWEVER, IF NEW INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE SUBMITTED.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER WHERE IT WAS INSPECTED. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS LOCATED IN THE LOGS, AND NO FAULTS WERE FOUND. THE DEVICE'S FILTERS, TUBING, AND MANUAL WERE REPLACED AND THE DEVICE SENT TO BE PUT BACK INTO LOAN STOCK. SOFTWARE V1.0.6.620. THE DEVICE PASSED ALL TESTS.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH." DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "NO FAULTS WERE FOUND. ZERO ERRORS LOCATED IN THE LOGS. THE DEVICE'S FILTERS, TUBING AND MANUAL WERE REPLACED. THE DEVICE TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO LOAN STOCK."
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE PASSED AWAY AND THE DEVICE WILL BE RETURNED TO THE MANUFACTURER. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE PASSED AWAY AND THE DEVICE WILL BE RETURNED TO THE MANUFACTURER. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND ERROR CODE E-370 (ERR_FR_ERROR_MISC) A CONTINUE LEVEL-ERROR NOT DUE TO ANY COMPONENT FAILURE. THERE WAS NO VALID COMPLAINT TO BE REPRODUCED. THE DEVICE WAS IN GOOD CONDITION AND WAS WORKING FINE. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER UPDATED BOX H IN THIS REPORT.
REMSTAR AUTO A-FLEX | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND VISIBLE FOAM PARTICLES WERE NOT OBSERVED. ADDITIONALLY, THE PATIENT¿S COMPLAINT WAS NOT CONFIRMED, AND THE DEVICE WAS SCRAPPED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION PATENT WAS DECEASED. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Diaphoresis; Sepsis; Fever; Low Blood Pressure/ Hypotension; Tachycardia; Bowel Perforation; Respiratory Failure; Abdominal Pain; Septic Shock; Elevated ketones/Diabetic Ketoacidosis; Peritonitis | Adverse Event Without Identified Device or Use Problem | ON 19/NOV/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE HOSPITALIZED ON (B)(6) 2024 DUE TO SEPSIS CAUSED BY PERITONITIS. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. MEDICAL RECORDS REVEALED THE PATIENT PRESENTED TO THE EMERGENCY ROOM (ER) ON (B)(6) 2024 WITH SYMPTOMS OF SEPTIC SHOCK INCLUDING: FEVER, DIAPHORESIS, ABDOMINAL PROTUBERANCE/TENDERNESS, TACHYCARDIA (HEART RATE = 124), HYPOTENSION (BLOOD PRESSURE = 98/58), TACHYPNEIC (RATE = 35), ABDOMINAL PAIN (ONSET - 17/NOV/2024), AND CLOUDY EFFLUENT (ONSET - 18/NOV/2024). THE PATIENTS¿ POINT OF CONTACT (POC) REPORTED THE PATIENT RECENTLY UNDERWENT A DIAGNOSTIC LAPAROSCOPY (DATE NOT PROVIDED) FOR THE EVALUATION OF HIS PD CATHETER (NOT A FRESENIUS PRODUCT) AND HAS NOT ¿ENTIRELY BEEN HIMSELF EVER SINCE.¿ HEMATOLOGICAL STUDIES REVEALED THE PATIENT WAS SUFFERING FROM SEVERE METABOLIC ACIDOSIS (LACTIC ACID = 8.6 MMOL/L), ELEVATED LIVER FUNCTION TESTS, AND LEUKOCYTOSIS. THE PATIENT WAS INITIALLY TREATED WITH 2.0 LITERS OF INTRAVENOUS (IV) NORMAL SALINE, WHICH IMPROVED THE LACTIC ACID LEVEL (5.6 MMOL/L) AND BLOOD PRESSURE, HOWEVER THE PATIENT¿S HYPOTENSION RETURNED AND REQUIRED THE ADMINISTRATION OF NOREPINEPHRINE. A PERITONEAL EFFLUENT CELL COUNT WAS ALSO PERFORMED AND WAS POSITIVE FOR GRAM-POSITIVE COCCI, AND A COMPUTED TOMOGRAPHY (CT) SCAN DISCOVERED A BOWEL PERFORATION, SMALL BOWEL ISCHEMIA, ASCITES, PORTAL VENOUS GAS, PNEUMATOSIS OF THE SMALL INTESTINE, AND ALTHOUGH THE TIMELINE FOR COLLECTION IS UNKNOWN, PERITONEAL CULTURES WERE POSITIVE FOR MULTIPLE BACTERIA (ORGANISMS NOT PROVIDED). THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) AND WAS INTUBATED SHORTLY THEREAFTER. THE PATIENT RECEIVED IV FLUIDS, VANCOMYCIN, ZOSYN, AND 3 DIFFERENT VASOPRESSORS (VASOPRESSIN, ANGIOTENSIN II, NOREPINEPHRINE) DUE TO SEVERE HYPOTENSION. A SURGICAL CONSULTATION WAS PLACED FOR THE BOWEL PERFORATION, AND AFTER CONSULTING WITH THE GENERAL SURGEON, THE PATIENT/FAMILY CHOSE TO WITHDRAW CARE AND PROVIDE COMFORT MEASURES ONLY (HOSPICE CARE) DUE TO HIS POOR PROGNOSIS (MULTISYSTEM ORGAN FAILURE, ACUTE RESPIRATORY FAILURE, SEPTIC SHOCK, BOWEL PERFORATION/ISCHEMIA, METABOLIC ENCEPHALOPATHY). LIFE SUPPORT WAS REMOVED LATER THE SAME DAY, AND THE PATIENT EXPIRED AT 2122 ON (B)(6) 2024 WITH FAMILY PRESENT. THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED APPROXIMATELY 24 HOURS PRIOR TO HIS PASSING. PER THE MEDICAL EXAMINER, THE PATIENT¿S DEATH WAS DUE TO NATURAL CAUSES AND NO AUTOPSY WAS PERFORMED. THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS SEPTICEMIA, SECONDARY TO A BOWEL PERFORATION AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. THE PATIENT¿S LIBERTY SELECT CYCLER HAS NOT BEEN RETURNED TO THE MANUFACTURER AT THE TIME OF THIS INVESTIGATION. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXIST BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF BOWEL PERFORATION/ISCHEMIA, SEPTIC PERITONITIS, MULTISYSTEM ORGAN FAILURE, ACUTE RESPIRATORY FAILURE, SEPTIC SHOCK, AND METABOLIC ENCEPHALOPATHY (CHARACTERIZED BY MENTAL STATUS CHANGES, FEVER, ABDOMINAL PROTUBERANCE/TENDERNESS, DIAPHORESIS, TACHYCARDIA, HYPOTENSION, TACHYPNEA, CLOUDY PERITONEAL EFFLUENT FLUID) AND DEATH, WHICH REQUIRED HOSPITALIZATION, ICU ADMISSION, AND INTUBATION. THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED APPROXIMATELY 24 HOURS PRIOR TO HIS PASSING, AND PER THE MEDICAL EXAMINER, THE PATIENT¿S DEATH WAS DUE TO NATURAL CAUSES. THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS SEPTICEMIA, SECONDARY TO A BOWEL PERFORATION, AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. SEPTICEMIA IS A LIFE-THREATENING CONDITION WITH A MORTALITY RATING OF 12-22% AND IS PRIMARILY CAUSED BY SEVERE BACTERIAL INFECTIONS. MEDICAL DOCUMENTATION CONFIRMED THE DECISION WAS MADE TO TRANSITION THE PATIENT INTO HOSPICE CARE ON (B)(6) 2024 AFTER A COMPLICATED HOSPITAL COURSE, AND THE PATIENT EXPIRED AS A RESULT LATER THE SAME DAY. HOSPICE CARE IS CONSIDERED MEDICAL INTERVENTION, WITH THE EXPECTATION BEING THE PATIENT WILL EXPIRE AS A RESULT. THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Insufficient Information | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR SURGERY OR INTERVENTION - UNPLANNED TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 77 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 4.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR COMPLICATION - MAJOR TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 6 VASCULAR COMPLICATION - MAJOR AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 75 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 4 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Insufficient Information | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR SURGERY OR INTERVENTION - UNPLANNED TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 11 VASCULAR SURGERY OR INTERVENTION - UNPLANNED AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 94 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 5 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 11 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 4.73. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR COMPLICATION - MAJOR TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 18 VASCULAR COMPLICATION - MAJOR AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 98 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 10 FEMALE(S) AND 8 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 18 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 42.94. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - HEMATOMA AT ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 2 BLEEDING - HEMATOMA AT ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 4 BLEEDING - ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 79 - 89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 3 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES (B)(4) BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.5. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR COMPLICATION - MAJOR TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 27 VASCULAR COMPLICATION - MAJOR AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 46 - 96 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 13 FEMALE(S) AND 14 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 27 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 41.78. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Retroperitoneal Hemorrhage | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - RETROPERITONEAL TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - RETROPERITONEAL AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Unspecified Vascular Problem | Insufficient Information | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR SURGERY OR INTERVENTION - UNPLANNED TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 19 VASCULAR SURGERY OR INTERVENTION - UNPLANNED AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 59 - 96 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 10 FEMALE(S) AND 9 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 19 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.26. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - HEMATOMA AT ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 53 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 8 BLEEDING - ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 59 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 4 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER (B)(4) AND SUMMARIZES 8 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.5. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4) . THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | Hematoma | Adverse Event Without Identified Device or Use Problem | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR BLEEDING - HEMATOMA AT ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE MITRAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 6.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
DREAMSTATION AUTO CPAP | K131982 | Dizziness; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, CANCER, DEATH AND BRAIN GLIOBLASTOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN BOX E. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. IN BOX E: REPORTER COUNTRY HAS BEEN CORRECTED.
REMSTAR PRO C-FLEX+ | K091319 | Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, KIDNEY DISEASE/ TOXICITY AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Headache; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, HYPERSENSITIVITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A LETTER RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem; No Apparent Adverse Event | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. EVALUATION OF THE DEVICE INDICATED THE DEVICE PASSED AND ALL TESTING WITHOUT ANY FAULTS FOUND. BASED ON THE SERVICE EVALUATION OF THE DEVICE, THERE IS NO EVIDENCE INDICATING A REPORTABLE MALFUNCTION OCCURRED OR THAT THERAPY WAS AFFECTED. THE INVESTIGATION FINDINGS DID NOT UNCOVER ANY DETAILS THAT WOULD CHANGE OR MODIFY THE RISK OF THE DEVICE.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A LETTER RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS FOUND DECEASED BY A SECURITY GUARD IN THE PATIENT'S APARTMENT COMPLEX. THE GUARD HEARD THE CYCLER ALARMING AND CONDUCTED A WELFARE CHECK ON THE PATIENT. THE PATIENT WAS ALREADY DECEASED WHEN FOUND. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER. THE CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN IF AN AUTOPSY WILL BE PERFORMED. THE PDRN STATED THAT THE PATIENT¿S DAUGHTER HAS NOT ANSWERED THEIR INQUIRIES. NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER. IT WAS REPORTED THAT PRIOR TO THE PATIENT¿S DEATH, THE PATIENT WAS FEELING WEAK AND FATIGUED. THE PATIENT HAD AN EVENT OF COVID AND PERITONITIS IN OCTOBER AND NEVER FULLY RECOVERED FROM FEELING SICK. NO ADDITIONAL INFORMATION WAS AVAILABLE. | THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. THE TEACH PUMPS TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. THERE WERE VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY AND THE BOTTOM COVER. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S). UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
REMSTAR AUTOA-FLEX | K131982 | Hypersensitivity/Allergic reaction; Eye Irritation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, HYPERSENSITIVITY, LUNG DISEASE, DEATH, OTHER AND LOST SENSE OF TASTE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DIED IN (B)(6) 2024. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER WHERE IT WAS INSPECTED. EVALUATION RESULTS DETERMINED NO ERRORS WERE FOUND. THE TUBING, MANUAL, AND FILTERS WERE REPLACED. ADDITIONALLY, THE DEVICE WAS CLEANED AND PASSED ALL TESTS.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEWED BY THE CLINICAL EXPERT AND THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Headache; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DIZZINESS AND/OR HEADACHE, CANCER. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY AND A RETURN LABEL WAS BEING REQUESTED. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT YET BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY AND A RETURN LABEL WAS BEING REQUESTED. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND NO ERRORS PRESENT. THERE WAS NO EVIDENCE OF FOAM PARTICLES OR DEGRADATION. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT FOR THE DEVICE EVALUATION.
REMSTAR PRO C-FLEX+ | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WERE 24 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REP DREAMSTATION AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTOA-FLEX | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K091319 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTOA-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973, AND Z-1974.
REMSTAR AUTOA-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY WHILE WAITING FOR THE REPLACEMENT DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973, AND Z-1974.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT TO ENROLL RECALLED DEVICE FOR THE CLASS ACTION LAWSUIT SETTLEMENT CLAIM. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Insufficient Information; No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT HER HUSBAND'S NAME TO BE REMOVED FROM EMAIL AND RECALL LIST. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT HER HUSBAND'S NAME TO BE REMOVED FROM EMAIL AND RECALL LIST. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE CORRECT B5 SHOULD BE- THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT HER HUSBAND'S NAME TO BE REMOVED FROM EMAIL AND RECALL LIST. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. OUTCOMES ATTRIBUTED TO AE SECTION HAS BEEN CORRECTED FROM OTHER TO DEATH. APPROPRIATE CODE UPDATED/CORRECTED. IN THE INITIAL REPORT IT WAS WRONGLY REPORTED UNDER RECALL, CORRECTED IN THIS FOLLOW-UP REPORT. THIS REPORT IS NOT A PART OF RECALL.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY 2 YEARS AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K091319 | Heart Failure/Congestive Heart Failure; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CONGESTED HEART FAILURE, BREATHING ISSUES WHICH LED TO THE DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DREAMSTATION AUTO CPAP DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. THE DEVICE WAS DISCARDED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY (B)(6) 2022. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE IS NOT AVAILABLE TO BE RETURNED FOR INVESTIGATION. THE USER'S WIFE STATES SHE NO LONGER HAS CPAP DEVICE AS SOMEONE CAME AND PICKED IT UP BEFORE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Insufficient Information | Degraded | THE MANUFACTURER BECAME AWARE THAT A USER OF THE REP DREAMSTATION AUTO CPAP HAD STATES DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K183387 | Insufficient Information | Device Alarm System | IT WAS REPORTED THE PATIENT PASSED AWAY AND THE CUSTOMER REQUESTED ASSISTANCE IN REVIEWING THE LOGS. THE STAFF REPORTED THEY DID NOT SEE A YELLOW LEADS OFF ALARM AND THE PATIENT PASSED AWAY. THE SETTINGS AND AUDIT LOGS WERE REVIEWED WITH THE NURSE MANAGER, REVEALING THE LEADS OFF ALARM WAS SET TO YELLOW AND INOPS WERE SET TO RE-ALARM. A YELLOW ALARM FOR LEADS OFF WAS GENERATED AND ACKNOWLEDGED AND RE-ALARMED FIVE TIMES. NO DEVICES WERE TAKEN OUT OF SERVICE AND THE MANAGER DID NOT FEEL THE DEVICE WAS AT FAULT. | THE REMOTE CLINICAL SUPPORT SPECIALIST REVIEWED THE AUDIT LOGS AND DEVICE SETTINGS WITH THE NURSE MANAGER, AND CONFIRMED THE LEADS OFF ALARM WAS SET TO YELLOW AND INOPS WERE SET TO RE-ALARM. A YELLOW ALARM FOR LEADS OFF WAS NOTED IN THE AUDIT LOGS AND IT WAS ACKNOWLEDGED AND RE-ALARMED 5 TIMES. THE MANAGER DID NOT FEEL THE MONITORING SYSTEM WAS AT FAULT; THEREFORE, THE DEVICE REMAINS IN SERVICE. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED WE WERE UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DAVINCI XI | K131861 | Insufficient Information | Insufficient Information | IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNEVENTFUL DA VINCI-ASSISTED INGUINAL HERNIA REPAIR WITHOUT ANY INTRA-OPERATIVE COMPLICATIONS. THE PATIENT WAS NOTED TO HAVE EXTENSIVE ADHESIONS AND ADHESIOLYSIS WAS PERFORMED OR BEDSIDE BEFORE PLACING PORTS FOR THE ROBOTIC PROCEDURE. POST-OPERATIVELY, THE PATIENT WAS DISCHARGED HOME AS PLANNED. THE SURGEON INFORMED THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE (CSR) THAT SEVERAL DAYS AFTER BEING DISCHARGED, THE PATIENT FELL WHILE AT HOME. THE PATIENT WAS ADMITTED TO A DIFFERENT HOSPITAL THAN WHERE THE DA VINCI PROCEDURE WAS PERFORMED AND EXPIRED SEVERAL DAYS AFTER BEING HOSPITALIZED. THE SURGEON OF THE DA VINCI PROCEDURE DID NOT KNOW THE PATIENT'S CAUSE OF DEATH AND STATED THAT THE FALL EVENT WAS NOT RELATED TO THE DA VINCI ROBOT. THE SURGEON STATED THAT IT WAS POSSIBLE THE PATIENT SUSTAINED A LIVER LACERATION AS A RESULT OF THE FALL; HOWEVER, IT WAS NOT PROVIDED HOW THAT WAS THEORIZED. THE SURGEON DECLINED TO PROVIDE ANY ADDITIONAL INFORMATION, INCLUDING THE NAME OF THE TREATING PHYSICIAN AT THE OTHER HOSPITAL. | NO SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE WAS PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. ON 11/19/2024, THE SURGEON STATED THAT THE PROCEDURE HAD BEEN PERFORMED ¿WITHIN THE LAST 2 MONTHS¿. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT AN INGUINAL HERNIA REPAIR AND FELL AT HOME A FEW DAYS LATER, WAS ADMITTED TO THE HOSPITAL, AND EVENTUALLY DIED A FEW DAYS LATER FOR REASONS THAT ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | ON 4/DEC/2024, IT WAS REPORTED THAT THIS PATIENT ON PERITONEAL DIALYSIS (PD) EXPIRED WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE STATED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF END STAGE RENAL DISEASE (ESRD), ATHEROSCLEROTIC HEART DISEASE WITH PREVIOUS MYOCARDIAL INFARCTION (UNRELATED TO DIALYSIS), HYPERLIPIDEMIA, DIABETES MELLITUS AND ULCERATIVE COLITIS. ADDITIONALLY, THE NURSE REPORTED THAT THE PATIENT WAS IMMOBILE AND IN POOR HEALTH RECEIVING DIALYSIS TREATMENT IN A NURSING HOME. PER THE NURSE, THE PATIENT ALSO HAD AN INDWELLING FOLEY CATHETER AND THAT THE PATIENT HAD SYMPTOMS OF A URINARY TRACT INFECTION SINCE (B)(6) 2024. HOWEVER, IT WAS REPORTED THAT THE PATIENT WAS NOT ON ANTIBIOTIC THERAPY. THE PD NURSE STATED THAT ON (B)(6) 2024 THE PATIENT BEGAN PD TREATMENT WITH THE FRESENIUS CYCLER (SERIAL NUMBER (B)(6)) WITHOUT ANY ISSUES IN THE NURSING HOME. AT APPROXIMATELY 5:30AM (ON (B)(6) 2024), THE PATIENT WAS ALIVE RECEIVING HER DIALYSIS OBSERVED ON MORNING NURSING CHECK. SUBSEQUENTLY, AT APPROXIMATELY 8:30 AM, THE PATIENT WAS FOUND UNRESPONSIVE. THE NURSE STATED THE PATIENT WAS IN THE LAST DRAIN OF PD TREATMENT AND IT IS UNKNOWN IF THE TREATMENT WAS COMPLETED. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION (CPR). HOWEVER, CPR WAS NOT SUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED IN THE NURSING HOME (NOT BROUGHT TO THE HOSPITAL). THE NURSE REPORTS THAT THE CAUSE OF DEATH HAS NOT BEEN FIRMLY ESTABLISHED. THE PATIENT¿S ESRD DEATH FORM IS NOT AVAILABLE. HOWEVER, THE NURSE INDICATED THAT IT IS SUSPECTED THE PATIENT MAY HAVE BEEN SEPTIC DUE TO CONCURRENT URINARY TRACT INFECTION WITH THE INDWELLING FOLEY CATHETER. THE NURSE CONFIRMED THERE WERE NO ISSUES WITH PD TREATMENT OR PRODUCT IN RELATION TO THE DEATH. THE CYCLER IS IN THE PD CLINIC PENDING RETURN TO MANUFACTURER. | ADDITIONAL INFORMATION: B7.
ALARIS SYSTEM | K133532 | Insufficient Information | Inaccurate Delivery; Use of Device Problem | IT WAS REPORTED THAT THE PATIENT RECEIVED PCA (PATIENT-CONTROLLED ANALGESIA) DOSE "OUTSIDE THE GUARDRAILS LIMIT" AND THE EVENT ALLEGEDLY RESULTED TO DEATH. PER REPORT, THE PCA INFUSION SET UP WAS NOT CONNECTED TO AN END-TIDAL CO2 (ETCO2) MONITOR AT THE TIME OF THE EVENT THEREFORE THE PATIENT'S ETCO2 LEVEL WAS NOT MONITORED. THIS WAS REPORTED TO A BD FIELD SPECIALIST WHILE THEY WERE ON THE CUSTOMER'S FACILITY WORKING ON A PROJECT. PER REPORT, THE FACILITY HAS SINCE CHANGED THEIR POLICY REQUIRING USE OF ETCO2 MONITOR WHENEVER PCA MODULE IS IN USE. PER CUSTOMER (BIOMEDICAL MANGER), "I HAVE BEEN TOLD THAT THIS HAS BEEN RESOLVED, AND NO FURTHER INVESTIGATION IS NEEDED." MULTIPLE COMMUNICATIONS HAVE BEEN MADE REQUESTING FOR ADDITIONAL INFORMATION AND PRODUCT RETURN. NO RESPONSE HAS BEEN PROVIDED TO DATE. | OMIT: C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. CORRECTION: DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES.
ALARIS SYSTEM | K043299 | Insufficient Information | Inaccurate Delivery; Use of Device Problem | IT WAS REPORTED THAT THE PATIENT RECEIVED PCA (PATIENT-CONTROLLED ANALGESIA) DOSE "OUTSIDE THE GUARDRAILS LIMIT" AND THE EVENT ALLEGEDLY RESULTED TO DEATH. PER REPORT, THE PCA INFUSION SET UP WAS NOT CONNECTED TO AN END-TIDAL CO2 (ETCO2) MONITOR AT THE TIME OF THE EVENT THEREFORE THE PATIENT'S ETCO2 LEVEL WAS NOT MONITORED. THIS WAS REPORTED TO A BD FIELD SPECIALIST WHILE THEY WERE ON THE CUSTOMER'S FACILITY WORKING ON A PROJECT. PER REPORT, THE FACILITY HAS SINCE CHANGED THEIR POLICY REQUIRING USE OF ETCO2 MONITOR WHENEVER PCA MODULE IS IN USE. PER CUSTOMER (BIOMEDICAL MANGER), "I HAVE BEEN TOLD THAT THIS HAS BEEN RESOLVED, AND NO FURTHER INVESTIGATION IS NEEDED." MULTIPLE COMMUNICATIONS HAVE BEEN MADE REQUESTING FOR ADDITIONAL INFORMATION AND PRODUCT RETURN. NO RESPONSE HAS BEEN PROVIDED TO DATE. | OMIT: C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. CORRECTION: DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 01/28/2025 AND H11 IS UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED BUT THERE WAS NO MENTION WHETHER THE DEVICE WILL BE AVAILABLE FOR EVALUATION, THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. AFTER DEVICE RETURN AND COMPLETION OF EVALUATION, A FOLLOW-UP REPORT WILL BE FILED. SECTION G AND H IS UPDATED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION PATENT PASSED AWAY. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. | THE MANUFACTURER HAS PREVIOUSLY FILED THIS REPORT AS REP DEVICE, BUT DEVICE INFORMATION HAS BEEN UPDATED. HENCE A CORRECTION REPORT HAS BEEN FILED. BOX D HAS BEEN UPDATED AND CORRECTED.
DREAMSTATION | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE PASSED AWAY AND THE PATIENT¿S WIFE WANTS TO RETURN IT. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE PATIENT HAS ALLEGED PATIENT DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REP DREAMSTATION AUTO BIPAP. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Hypoxia | Degraded | THE MANUFACTURER WAS CONTACTED REGARDING THE VOLUNTARY FIELD SAFETY NOTICE/RECALL CONCERNING THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER HAS BEEN INFORMED THAT A USER PASSED AWAY DUE TO A LACK OF OXYGEN TO THE BRAIN, WHICH WAS ALLEGEDLY CAUSED BY NOT USING THE AFFECTED DEVICE. THE USER HAD NOT BEEN USING THE DEVICE FOR TWO YEARS AND PASSED AWAY IN 2023. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K091319 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
PATIENT INFORMATION CENTER IX | K211900 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THE CUSTOMER WANTED TO KNOW IF THE EQUIPMENT HAS AN OPTIONAL FUNCTION FOR EGC ANALYSIS AND ALARMING DURING POSTOPERATIVE CARE AND CARDIAC ARREST. IT WAS INDICATED THAT A PATIENT DEATH OCCURRED WHILE THE PATIENT WAS CONNECTED TO THE MX40 WHICH WAS CONNECTED TO THE PIC IX. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE CONTACTED RECALL SUPPORT AND STATED THAT HER HUSBAND HAD A RECALLED PHILIPS DEVICE AND NEVER GOT A REPLACEMENT FOR IT. SHE SAID THE PATIENT HAS BEEN DECEASED FOR 2 YEARS AND DOES NOT NEED A REPLACEMENT ANY LONGER. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY AND THE PATIENT¿S WIFE WOULD LIKE THE LETTERS TO STOP. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THIS ISSUE HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THIS ISSUE HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AS PART OF THE COMPLAINT HANDLING PROCESS, ATTEMPT TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION HAS BEEN UNSUCCESSFUL BECAUSE THERE IS NO CUSTOMER CONTACT INFORMATION AVAILABLE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI XI | K131861 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING A DA VINCI ASSISTED URETERAL REIMPLANTATION PROCEDURE, AN UNSPECIFIED ARTERY WAS DAMAGED AND THE PROCEDURE WAS CONVERTED TO OPEN; THE PATIENT EXPIRED IN THE OPERATING ROOM. THE PATIENT HAD A HISTORY OF RETROPERITONEAL FIBROSIS. THE SITE ROBOTICS COORDINATOR INFORMED AN INTUITIVE CLINICAL REPRESENTATIVE THAT THE PROCEDURE WAS NEAR COMPLETION WHEN THE COMPLICATION OCCURRED. IT WAS REPORTED THAT THE ARTERY WAS DAMAGED WHEN THE SURGEON WAS ATTEMPTING TO DISSECT VERY DIFFICULT ADHESIONS THAT WERE STUCK TO THE URETER. A DA VINCI GRASPING INSTRUMENT WAS USED TO CLAMP THE VESSEL TO CONTROL THE BLEEDING WHILE THE PROCEDURE WAS CONVERTED TO OPEN. THE PATIENT ULTIMATELY EXPIRED IN THE OPERATING ROOM. THE ROBOTICS COORDINATOR STATED THAT THE COMPLICATIONS WERE NOT RELATED TO THE DA VINCI PRODUCTS IN USE. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, LARGE NEEDLE DRIVER, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A BLEEDING EVENT FROM AN ARTERY DURING A DIFFICULT DISSECTION NEAR THE URETER. WHILE THE SPECIFIC VESSEL IS NOT MENTIONED, THE MOST LIKELY MAJOR VESSELS IN THIS AREA INCLUDE THE COMMON AND EXTERNAL ILIACS. OTHER POSSIBLE LARGE VESSELS MAY INCLUDE THE INTERNAL ILIACS, AORTA, AND SACRAL VESSELS. THERE WAS DIFFICULTY DURING DISSECTION DUE TO PATIENT DISEASE AND ADHESIONS WHICH MAY HAVE CONTRIBUTED TO THIS INJURY. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | Air Embolism; Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PATIENT EXPIRED POST-OPERATIVELY AFTER A 5.5 HOUR DA VINCI-ASSISTED PARTIAL HEPATECTOMY - WEDGE RESECTION. A SURGEON INFORMED THE INTUITIVE CLINICAL SALES REPRESENTATIVE THAT THE PATIENT OUTCOME AND POST-OPERATIVE COMPLICATIONS WERE NOT RELATED TO THE DA VINCI PRODUCTS; THERE WERE NO ISSUES WITH THE DA VINCI PRODUCTS USED DURING THE PROCEDURE. ONE ASSISTING PHYSICIAN REPORTED THAT AN UNSPECIFIED VEIN WAS DAMAGED WHILE REMOVING A PORTION OF THE LIVER. THE ROBOTIC PORTION OF THE PROCEDURE WAS NOT YET COMPLETE WHEN THE INJURY OCCURRED. THE ESTIMATED BLOOD LOSS WAS UNKNOWN; BLOOD TRANSFUSIONS WERE ADMINISTERED; IT WAS UNKNOWN WHAT OTHER INTERVENTIONS WERE PERFORMED. IT WAS REPORTED THAT THE PATIENT ALSO EXPERIENCED A GAS EMBOLISM; TREATMENT FOR THE GAS EMBOLISM WAS UNKNOWN. THE PHYSICIAN DID NOT KNOW IF THE PATIENT¿S DETERIORATING POST-OPERATIVE CONDITION WAS RELATED TO THE INTRA-OPERATIVE VASCULAR INJURY, OR THE GAS EMBOLISM. ANOTHER ASSISTING SURGEON REPORTED THAT THERE WAS NO UNEXPECTED VASCULAR DAMAGE DURING THE PROCEDURE, THE BLEEDING WAS WITHIN THE EXPECTED RANGE FOR THE PROCEDURE, AND THE BLOOD TRANSFUSIONS ADMINISTERED WERE NOT DUE TO UNEXPECTED BLEEDING. THE PATIENT DEVELOPED A POST-OPERATIVE GAS EMBOLISM, BUT THE DETAILS OF THE PATIENT¿S POST-OPERATIVE COURSE WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND THAT NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THIS REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MONOPOLAR CURVED SCISSORS, MEDIUM-LARGE CLIP APPLIER, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, LARGE SUTURE CUT NEEDLE DRIVER, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING A PARTIAL HEPATECTOMY. AN INJURY OCCURRED WHILE EXTRACTING THE SPECIMEN. THE BLEEDING THAT OCCURRED AND THE IMPACT IT MAY HAVE HAD IS UNCLEAR, AND THERE ARE SOMEWHAT CONFLICTING ACCOUNTS AS TO THE DETAILS OF THIS BLEED. IT IS UNCLEAR HOW OR IF THAT EVENT CONTRIBUTED TO THE EVENTUAL DEMISE OF THE PATIENT. WHILE THE CAUSE OF DEATH IS ALSO NOT CONFIRMED, THERE WAS SUGGESTION THAT A GAS EMBOLISM MAY HAVE OCCURRED. GAS EMBOLISM CAN OCCUR AT THE TIME OF MINIMALLY INVASIVE SURGERY DUE TO INSUFFLATION OF CO2. THERE ARE NO DETAILS AS TO HOW A GAS EMBOLISM OCCURRED NOR THE DETAILS OF HOW IT WAS DIAGNOSED. IT IS ALSO UNKNOWN IF OR HOW THIS MAY HAVE CONTRIBUTED TO THE EVENTUAL DEATH OF THE PATIENT. THERE IS NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, CANCER, AND DEATH. NO FURTHER CLINICAL DETAILS OR MEDICAL INTERVENTION WERE REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. THE MANUFACTURER RECEIVED NEW AND RELEVANT INFORMATION ON 02/14/2024. THE DEVICE WAS RETURNED TO THE SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE, THERE WAS NO VISIBLE FOAM PARTICLES FOUND. DUST CONTAMINATION FOUND DURING THE DEVICE EVALUATION. THE DEVICE WAS REPAIRED. SECTION G3 HAS BEEN UPDATED. IN ADDITION, APPROPRIATE CODE UPDATED/CORRECTED.
REMSTAR PLUS | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Insufficient Information | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | Cancer | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION CPAP PRO DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Unspecified Heart Problem; Unspecified Mental, Emotional or Behavioural Problem; Unspecified Respiratory Problem; Cancer; Septic Shock | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PANCREATIC AND PROSTATE CANCER, CORONARY ARTERY DISEASE, ACUTE HYPOXIC RESPIRATORY FAILURE WITH ASSOCIATED ALTERED MENTAL STATUS, SEPTIC SHOCK AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE IS NO REPORT OF MEDICAL INTERVENTION BEING REQUIRED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE MANUFACTURER SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.
DAVINCI XI | K131861 | Insufficient Information | Insufficient Information | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY PROCEDURE, THE PATIENT EXPIRED. THE PROCEDURE DATE WAS UNKNOWN, AND THE CAUSE OF DEATH WAS NOT PROVIDED. THE DA VINCI PROCEDURE WAS CONVERTED TO OPEN; THE SITE OPERATING ROOM MANAGER REPORTED THAT THE CONVERSION WAS NOT RELATED TO A DEVICE ISSUE OR A DA VINCI PRODUCT MALFUNCTION. THE REASON FOR THE CONVERT TO OPEN WAS NOT PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED; NO RESPONSE HAS BEEN RECEIVED. | NO SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE WAS PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A ROBOTIC CHOLECYSTECTOMY WHICH REQUIRED CONVERSION TO AN OPEN PROCEDURE. NO INFORMATION AS TO THE REASONS FOR THE CONVERSION WERE PROVIDED. THE PATIENT SUBSEQUENTLY DIED BUT NO DETAILS AS TO THE EVENTS THAT LED TO THEIR DEATH OR HOW THEY DIED WERE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
REMSTAR AUTO A-FLEX | K091319 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Asthma; Insufficient Information; Dizziness; Nausea; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), CANCER, OTHER (UNSPECIFIED EVENT) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE
DREAMSTATION AUTO CPAP | K131982 | Vomiting; Dizziness; Nausea; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGATION OF ASTHMA (NEW OR WORSENING); LIVER DISEASE/TOXICITY; LUNG DISEASE; DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Insufficient Information | Degraded | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER BECAME DUE TO A USER OF DREAMSTATION CPAP. DAUGHTER CALLED IN STATING PATIENT PASSED AWAY LAST (B)(6) 2023. DID NOT PASS DUE TO THE USE OF CPAP. SHE WANTS TO STOP ALL MAIL. DON'T KNOW WHERE THE CPAP DEVICE WENT THAT WAS RECALLED AND WILL NOT BE RETURNING FOR EVALUATION. REPLACEMENT DEVICE WAS DONATED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | Insufficient Information; Unspecified Heart Problem | Degraded | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | IN THE PREVIOUS REPORT PATIENT OUTCOME GRID WAS NOT ADDED. IN THIS REPORT PATIENT OUTCOME CODE GRID (BOX H) HAS BEEN UPDATED OR CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY AND HE HAD LUNG CANCER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K091319 | Eye Irritation; Asthma; Dizziness; Pulmonary Dysfunction; Inflammation; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Sore Throat; Cough | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF KIDNEY DISEASE/TOXICITY, CANCER, DEATH, SORE THROAT AND COUGH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN BOX B. AFTER REVIEW, IT WAS DETERMINED THAT BOX B SHOULD BE CORRECTED. OUTCOMES ATTRIBUTED TO AE HAS BEEN CORRECTED AS AND REPORTED AS BELOW.
REMSTAR PRO C-FLEX+ | K131982 | Vomiting; Eye Irritation; Dizziness; Nausea; Pulmonary Dysfunction; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache | Material Integrity Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS HEADACHE, NAUSEA, VOMITING, KIDNEY DISEASE REDUCED CARDIOPULMONARY RESERVE AND DEATH PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Eye Irritation; Dizziness; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Insufficient Information; Headache | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE (UNSPECIFIED), DEATH AND OTHER (UNSPECIFIED EVENT). THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ALARIS SYSTEM | K133532 | Insufficient Information | Break; Inaccurate Delivery; Use of Device Problem; Contamination | IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON 06 DECEMBER 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915. | OMIT : G07001 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. ADDITIONAL INFORMATION: IMDRF ANNEX G CODE.
ALARIS SYSTEM | K133532 | Insufficient Information | Corroded; Inaccurate Delivery; Use of Device Problem; Contamination | IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON 06 DECEMBER 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915. | OMIT: G07001 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. ADDITIONAL INFORMATION: IMDRF ANNEX G CODE.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE CALLER INDICATED THAT THE DEVICE IS NO LONGER IN THEIR POSSESSION. NO ADDITIONAL DETAILS WERE PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | Sleep Dysfunction | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION CPAP PRO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S WIFE STATES ON BEHALF OF HER LATE HUSBAND THAT HER HUSBAND HAD SLEEP APNEA, AND HE PASSED AWAY TWO YEARS AGO. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | IN THIS REPORT, SECTION BOX H: DEVICE MANUFACTURERS (EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID) HAS BEEN UPDATED/CORRECTED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH." THIS ISSUE HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS LOCATED IN THE LOGS, AND NO FAULTS WERE FOUND. THE DEVICE'S FILTERS, TUBING, AND MANUAL WERE REPLACED AND THE DEVICE SENT TO BE PUT BACK INTO LOAN STOCK.
PHILIPS CPAP DEVICE | K131982 | Insufficient Information; No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER BECAME AWARE THAT A USER OF THE PHILIPS CPAP DEVICE HAD PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | (DEVICE) PROBLEM CODE GRID AND PATIENT OUTCOME CODE GRID HAVE BEEN CORRECTED IN THIS FOLLOW-UP REPORT.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information; No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: PATIENT OUTCOME CODE GRID, EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED. BOX B: ADVERSE EVENT/PRODUCT PROBLEM HAS BEEN UPDATED.
DREAMSTATION CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Skin Inflammation/ Irritation; Eye Irritation; Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LIVER DISEASE/TOXICITY, AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PHILIPS CPAP DEVICE | K131982 | Eye Irritation; Appropriate Term / Code Not Available; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATENT HAS EXPIRED DUE TO EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER AND OTHER PANCREATIC ISSUE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Cancer; Liver Damage/Dysfunction; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information; No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H6: PATIENT OUTCOME CODE GRID WAS CORRECTED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION AND OTHER (UNSPECIFIED INFORMATION). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Asthma; Unspecified Heart Problem; Dizziness; Nausea; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA / VOMITING, ASTHMA, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION AUTO BIPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
REMSTAR AUTOA-FLEX | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. IT IS UNKNOWN IF THE PATIENT WAS USING THE DEVICE AT THE TIME OF DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Thromboembolism | Output Problem | REFER TO H11 FOR FOLLOW-UP INFORMATION. | THIS PATIENT'S SURGICAL PROCEDURE WAS A SCHEDULED NON-ROBOTIC SURGERY; THE REPORTED PATIENT EVENT WAS ALLEGED TO BE INDIRECTLY RELATED BY THE SURGEON WHO REPORTED THAT THE PATIENT SUFFERED A PROBABLE THROMBOEMBOLIC ACCIDENT EVENT RELATED TO INTERRUPTION OF THE THE PATIENT¿S ANTICOAGULANT TREATMENT, THEN RESUMED FOLLOWING THE SURGEON'S EARLIER DA VINCI IESU ERBE EVENT, LEADING TO POSTPONEMENT OF THEIR SURGERY. BASED ON THE INVESTIGATION, THERE IS NO INDICATION THAT THE DEVICE DIRECTLY CAUSED THE REPORTED ADVERSE EVENT; THE CASCADE LED TO THE SURGEON'S SUBSEQUENT PROCEDURE(S) BEING POSTPONED BY THE SURGEON. PLEASE SEE SECTION H10 RELATED MEDWATCH REPORT FOR THE IESU ERBE DEVICE USED DURING A DIFFERENT PATIENT PROCEDURE. AN INTUITIVE SURGICAL, INC. (ISI) FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER SITE TO INVESTIGATE THE SURGEON'S UNRELATED EARLIER ROBOTIC CASE WHERE THE ERBE GENERATOR WAS REPLACED AS A RESULT OF MONOPOLAR POWER ACTIVATION PROBLEMS AND C-35 ERRORS NOTED DURING TESTING. ISI DID RECEIVE THE IESU ERBE GENERATOR FOR FAILURE ANALYSIS. DURING THE POWER-ON TEST, IESU CAUTERY ACTIVATION, AND VISUAL INSPECTION, THE ERROR WAS FOUND, CONFIRMING THE FAULT OCCURRED IN THE FIELD. UPON VISUAL INSPECTION, NO ISSUES WERE FOUND THAT WOULD BE RELATED TO THE REPORTED EVENT. THE COMPLAINT WAS CONFIRMED BASED ON THE FIELD EVALUATION AND FAILURE ANALYSIS, WHICH INDICATES THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE CUSTOMER REPORTED ISSUE. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THE FOLLOWING POSSIBLY RELATED SYSTEM ERRORS: C-34 - ENERGY ACTIVATION HALTED BY AN INSTRUMENT OR INSTRUMENT CABLE CONNECTED TO THE ERBE WHILE USING THE COAG FUNCTION. THE PROBABLE CAUSE OF THE ERRORS IS ATTRIBUTED TO A FAULTY ELECTRICAL COMPONENT INSIDE THE IESU. ONE ERROR INDICATES A LOWER VOLTAGE THAN EXPECTED DURING ENERGY ACTIVATION. THE OTHER ERROR INDICATES A HIGHER CURRENT MEASUREMENT VALUE THAN EXPECTED WHILE THE UNIT IS IN AN OFF STATE. ADDITIONAL B, C, D IMDRF ANNEX CODES ADDED REFLECT THE ISSUE THAT OCCURRED WITH THE IESU ERBE DURING A DIFFERENT SURGERY, RESULTING IN THIS PATIENT'S NON-ROBOTIC SURGERY BEING POSTPONED. C20 - NO FINDINGS AVAILABLE AND D15 - CAUSE NOT ESTABLISHED REFER TO THE IMPACT FOR THE PATIENT IN THIS REPORT.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information; Pulmonary Dysfunction; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, KIDNEY DISEASE/TOXICITY AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Vomiting; Dizziness; Nausea; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, VOMITING, LIVER DISEASE/TOXICITY AND CANCER, DEATH AND NAUSEA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Eye Irritation; Asthma; Unspecified Heart Problem; Pulmonary Dysfunction; Unspecified Respiratory Problem; Pneumonia | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), REDUCED CARDIOPULMONARY RESERVE, DEATH, HEART CONDITION, AND PNEUMONIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR AUTO A-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | IN PREVIOUS REPORT, THE MANUFACTURER REPORTED INCORRECT INFORMATION IN BOX B: COMMON DEVICE NAME. AFTER REVIEWED, IT WAS CAPTURED CORRECTLY. IN BOX B: COMMON DEVICE NAME HAS BEEN UPDATED.
DREAMSTATION CPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Dizziness; Nausea; Pulmonary Dysfunction; Inflammation; Skin Inflammation/ Irritation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, THROAT, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information; Dizziness; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Pulmonary Dysfunction; Respiratory Failure; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Inflammation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND ESOPHAGEAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Upper Respiratory Tract Irritation; Cough; Unspecified Kidney or Urinary Problem; Sinusitis; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, COUGH, SINUS INFECTION AND UPPER AIRWAY IRRITATION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT PASSED AWAY. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Asthma; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), LUNG DISEASE AND DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Hypersensitivity/Allergic reaction; Eye Irritation; Dizziness; Pulmonary Dysfunction; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/ OR HEADACHE, HYPERSENSITIVITY, LIVER DISEASE/ TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, LUNG CANCER AND LIVER CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Bronchitis; Pulmonary Dysfunction; Pneumonia | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CHRONIC BRONCHITIS, RECURRING PNEUMONIA, COPD. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Skin Inflammation/ Irritation; Eye Irritation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ESOPHAGEAL ADAMOU CARCINOMA, EYE IRRITATION, SKIN IRRITATION CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Eye Irritation; Dizziness; Unspecified Respiratory Problem; Insufficient Information; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Hypersensitivity/Allergic reaction; Eye Irritation; Dizziness; Pulmonary Dysfunction; Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Insufficient Information; Headache; Epistaxis | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, DEATH, NOSE BLEEDS, LUNG INFLAMMATION AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Eye Irritation; Dizziness; Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Insufficient Information; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Unspecified Kidney or Urinary Problem; Respiratory Failure; Eye Irritation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, KIDNEY DISEASE/TOXICITY AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Dizziness; Inflammation; Unspecified Respiratory Problem; Cancer; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Myocardial Infarction | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING HEART ATTACK AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Vomiting; Eye Irritation; Insufficient Information; Dizziness; Nausea; Inflammation; Skin Inflammation/ Irritation; Breast Cancer; Unspecified Respiratory Problem; Cancer; Headache; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER, OTHER (UNSPECIFIED EVENT), BREAST, MULTIPLE MYELOMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Inflammation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, LUNG DISEASE, CANCER AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Skin Inflammation/ Irritation; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, CANCER, DEATH, CANCER, BLADDER AND PROSTATE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Dyspnea | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND SHORTNESS OF BREATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Cancer; Unspecified Respiratory Problem | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED/RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF DEATH, LUNG DISEASE AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H: TYPE OF REPORTED COMPLAINT WAS CAPTURED SERIOUS INJURY IN PREVIOUS REPORT. TYPE OF REPORTED COMPLAINT WAS CORRECTED.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Asthma; Cancer; Unspecified Kidney or Urinary Problem; Lymphoma | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, LYMPHOMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Headache; Dizziness; Unspecified Kidney or Urinary Problem; Inflammation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Adult Respiratory Distress Syndrome | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND ARDS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Headache; Dizziness | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, DIZZINESS AND HEADACHE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Pulmonary Dysfunction; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Hypersensitivity/Allergic reaction; Vomiting; Eye Irritation; Dizziness; Nausea; Pulmonary Dysfunction; Inflammation; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, REDUCED PULMONARY RESERVE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Heart Failure/Congestive Heart Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Gastrointestinal Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING STOMACH ISSUES, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH, THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Insufficient Information; Appropriate Term / Code Not Available; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, APPENDIX, COLON, LIVER PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, METASTATIC LUNG CANCER AND BRAIN AND BONE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | Hypothyroidism | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING THYROID AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Vomiting; Eye Irritation; Asthma; Nausea; Skin Inflammation/ Irritation; Unspecified Respiratory Problem; Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE PASSED AWAY. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY, AND THE DEVICE WAS RETURNED. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | Unspecified Kidney or Urinary Problem; Myocardial Infarction; Unspecified Hepatic or Biliary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY,DEATH,KIDNEY DAMAGE, HEART ATTACK. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Vomiting; Asthma; Dizziness; Nausea; Cardiovascular Dysfunction/Insufficiency; Unspecified Respiratory Problem; Headache | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA /VOMITING, ASTHMA (NEW OR WORSENING), REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER, DEATH PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PATIENT INFORMATION CENTER IX | K183387 | Cardiac Arrest | Defective Alarm; No Audible Alarm | IT WAS REPORTED THE CENTRAL STATION DID NOT GENERATE AN ALARM WHEN THE PATIENT WENT INTO CARDIAC ARREST AND PASSED AWAY. THE CUSTOMER REQUESTED ASSISTANCE IN DETERMINING WHETHER ALARMS WERE TRIGGERED AT THE TIME OF THE EVENT. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | PATIENT INVOLVEMENT WAS REPORTED, THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. THE DEVICE, INCLUDING THE ASSOCIATED BEDSIDE DEVICE(S), WERE TESTED AND DETERMINED TO BE FUNCTIONING AS DESIGNED. IT WAS ALSO INDICATED DURING TESTING THAT THE PIC IX SPEAKER WAS PUSHED TO THE BACK OF THE SHELF. THE PROVIDED LOGS WERE REVIEWED BY A PHILIPS PRODUCT SUPPORT ENGINEER AND CLINICAL SPECIALIST. IT WAS DETERMINED THAT THE DEVICE DID ALARM FOR ASYSTOLE DURING THE INDICATED TIMEFRAMES. THE DEVICE WAS TESTED AND DETERMINED TO BE FUNCTIONING AS DESIGNED. REVIEW OF THE LOGS ALSO INDICATED THE DEVICE WAS ALARMING AS EXPECTED. IT WAS CONFIRMED THAT THE DEVICE WAS TESTED AND IS STILL IN CLINICAL USE. REPORTING INSTITUTION PHONE#: (B)(6).
DREAMSTATION CPAP | K131982 | Vomiting; Dizziness; Nausea; Inflammation; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem; Headache; Chronic Obstructive Pulmonary Disease (COPD) | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, AND/OR HEADACHE, NAUSEA / VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | BOX H WAS UPDATED FROM SERIOUS INJURY TO DEATH.
DREAMSTATION CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Inflammation | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Chronic Obstructive Pulmonary Disease (COPD) | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER AND COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAD BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Skin Inflammation/ Irritation; Pulmonary Dysfunction | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, SKIN IRRITATION, REDUCED CARDIOPULMONARY RESERVE AND 2 SURGERIES. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Pneumonia; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED TO FRESENIUS THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP DOCUMENTATION RECEIVED FROM THE PATIENT¿S HOME PROGRAM FACILITY ADMINISTRATOR (FA) CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2025 WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. PER THE FA, THE PATIENT¿S CAUSE OF DEATH IS UNCONFIRMED DUE TO THE PENDING AUTOPSY RESULTS, HOWEVER IT WAS LIKELY A CARDIAC ARREST. ALTHOUGH THE SPECIFICS ARE LARGELY UNKNOWN, IT APPEARS THE PATIENT COMPLETED THEIR LAST CCPD TREATMENT ON (B)(6) 2025 WITHOUT REPORTED ISSUE. PER THE FA¿S RESPONSE, THERE WAS NO ALLEGATION THE SERIOUS ADVERSE EVENTS WERE RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN PD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF A PROBABLE CARDIAC ARREST AND DEATH, AS THE PATIENT WAS CONNECTED TO THE DEVICE WHEN SHE EXPIRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN (AUTOPSY RESULTS AND DEATH CERTIFICATE UNAVAILABLE); THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON FOLLOW-UP DOCUMENTATION FROM THE HOME PROGRAMS¿ FA, THERE WAS NO ALLEGATION OR INDICATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) WAS DEFICIENT OR MALFUNCTIONED IN ANY WAY. CARDIAC DISEASE IS THE LEADING CAUSE OF DEATH IN DIALYSIS PATIENTS, AND MORTALITY RATES ARE UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION PRESENTLY AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN PATIENT¿S CARDIAC ARREST AND DEATH. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS CONNECTED TO THE DEVICE WHEN SHE EXPIRED, THE PENDING AUTOPSY REPORT, AND NO CONFIRMED CAUSE OF DEATH; THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE LIBERTY SELECT CYCLER FROM HAVING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS EXPERIENCED BY THE PATIENT. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Bronchitis | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CHRONIC BRONCHITIS AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING COLON AND ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S SISTER HAS REQUESTED A RETURN LABEL FOR THE DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION, BUT THE EVALUATION IS NOT YET COMPLETE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE DEVICE IS NO LONGER AVAILABLE FOR RETURN, AS IT WAS RECYCLED FOLLOWING THE USER'S PASSING. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT YET BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE USER'S WIFE HAS STATED THAT SHE IS UNSURE OF THE DEVICE'S CURRENT LOCATION AND IS UNABLE TO RETURN IT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION DUE TO A USER OF DREAMSTATION AUTO CPAP. WIFE CALLED IN STATING PATIENT PASSED AWAY LAST 2022. SHE SAID THAT THE DEVICE HAS BEEN RETURNED AND NO LONGER HAS THE UNIT OR TRACKING DETAILS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE PREVIOUSLY CALLED TO ASK FOR ALL COMMUNICATION TO STOP, BUT SHE IS STILL RECEIVING EMAILS AND LETTERS. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR THE PATIENT TO BE REMOVED FROM THE MAILING LIST. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY THREE YEARS AGO. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS BEEN MISPLACED AND WILL NOT BE RETURNED FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DS2ADV AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND NO ISSUES WERE IDENTIFIED DURING THE ASSESSMENT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS BEING USED IN A NURSING HOME, AND A FAMILY MEMBER INDICATED THAT, HAD THEY STILL HAD POSSESSION OF IT, THEY WOULD HAVE RETURNED IT FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE USER PASSED AWAY YEARS AGO. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS GIVEN TO A FRIEND AND IS NOT AVAILABLE TO BE RETURNED FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY LAST YEAR. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO BIPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PATIENT WITH BREAST CANCER UNDERWENT A DA VINCI-ASSISTED BILATERAL NIPPLE-SPARING MASTECTOMY WITH IMMEDIATE RECONSTRUCTION IN (B)(6) 2018. THE PATIENT PASSED AWAY 2 YEARS POST-PROCEDURE. THE REPORTER DID NOT PROVIDE A CAUSE OF DEATH. THERE WAS NO REPORT THAT ANY DA VINCI DEVICE MALFUNCTION OR INTRAOPERATIVE ISSUES OCCURRED DURING THE PROCEDURE. REQUESTS FOR ADDITIONAL INFORMATION FROM THE OPERATING SURGEON HAVE BEEN SENT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE DEVICE LOG FOUND THAT THE FOLLOWING INSTRUMENTS WERE USED IN THE PROCEDURE: ENDOSCOPE, TWO MONOPOLAR CURVED SCISSORS, TWO FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. ALL INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES WITH THE EXCEPTION OF TWO OF THE INSTRUMENTS; ONE SINGLE-USE INSTRUMENT, AND ONE INSTRUMENT THAT HAD REACHED ITS FINAL LIFE USE LIMIT WITH THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. A PATIENT DEATH WAS REPORTED 2 YEARS AFTER A BILATERAL NIPPLE SPARING MASTECTOMY. NO OTHER INTRAOPERATIVE OR POST OPERATIVE DETAILS ARE PROVIDED. HOW THEY DIED OR THE EVENTS LEADING TO THEIR DEATH ARE NOT PROVIDED. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE DEVICE LOG FOUND THAT THE FOLLOWING INSTRUMENTS WERE USED IN THE PROCEDURE: ENDOSCOPE, TWO MONOPOLAR CURVED SCISSORS, TWO FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. ALL INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES EXCEPT FOR TWO OF THE INSTRUMENTS THAT WERE EXPIRED AFTER THIS PROCEDURE AND SINGLE-USE INSTRUMENT. SECTION H10 OF THIS REPORT (RELATED REPORT NUMBER) FOR THE RECEIVED VOLUNTARY MEDWATCH REPORT NUMBER (MW5164582). SECTION E: THE USER FACILITY WHERE THE SURGERY WAS PERFORMED IS DOCUMENTED. THE REPORTER'S NAME AND CORRESPONDING INFORMATION IS NOT AFFILIATED WITH THE USER FACILITY.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFICIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Unspecified Kidney or Urinary Problem; Liver Failure | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER AND KIDNEY SHUTDOWN AND DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Bowel Perforation | Adverse Event Without Identified Device or Use Problem | REFER TO H11 FOR FOLLOW-UP INFORMATION. | ADDITIONAL INFORMATION RECEIVED FROM THE HOSPITAL RISK MANAGEMENT DEPARTMENT STATED, "[THE INFORMATION SUMMARY YOU PROVIDED] "IS NOT ACCURATE. THERE WAS NO INJURY TO A PATIENT INVOLVING DA VINCI EQUIPMENT." NO FURTHER INFORMATION WAS PROVIDED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT'S WIFE NO LONGER HAS THE DEVICE AS HER HUSBAND PASSED AWAY TWO YEARS AGO. NO ADDITIONAL DETAILS HAVE BEEN PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT'S WIFE CONTACTED US TO REPORT THAT SHE HAD RECEIVED A LETTER REGARDING A REPLACEMENT DEVICE FOR HER HUSBAND, WHICH HE NEVER RECEIVED. SHE CLARIFIED THAT HER HUSBAND HAS PASSED AWAY AND NO LONGER REQUIRES A REPLACEMENT. NO FURTHER DETAILS HAVE BEEN PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | ON (B)(6) 2025, A HOME CARE FACILITY ADMINISTRATOR AND PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PDRN, IT WAS REPORTED THIS PATIENT EXPIRED ON (B)(6) 2025 AT HOME DUE TO A CARDIAC ARREST, ATTRIBUTED TO MULTIMORBIDITY THAT INCLUDED A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE. IT WAS EXPLAINED THAT THE PATIENT HAD BEEN IN AND OUT OF THE HOSPITAL IN RECENT MONTHS FOR A DECLINE IN OVERALL HEALTH AND FOR REASONS UNRELATED TO PD THERAPY OR THE USE OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). THE PATIENT WAS FOUND UNRESPONSIVE BY FAMILY ON THE MORNING OF (B)(6) 2025 WHILE STILL CONNECTED TO HER CYCLER. EMERGENCY SERVICES WERE NOT CONTACTED, AND THE PATIENT WAS PRONOUNCED DECEASED IN HER HOME WITHOUT TRANSPORT TO THE HOSPITAL. IT WAS CONFIRMED THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO HER DEATH. THE CAUSE OF THIS PATIENT¿S CARDIAC ARREST AND SUBSEQUENT DEATH CAN BE ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE WITH MULTIMORBIDITY. THERE WAS NO INDICATION OF AN ASSOCIATED DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS CONFIRMED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOVASCULAR DISEASE. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. AT THIS TIME, THE DEVICE WAS RETURNED AND A FOLLOW-UP WILL BE FILED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED.
ALARIS SYSTEM | K133532 |  |  | IT WAS REPORTED THAT ON (B)(6) 2024, AN UMBILICAL ARTERIAL CATHETER (UAC) AND UMBILICAL VENOUS CATHETER (UVC) WERE PLACED ON A 27-WEEK-OLD INFANT SHORTLY AFTER ARRIVAL TO THE NEONATAL INTENSIVE CARE UNIT (NICU). AN INFUSION PUMP WAS USED WITH TUBING REF# (B)(4) FOR BOTH UAC AND UVC LINES. THE UAC LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE REPORTEDLY PRIMED AND FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. IT WAS REPORTED THAT THE INFANT'S CONDITION "ABRUPTLY DETERIORATED WITHIN MINUTES"; "HEART RATE DROPPED"; "CYANOSIS WAS NOTED". ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH REPORTEDLY SHOWED "AIR IN THE LEFT VENTRICLE OF THE HEART." THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2 TO 3ML. THE PATIENT REPORTEDLY DIED 8 DAYS LATER. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "NO AUDIBLE ALARMS SOUNDED, PER STAFF IN THE ROOM. THE LOCATION OF THE AIR WAS DISTAL TO THE PUMP. PUMP FAILURE WAS NOT A CONCERN. PUMPS WERE EVALUATED BY AGILITI BIOMED AND NO DEFICIENCIES WERE FOUND..." THE CUSTOMER (REGISTERED NURSE QUALITY IMPROVEMENT: RL ADMIN LITE & MEDSUN REPORTER) STATED THAT THE PRIMARY AND SECONDARY CAUSES OF DEATH WERE "COMPLICATIONS OF PREMATURE BIRTH ¿ 27 WEEKS AND 4 DAYS. ANOXIC BRAIN INJURY, MULTIPLE ORGAN DYSFUNCTION." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DREAMSTATION CPAP | K131982 |  |  | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE AND DECEASED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 |  |  | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
RENTAL, DREAMSTATION ST30, GB | K102465 |  |  | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K102465 |  |  | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K183387 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | PHILIPS RECEIVED A COMPLAINT ON A PATIENT INFORMATION CENTER IX. IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN OBTAINING THE CLINICAL AUDIT LOGS FOLLOWING A PATIENT INCIDENT. IT WAS INDICATED THE PATIENT DEVELOPED AN SPO2 ISSUE AND PASSED AWAY; HOWEVER, IT WAS ALSO INDICATED THERE WAS NO IMPACT TO THE PATIENT. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. | A PHILIPS TECHNICAL CONSULTANT (TC) SPOKE TO THE CUSTOMER BIOMEDICAL ENGINEER (BIOMED) MANAGER. THE BIOMED CONTACTED THE NURSE/CLINICAL QUALITY IMPROVEMENT SPECIALIST WHO STATED THEY WERE REQUESTING THE LOGS FILES TO CONFIRM THAT THERE WERE NO DELAYS IN NOTIFICATION OR TREATMENT OF THE PATIENT AND TO HOPEFULLY VERIFY THAT THE PATIENT WAS ON CONTINUOUS PULSE OXIMETRY IF POSSIBLE. THE BIOMED LATER CONFIRMED WITH THE TC THAT THERE WAS NO DELAY; THEY JUST WANTED TO SEE WHAT O2 SENSOR WAS READING AT THE MOMENT. BASED ON THE INFORMATION AVAILABLE, THERE WAS NO ALLEGED PRODUCT FAILURE; THE CUSTOMER JUST NEEDED THE LOG FILES FOR REVIEW. THIS COMPLAINT IS DEEMED NOT REPORTABLE WITH PHILIPS DUE THERE IS NO ALLEGATION OF PRODUCT DEFICIENCY OR OTHER INDICATION THAT A PRODUCT PERFORMANCE ISSUE OR MALFUNCTION HAS OCCURRED. THE CUSTOMER HAS REQUESTED ASSISTANCE FROM PHILIPS ONLY.
PATIENT INFORMATION CENTER IX | K153702 | Insufficient Information | No Audible Alarm | IT WAS REPORTED THE PATIENT PASSED AWAY WHILE CONNECTED TO THE MONITOR AND THE CUSTOMER REQUESTED ASSISTANCE WITH THE ALARM LOGS AND CONFIGURATION SETTINGS. THE HOSPITAL DID NOT SHARE THE CAUSE OF DEATH AS THEY FELT IT WAS PROTECTED HEALTH INFORMATION (PHI). BIOMED INDICATED THE BEDSIDE MONITOR ALARMED AS EXPECTED BUT THE CLINICAL MANAGER WANTED TO REVIEW THE AUDIT LOG, SPECIFICALLY RELATED TO SPO2 SENSOR OFF INOP TECHNICAL ALARMS, SPO2 SENSOR OFF CONFIGURATION, AND ECG AND HEART RATE EXTREME ALARMS. GOOD FAITH EFFORTS (GFE) CONFIRMED THE DEVICE DID NOT CONTRIBUTE TO THE DEATH SINCE THE ISSUE IN QUESTION WAS SPO2 SENSOR OFF. THE SYSTEM GENERATED EKG ALARMS AND OTHER ALARMED APPROPRIATELY. | A PHILIPS REMOTE APPLICATION SPECIALIST (RAS) INTERVIEWED THE CUSTOMER, AND THEY REPORTED ON (B)(6) 2024 (TIME 2200 TO (B)(6) 2024 1:30AM) THEIR PATIENT (ICU BED239) PASSED AWAY. THE CUSTOMER STATED THAT THE PATIENT'S BEDSIDE MONITOR DID ALARM AS EXPECTED; HOWEVER, THE CLINICAL MANAGER (CM) HAD INQUIRIES REGARDING THE ALARM LOG AND THE APPLICATION CONFIGURATION SETTINGS. THE CM HAD THE FOLLOWING QUESTIONS: 1) THE CLINICAL AUDIT LOG DOES NOT SHOW THE "SPO2 SENSOR OFF" INOP ALARM EVENT WHY? 2) IS IT POSSIBLE TO CHANGE THE "SPO2 SENSOR OFF" ALARM SEVERITY CONFIG SETTING? 3) HOW IT THE ECG/HR EXTREME ALARM BEHAVIORS CONFIGURED? THE PHILIPS RAS REVIEWED FOLLOWING QUESTIONS AND NOTED A CENTRAL STATION WITH SOFTWARE B.02.X CANNOT LOG A "SENSOR OFF INOP" ALARM AS AN EVENT IN THE CLINICAL AUDIT LOG. THE CLINICAL AUDIT CAN LOG THE FOLLOWING INOPS EVENTS: ALL SERVER YELLOW OR RED INOP MESSAGES. ALL BATTERY INOP MESSAGES. ECG LEADS OFF. TRANSMITTER OFF. NO SPO2T. BATT LOW. THE CENTRAL STATION SERVER WITH SOFTWARE C AND GREATER WILL TRACK ALL THE TECHNICAL INOP ALARMS EVENT ON THE AUDIT LOG. THE RAS EXPLAINED THE SPO2 INOP TECHNICAL ALARMS ARE NOT CONFIGURABLE. THESE ALARMS CANNOT BE CONFIGURED TO YELLOW OR RED ALARM SEVERITY. THE CENTRAL STATION LOCAL SERVER INOPERATIVE SOUND VOLUME CAN BE SET TO BE LOUDER THAN THE CURRENT VOLUME. THE RAS EMAILED REFERENCES TO THE CUSTOMER, AND THEY CONFIRMED THE MONITOR GENERATED THE PATIENT'S HEART RATE LOW ALARMS AND THE ALARMS WERE SILENCED. CENTRAL STATION LOGS WERE REQUESTED BUT COULD NOT BE PROVIDED. THE CUSTOMER NOTED THE SYSTEM WORKED AS DESIGNED SINCE SPO2 SENOR BEING OFF WAS A BLUE ALARM AND DOES NOT RE-ALARM LIKE YELLOW AND RED ALARMS. THE QUESTIONS FROM THE CLINICAL MANGER WERE ANSWERED TO THEIR SATISFACTION AS THEY NOW UNDERSTAND THAT SPO2 SENSOR ALARM IS A BLUE ALARM. NO FURTHER INVESTIGATION OR ACTION IS WARRANTED AT THIS TIME.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. THE EXECUTOR OF THE ESTATE CONTACTED US, MENTIONING THAT SHE HAD RECEIVED A LETTER FROM PHILIPS REGARDING A REPLACEMENT DEVICE. HOWEVER, SINCE THE PATIENT HAD ALREADY PASSED AWAY, SHE WAS UNCERTAIN WHETHER THE DEVICE WAS IN THEIR POSSESSION. SHE STATED THAT SHE WOULD FOLLOW UP IF THE DEVICE IS LOCATED SO THAT IT CAN BE RETURNED. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS INVOLVED IN THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE REPORTED RECEIVING A LETTER REGARDING THE RETURN OF HER HUSBAND'S DEVICE, BUT SINCE HE PASSED AWAY TWO YEARS AGO, SHE IS UNSURE OF THE DEVICE'S WHEREABOUTS OR WHAT HAPPENED TO IT. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S HUSBAND CONTACTED US, EXPRESSING HIS INTENTION TO RETURN THE DEVICE BUT MENTIONED HE WAS UNFAMILIAR WITH HOW TO USE FEDEX. HE WAS REFERRED TO THE DME FOR FURTHER ASSISTANCE. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ALARIS SYSTEM | K133532 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A TECH CALLED IN TO REPORT A PATIENT INCIDENT WHERE IT WAS NOTED THAT THE PATIENT PASSED AWAY. AT THE TIME, THE TECH DIDN'T HAVE ALL THE DEVICES INVOLVED AND SAID THEY WOULD CALL BACK WITH THE OTHER DEVICE INFORMATION AND WHETHER THEY WANTED THE INCIDENT INVESTIGATED. HOWEVER, IT WAS NOTED THAT THE TECH NEVER CALLED BACK. AFTER NUMEROUS ATTEMPTS OF REACHING OUT TO THE TECH, THEY NEVER RESPONDED. | THE ACTUAL DATE OF EVENT IS UNKNOWN. THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
PATIENT INFORMATION CENTER IX | K183387 | Cardiac Arrest | Defective Alarm | IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN DETERMINING WHETHER AN ALARM WAS ACTIVELY ACKNOWLEDGED RELATED TO A PATIENT DEATH. IT WAS INDICATED THE ALARMS FOR 'ECG LEADS OFF' WERE SILENCED, WHICH LED TO A DELAY IN CARE. THE PATIENT WAS FOUND AND SUCCESSFULLY RESUSCITATED BUT LATER PASSED AWAY; HOWEVER, THE CAUSE OF DEATH REMAINS UNKNOWN. NO OTHER CLINICAL INFORMATION WAS PROVIDED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL SPECIALIST REVIEWED THE LOGS PROVIDED. THE CLINICAL AUDIT LOG PROVIDES A CHRONOLOGICAL RECORD OF THE ALARMS AND ACTIONS. THE CLINICAL AUDIT LOG SHOWS FREQUENT SPO2 INOPS ALARM DURING THE TIME IN QUESTION, SUCH AS SPO2 SENSOR OFF, SPO2 NO PULSE, SPO2 PULSE? AND SPO2 NOISY SIGNAL. THESE ALARMS WERE ACKNOWLEDGED AT THE PIC IX. INOPS ARE TECHNICAL ALARMS, THEY INDICATE THAT THE MONITOR CANNOT MEASURE OR DETECT ALARM CONDITIONS RELIABLY. IF AN INOP INTERRUPTS MONITORING AND ALARM DETECTION (FOR EXAMPLE, LEADS OFF), THE MONITOR PLACES A QUESTION MARK IN PLACE OF THE MEASUREMENT NUMERIC AND AN AUDIBLE INDICATOR TONE WILL BE SOUNDED. THESE INOPS APPEAR IN THE CLINICAL AUDIT LOG. CANNOT ANALYZE ECG INOPS ARE ALSO SHOWN IN THE LOG STARTING AT 07:05:45 AND TRANSITED TO !!ECG LEADS OFF AT 7:33:52. CANNOT ANALYZE ECG INOP IS GENERATED IF 20 OF THE LAST 30 SECONDS ARE CLASSIFIED AS EITHER NOISY OR QUESTIONABLE. DURING THIS INOP/TECHNICAL CONDITION, THE ARRHYTHMIA ANALYSIS CONTINUES AND AN ALARM WILL BE ANNOUNCED IF AN ALARM CONDITION IS MET. HOWEVER, THE CANNOT ANALYZE ECG INOP/TECHNICAL ALARM INDICATES THAT THE EFFECTIVENESS OF THE ARRHYTHMIA MONITORING FOR THE PATIENT IS COMPROMISED AND A QUICK RESPONSE TO THIS ALARM IS IMPORTANT.!!! ECG LEADS OFF INDICATES THAT THE REQUIRED ECG LEADS ARE NOT AVAILABLE. AT 7:37:02, ALL ALARMS WERE PAUSED AT THE MONITOR. IN ADDITION, TESTS CARRIED OUT ON SITE CONFIRMED THAT THE DEVICE WAS FUNCTIONING WITHIN MANUFACTURER'S SPECIFICATIONS AND ARE READY FOR CLINICAL USE. BASED ON A REVIEW OF THE INFORMATION PROVIDED AND THE TESTING CONDUCTED ON SITE, THE PIC IX DEVICE WAS FUNCTIONING AS DESIGNED.
PATIENT INFORMATION CENTER IX | K183387 | Ventricular Fibrillation | Defective Alarm | IT WAS REPORTED THAT A (B)(6) Y/O MALE PATIENT WAS BEING MONITORED WHEN INSTANCES OF V-FIB DID NOT ALARM ON THE MONITOR AND THE PATIENT PASSED AWAY. THE PATIENT HAD AN IMPLANTED ICD/PACEMAKER, WHICH ATTEMPTED TO PACE UNSUCCESSFULLY. THE EVENT OCCURRED IN THE MED SURG UNIT AND THE SPEAKERS WERE TESTED FOR FUNCTIONALITY. THE REMAINING QUESTIONS WERE RESPONDED TO WITH ASKU (ASKED BUT UNKNOWN), NOT ALL CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. AT THIS TIME IT CANNOT BE DETERMINED IF THIS DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. | DIAGNOSTIC/FUNCTIONAL TESTING WAS PERFORMED PER A PHILIPS FIELD SERVICE ENGINEER (FSE) ONSITE. THE FSE COLLECTED LOG FILES AND DATA FOR THE EVENT AND VERIFIED THE INTELLIVUE MX40 AND THE PATIENT INFORMATION CENTER IX (PIC IX) ENTERPRISE SYSTEM FOR PROPER OPERATION. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION TO THE PHILIPS SENIOR PRODUCT SUPPORT ENGINEER AND PHILIPS CLINICAL PRODUCT SPECIALIST TO VERIFY IF THE (B)(6) MX40 SN: (B)(6) PATIENT INFORMATION CENTER IX FUNCTIONED AS DESIGNED AND DID ALARM. THE INTELLIVUE MX40 AND THE PATIENT INFORMATION CENTER IX (PIC IX) DEVICE WERE CONFIRMED TO HAVE ALARMED FOR THE PATIENT¿S ECG CHANGES (BOTH YELLOW AND RED LEVEL ALARMS). THE EQUIPMENT WAS WORKING AS DESIGNED, CONFIGURED, AND OPERATED. THE UNACKNOWLEDGED ASYSTOLE ALARM SUPPRESSED VF/VT DETECTION. NO PRODUCT MALFUNCTION WAS DETECTED. THE RESULTS OF THE CLINICAL AUDIT LOG REVIEW SHOW: THE AUDIT LOGS SHOW THAT FOR THE 1-HOUR TIMEFRAME ON (B)(6)2025, FROM 13:00:00 UNTIL 14:00:00: A RED LEVEL CARDIAC ALARM WAS ACTIVE FOR 0:41:36 (41 MINUTES AND 36 SECONDS: ASYSTOLE FOR 40 MINUTES AND 13 SECONDS, VENTRICULAR FIBRILLATION/TACHYCARDIA FOR 1 MINUTE AND 23 SECONDS). A YELLOW LEVEL CARDIAC ALARM (IRREGULAR HR OR PAUSE) WAS ACTIVE FOR 15 MINUTES. THE AUDIT LOGS ALSO REVEAL: ISSUES IMPACTING ECG ANALYSIS: CANNOT ANALYZE ECG INOPERATIVE CONDITION FOR 27 MINUTES AND 29 SECONDS. LL LEAD OFF INOPERATIVE CONDITION FOR 20 MINUTES AND 26 SECONDS. ECG LEADS OFF (NO MONITORING POSSIBLE) FOR 3 MINUTES AND 26 SECONDS. BASED ON THE INFORMATION AVAILABLE AND THE RESULT OF ADDITIONAL ANALYSIS CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS DUE TO USER KNOWLEDGE. THE REPORTED PROBLEM WAS NOT CONFIRMED. ANALYSIS WAS PERFORMED AND INVESTIGATION HAS BEEN COMPLETED. INFORMATION WAS PROVIDED TO THE CUSTOMER OF THE ANALYSIS FINDINGS AND CONFIRMED THE DEVICE WAS WORKING TO ITS SPECIFICATION. THERE IS NO DEVICE MALFUNCTION AND THE REPORTED ISSUE WAS DUE TO USER KNOWLEDGE. THE DEVICE REMAINS AT THE CUSTOMER SITE. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. - PLEASE SEE SPLIT CASE FOR INTELLIVUE MX40 (PRODUCT PROBLEM) REPORTED UNDER MFR REPORT NUMBER 1218950-2025-000075.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE PATIENT OF THE REP DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY AND THE PATIENT'S WIFE NEEDED A RETURN LABEL TO SEND BACK THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION AUTO CPAP DEVICE WAS NO LONGER AVAILABLE TO RETURN AND HAD BEEN DONATED. ACCORDING TO THE PATIENT'S DAUGHTER, HER FATHER PASSED AWAY LAST YEAR AND STOPPED USING THE DEVICE PRIOR TO PASSING. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. THE PATIENT'S DAUGHTER REQUESTED THAT NO FURTHER CORRESPONDENCE BE SENT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DREAMSTATION AUTO CPAP DEVICE. THE CALLER STATED THE DEVICE WAS DISCARDED AND CANNOT BE RETURNED. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
ION | K182188 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED A CARDIAC ARREST AND THE PROCEDURE WAS ABORTED. ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS ADMINISTERED, RESULTING IN THE RETURN OF SPONTANEOUS CIRCULATION (ROSC). HOWEVER, THE PATIENT CODED FIVE MORE TIMES IN THE PROCEDURE ROOM. EACH TIME, ACLS WAS PERFORMED UNTIL THE PATIENT WAS STABLE ENOUGH FOR TRANSFER TO THE INTENSIVE CARE UNIT. THE PATIENT CODED AGAIN IN THE ICU AND AT AN UNSPECIFIED TIME LATER PASSED AWAY. THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE (ETA) WAS PRESENT DURING THE ION PROCEDURE. THE PHYSICIAN INFORMED THE ETA THAT THE CARDIAC ARREST AND THE PATIENT OUTCOME WERE NOT RELATED TO THE ION SYSTEM AND WOULD LIKELY HAVE OCCURRED WITH ANOTHER MODALITY. THE PHYSICIAN ATTRIBUTED THE EVENT TO THE PATIENT'S EXTENSIVE MULTIPLE COMORBID CONDITIONS; THE PATIENT'S MEDICAL HISTORY WAS NOT PROVIDED. THERE WAS NO DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. | THE PHYSICIAN STATED THAT THE PATIENT OUTCOME WAS RELATED TO EXTENSIVE, MULTIPLE CO-MORBIDITIES. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND THERE WERE NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH EXTENSIVE MEDICAL COMORBIDITIES NOT OTHERWISE SPECIFIED UNDERWENT AN ION LUNG BIOPSY. DURING THE PROCEDURE THE PATIENT SUFFERED MULTIPLE EPISODES OF CARDIAC ARREST WITH EVENTUAL ROSC THEN TRANSFERRED TO THE ICU WHERE THE PATIENT ARRESTED AGAIN AND ULTIMATELY DIED. THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT OR ACCESSORIES. THE AVAILABLE DATA SUGGESTS THE ADVERSE EVENT WAS RELATED TO THE PROCEDURE IN A PATIENT WITH SIGNIFICANT UNDERLYING CO-MORBIDITIES. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. ---FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). ---OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. ---FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. ---KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. ---BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. FOR SECTION B2: DATE OF DEATH IS ESTIMATED AS SAME DAY AS PROCEDURE(B)(6) 2025 OR FOLLOWING DAY (B)(6) 2025. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
DAVINCI XI | K131861 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LEFT ADRENALECTOMY PROCEDURE, BLEEDING OCCURRED AND THE PROCEDURE WAS CONVERTED TO OPEN. IT WAS REPORTED THAT THE PATIENT EXPIRED. IT IS UNKNOWN IF THE PATIENT EXPIRED INTRA-OPERATIVELY OR POST-OPERATIVELY. THE CUSTOMER REPORTED THAT THIS EVENT HAD NOTHING TO DO WITH THE DA VINCI PRODUCTS USED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, PERMANENT CAUTERY HOOK, SMALL GRASPING RETRACTOR, CADIERE FORCEPS, FENESTRATED BIPOLAR FORCEPS, LONG BIPOLAR GRASPER, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN ADRENALECTOMY. NEITHER THE DIAGNOSIS, NOR COMORBIDITIES, DETAILS OF THE CASE, POST OPERATIVE COURSE, OR CAUSE OF DEATH ARE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY AT A NURSING HOME FACILITY WHILE CONNECTED TO THE PD CLINIC¿S LIBERTY SELECT CYCLER. THE DEATH WAS REPORTEDLY NOT RELATED TO PD THERAPY. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION, AND THUS FAR NO FURTHER DETAILS HAVE BEEN PROVIDED. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS (PD) UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT DEATH. HOWEVER, THERE WAS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE CYCLER. ADDITIONALLY, THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THE EVENT. THE INFORMATION INITIALLY REPORTED STATED THAT ALTHOUGH THE PATIENT WAS CONNECTED TO THE CYCLER AT THE TIME OF DEATH, THE DEATH WAS NOT PD RELATED. ALTHOUGH MAINTENANCE DIALYSIS PREVENTS DEATH FROM UREMIA, MORTALITY AMONG PATIENTS WITH END-STAGE KIDNEY DISEASE (ESKD) REMAINS HIGH. AMONG PATIENTS WITH ESKD, CARDIOVASCULAR DISEASE ACCOUNTS FOR APPROXIMATELY 50 PERCENT OF DEATHS. BASED ON THE LIMITED INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN NOTIFIED REGARDING THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S MOTHER REPORTED THAT HER DAUGHTER PASSED AWAY ON (B)(6) 2021, AND HAS REQUESTED TO BE REMOVED FROM ALL FURTHER COMMUNICATION. THE MOTHER ALSO INDICATED THAT SHE DOES NOT HAVE THE DEVICE AVAILABLE FOR RETURN. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION DUE TO DEATH HOWEVER NO OTHER MEDICAL OR CLINICAL INFORMATION WERE REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A PERITONEAL DIALYSIS (PD) PATIENT CONTACT REPORTED TO FRESENIUS THAT THE PATIENT EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS REPORTED THAT THE PATIENT'S DEATH WAS NOT PD RELATED. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. ADDITIONAL INFORMATION WAS OBTAINED DURING FOLLOW-UP WITH THE PATIENT'S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). PER THE PDRN, THE PATIENT EXPIRED ON (B)(6) 2025. THE PATIENT HAD COMPLETED TREATMENT AND WAS STILL ATTACHED TO THE CYCLER. IT WAS EXPLAINED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF CARDIOVASCULAR DISEASE WITH PREVIOUS STROKE (UNRELATED TO PD THERAPY) AND WAS DECLINING IN OVERALL HEALTH. THE NURSE INDICATED THAT THE PATIENT EXPIRED FROM CARDIAC ARREST WHICH WAS ATTRIBUTED TO THE PATIENT¿S PRE-EXISTING CARDIOVASCULAR COMORBID CONDITIONS. THE PDRN STATED THERE WERE NO ISSUES WITH THE PD TREATMENT AND STATED SEVERAL TIMES THAT THE DEATH WAS NOT RELATED TO PRODUCT OR PD THERAPY. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXITS BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH FROM CARDIAC ARREST. HOWEVER, THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER OR ANY ISSUES WITH PD TREATMENT. THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED PRE-EXISTING CARDIAC COMORBIDITIES AND STROKE (UNRELATED TO DIALYSIS) WITH REPORT OF DECLINE IN OVERALL HEALTH. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT'S PRE-EXISTING CARDIAC COMORBID CONDITIONS AND STROKE PLACE THAT PATIENT AT HIGHER RISK OF MORTALITY. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DAVINCI XI | K131861 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PATIENT INJURY OCCURRED WHILE PORTS WERE BEING PLACED FOR A DA VINCI-ASSISTED SACROCOLPOPEXY PROCEDURE. ALL PORTS WERE PLACED AND VISUALIZATION WAS CONFIRMED. BEFORE THE DA VINCI SYSTEM WAS DOCKED TO THE PATIENT, THE ANESTHESIOLOGIST OBSERVED THAT THE PATIENT'S BLOOD PRESSURE WAS DROPPING. THE SURGEON DID NOT VISUALIZE ANY BLEEDING AT THE TIME, BUT CONVERTED THE PROCEDURE TO OPEN AND REQUESTED MULTIPLE GENERAL SURGEONS TO ASSIST. THE ROBOTICS COORDINATOR REPORTED THAT THE LOCATION OF THE BLEEDING WAS NEVER FOUND. NO FURTHER INFORMATION WAS PROVIDED REGARDING WHAT OCCURRED AFTER CONVERTING TO OPEN. THE ROBOTICS COORDINATOR STATED THE THEORY WAS THAT THE VERESS NEEDLE USED DURING TROCAR PLACEMENT DAMAGED RETROPERITONEAL ANATOMY. ATTEMPTS WERE MADE TO SPEAK WITH THE SURGEON, HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. | NO SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE AS THE DA VINCI SYSTEM WAS NEVER DOCKED / USED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT ACCORDING TO THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, THE PATIENT IN THIS REPORT IS SPECULATED TO HAVE DIED POSSIBLY AS A RESULT OF A VERESS NEEDLE INJURY WHICH IS A THIRD PARTY DEVICE. THE INTUITIVE SURGICAL DA VINCI ROBOT WAS NEVER DOCKED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED AS SAME DAY OF SURGICAL PROCEDURE.
ION | K182188 | Hemorrhage/Blood Loss/Bleeding; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE. THE PHYSICIAN OBTAINED BIOPSY SAMPLES FROM THE RIGHT LOWER LOBE TARGET LESION. THERE WERE NO ISSUES WITH THE ION SYSTEM. AFTER COMPLETING THE ION PORTION OF THE PROCEDURE, THE ION SCOPE WAS REMOVED AND MANUAL BRONCHOSCOPY WAS PERFORMED. BLEEDING WAS NOTED AT THE BIOPSY SITE WHILE OBTAINING A BRONCHOALVEOLAR LAVAGE SAMPLE. THE ESTIMATED BLOOD LOSS WAS 60ML. THE BLOOD ENTERED THE ENDOTRACHEAL TUBE (ETT) AND THE PATIENT SUFFERED A CARDIAC ARREST, PROMPTING ADVANCED CARDIAC LIFE SUPPORT INTERVENTIONS. THE PATIENT WAS REVIVED AND HEMOSTASIS WAS ACHIEVED. THE ETT WAS REPLACED WITH A DOUBLE LUMEN ETT AND INTRAVENOUS TRANEXAMIC ACID AND AN EPINEPHRINE DRIP WERE ADMINISTERED. THE PATIENT EXPERIENCED A SECOND CARDIAC ARREST 2 DAYS POST-PROCEDURE AND EXPIRED. THE PHYSICIAN REPORTED THAT THE BLEEDING AND PATIENT OUTCOME WERE DUE TO SIGNIFICANT RISKS ASSOCIATED WITH THE BIOPSY PROCEDURE AND COMPOUNDED BY THE PATIENT'S SEVERE PREEXISTING COMORBID CONDITIONS, INCLUDING HEART FAILURE WITH REDUCED EJECTION FRACTION, PULMONARY HYPERTENSION, AND SEVERE COPD. | REVIEW OF THE ION SYSTEM LOGS FOUND NO RELATED ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY; THE ION PORTION OF THE CASE WAS COMPLETED, THE ION SCOPE WAS REMOVED AND MANUAL BRONCHOSCOPY WAS PERFORMED. BLEEDING WAS NOTED WHILE OBTAINING A BRONCHOALVEOLAR LAVAGE SAMPLE. THIS WAS COMPLICATED BY A CARDIAC ARREST. RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED WITH RESUSCITATIVE EFFORTS AND THE PATIENT WAS HOSPITALIZED IN THE CCU. THE PATIENT SUFFERED ANOTHER ARREST 2 DAYS AFTER THE PROCEDURE AND DIED. THERE IS NO ALLEGATION THAT A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT, OR ACCESSORY OCCURRED. GIVEN THE AVAILABLE DATA THE REPORTED EVENT WAS LIKELY PROCEDURE-RELATED AND NOT DEVICE-RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW-RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE-CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF RELATIVELY SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE PATIENT OF THE DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION AND THE DEVICE IS CURRENTLY REGISTERED TO A NEW USER. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE, EVALUATION RESULT CODE, COMPONENT CODE AND CONCLUSION CODE HAS BEEN UPDATED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS INSPECTED AND THE UNIT PASSED THE FUNCTIONAL TEST. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE SERVICE CENTER AND ERROR CODE FOUND. ADDITIONALLY, CUSTOMER COMPLAINT WAS NOT REPRODUCED, AND THE DEVICE WAS SCRAPPED. THERE WAS AN ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Insufficient Information | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS NOTIFIED THAT A PATIENT WHO HAD BEEN USING A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER, STATING THAT HER HUSBAND HAD PASSED AWAY AND THAT SHE NO LONGER REQUIRED THE DEVICE. SINCE THE DEVICE HAD BEEN USED, THE RESPIRATORY THERAPIST INFORMED THE CALLER THAT IT WAS NOT NECESSARY TO RETURN THE DEVICE. THE CALLER INQUIRED ABOUT THE POSSIBILITY OF FINANCIAL COMPENSATION, AS SHE HAD HEARD ABOUT A SETTLEMENT. UPON REVIEWING THE X PRIORITY NOTIFICATION, THE RESPIRATORY THERAPIST CONFIRMED THAT THE AFFECTED DEVICE HAD NOT BEEN RETURNED. THE CALLER INDICATED THAT SHE DID NOT HAVE ANY TRACKING INFORMATION FOR THE DEVICE, EXPLAINING THAT THE DURABLE MEDICAL EQUIPMENT (DME) PROVIDER HAD TAKEN POSSESSION OF IT. THE CALLER WAS SUBSEQUENTLY INFORMED THAT PHILIPS' ONLY REMEDIATION FOR THE AFFECTED DEVICE WAS A REPLACEMENT UNIT, WHICH THE CALLER ACKNOWLEDGED AND UNDERSTOOD. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION OR ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS NOTIFIED THAT A PATIENT WHO HAD BEEN USING A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER, STATING THAT HER HUSBAND HAD PASSED AND REQUESTED A REFUND FOR THE DEVICE. IN RESPONSE, THE MANUFACTURER CLARIFIED THAT DEVICES ARE BEING PROVIDED TO PATIENTS THROUGH THE REPAIR AND REPLACEMENT PROGRAM, AND THAT FINANCIAL COMPENSATION, REIMBURSEMENT, OR PAYMENT IS NOT BEING OFFERED. ADDITIONALLY, THE WIFE WAS PROVIDED WITH THE CONTACT INFORMATION (PHONE NUMBER AND EMAIL ADDRESS) OF A THIRD-PARTY SETTLEMENT GROUP, WHICH HAS BEEN RETAINED TO ADMINISTER AN ECONOMIC LOSS SETTLEMENT, SHOULD SHE HAVE ANY FURTHER QUESTIONS. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION OR ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) REPORTED THE PATIENT EXPIRED PEACEFULLY IN HIS SLEEP DUE TO A CARDIAC ARREST ON (B)(6) 2025. THE PDRN STATED THE PATIENT'S TREATMENT WAS COMPLETED WITHOUT ANY KNOWN ALARMS/ISSUES (VERIFIED THROUGH TREATMENT DATA). THE PDRN CONFIRMED THE PATIENT'S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT'S ADVANCED AGE AND CARDIAC COMORBID CONDITIONS. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT'S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE DEFINITIVE CAUSE OF THE PATIENT¿S CARDIAC ARREST IS UNKNOWN; THEREFORE, CAUSALITY CANNOT FIRMLY BE ESTABLISHED. HOWEVER, THE PATIENT¿S PDRN REPORTED THE PATIENT WAS ELDERLY AND POSSESSED A SIGNIFICANT CARDIAC HISTORY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
PATIENT INFORMATION CENTER IX | K183387 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | PHILIPS RECEIVED A COMPLAINT ON THE PATIENT INFORMATION CENTER IX INDICATING THAT THE CUSTOMER REQUESTED INFORMATION ON WHETHER THE DEVICE HAD AN 'OPTIONAL FUNCTION' FOR EGC ANALYSIS AND ALARMS DURING POSTOPERATIVE CARE AND CARDIAC ARREST. PRIOR TO THE ARREST, THE PATIENT WAS PACED AT 80 BPM WITH AN EXTERNAL PACEMAKER VIA THE JUGULAR VEIN WITH A PLANNED INSERTION OF AN IMPLANTABLE PACEMAKER. THE PATIENT HAD AN UNWITNESSED CARDIAC ARREST, WAS FOUND UNCONSCIOUS, AND CPR WAS INITIATED. THE DEFIBRILLATOR DISPLAYED THE PACED RHYTHM BUT THERE WAS NO PULSE. AN ULTRASOUND WAS PERFORMED, WHICH REVEALED CARDIAC STANDSTILL, AND THE PATIENT WAS PRONOUNCED DEAD. IT WAS CLARIFIED THAT THE DEVICE FUNCTIONED AS INTENDED AND THAT THERE WAS AN ASYSTOLE ALARM GENERATED ON THE PIC IX WHICH NOTIFIED THE NURSING STAFF. THE ALARM WAS SILENCED AT THE BEDSIDE 2 MINUTES LATER. THE RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL, AND THE PATIENT PASSED AWAY. THE DEATH WAS UNRELATED TO THE DEVICE AND THERE WAS NO DEVICE MALFUNCTION. | THIS COMPLAINT IS A SUPPLEMENTAL TO (9610816-2024-001005) DUE TO INCORRECT REGISTRATION NUMBER USED. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION, AND A CLINICAL APPLICATION SPECIALIST REVIEWED THE EVENT, REVEALING THE ECG RHYTHM AT 1826 IS PRESENT BUT THEN BECOMES VERY CHAOTIC OVER THE NEXT FEW SECONDS, AND TRANSITIONS TO ECG LEADS OFF WITH NO DATA STORED AFTER THAT TIME. THE LOGS ALSO INDICATE THE PATIENT'S DATA WAS ACCESSED BY A USER FIVE TIMES BETWEEN 1830 AND 1900. THE DEVICE WAS FOUND TO BE OPERATING PER SPECIFICATIONS, INCLUDING ALARMS AND CAREEVENT NOTIFICATIONS. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE DEVICE WAS FUNCTIONING AS INTENDED, AND THERE WAS NO TROUBLE FOUND WITH THE DEVICE. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. SECTION E REPORTING INSTITUTION PHONE #: (B)(6), SECTION E REPORTER PHONE #: (B)(6).
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A PERITONEAL DIALYSIS (PD) PATIENT'S DAUGHTER REPORTED THAT THIS PATIENT ON PD EXPIRED WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. HOWEVER, THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE STATED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF END STAGE RENAL DISEASE (ESRD) WITH ADDITIONAL CARDIAC COMORBIDITIES INCLUDED AORTIC VALVE STENOSIS AND HYPERTENSION. PER THE NURSE, THE PATIENT EXPIRED ON (B)(6) 2025. THE NURSE STATED THAT THE PATIENT¿S DAUGHTER REPORTED THAT THE PATIENT WAS ATTACHED TO THE FRESENIUS CYCLER. HOWEVER, NO TREATMENT DATA WAS AVAILABLE AND IT WAS UNKNOWN IF THE PATIENT WAS IN ACTIVE TREATMENT WHEN SHE EXPIRED. THE NURSE STATED THE PATIENT WAS AWAKE AND WELL AT 8AM. HOWEVER, SUBSEQUENTLY, AT APPROXIMATELY 10AM THE PATIENT WAS FOUND UNRESPONSIVE (BY DAUGHTER). THE NURSE STATED THAT EMERGENCY MEDICAL SERVICES WAS CALLED. THE PATIENT WAS A DNR AND NOT BROUGHT TO THE HOSPITAL; PRONOUNCED DECEASED IN THE HOME. THE NURSE STATED THERE WERE NO REPORTED ISSUES WITH THE PATIENT¿S PD TREATMENT OR PRODUCT IN RELATION TO THE DEATH. THE NURSE STATED THAT THE PATIENT¿S NEPHROLOGIST SUSPECTS THE CAUSE OF DEATH WAS FROM CARDIAC DUE TO NATURAL CAUSES FROM PRE-EXISTING COMORBID CONDITIONS. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE. THE FRESENIUS CYCLER WAS PENDING RETURN TO MANUFACTURER AT THE TIME FOLLOW-UP WAS COMPLETED. | CLINICAL REVIEW: TEMPORAL RELATIONSHIP EXISTS BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. HOWEVER, THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER OR ANY ISSUES WITH PD TREATMENT. IT WAS REPORTED THAT THE SUSPECTED CAUSE OF DEATH WAS CARDIAC ARREST. THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED ESRD WITH PRE-EXISTING CARDIAC COMORBIDITIES THAT INCLUDED AORTIC VALVE STENOSIS AND HYPERTENSION. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT¿S PRE-EXISTING CARDIAC COMORBID CONDITIONS INCREASE RISK OF MORTALITY. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A PATIENT WITH A RENTAL DREAMSTATION ST30 GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINIC REACHED OUT TO THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. A FOLLOW-UP REPORT WILL BE FILED IF ANY FURTHER RELEVANT INFORMATION BECOMES AVAILABLE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. | THE MANUFACTURER PREVIOUSLY REPOTED THAT THEY RECEIVED INFORMATION ALLEGING THAT A PATIENT WITH A RENTAL DREAMSTATION ST30 GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINIC REACHED OUT TO THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. THE DEVICE UNDERWENT A REPAIR EVALUATION BY A THIRD-PARTY SERVICE CENTER, WHERE IT WAS VISIBLY INSPECTED, TESTED, AND SUBSEQUENTLY PASSED, WITH NO FAULTS FOUND. THE UNIT WAS THEN RETURNED TO LOAN STOCK . NO FURTHER INVESTIGATION IS NECESSARY. IF WE RECEIVE ANY RELEVANT INFORMATION, A FOLLOW UP REPORT WILL BE SUBMITTED.
ALARIS SYSTEM | K211218 | Insufficient Information | Pumping Stopped; Use of Device Problem; Contamination | IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY?...THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. " | OMIT : B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE, ADDITIONAL INFORMATION : DEVICE AVAILABLE FOR EVAL?, RETURNED TO MANUFACTURER ON, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, G, B, C, D CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE CUSTOMER SOUGHT INFORMATION AROUND THE MEDICATION INFUSION DETAILS AND WHETHER THERE WAS ANY TIME WHERE THE MEDICATIONS WERE NOT INFUSING. REVIEW OF THE DEVICE LOGS IDENTIFIED INSTANCES OF OCCLUDED PATIENT SIDE ALARMS, BUT THERE WERE NO INDICATIONS OF DEVICE MALFUNCTIONS OR INTERRUPTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT¿S CONDITION. INTERRUPTIONS IN INFUSION WERE NOT OBSERVED OR DURING DEVICE TESTING. INSPECTION AND TESTING OF THE SUSPECT PUMP MODULES AND PCU DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULES WERE INFUSING WITHIN SPECIFICATIONS. THE PCU PASSED PREVENTATIVE MAINTENANCE TESTING AND WAS OPERATING AS INTENDED. REVIEW OF THE SUSPECT PCU BATTERY LOGS DID NOT SHOW ANY EVENTS OR MALFUNCTIONS THAT COULD HAVE CONTRIBUTED TO THE REPORTED INCIDENT. THERE WERE NO INSTANCES OF THE SYSTEM POWERING OFF DURING THE REPORTED INCIDENT TIME. THE INITIAL INFUSION PROGRAMMING OF NOREPINEPHRINE (DRUGID = 308), AT A CONCENTRATION OF 4 MG / 250 ML, WAS RECORDED ON 7FEB2025 AT 1:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6). THE INFUSION RATE WAS SET AT 41.25 ML, THE VOLUME TO BE INFUSED (VTBI) WAS 100 ML. A BASIC INFUSION WAS PROGRAMMED ON (B)(6) 2025 AT 3:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6) AT A RATE OF 100 ML/HR AND A VTBI OF 950 ML. A THIRD PUMP MODULE (SN (B)(6), NOT RECEIVED) WAS IDENTIFIED TO BE IN USE DURING THE REPORTED INCIDENT TIME. A BASIC INFUSION WAS PROGRAMMED AND STARTED ON (B)(6) 2025 AT 1:47 PM, WITH A RATE OF 999 ML/HR AND A VTBI OF 950 ML. AT 2:40 PM, THE CONCOMITANT PUMP MODULE SN (B)(6) WAS CHANNELED OFF. THEN AT 2:59 PM AN INFUSION FOR THE DRUG CEFTRIAXONE (DRUGID = 131), AT A CONCENTRATION OF 1000 MG / 50 ML, WAS PROGRAMMED AND STARTED. THE INFUSION RATE WAS 100 ML/HR AND THE VTBI WAS 50 ML. THE DOSE FOR THE INFUSION OF NOREPINEPHRINE ON THE PUMP MODULE SN (B)(6) WAS CHANGED TO 0.4 MCG/KG/HR AT 3:18 PM AND THEN CHANGED AGAIN TO 0.22 MCG/KG/HR AT 3:20 PM. THE DOSE WAS CHANGED TO 0.25 MCG/KG/HR AT 4:28 PM. AN AIR-IN-LINE ALARM WAS OBSERVED ON THE CONCOMITANT PUMP MODULE SN (B)(6), LASTING 2 MINUTES BEFORE THE UNIT WAS CHANNELED OFF. THE TOTAL VOLUME INFUSED (TVI) WAS 902.015 ML. FIVE (5) OCCLUDED PATIENT SIDE ALARMS WERE OBSERVED BETWEEN 1:13 PM AND 4:27 PM ON THE PUMP MODULE SN (B)(6). THE ALARM AT 2:00 PM LASTED 8 MINUTES BEFORE THE INFUSION WAS RESTARTED. THEN, AT 3:15 PM, THE ALARM LASTED 2 MINUTES AND 48 SECONDS BEFORE THE INFUSION WAS RESUMED. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 3:46 PM, BUT WAS THEN POWERED ON AGAIN AT 4:24 PM AND AN INFUSION OF NOREPINEPHRINE AT A DOSE OF 0.22 MCG/KG/HR AND A VTBI OF 200 ML WAS PROGRAMMED AND STARTED AT 4:26 PM. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:29 PM. THE TVI WAS RECORDED AT 101.491 ML. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:30 PM. THE TVI WAS RECORDED AT 127.003 ML. IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF THE PCU EVENT LOG. IT IS ALSO NOT POSSIBLE TO DETERMINE WHAT THE FLUID IS THAT IS CONTAINED WITHIN THE IV CONTAINER. THIS IS NOT A FUNCTION OF THE PCU EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED AND FLUID SELECTED. THE DEVICE WAS REPORTEDLY IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INTERRUPTION OF INFUSIONS IS DUE TO THE USER POWERING OFF THE PUMP MODULES. IT WAS NOTED THAT THE SUSPECT PUMP MODULES SN (B)(6) AND SN (B)(6) WERE POWERED OFF ON (B)(6) 2025 AT 4:29 PM AND 4:30 PM, RESPECTIVELY. THE CONCOMITANT PUMP MODULE SN (B)(6) WAS POWERED OFF AT 3:28 PM. NO FUNCTIONALITY ISSUES OR ANOMALIES WERE OBSERVED WITH THE DEVICES THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULES WERE FOUND TO BE INFUSING WITHIN SPECIFICATION. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A1801, G04037, C0601 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE.
ALARIS SYSTEM | K211218 | Insufficient Information | Pumping Stopped; Use of Device Problem; Contamination | IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. " | THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K211218 | Insufficient Information | Pumping Stopped; Use of Device Problem; Contamination | IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR 07 FEBRUARY 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED." | THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Dyspnea; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED THAT PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PDRN. THE PATIENT WAS IN PD TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6) 2025. THE PATIENT EXPERIENCED SHORTNESS OF BREATH. THE PATIENT¿S PARTNER CALLED 911 AND DISCONNECTED THE PATIENT FROM THE TREATMENT. EMERGENCY MEDICAL SERVICES (EMS) ARRIVED TO FIND THE PATIENT IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED. RESUSCITATIVE EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. NO AUTOPSY WAS CONDUCTED. THE CAUSE OF DEATH IS DOCUMENTED AS CARDIAC ARREST, CAUSE UNKNOWN. NO ISSUES WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT WERE REPORTED PRIOR TO THE PATIENT¿S DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. | PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST ¿ PASSED. SYSTEM AIR LEAK TEST ¿ PASSED. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DREAMSTATION | K131982 | High Blood Pressure/ Hypertension; Dyspnea; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION DEVICE. THERE IS AN ALLEGATION THAT THE DEVICE WAS NOT WORKING CORRECTLY, AND THE PATIENT WAS HAVING PROBLEMS BREATHING. THE PATIENT WAS TAKEN TO AN EMERGENCY ROOM WHERE THEY PASSED FROM HYPERTENSION AND CARDIAC ARREST. THE PATIENT¿S FAMILY ALSO REPORTED THAT THE UNIT WAS BEING CLEANED WITH AN OZONE-BASED DISINFECTION DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION ST30 DEVICE. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE OPERATED AS DESIGNED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K091319 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER HAS BEEN INFORMED OF THE RECENT PASSING OF A PATIENT WHO WAS USING A REMSTAR AUTO A-FLEX DEVICE. ACCORDING TO THE PATIENT'S WIFE, HER HUSBAND RECENTLY PASSED AWAY, AND SHE WAS UNABLE TO LOCATE THE DEVICE. SHE HAS INDICATED THAT SHE WILL RETURN IT IF SHE COMES ACROSS IT WHILE SORTING THROUGH HIS BELONGINGS. NO FURTHER DETAILS HAVE BEEN PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE HAD ANY ROLE IN OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER HAS BEEN NOTIFIED OF THE RECENT PASSING OF A PATIENT WHO WAS USING A DREAMSTATION CPAP PRO DEVICE. THE PATIENT'S DAUGHTER HAS REQUESTED INFORMATION REGARDING THE RECALL AND WHETHER THE RECALLED DEVICE COULD BE LINKED TO HER MOTHER'S CANCER DIAGNOSIS. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING EVALUATION, NO ERROR WAS FOUND IN THE DEVICE LOG, AND THE DEVICE WAS AWAITING SCRAPPING CONFIRMATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION CPAP | K131982 | Cancer; Insufficient Information; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE CANCER, AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION STATING A PATIENT SADLY PASSED AWAY. THE PATIENT WAS USING A BIPAP AVAPS30 DEVICE, ALTHOUGH NO INFORMATION WAS PROVIDED TO VERIFY IF THE DEVICE WAS IN USE WHEN THE PATIENT PASSED. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. ALSO, PER REPORT OF THE EVENT, "THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH." THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE AND DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM IF ANY MALFUNCTION OCCURRED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Peritonitis | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON 06/MAR/2025 REPORTEDLY DUE TO PERITONITIS AND A CLOSTRIDIUM DIFFICILE (C-DIFF) INFECTION. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, HOSPITAL RECORDS, TREATMENT RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, AUTOPSY REPORT, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: IT IS UNKNOWN IF A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF PERITONITIS, A C-DIFF INFECTION, AND DEATH. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. ACCORDING TO THE INITIAL REPORTING, THE PATIENT¿S DEATH WAS ATTRIBUTED TO THE PERITONITIS AND C-DIFF INFECTIONS. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP INFORMATION PRECLUDED A MORE COMPREHENSIVE CLINICAL ASSESSMENT. THOSE INDIVIDUALS UNDERGOING PD FOR RRT (MANUAL OR CYCLER BASED) ARE AT HIGH RISK FOR INFECTIONS OF THE PERITONEUM. ADDITIONALLY, RISK FACTORS ASSOCIATED WITH THE DEVELOPMENT OF C. DIFFICILE INCLUDE LONGER DURATIONS OF ANTIBIOTIC ADMINISTRATION REQUIRED FOR ESRD PATIENTS TO CLEAR THE INFECTION. LASTLY, THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER AND LIBERTY CYCLER SET CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE LACK OF PATIENT DATA, TREATMENT DATA, CLINICAL FOLLOW-UP, DISCHARGE SUMMARY, DEATH CERTIFICATE, ESRD DEATH NOTIFICATION AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
ION | K182188 | Stroke/CVA | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNEVENTFUL ION ENDOBRONCHIAL LUNG BIOPSY WITH MEDIASTINAL STAGING PROCEDURE. THE PHYSICIAN REPORTED THAT THERE WERE NO INTRA-PROCEDURAL COMPLICATIONS OR MALFUNCTIONS WITH THE ION SYSTEM. APPROXIMATELY 30 MINUTES POST-PROCEDURE, THE PATIENT SHOWED ACUTE STROKE SYMPTOMS OF LEFT-SIDED WEAKNESS AND GAZE DEVIATION TREATED WITH TISSUE PLASMINOGEN ACTIVATOR (TPA) THAT RESULTED IN EPISODES OF HEMOPTYSIS. THE PHYSICIAN STATED THAT THE PATIENT SUFFERED AN ISCHEMIC STROKE; UNSPECIFIED HEAD IMAGES DID NOT SHOW ANY ACUTE CHANGES. THE PATIENT EXPIRED SHORTLY THEREAFTER. THE CAUSE OF THE STROKE WAS REPORTED TO BE RELATED TO THE PATIENT¿S MEDICAL HISTORY, INCLUDING SUBSTANTIAL CARDIOVASCULAR DISEASE WITH MULTIPLE STENTS AND MULTIPLE STROKES TREATED WITH ASPIRIN AND PLAVIX THERAPY. THE PLAVIX HAD BEEN HELD FOR 5 DAYS PRIOR TO THE PROCEDURE, BUT THE ASPIRIN THERAPY WAS CONTINUED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT AN ELDERLY PATIENT WITH MULTIPLE MEDICAL CO-MORBIDITIES INCLUDING MULTIPLE PRIOR STROKES, CARDIOVASCULAR DISEASE REQUIRING MULTIPLE VASCULAR STENTS UNDERWENT AN ION LUNG BIOPSY. PLAVIX WAS HELD FOR 5 DAYS IN PREPARATION FOR THE BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS TRANSPORTED TO THE PACU. ABOUT 30 MINUTES LATER THE PATIENT DEVELOPED A LEFT SIDED WEAKNESS AND GAZE DEVIATION PROMPTING TPA TO TREAT AN ACUTE STROKE. WORKUP WAS CONSISTENT WITH AN ISCHEMIC STROKE AND THE PATIENT DIED LATER THE SAME DAY. BASED ON THE AVAILABLE DATA THE EVENT WAS PROCEDURE RELATED IN A PATIENT WITH MULTIPLE PRIOR STROKES WHOSE PLAVIX HAD BEEN STOPPED FOR THE BIOPSY AND WAS NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 5 (0.02%) TRANSIENT ISCHEMIC ATTACKS AND 4 (0.02%) TOTAL DEATHS BUT NO STROKES. ANOTHER MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO STROKES AND NO DEATHS ASSOCIATED WITH 581 CASES. A RECENT PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 BRONCHOSCOPIES REPORTED 1 ASSOCIATED DEATH (0.08%) AND NO STROKES. ANOTHER PROSPECTIVE SERIES OF 415 ION PROCEDURES REPORTED 1 CVA (0.2%). THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT OR ACCESSORIES ASSOCIATED WITH THE REPORTED COMPLICATION. THE RISK OF STROKE ASSOCIATED WITH GENERAL ANESTHESIA HAS BEEN REPORTED TO VARY FROM 0.1-1.9% IN NON-CARDIAC, NON-NEUROLOGIC, AND NON-MAJOR SURGERY AND AS HIGH AS 10% IN THE SETTING OF BRAIN OR HIGH-RISK CARDIOVASCULAR SURGERY. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BATEMAN BT, SCHUMACHER HC, WANG S, SHAEFI S, BERMAN MF. PERIOPERATIVE ACUTE ISCHEMIC STROKE IN NONCARDIAC AND NONVASCULAR SURGERY: INCIDENCE, RISK FACTORS, AND OUTCOMES. ANESTHESIOLOGY. 2009. SELIM M. PERIOPERATIVE STROKE. NEW ENGLAND JOURNAL OF MEDICINE. 2007. SECTION A2: AGE REPORTED AS "IN THEIR 80'S."
REMSTAR AUTO A-FLEX | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE. THE PATIENT HAS PASSED AWAY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H UPDATED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S WIFE HAS ALLEGED HER HUSBAND PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN UNABLE TO CONFIRM FOAM PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION. THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. ADDITIONALLY, NO FAULTS WERE FOUND AND ALL TESTS WERE PASSED. | THE BECOME AWARENESS DATE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED/CORRECTED IN THIS REPORT. THE DEVICE INFORMATION HAS ALSO BEEN UPDATED/CORRECTED. IN THIS REPORT, BOX D, G, H HAS BEEN UPDATED/CORRECTED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Arrhythmia; Malaise | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THEIR LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). IT WAS CONFIRMED THE PATIENT WAS IN ACTIVE TREATMENT UTILIZING THE LIBERTY SELECT CYCLER AT THE TIME OF THE EVENT. IT IS UNKNOWN WHAT PHASE AND CYCLE THE PATIENT WAS IN. THE PATIENT CALLED THEIR SPOUSE INTO THE BEDROOM AND STATED THEY WERE NOT FEELING WELL. THE PATIENT WAS SITTING ON THE EDGE OF THE BED AND BECAME QUIET. THE PATIENT THEN BECAME UNRESPONSIVE. 911 WAS CALLED. THE 911 OPERATOR INSTRUCTED THE SPOUSE TO PLACE THE PATIENT ON THE FLOOR. DESPITE THE PATIENT BEING A DO NOT RESUSCITATE (DNR) ORDER, THE SPOUSE BEGAN CARDIOPULMONARY RESUSCITATION (CPR). EMERGENCY MEDICAL SERVICES (EMS) ARRIVED AND CONTINUED TO PERFORM RESUSCITATIVE EFFORTS. PER THE PATIENT¿S SPOUSE, THEY ADMINISTERED CPR AND MEDICATIONS (UNSPECIFIED). RESUSCITATION EFFORTS WERE UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED AT HOME. NO AUTOPSY WAS COMPLETED. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. THE PATIENT¿S CAUSE OF DEATH IS REPORTEDLY CARDIAC ARRHYTHMIA. THE PATIENT DID HAVE A HISTORY OF UNSPECIFIED CARDIAC ARRHYTHMIA. THERE WAS NO REPORTED ISSUE WITH THE PATIENT¿S TREATMENT, OR THE LIBERTY SELECT CYCLER IN RELATION TO THE EVENT. THE LAST REPORTED ISSUES WERE DRAIN COMPLICATIONS WHICH WERE SUBSEQUENTLY FOUND TO BE CAUSED BY SEVERE CONSTIPATION AS A RESULT OF THE PATIENT BEING BEDRIDDEN. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE REPORTED CAUSE OF DEATH WAS CARDIAC ARRHYTHMIA. CARDIAC DISEASE IS THE MAJOR CAUSE OF DEATH, ACCOUNTING FOR APPROXIMATELY 37 PERCENT OF ALL-CAUSE MORTALITY IN PATIENTS RECEIVING EITHER HEMODIALYSIS OR PERITONEAL DIALYSIS WITH THE LARGEST SPECIFIC CAUSE ATTRIBUTED TO ARRHYTHMIC MECHANISMS OR SUDDEN CARDIAC ARREST (SCA). THE PATIENT REPORTEDLY HAD A HISTORY OF UNSPECIFIED CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED TO THE CLINIC. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, TREATMENT RECORD, HOSPITAL RECORDS, DEATH CERTIFICATE) WERE UNSUCCESSFUL. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER AND SHOWED BEZEL DAMAGE. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. AN AS-RECEIVED WITH REDUCED DWELL SIMULATED TREATMENT WAS PERFORMED ON THE CYCLER AND PASSED. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST AND VALVE ACTUATION TEST. AN INVESTIGATION OF THE CYCLER MUSHROOM HEADS VERIFIED THAT THE SURFACE CONDITIONS AND ALIGNMENTS WERE WITHIN SPECIFICATION. AN INTERNAL VISUAL INSPECTION IDENTIFIED EVIDENCE OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE DRIED FLUID COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DREAMSTATION ST30 | K102465 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | CORRECTION TO THE PREVIOUSLY REPORTED INFORMATION: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS INSPECTED AND NO FAULT FOUND. ADDITIONALLY, ALL TESTS PASSED. THE "PROBLEM CODE GRID" HAS BEEN CORRECTED IN THIS REPORT. IN THIS REPORT, BOX H HAS BEEN UPDATED/CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN THIS REPORT IN BOX G: DATE RECEIVED BY MFG, IN BOX H: EVAL METHOD CODE, EVAL RESULT CODE, CONCLUSION CODE HAS BEEN UPDATED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A HOME THERAPY PROGRAM MANAGER AND PERITONEAL REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO A CARDIAC ARREST. IT WAS AFFIRMED THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF THE CARDIAC ARREST AND HER SUBSEQUENT DEATH. THE PATIENT¿S CARDIAC ARREST WAS ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOPULMONARY DISEASE THAT WAS UNRELATED TO CCPD THERAPY. IT WAS EXPLAINED THAT THE PATIENT WAS FOUND UNRESPONSIVE AND PULSELESS ON THE MORNING OF (B)(6) 2025. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS NOT TRANSPORTED TO THE HOSPITAL AND PRONOUNCED DECEASED IN HER HOME. IT WAS CONFIRMED THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | CORRECTION PROVIDED IN G4. ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL NO INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A PERITONEAL DIALYSIS (PD) NURSE REPORTED THAT A PD PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. REPORTEDLY, THE CYCLER WAS NOT ACTIVELY RUNNING. THE PATIENT¿S TREATMENT HAD BEEN COMPLETED. THE PATIENT WAS FOUND BY SPOUSE DECEASED IN THE BEDROOM AT APPROXIMATELY 7AM. IN A FOLLOW UP, ADDITIONAL DETAILS VERIFIED THAT THE PATIENT WAS FOUND DECEASED AT APPROXIMATELY 7AM ON (B)(6) 2025. THE PATIENT¿S SPOUSE FOUND PATIENT COLD AND PULSELESS. THE SPOUSE ADMINISTERED CARDIOPULMONARY RESUSCITATION (CPR) UNTIL THE RESCUE SQUAD ARRIVED. THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME AND TRANSPORTED TO THE FUNERAL HOME. THE PATIENT¿S TREATMENT WAS RECORDED AS COMPLETED AT 5:29AM, HOWEVER; THE PATIENT HAD NOT DISCONNECTED FROM THE LIBERTY SELECT CYCLER. THE EXACT TIME OF DEATH WAS NOT REPORTED. NO ISSUES WERE REPORTED WITH THE CYCLER OR THE PATIENT¿S TREATMENT. THERE WERE 0 DRAIN ALARMS AND 0 FILL ALARMS PER THE TREATMENT SHEET. THE MACHINE HAD A HISTORY OF 5 ALARMS (UNKNOWN TYPE OR DATE). THE PATIENT HAD REPORTED SLOW DRAINS PREVIOUSLY. THE CAUSE OF DEATH WAS DOCUMENTED AS CARDIAC ARREST PER THE DOCTOR. THE DEATH WAS NOT RELATED TO ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY. | ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL ANALYSIS. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT ANY FAILURES OR PROBLEMS. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. SYSTEM AIR LEAK TEST PASSED. TEACH PUMP TEST PASSED. VALVE ACTUATION TEST PASSED. AN INTERNAL VISUAL OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DREAMSTATION | K131982 | Fibrosis | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY DUE TO PULMONARY FIBROSIS. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE SPOUSE OF THE PATIENT REPORTED THE PATIENT WAS DECEASED. THE REPORTER INQUIRED ABOUT THE RETURN AND/OR DISPOSAL OF THE DEVICE. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE REPORTER. THE REPORTER WAS ADVISED IT WAS SAFE TO DISPOSE OF THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAD PASSED AWAY. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | Unspecified Heart Problem | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION DEVICE HAS PASSED DUE TO AN UNKNOWN HEART CONDITION. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT¿S SON HAD ARISEN TO USE THE RESTROOM AT APPROXIMATELY 1:30 AM, AND NOTED HIS MOTHER WAS REPORTEDLY SLEEPING COMFORTABLY. HOWEVER, AT APPROXIMATELY 6:30 AM THE PATIENT¿S SON DISCOVERED THE PATIENT HAD EXPIRED WHEN HE WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 THROUGH (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HER SIGNIFICANT CARDIAC HISTORY (CORONARY ARTERY DISEASE, STENOSED AORTIC VALVE NEEDING REPLACEMENT). THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. THE PDRN REPORTED THE CAUSE OF THE SERIOUS ADVERSE EVENTS WAS THE PATIENT¿S SIGNIFICANT CARDIAC COMORBIDITIES. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. THE PDRN SERIOUS ADVERSE EVENTS WERE REPORTEDLY UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
ION | K182188 | Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE THE PATIENT EXPIRED. THE ENDOLUMINAL SALES ASSOCIATE (ESA) REPORTED THAT UPON ARRIVAL AT THE HOSPITAL PROCEDURE ROOM, THEY SAW THAT THE ION SYSTEM WAS UNDOCKED AND BLOOD WAS COMING OUT OF THE ENDOTRACHEAL TUBE. THE PHYSICIAN HAD PLACED A BRONCHOSCOPE, WAS ADMINISTERING SALINE AND EPINEPHRINE VIA THE ENDOTRACHEAL TUBE, AND WAS ATTEMPTING TO STOP THE BLEEDING VIA SUCTIONING THE AIRWAY. THE ESA WAS LATER INFORMED THAT THE PATIENT HAD EXPIRED. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN WERE MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | AN ION SYSTEM LOG REVIEW COULD NOT BE PERFORMED AS THE LOGS ARE NOT AVAILABLE. A DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE WITH UNKNOWN MEDICAL CO-MORBIDITIES UNDERWENT AN ION BRONCHOSCOPY FOR BIOPSY. DURING THE PROCEDURE BLEEDING WAS NOTED IN THE ET TUBE PROMPTING TRANSITION TO MANUAL BRONCHOSCOPY TO INSTILL EPINEPHRINE, SALINE AND SUCTIONING. THE PATIENT DIED DURING THE PROCEDURE. THERE WAS NO ALLEGATION OF MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE REPORTED EVENT WAS LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. ATTEMPTS AT OBTAINING FURTHER INFORMATION HAS BEEN UNSUCCESSFUL. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DAVINCI XI | K131861 | Respiratory Acidosis; Low Blood Pressure/ Hypotension | Adverse Event Without Identified Device or Use Problem | A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE STUDY TO COMPARE THE OUTCOMES OF 96 PATIENTS UNDERGOING ROBOTIC-ASSISTED RADICAL CYSTECTOMY (RARC) WITH URINARY DIVERSION FOR BLADDER CANCER. THE STUDY AIMED TO EVALUATE THE POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF 47 SINGLE-PORT (SP) VS 49 MULTIPORT (MP) ROBOTIC-ASSISTED RADICAL CYSTECTOMY. THE STUDY REPORTED THAT IN THE MP COHORT, A MORTALITY WAS NOTED WITH THE PATIENT EXPIRING ON POSTOPERATIVE DAY 2 SECONDARY TO SEVERE ACIDOSIS AND HYPOTENSION. THE STUDY CONCLUDED THAT THE SP ROBOTIC PLATFORM OFFERS A SAFE ALTERNATIVE TO THE MP ROBOTIC PLATFORM IN PERFORMING RARC WITH URINARY DIVERSION. NO CLINICALLY SIGNIFICANT DIFFERENCE WAS DEMONSTRATED IN PERIOPERATIVE OR ONCOLOGIC OUTCOMES. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS REPORTED, NOR DID THE AUTHORS ALLEGED THAT ANY INTUITIVE PRODUCTS CAUSED OR CONTRIBUTED TO THE DEATHS. THE DESIGNATED AUTHOR CONTACT HAS NOT RESPONDED TO A REQUEST FOR ADDITIONAL INFORMATION. | THIS MEDWATCH REPORT REPRESENTS THE DA VINCI MULTIPORT (MP) REPORTED DEATHS. SEE SECTION H10 FOR THE ASSOCIATED DA VINCI SINGLE PORT (SP) DEATH REPORTS. AN INVESTIGATION COULD NOT BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATES OR DA VINCI SYSTEM IDENTIFIERS. THERE WERE NO DEVICE ISSUES DOCUMENTED IN THE ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THIS ARTICLE DISCUSSED A SERIES OF PATIENTS COMPARING MULTIPORT AND SINGLE PORT ROBOTIC SURGERY WHO UNDERWENT RADICAL CYSTECTOMY. THIS REPORT IS REGARDING A MULTIPORT PATIENT WHO DIED ON POST OPERATIVE DAY 2 SECONDARY TO SEVERE ACIDOSIS AND HYPOTENSION. NO INTRAOPERATIVE DETAILS ARE PROVIDED. HOW THEY DIED AND THE CIRCUMSTANCES WHICH LED TO THEIR DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. CITATION: FANG, A. M., HAYEK, O., KAYLOR, J. M., PEYTON, C. C., FERGUSON, J. E., NIX, J. W., & RAIS BAHRAMI, S. (2024). POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF SINGLE-PORT VS MULTIPORT ROBOTIC-ASSISTED RADICAL CYSTECTOMY. JOURNAL OF ENDOUROLOGY, 38(5), 438¿443. HTTPS://DOI.ORG/10.1089/END.2023.0553. SECTION A2 AGE: THE STUDY POPULATION AGE RANGE FOR DA VINCI MP WAS 67.0 +/- 9.4 YEARS. MEAN AGE 67 YEARS. SECTION A3 GENDER: THE STUDY POPULATION FOR DA VINCI MP WAS MALE - 41; FEMALE - 8. SECTION A4 WEIGHT: THE STUDY POPULATION FOR DA VINCIE MP WAS BMI OF 30.3+/-6.5 KG/M2. SECTION B3 EVENT DATE IS UNKNOWN. THE DATE REFLECTS THE DATE THE ARTICLE WAS PUBLISHED - MAY 2024.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS OR FAULTS FOUND. THE DEVICE IS CURRENTLY AWAITING CONFIRMATION FOR RETURN TO LOAN OR TO BE SCRAPPED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. IN THE PREVIOUS FILE THE REPORT WAS SUBMITTED FOR AN INITIAL WHICH IS UPDATED TO THE FINAL IN THIS REPORT. ADDITIONALLY, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID WERE ALSO UPDATED IN THIS REPORT.
PATIENT INFORMATION CENTER IX | K211900 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THE PATIENT PASSED WAY WHILE CONNECTED TO THE MONITOR IN THE TIMEFRAME OF (B)(6) 2025. WHEN THE CUSTOMER WAS ATTEMPTING TO COLLECT THE LOGS ON (B)(6) 2025, THE APPLICATION SHUT DOWN; THEREFORE, THE CUSTOMER WAS REQUESTING ASSISTANCE IN DOWNLOADING THE LOGS. IN ADDITION, IT WAS INDICATED THE CUSTOMER ALLEGED NO FAILURE OR FAULT OF THE DEVICE AS IT RELATES TO THE PATIENT DEATH. A REVIEW OF THE LOGS BY PRODUCT SUPPORT ENGINEERING CONFIRMED THE DEVICE WAS FUNCTIONING AS EXPECTED DURING THE EVENT TIMEFRAME. | THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP | K200480 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CALLED AND WANTED TO KNOW IF SHE NEEDED TO RETURN THE REPLACEMENT DEVICE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S WIFE WAS INFORMED SHE COULD RETURN THE DEVICE IF SHE WANTED TO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAD PASSED AWAY. THE PATIENT'S WIFE CALLED AND WANTED TO KNOW IF SHE NEEDED TO RETURN THE REPLACEMENT DEVICE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S WIFE WAS INFORMED SHE COULD RETURN THE DEVICE IF SHE WANTED TO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H CODING IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A PATIENT'S SON STATING THE PATIENT DID NOT RECEIVE THE DEVICE. THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE TOTAL BLOWER HOURS WERE 1452.3 WITH NO ERRORS OR REPORTED FAULTS. THE BLOWER BOX ASSEMBLY WAS CONTAMINATED AND REPLACED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED AS PREVENTIVE MAINTENANCE. IN ADDITION, THE SD CARD WAS FITTED. THE DEVICE WAS CLEANED AND TESTED. THE DEVICE PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION CPAP PRO | K131982 | Insufficient Information; Sleepiness/Drowsiness/Somnolence; Hypoxia | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS REPORTED AS BEING RETURNED TO THE DME TWO YEARS AGO, HOWEVER THE DEVICE WAS NEVER RECEIVED BY THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS REPORTED AS BEING RETURNED TO THE DME TWO YEARS AGO, HOWEVER THE DEVICE WAS NEVER RECEIVED BY THE MANUFACTURER. THE MANUFACTURER RECEIVED ADDITIONAL INFORMATION FROM THE PATIENT'S DAUGHTER REGARDING THE EL SETTLEMENT PAYMENT IN REFERENCE TO THE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. SHE EXPRESSED, "MY DAD SHOULD NOT HAVE PASSED IF HE RECEIVED HIS NEW MACHINE ON TIME. WE TRIED FOR MONTHS." THE DAUGHTER SHARED THAT HER FATHER HAD BEEN USING THE CPAP DEVICE FOR APPROXIMATELY 4 OR 5 YEARS BEFORE HIS DEATH. SHE EXPLAINED THAT HER FATHER REGULARLY CLEANED HIS DEVICE ACCORDING TO THE INSTRUCTIONS, USING THE RECOMMENDED CLEANING SOLUTIONS AND WATER. WHILE SHE WAS UNSURE IF HE USED A SOCLEAN OR A SIMILAR DEVICE, SHE NOTED THAT HE ALWAYS KEPT UP WITH THE LATEST EQUIPMENT, SO IT IS POSSIBLE HE DID. THE DAUGHTER FURTHER MENTIONED THAT HER FATHER NEVER APPEARED WELL-RESTED, OFTEN REQUIRING OXYGEN AT 3 OR 4 LITERS AND USED THE CPAP MACHINE AT NIGHT. SHE RECALLED THAT HE HAD USED THE CPAP THE NIGHT BEFORE HIS PASSING BUT WAS STRUGGLING TO GET IT TO WORK PROPERLY. UPON WAKING, HIS OXYGEN LEVELS WERE IN THE 70S, AND HE HAD EXPERIENCED DIFFICULTIES WITH THE CPAP DEVICE OVER THE PAST FEW NIGHTS. SHE STATED THAT HER FATHER ALWAYS SEEMED TIRED AND GROGGY, AND ALTHOUGH HE INTENDED TO CONTACT PHILIPS ABOUT HIS DEVICE, SHE BELIEVES HE NEVER MANAGED TO DO SO BEFORE HE FELL ILL AND PASSED AWAY. THE DAUGHTER DESCRIBED A PARTICULARLY DIFFICULT NIGHT BEFORE HIS PASSING, WHEN HER FATHER HAD TROUBLE USING THE MACHINE, COULD NOT SLEEP WITH IT, AND FOUND IT MALFUNCTIONING. SHE ALSO REPORTED THAT HE RELIED ON OXYGEN, THE CPAP MACHINE, AND A NEBULIZER. THE DAUGHTER STATED HER FATHER PASSED AWAY FROM COPD AND "ALL THAT STUFF" RESTRICTIVE AIRWAYS, CHRONIC COPD, THE WHOLE REASON HE WAS ON THE NEBULIZER AND OXYGEN. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX A AGE AND SEX IN THIS REPORT. THE MANUFACTURER HAS UPDATED BOX B DATE OF DEATH AND OTHER RELEVANT HISTORY IN THIS REPORT. THE MANUFACTURER HAS UPDATED PRODUCT PROBLEM OR ADVERSE EVENT TO BOTH. THE PREVIOUSLY SUBMITTED REPORT WAS FILED AS AN ADVERSE EVENT ONLY. THE MANUFACTURER HAS UPDATED BOX E INITIAL REPORTER IN THIS REPORT. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | ON (B)(6) 2025 IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY THAT MORNING WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. ADDITIONAL INFORMATION WAS PROVIDED THROUGH FOLLOW-UP WITH THE PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS FOUND DECEASED BY THEIR SPOUSE ON (B)(6) 2025. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS UNKNOWN WHAT PHASE AND CYCLE OF THE PATIENT¿S TREATMENT THEY WERE IN AT THE TIME OF DEATH. NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT. THE PATIENT¿S SPOUSE REPORTED THAT EVERYTHING WAS FINE PRIOR TO THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. IT WAS REPORTED THAT CARDIAC ISSUES ARE SUSPECTED. THE DEATH IS NOT RELATED TO THE USE OF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. ALTHOUGH NO CAUSE OF DEATH WAS DOCUMENTED, CARDIAC ISSUES WERE SUSPECTED. CARDIOVASCULAR DISEASE IS THE LEADING CAUSE OF DEATH IN PATIENTS WITH CHRONIC KIDNEY DISEASE (CKD) WITH AN ESTIMATED 40%-50% OF ALL DEATHS IN PATIENTS WITH CKD STAGE 4 AND 5 BEING ATTRIBUTABLE TO CARDIOVASCULAR-RELATED MORTALITY. THIS PATIENT HAD SEVERAL COMORBID CARDIAC DISEASES INCLUDING CHRONIC SYSTOLIC HEART FAILURE. BASED ON THE AVAILABLE INFORMATION AND NO EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. IF ADDITIONAL INFORMATION IS OBTAINED ABOUT THIS EVENT, THE PMS CLINICAL EXPERT WILL REVIEW. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING EVALUATION AN ERROR INDICATING A REBOOT OF THE CENTRAL PROCESSING UNIT (CPU). PER THE DREAMSTATION SERVICE AND TECHNICAL REFERENCE MANUAL, THIS LEVEL ERROR DOES NOT AFFECT DEVICE FUNCTIONALITY. THE SERVICE TECHNICIAN RECOMMENDED THE DEVICE BE SCRAPPED. NO FURTHER INVESTIGATION DETAILS WERE PROVIDED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ION | K182188 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A CLINICAL STUDY REGISTRY PATIENT UNDERWENT AN UNEVENTFUL 19-MINUTE ION-ASSISTED LUNG BIOPSY. THE PROCEDURE WAS COMPLETED AS PLANNED, AND THE PATIENT WAS DISCHARGED THE SAME DAY. THERE WERE NO REPORTED DEVICE ISSUES OR COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. THE PATIENT¿S SPOUSE INFORMED THE PHYSICIAN THAT THE PATIENT EXPIRED 18 DAYS POST-PROCEDURE. THE REGISTRY STUDY PHYSICIAN DEEMED THE DEATH TO BE NOT RELATED TO THE PROCEDURE, NOT RELATED TO THE ION SYSTEM OR ACCESSORIES, AND PROBABLY RELATED TO THE PATIENT¿S PRE-EXISTING CO-MORBIDITIES, INCLUDING AN 80 PACK-YEAR HISTORY OF SMOKING. NO ADDITIONAL INFORMATION WAS AVAILABLE. | THE REGISTRY STUDY PHYSICIAN REPORTED THAT THE EVENT WAS PROBABLY RELATED TO THE PATIENT¿S EXISTING CONDITION(S), NOT RELATED TO THE PROCEDURE, NOT RELATED TO ION SYSTEM/ACCESSORIES. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WAS RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT UNDERWENT AN ION ENDOLUMINAL BIOPSY AS PART OF A CLINICAL STUDY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS DISCHARGED HOME THE SAME DAY. THE BIOPSY CONFIRMED A SQUAMOUS CELL CANCER. THE PATIENT DIED 18 DAYS AFTER THE PROCEDURE. THE DEATH WAS ATTRIBUTED TO THEIR MULTIPLE UNDERLYING MEDICAL CO-MORBIDITIES INCLUDING CHRONIC DEMYELINATING POLYNEUROPATHY, HEPATITIS B, IRON DEFICIENCY ANEMIA, COPD/EMPHYSEMA AND MODERATE PULMONARY FIBROSIS. BASED ON THE AVAILABLE DATA THE DEATH WAS DUE TO UNDERLYING MEDICAL CONDITIONS AND ALTHOUGH THERE IS NO CLEAR EVIDENCE OF ASSOCIATION WITH THE INDEX PROCEDURE, THE POSSIBILITY THAT THE EVENT WAS PROCEDURE RELATED CANNOT BE EXCLUDED. THE DEATH WAS IDENTIFIED AS PART OF A CLINICAL STUDY. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem; Degraded | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE WAS RECEIVED BY A THIRD-PARTY SERVICE CENTER, BUT THE EVALUATION HAS NOT YET BEGUN. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | THIS REPORT IS BEING SUBMITTED TO PROVIDE THE DEVICE'S THIRD-PARTY EVALUATION RESULTS. IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE RENTAL DREAMSTATION ST30 DEVICE WAS RETURNED AND EVALUATED BY A THIRD-PARTY SERVICE CENTER. THE DEVICE FOUND THAT NO FAULTS WERE REPORTED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED FOR PREVENTATIVE MAINTENANCE, AND THE ACCESSORY FLIP MODULE DOOR WAS MISSING FROM THE DEVICE. THE DEVICE WAS CLEANED AND TESTED, AND ALL TESTS PASSED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Bowel Perforation | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY DUE TO BOWEL PERFORATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION A USER OF A DREAMSTATION CPAP PRO PASSED AWAY THREE YEARS AGO. THE PATIENT'S SISTER STATES THEY ARE RECEIVING LETTERS REGARDING THE DEVICE AND THE PATIENT WANTED THE LETTERS TO STOP BEFORE PASSING AWAY. THE PATIENT'S SISTER COULD NOT RECALL WHAT HAPPENED WITH THE RECALLED DEVICE OR THE REPLACEMENT DEVICE. IT WAS NOT RECEIVED BY THE MANUFACTURER. THE CAUSE OF DEATH WAS NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION VIA EMAIL TO THE NHS (NATIONAL HEALTH SERVICE) (SBS.APINVOICING@NHS.NET) ON 06/30/2025, 07/03/2025, 07/09/2025. NO RESPONSE WAS RECEIVED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K091319 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING USER OF UNKNOWN DEVICE PASSED AWAY. THE PATIENT'S WIFE REPORTED THEY NEED TO RETURN THE PATIENT'S REPLACEMENT DEVICE SINCE THE PATIENT HAD PASSED AWAY BEFORE HE RECEIVED THE REPLACEMENT. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THIS REPORT WILL CAPTURE THE DEATH OF THE PATIENT ON THE DEVICE THAT WAS RECALLED SINCE THE PATIENT DID NOT USE THE REPLACEMENT DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING USER OF UNKNOWN DEVICE PASSED AWAY. THE PATIENT'S WIFE REPORTED THEY NEED TO RETURN THE PATIENT'S REPLACEMENT DEVICE SINCE THE PATIENT HAD PASSED AWAY BEFORE HE RECEIVED THE REPLACEMENT. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THIS REPORT WILL CAPTURE THE DEATH OF THE PATIENT ON THE DEVICE THAT WAS RECALLED SINCE THE PATIENT DID NOT USE THE REPLACEMENT DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE WAS IN GOOD CONDITION AND PASSED EVALUATION. THE CUSTOMER COMPLAINT WAS NOT ADDRESSED. THE DEVICE WAS SCRAPPED.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION VIA EMAIL TO THE NHS (NATIONAL HEALTH SERVICE) (SBS.APINVOICING@NHS.NET) ON 06/30/2025, 07/03/2025, 07/09/2025. NO RESPONSE WAS RECEIVED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | Cancer | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION DEVICE DEVELOPED MELANOMA AND HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Myocardial Infarction; Pulmonary Embolism | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED VENTRAL HERNIA ¿ COMPLEX TRANSVERSUS ABDOMINUS RELEASE (TAR) REPAIR, THE PATIENT WAS KEPT OVERNIGHT FOR OBSERVATION AND LABORATORY RESULTS. THE SURGEON INFORMED THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE (CSR) THAT THE PATIENT WAS DOING WELL POST-OPERATIVELY; HOWEVER, EXPIRED OVER NIGHT DUE TO UNKNOWN CAUSES. THE SURGEON REPORTED THAT THE CAUSE OF DEATH WAS LIKELY A HEART ATTACK OR A PULMONARY EMBOLISM. THE SURGEON SPECIFIED THAT THE PATIENT OUTCOME WAS NOT CAUSED BY THE DA VINCI-ASSISTED PROCEDURE AND THERE WERE NO DA VINCI ISSUES DURING THE PROCEDURE. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE REQUESTED; NO RESPONSE WAS RECEIVED. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, MARYLAND BIPOLAR FORCEPS, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING OR SUBSEQUENT TO THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED UNEXPECTEDLY AFTER AN UNEVENTFUL HERNIA REPAIR. NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT HAS BEEN MADE AND THE CASE HAD NO ISSUES. THE SURGEON STATED THE PATIENT MAY HAVE DIED OF A MYOCARDIAL INFARCTION OR PULMONARY EMBOLISM BUT NO SPECIFIC INFORMATION ALLUDING TO THESE CAUSES IS PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED - SURGEON REPORTED THE PATIENT EXPIRED DURING THE NIGHT, TIME NOT PROVIDED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY AND THE SON WOULD LIKE TO HAVE THE FATHER'S INFORMATION CLEARED AND HIS ENTERED SO HE COULD USE THE DEVICE. THE RESPIRATORY THERAPIST DIRECTED THE SON TO THE DME FOR THOSE CHANGES. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION AS THE SON WOULD LIKE TO USE THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION ST30. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT OF THE DREAMSTATION AUTO BIPAP DEVICE PASSED AWAY IN 2021 WHILE ON HOSPICE CARE. THE PATIENT WAS NOT USING THE DEVICE TOWARD THE END OF HER LIFE, AND THE DAUGHTER CONTACTED PHILIPS TO RECYCLE THE DEVICE. A SHIPPING LABEL WAS CREATED, AND THE DEVICE WAS RETURNED TO PHILIPS. THE DEVICE WAS EVALUATED, AND NO VISIBLE FOAM DEGRADATION OR PROBLEMS WERE FOUND. THERE IS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT PASSED AWAY ONE MONTH AGO AND THE REP RECRT DREAMSTATION AUTO CPAP DEVICE WAS THROWN AWAY AND THE PATIENT "WOULD NOT HAVE BEEN USING THIS DEVICE". A REPLACEMENT FOR THE DEVICE WAS DECLINED. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
ALARIS SYSTEM | K133532 | Insufficient Information | Corroded; Excess Flow or Over-Infusion; Fracture; Contamination | IT WAS INITIALLY REPORTED THE DEVICE HAD AN EVENT WHERE THE PUMP DELIVERED AN INCORRECT DOSE. THE DEVICE WAS INVESTIGATED BY THE FACILITY'S BIOMEDICAL ENGINEERING AND PER REPORT, IT "PASSED PREVENTATIVE MAINTENANCE AND RETURNED BACK TO SERVICE". THERE WAS PATIENT INVOLVEMENT HOWEVER OUTCOME IS UNKNOWN. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT INVOLVED A "COMFORT CARE PATIENT" ON A CONTINUOUS MORPHINE DRIP INFUSING AT 5MG/HR. FROM A 100ML (CONCENTRATION: 1MG/1ML) BAG. THE INFUSION WAS INITIATED AT 12:14. AT 14:09 THE NURSE RESPONDED TO THE ROOM AND FOUND THAT THE BAG WAS EMPTY. THE BAG AND TUBING WERE INSPECTED BY CUSTOMER; NO LEAKS WERE NOTED PER REPORT. IT WAS REPORTED THE PATIENT LATER PASSED AWAY. THE PATIENT'S FAMILY REQUESTED NO FURTHER INTERVENTIONS. THE EVENT OCCURRED ON 11 JANUARY 2025. PER CUSTOMER, THE EVENT WAS NOT IMMEDIATELY REPORTED TO BD AS THEY "DID NOT HAVE REASON TO BELIEVE THERE WAS AN EQUIPMENT MALFUNCTION." THE CUSTOMER, AT THE TIME, BELIEVED THAT THE ISSUE WAS DUE TO "DRUG DIVERSION". HOWEVER, THE SAME DEVICE HAS BEEN INVOLVED IN ANOTHER (CAPTURED UNDER SEPARATE RECORD MFR REPORT # 2016493-2025-77596) AND THIS PROMPTED THE CUSTOMER TO REPORT THE EVENT TO BD. | THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP. THE SPOUSE OF THE PATIENT REPORTED THE PATIENT HAS SADLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP. IT HAS BEEN REPORTED THE PATIENT HAS SADLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Insufficient Information; Exsanguination | Insufficient Information; Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING A DA VINCI ASSISTED NEPHRECTOMY, THERE WAS AN UNSPECIFIED SURGICAL COMPLICATION RESULTING IN MAJOR HEMORRHAGING. THE SURGEON REPORTED THAT THE BLEEDING REQUIRED A CONVERSION TO OPEN LAPAROTOMY AND AN OPEN THORACOTOMY. DESPITE ALL RESUSCITATIVE EFFORTS THE PATIENT EXPIRED IN THE OPERATING ROOM. THE SURGEON REPORTED THAT THE SURGICAL COMPLICATION WAS NOT RELATED TO THE ROBOTIC SYSTEM AND THERE WAS NO MALFUNCTION OF THE SYSTEM OR INSTRUMENTS. THE STATED CAUSE OF DEATH WAS MAJOR HEMORRHAGE. ADDITIONAL EVENT INFORMATION WAS REQUESTED BUT WAS NOT PROVIDED BY THE SURGEON. | THE FOLLOWING CONCOMITANT DEVICES WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, VESSEL SEALER EXTEND. A SITE REVIEW FOUND THAT NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THE PRODUCTS USED. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICES USED DURING THE EVENT FOUND NO RELATED NON-CONFORMANCES. DIGITAL SIGNAL PROCESSOR (DSP) LOGS SHOW THAT THE VESSEL SEALER EXTEND (VSE) INSTRUMENT WAS INSTALLED 3 TIMES AND PASSED HOMING 3 TIMES. ALL CUT COMPLETE AND SEAL EVENTS WERE NOTED WITH NO ERRORS. THE MASTER SUPERVISORY CONTROLLER (MSC) LOGS DID NOT SHOW ANY VSE RELATED ERRORS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT, WITH A SIGNIFICANT COMORBID DIAGNOSIS OF XANTHOGRANULOMATOUS PYELONEPHRITIS, WAS UNDERGOING A LEFT NEPHRECTOMY WHEN SIGNIFICANT BLEEDING OCCURRED. WHAT BLED AND HOW THEY GOT INTO THAT BLEEDING WAS NOT PROVIDED. THE PROCEDURE WAS CONVERTED TO AN OPEN LAPAROTOMY AND ALSO REQUIRED A THORACOTOMY WHICH SUGGESTS THEY HAD TO EMERGENTLY GET CONTROL OF THE AORTA. HOW THESE EVENTS UNFOLDED ARE NOT PROVIDED, HOWEVER, THE PATIENT EVENTUALLY DIED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
ION | K182188 | Low Blood Pressure/ Hypotension; Hemodynamic instability | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT'S BLOOD PRESSURE DROPPED SIGNIFICANTLY, CAUSING HEMODYNAMIC INSTABILITY LEADING TO DEATH. THE PROCEDURE INVOLVED ION NAVIGATION AND RADIAL ENDOBRONCHIAL ULTRASOUND (REBUS), IT WAS INTERRUPTED BEFORE THE BIOPSY COULD BE PERFORMED. APPROXIMATELY 10-15 MINUTES INTO NAVIGATION, THE ANESTHESIOLOGIST INFORMED THE PROCEDURE TEAM OF THE PATIENT¿S ¿DANGEROUSLY LOW¿ BLOOD PRESSURE. THE ION SYSTEM WAS UNDOCKED, AND CARDIOPULMONARY RESUSCITATION WAS INITIATED. THE PATIENT WAS PRONOUNCED DEAD APPROXIMATELY 45 MINUTES LATER. THERE WAS NO REPORTED DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN WERE MADE; NO RESPONSE HAS BEEN RECEIVED. | A REVIEW OF THE ION SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND NO RELATED SYSTEM ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION BRONCHOSCOPY, AND DURING NAVIGATION AND REBUS IMAGING, PRIOR TO BIOPSY, THE PATIENT DEVELOPED HYPOTENSION. EFFORTS AT RESUSCITATION WERE UNSUCCESSFUL AND THE PATIENT DIED. THERE WERE NO REPORTED MALFUNCTIONS OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. NO FURTHER DATA IS AVAILABLE DESPITE MULTIPLE EFFORTS. BASED ON THE AVAILABLE DATA THE EVENTS WERE PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. .
ION | K182188 | Insufficient Information; Pulmonary Embolism; Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, A PATIENT EXPERIENCED SIGNIFICANT BLEEDING, PULMONARY EMBOLISM, CARDIAC ARREST AND DEATH. THE INTUITIVE ENDOLUMINAL SALES MANAGER (ESM) WAS INFORMED BY HOSPITAL PERSONNEL THAT DURING THE BIOPSY PROCEDURE, A THIRD PARTY CRYOPROBE WAS USED TO COLLECT THE BIOPSY SAMPLES; THE BIOPSY SPECIMEN APPEARED SUGGESTIVE OF BLOOD VESSEL INVOLVEMENT. FOLLOWING THE BIOPSY, THE PATIENT HAD SIGNIFICANT BLEEDING, PROMPTING A BRONCHOSCOPIC AIRWAY INSPECTION AND SUCTIONING OF THE BLEEDING. CONE BEAM CT SCAN SHOWED A PULMONARY EMBOLISM. SHORTLY THEREAFTER, THE PATIENT SUFFERED CARDIAC ARREST AND REQUIRED RESUSCITATION, INCLUDING TRANSFUSION OF BLOOD PRODUCTS. THE PATIENT WAS STABILIZED; HOWEVER, EXPIRED AT AN UNSPECIFIED LATER TIME. THE ESM WAS INFORMED THAT THE ADVERSE EVENTS WERE NOT ATTRIBUTED TO THE ION SYSTEM, AND NO DEVICE MALFUNCTION WAS IDENTIFIED. THE COMPLICATIONS WERE BELIEVED TO BE RELATED TO THE CRYOPROBE BIOPSY THAT INDICATED VESSEL INVOLVEMENT. ADDITIONAL INFORMATION WAS REQUESTED FROM THE RESPIRATORY MEDICINE PHYSICIAN WHO STATED THE FACTS PROVIDED WERE INCORRECT, AND THAT THE MATTER IS CONFIDENTIAL AND INTERNAL REVIEW IS IN PROGRESS. NO CLARIFICATION OF THE INCORRECT FACTS WAS PROVIDED. | SYSTEM LOGS ARE NOT AVAILABLE FOR REVIEW. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY WHERE A CRYO-PROBE WAS USED TO OBTAIN BIOPSY SAMPLES. DURING THE PROCEDURE, AFTER BIOPSIES WERE OBTAINED, SIGNIFICANT BLEEDING WAS NOTED AND A PULMONARY EMBOLISM WAS DIAGNOSED ON CONE BEAM CT. THE PATIENT FURTHER DECOMPENSATED WITH A CARDIAC ARREST. DESPITE AGGRESSIVE INTERVENTIONS AND EFFORTS AT RESUSCITATION, THE PATIENT ULTIMATELY DIED. BASED ON THE AVAILABLE DATA THE BLEEDING, PULMONARY EMBOLISM, AND DEATH POST BIOPSY WERE LIKELY PROCEDURE RELATED. THERE IS NO COMPELLING EVIDENCE THAT THE EVENTS WERE DEVICE RELATED. OF NOTE, THE PHYSICIAN REPORTED DISCREPANCIES IN THE DETAILS BUT DECLINED TO PROVIDE FURTHER INFORMATION. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW-RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE-CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. THE RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). THE PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). THE RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). THERE WERE NO REPORTED EVENTS OF VENOUS THROMBOEMBOLISM (VTE) INCLUDING DEEP VENOUS THROMBOSIS OR PULMONARY EMBOLISM. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS AND NO VTE. HOWEVER, IT IS POSSIBLE THAT IMMOBILITY DURING GENERAL ANESTHESIA AND BLEEDING WITH SUBSEQUENT CLOTTING ASSOCIATED WITH LUNG BIOPSIES MAY CONFER SOME DEGREE OF INCREASED RISK FOR VTE EVENTS. LIKELY OF MORE RELEVANCE, IT IS ESTIMATED THAT CANCER CONFERS A 4-7X HIGHER RISK FOR VTE. THE HIGHER RATE OF VTE HAS BEEN REPORTED TO RANGE BETWEEN 3-13.9% IN PATIENTS WITH LUNG CANCER. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLOM JW, DOGGEN CJ, OSANTO S, ROSENDAAL FR. MALIGNANCIES, PROTHROMBOTIC MUTATIONS, AND THE RISK OF VENOUS THROMBOSIS. JAMA. 2005. BAGCHI A, KHAN MS, SARASWAT A, ANSARI A, NAI Q, IYER V, HAMOUDA D, KHUDER S, VERGHESE C. INCREASED INCIDENCE OF THROMBOTIC COMPLICATIONS WITH NON-SMALL CELL LUNG CANCER NECESSITATES CONSIDERATION OF PROPHYLACTIC ANTICOAGULATION IN YOUNG INDIVIDUALS. CUREUS. 2021. CHEW HK, DAVIES AM, WUN T, HARVEY D, ZHOU H, WHITE RH. THE INCIDENCE OF VENOUS THROMBOEMBOLISM AMONG PATIENTS WITH PRIMARY LUNG CANCER. J THROMB HAEMOST. 2008. CONNOLLY GC, DALAL M, LIN J, KHORANA AA. INCIDENCE AND PREDICTORS OF VENOUS THROMBOEMBOLISM (VTE) AMONG AMBULATORY PATIENTS WITH LUNG CANCER. LUNG CANCER. 2012.
REMSTAR AUTOA-FLEX | K113068 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REMSTAR AUTOA-FLEX DEVICE HAS PASSED AWAY AND THE DAUGHTER WANTED TO RETURN THE DEVICE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY AND WIFE REPORTED THAT SHE RECEIVED A LETTER STATING TO RETURN THE RECALLED DEVICE. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND VISIBLE FOAM PARTICLES WERE NOT OBSERVED. ADDITIONALLY, THE PATIENT¿S COMPLAINT WAS NOT CONFIRMED, AND THE DEVICE WAS SCRAPPED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ALARIS SYSTEM | K133532 | Insufficient Information | Use of Device Problem | IT WAS REPORTED THAT THERE WAS AN OVER INFUSION OF HEPARIN TO A PATIENT ADMITTED WITH UPPER BACK PAIN, WHO "CODED" FOLLOWING ADMISSION AND WAS ADMITTED TO ICU WITH POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE HEPARIN (HEPARIN 25,000 UNITS/500ML) WAS INTENDED TO INFUSE AT A RATE OF 16ML/HOUR (800 UNITS/HOUR), HOWEVER THE MEDICATION BAG WAS FOUND TO BE EMPTY UNEXPECTEDLY QUICKLY. THE HEPARIN INFUSION WAS STARTED AT 0401, AND THE ENTIRE 500ML BAG WAS NOTED TO BE EMPTY AT 0520. THE PATIENT'S PARTIAL THROMBOPLASTIN TIME WAS 25.3 SECONDS ON THE INITIAL BLOOD DRAW AT 0120, THEN JUMPED TO >169 SECONDS FOLLOWING THE OVER INFUSION. THE PATIENT DID NOT RECEIVE THE ANTIDOTE BECAUSE THEIR BLOOD PRESSURE WAS TOO LOW, ACCORDING TO THE INTENSIVIST. THE PATIENT EXPIRED LATER THAT MORNING. THE CLINICIAN INDICATED THE CAUSE OF DEATH WAS MULTIFACTORIAL. THE CLINICIAN INDICATED THE PATIENT HAD A LARGE POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION, "CODED", ACHIEVED RETURN OF SPONTANEOUS CIRCULATION, AND WAS ON MULTIPLE VASOPRESSORS. A REPORT WAS RECEIVED FROM HEALTH CANADA'S CANADA VIGILANCE PROGRAM WHICH STATES, "THE PUMP WAS PROGRAMED FOR 800UTS/HR. AND SHOULD HAVE INFUSED AT 16ML/HR. BUT THE PATIENT RECEIVED AN ENTIRE BAG OF HEPARIN IN APPROXIMATELY 1 HOUR. THE CONSEQUENCES OF THAT WERE AN EXTREMELY HIGH PTT. THE PATIENT WAS NOT ABLE TO RECEIVE THE ANTIDOTE DUE TO CLINICAL CONDITION (LOW BP) AS DECIDED BY ICU. THE PATIENT PASSED BUT THE CAUSE OF THIS WAS MULTIFACTORIAL." | ADDITIONAL INFORMATION: OTHER RELEVANT HISTORY, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX B CODE, REMEDIAL ACTION REQUIRED, REMEDIAL ACTION # AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE REPORT OF AN OVER-INFUSION OF HEPARIN WAS CONFIRMED THROUGH LOG ANALYSIS AND ATTRIBUTED TO A USER PROGRAMMING ERROR, WHICH REVEALED A DISCREPANCY BETWEEN THE REPORTED HEPARIN CONCENTRATION OF 25000U/500 ML AND THE ENTERED CONCENTRATION OF 800 U/500 ML. THE CONCENTRATION ENTRY ERROR RESULTED IN A PUMP CALCULATED RATE OF 500 ML/H WHEN A DOSE OF 800 U/H WAS ENTERED. THE REPORT INDICATED THE DESIRED RATE WAS 16 ML/HOUR. THE EVENT LOG REVIEW COULD NOT CONFIRM THE REPORT THAT ON THE DATE (B)(6) 2024 AT THE TIME OF 0401 AN INFUSION HEPARIN WAS STARTED AND AT THE TIME OF 0520 THE HEPARIN BAG WAS FOUND EMPTY. THE INSPECTION AND TESTING PROCESS DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATIONS WITH NO OBSERVANCES OF UNREGULATED FLOW OCCURRING. THE SEAR WAS ALSO FOUND TO BE PROPERLY CLOSING THE ADMINISTRATION SET SAFETY CLAMP WHEN THE DOOR WAS OPENED. AFTER UNREGULATED FLOW TESTING, THE SUSPECT PUMP MODULE WAS PROGRAMMED TO PERFORM A ONE-HOUR SYSTEM LEVEL RATE ACCURACY TEST AT THE INCIDENT INFUSION RATE OF 16ML/HR AND TEST RESULTS MEASURED THE ACTUAL RATE AT 15.43ML/HR (-3.55% ERROR) INDICATING THE DEVICE IS WITHIN SPECIFICATION AFTER RATE CALIBRATION HAVING NO INCIDENTS OF UNREGULATED FLOW OCCURRING DURING THE TESTING. NO ERRORS WERE OBSERVED ON BOTH PCU AND SUSPECT PUMP MODULE ON THE REPORTED DATE OF THE EVENT. REVIEW OF THE SUSPECT PUMP MODULE ERROR LOG SHOWED NO MALFUNCTIONS RELEVANT TO THE FACILITY¿S COMPLAINT. ON THE DATE (B)(6) 2024 AT THE TIME OF 2:25 AM, THE CLINICIAN WAS RECORDED TO SELECT HEPARIN (ACS) (DRUGID=251) FROM GUARDRAIL DRUGS. THE HEPARIN SELECTION WAS A CUSTOM CONCENTRATION WITH THE CLINICIAN ENTERING A DRUG AMOUNT OF 800 UNITS AND A DILUENT VOLUME 500 ML. THE INTENDED CONCENTRATION HAD BEEN 25000U/500 ML. THE ENTRY ERROR RESULTED IN A SOFT GUARDRAIL CONCENTRATION ALERT WHICH THE USER ACCEPTED AND CONTINUED PROGRAMMING THE INFUSION. THE DOSE WAS ENTERED CORRECTLY AT 800 U/HR BUT BECAUSE OF THE CONCENTRATION ERROR THE INFUSION RATE CALCULATED TO 500 ML/H INSTEAD OF THE INTENDED 16 ML/H. THE INFUSION STARTED AT 2:27 AM BUT WAS PAUSED A SHORT TIME LATER BY THE CLINICIAN WITH A PRIMARY VOLUME INFUSED OF 4.171 ML. AT 2:31 AM THE SUSPECT WAS PROGRAMED TO CONTINUE THE INFUSION WITH A VOLUME TO BE INFUSED OF 495.004ML. THE INFUSION FINISHED AT 3:33 AM WITH A REPORTED PRIMARY VOLUME INFUSED OF 511.403ML ML AFTER INFUSING FOR 61 MINUTES. THE ADMINISTRATION SET WAS REQUESTED BUT WAS NOT RETURNED; THEREFORE, IT COULD NOT BE INSPECTED OR TESTED. ROOT CAUSE: THE ROOT CAUSE FOR THE REPORT OF OVER INFUSION OF HEPARIN WAS ABLE TO BE IDENTIFIED FROM THE LOG ANALYSIS AND IS BEING ATTRIBUTED TO USER PROGRAMMING ERROR. THE SUSPECT PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATION, AND THE SEAR WAS FOUND TO BE PROPERLY CLOSING THE SAFETY CLAMP WHEN THE ADMINISTRATION SET IS REMOVED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | A PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO AN UNREPORTED CAUSE. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO THEIR LIBERTY SELECT CYCLER AT THE TIME OF DEATH. THE PATIENT WAS FOUND BY FAMILY TO BE UNRESPONSIVE AND PULSELESS ON THE MORNING OF (B)(6) 2025. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS NOT TRANSPORTED AND PRONOUNCED DECEASED IN HOME. THE CAUSE OF DEATH WAS NOT REPORTED TO THE OUTPATIENT CLINIC; HOWEVER, THE PATIENT HAD A RECENT DECLINE IN HEALTH INVOLVING ADVANCED DEMENTIA THAT PREEXISTED HIS START ON PD THERAPY IN (B)(6) 2025. IT WAS CONFIRMED, THOUGH THE CAUSE OF DEATH WAS NOT REPORTED TO THE CLINIC, THE PATIENT¿S DEATH WAS DETERMINED TO BE UNRELATED TO PD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CANNOT BE DETERMINED; HOWEVER, THERE WAS NO INDICATION THE PATIENT¿S DEATH WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
REMSTAR PRO C-FLEX+ | K091319 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED DEVICE TURNS OFF RANDOMLY. THE PATIENT WAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE CUSTOMER ALLEGED DEVICE TURNS OFF RANDOMLY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. AFTER MULTIPLE ATTEMPTS, THE CUSTOMER CONFIRMED THAT THE DEVICE DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THIS REPORT AND NO INCREASED RISK TO THE INTENDED USER. BASED ON THE INFORMATION AVAILABLE, IT IS A NON-REPORTABLE COMPLAINT.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER AND INSPECTED. NO ERRORS WERE FOUND; THEREFORE, NO COMPONENTS WERE REPLACED TO CORRECT A MALFUNCTION. THE DEVICE PASSED ALL TESTS AND WAS RETURNED TO STOCK.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Thrombocytopenia; Hemorrhage/Blood Loss/Bleeding; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION B5 ADDITIONAL INFORMATION: PRIOR TO THE PATIENT'S CARDIAC ARREST, BLOOD WAS NOTED IN THE SURGICAL DRAIN; THERE WAS NO BLOOD IN THE DRAIN EARLIER IN THE DAY. AN EMERGENCY EXPLORATORY LAPAROTOMY WAS PERFORMED, REVEALING A HEMOPERITONEUM. THE BLEEDING SOURCE WAS IDENTIFIED AT THE JUNCTION BETWEEN THE PANCREATIC HEAD AND BODY, SPECIFICALLY AT THE SECTIONED RAW TRANSITION AREA OF THE "RESIDUAL PANCREATIC PARENCHYMA". THIS HEMORRHAGE WAS ATTRIBUTED TO THROMBOCYTOPENIA. A REVIEW OF THE ADDITIONAL EVENT INFORMATION PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED FOLLOWING A LATE BLEEDING EVENT SEVERAL DAYS AFTER A PANCREATICODUODENECTOMY. THE LOCATION OF THE BLEED WAS NOTED TO BE THE RAW SURFACE OF THE PANCREATIC TRANSECTION. THE CAUSE OF THE BLEED IS UNCLEAR ALTHOUGH IT WAS NOTED THE PATIENT HAD BECOME THROMBOCYTOPENIC AND MAY HAVE DEVELOPED COAGULOPATHY. THE REASON FOR THE COAGULOPATHY IS UNCLEAR. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
PATIENT INFORMATION CENTER IX | K211900 | Bradycardia; Asystole | No Audible Alarm | IT WAS REPORTED THE ALARM DID NOT SOUND AND THE INTENSIVE CARE UNIT (ICU) PATIENT PASSED AWAY. | A PHILIPS REMOTE SERVICE ENGINEER (RSE) SPOKE WITH THE CUSTOMER. THE EVENT TOOK PLACE ON (B)(6) 2025 AT BED D137. WHILE A SPECIFIC TIME FOR THE FAILURE WAS NOT PROVIDED IN THE ALLEGATION, THE REMOTE INVESTIGATION FOCUSED ON THE TIMELINE OF 13:22 THROUGH THE PATIENT DISCONNECTION FROM THE MONITOR AT 13:58. THE PATIENT WAS BEING MONITORED FOR ECG AND FOR SPO2. PER THE LOG REVIEW, 'DURING THE OBSERVATION PERIOD, A TOTAL OF EIGHT EXTREME BRADYCARDIA RED ALARMS WERE RECORDED, ACCOMPANIED BY SEVERAL YELLOW ALARMS, IN A CONTEXT MARKED BY REPEATED TECHNICAL ALERTS CONCERNING THE SPO2 SENSOR (SIGNAL DIFFICULT TO CAPTURE, AGITATED PATIENT, POORLY POSITIONED SENSOR). THE FIRST RED ALARMS OCCURRED FROM 13:22:23, WITH A SUCCESSION OF CRITICAL EVENTS UNTIL 13:26:02. NO MANUAL ACKNOWLEDGMENT WAS PERFORMED DURING THIS INITIAL SEQUENCE. THE FIRST MANUAL ACKNOWLEDGMENT OF A RED ALARM AT THE CENTRAL UNIT WAS RECORDED AT 13:27:41, FOR AN ALARM GENERATED AROUND 13:26. THE ASYSTOLE RED ALARMS BEGAN AROUND 13:29, WITH A MANUAL ACKNOWLEDGMENT AT THE CENTRAL UNIT 30 SECONDS LATER. OTHER REPETITIONS OF THE ASYSTOLE ALARM WERE OBSERVED, NOTABLY AT 1:32 P.M., 1:35 P.M., 1:39 P.M., 1:42 P.M. AND 1:46 P.M., WITH SUCCESSIVE ACKNOWLEDGMENTS AT THE CENTRAL UNIT BUT ALSO AT THE MONITOR IN THE PATIENT'S ROOM, THE FIRST BEING NOTED AT 1:40:54 AND THE SECOND AT 1:46 P.M. THE SPO2 SENSOR, OFTEN DISCONNECTED OR POORLY POSITIONED, COMPLICATED THE QUALITY OF MONITORING. THE EPISODE ENDED WITH THE PATIENT BEING DISCONNECTED AT 1:58 P.M., PUTTING AN END TO THE ASYSTOLE ALARMS.' THE SPO2 SENSOR, WHICH WAS OFTEN DISCONNECTED OR POORLY POSITIONED, COMPLICATED THE QUALITY OF THE MONITORING. THE EPISODE ENDED WITH THE PATIENT BEING DISCONNECTED AT 13:58, ENDING THE ASYSTOLE ALARMS. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL PRODUCT SPECIALIST (CPS) REVIEWED THE PROVIDED INFORMATION AND CONFIRMED THE RSE FINDINGS. THE PSE ALSO REVIEWED PIC IX TECHNICAL LOG DATA FROM THE "SURVEILLANCE" HOST. THE PSE NOTED SEVERAL THREE STAR (***) XBRADY ALARMS AND AN ASYSTOLIE SENT TO THE RECORDER. THE TIMESTAMPS CORRELATE WITH THE SOUND MANAGER ENTRIES FROM THE PIC, INDICATING THE RED ALARM SOUND BEING PLAYED AT THE CENTRAL STATION. THE REPORTED ISSUE WAS NOT CONFIRMED. THE CAUSE OF THE DEATH AND THE RELATION TO THE DEVICE REMAINS UNKNOWN DUE TO LIMITED INFORMATION FROM THE CUSTOMER. THE ORIGINAL EVENT DESCRIPTION INDICATES THE PATIENT WAS NOT TREATED IN TIME. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE A DEVICE MALFUNCTION WHICH CAUSED OR CONTRIBUTED TO THE REPORTED EVENT; HOWEVER, ALARM MANAGEMENT, CLINICAL WORKFLOW, OR OTHER USE ISSUE MAY HAVE BEEN FACTORS.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE HAD FOUND THE DEVICE AFTER THE PATIENT PASSED AWAY A FEW YEARS AGO AND WOULD LIKE IT REPLACED. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION IT IS OBSERVED THAT THERE ARE 0 ERRORS FOUND IN LOGS, NO FAULTS FOUND, AND ALL TESTING PASSED.
DAVINCI XI | K131861 | Septic Shock; Ischemia; Coagulation Disorder | Adverse Event Without Identified Device or Use Problem | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION B5 ADDITIONAL INFORMATION: THE RESECTIONS AND ANASTOMOSES WERE PERFORMED ROBOTICALLY USING A THIRD-PARTY STAPLER, LIGASURE, AND A DA VINCI ROBOTIC MARYLAND BIPOLAR FORCEPS. THE PANCREATICOJEJUNOSTOMY AND HEPATICOJEJUNOSTOMY WERE SUTURED, WHILE THE GASTROJEJUNOSTOMY WAS STAPLED. THE PROCEDURE WAS THEN CONVERTED FROM ROBOTIC TO LAPAROSCOPY PRIOR TO PERFORMING LYMPHADENECTOMY AND RESECTING THE SECOND MESENTERIC TUMOR. THIS DECISION WAS MADE DUE TO THE TUMOR'S LOCATION, WHICH WOULD HAVE NECESSITATED REPOSITIONING ALL ROBOTIC PORT SITES. DURING THE LYMPHADENECTOMY, APPROXIMATELY 700 ML OF RETROPERITONEAL BLEEDING OCCURRED, RESULTING IN THE PATIENT EXPERIENCING TACHYCARDIA AND HYPOTENSION. HEMOSTASIS WAS ACHIEVED, AND THE LYMPHADENECTOMY WAS SUCCESSFULLY COMPLETED. THE MESENTERIC TUMOR, INITIALLY IDENTIFIED AS A MESENTERIC LYMPH NODE, WAS RESECTED. PATHOLOGY CONFIRMED IT TO BE A NEUROENDOCRINE TUMOR CONCURRENT WITH DUODENAL CARCINOMA. THE RESECTION RESULTED IN APPROXIMATELY 1 LITER OF BLOOD LOSS DUE TO THE TUMOR'S LOCATION AND ADHERENCE. HEMOSTASIS WAS ACHIEVED USING A LIGASURE DEVICE AND SUTURING. THIRTY-SIX HOURS POST-OPERATIVELY, THE PATIENT DEVELOPED MESENTERIC ISCHEMIA, SECONDARY TO THE RESECTION OF THE NEUROENDOCRINE MESENTERIC TUMOR. A LAPAROSCOPIC PROCEDURE WAS PERFORMED FOR AN ENTERECTOMY AND AN ENTEROANASTOMOSIS USING A THIRD-PARTY STAPLER INSTRUMENT. THE PATIENT LATER DEVELOPED A PANCREATIC FISTULA WITH A SUBHEPATIC COLLECTION, WHICH REQUIRED TWO ADDITIONAL LAPAROSCOPIC PROCEDURES FOR REVISION AND DRAINAGE. ON POST-OPERATIVE DAY 19, THE PATIENT DEVELOPED SEPTIC SHOCK AND WAS EXPERIENCING A LEAKAGE OF SECRETIONS THROUGH ONE OF THE ROBOTIC PORT SITES. AN EXPLORATORY LAPAROTOMY FOUND A FISTULA HAD FORMED AT THE ENTERO-ETEROANASTOMOSIS. LATER THAT DAY, THE PATIENT¿S CONDITION FURTHER DETERIORATED, RESULTING IN BLOOD DYSCRASIA AND INSTABILITY. A SECOND EXPLORATORY LAPAROTOMY WAS PERFORMED REVEALING A HEMOPERITONEUM, THOUGH THE SOURCE OF THE BLEEDING COULD NOT BE IDENTIFIED. DESPITE THESE EFFORTS, THE SEQUENCE OF EVENTS LEAD TO THE PATIENT DYING THE NEXT MORNING, ON POST-OPERATIVE DAY 20. A REVIEW OF THE ADDITIONAL EVENT INFORMATION PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT HAD A COMPLEX AND CHALLENGING TUMOR IN WHICH THEY UNDERWENT AN INITIAL RESECTION WHICH INCORPORATED ROBOTIC INSTRUMENTATION AS WELL AS SEVERAL 3RD PARTY DEVICES. A MESENTERIC EXTENSION OF THE TUMOR IN THE FORM OF A LARGE LYMPH NODE NECESSITATED A BOWEL RESECTION WHICH WAS PERFORMED USING A 3RD PARTY DEVICE AND THAT BOWEL WENT ON TO BECOME ISCHEMIC REQUIRING ADDITIONAL EXPLORATION AND RESECTION. A LEAK ALSO DEVELOPED BETWEEN AN ANASTOMOSIS FOLLOWING THE SECOND OPERATION AND FURTHER COMPLICATIONS INCLUDING A FISTULA OCCURRED. THE ULTIMATE CAUSE OF DEATH WAS SEPTIC AND HYPOVOLEMIC SHOCK. HOW THESE COMPLICATIONS MAY OR MAY NOT HAVE BEEN RELATED TO THE INITIAL SURGERY AND ANY ROBOTIC DEVICE IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | Respiratory Arrest; Oversedation; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY PROCEDURE, THE PATIENT EXPIRED. THE SURGEON STATED THAT THE ROBOTIC PROCEDURE WAS COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS; THERE WERE NO DA VINCI-RELATED ERRORS OR ISSUES. POSTOPERATIVELY, WHILE THE PATIENT WAS RECOVERING IN THE POST ANESTHESIA CARE UNIT (PACU), NARCOTICS WERE ADMINISTERED FOR PAIN. DUE TO A MISADMINISTRATION OF NARCOTICS IN THE PACU, THE PATIENT EXPIRED. THE PARTNER OF THE SURGEON CLARIFIED THAT THE PATIENT WAS RECOVERING IN THE OUTPATIENT SURGICAL CENTER¿S PACU AND WAS ADMINISTERED FENTANYL. THE SURGICAL CENTER'S PACU DID NOT HAVE NARCAN AVAILABLE AT THE TIME OF THE CODE. PARAMEDICS WERE CALLED AND ADMINISTERED NARCAN UPON ARRIVAL. ADVANCED RESUSCITATION ACTIVITIES AND CHEST COMPRESSIONS WERE CONTINUED FOR TWO HOURS. AN INITIAL POST-PROCEDURE ULTRASOUND SHOWED NO SIGNS OF BLOOD WITHIN THE ABDOMEN. AT AN UNSPECIFIED TIME DURING RESUSCITATION, A REPEAT ULTRASOUND DID SHOW SIGNS OF BLOOD IN THE ABDOMEN. IT WAS PLANNED TO TRANSFER THE PATIENT TO THE HOSPITAL'S MAIN OPERATING ROOM FOR A REOPERATION DUE TO THE FINDINGS; HOWEVER, THE PATIENT EXPIRED IN THE PACU AND THE SUBSEQUENT PROCEDURE DID NOT OCCUR. IT WAS STATED THAT THE NEW BLEEDING WITHIN THE ABDOMEN WAS BELIEVED TO BE DUE TO THE CHEST COMPRESSIONS, POSSIBLY CRACKED RIBS. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES FOUND THERE WERE NO INTERRUPTIONS OR ERROR EVENTS IMPACTING ANY OF THE PROCEDURES. THE SYSTEM WAS FUNCTIONING TO MANUFACTURING SPECIFICATIONS. THE DAY AFTER THE EVENT, THE CUSTOMER REQUESTED A FIELD SERVICE ENGINEER (FSE) TO CHECK OUT THE SYSTEM EVEN THOUGH ¿THEY DID NOT BELIEVE THE ROBOT HAD ANYTHING TO DO WIGHT THE ADVERSE EVENT¿ TO VERIFY THE EQUIPMENT. THE FSE TESTED THE SYSTEM, INCLUDING THE E-100 AND ERBE GENERATORS WITH NO ISSUES FOUND. ELECTRICAL SAFETY TESTING OF ALL THE CARTS AND GENERATORS WAS ALSO PERFORMED, FOLLOWED BY A SECOND SYSTEM TEST DRIVE. NO ISSUES WERE FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, PERMANENT CAUTERY HOOK, MARYLAND BIPOLAR FORCEPS, CADIERE FORCEPS, PROGRASP FORCEPS, MEDIUM-LARGE CLIP APPLIERS X2. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING OR SUBSEQUENT TO THIS PROCEDURE. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICES USED DURING THE REPORTED EVENT FOUND NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A CHOLECYSTECTOMY, WHERE NO INTRAOPERATIVE COMPLICATIONS OR ISSUES WERE REPORTED. POSTOPERATIVELY, THE PATIENT RECEIVED FENTANYL AND CODED. THE TEMPORAL RELATIONSHIP OF THE CARDIOVASCULAR COLLAPSE SUGGESTS THE EVENT WAS PRECIPITATED BY THE NARCOTIC AS AN INITIAL ULTRASOUND SHOWED NO EVIDENCE OF FLUID IN THE ABDOMEN. ALTHOUGH A SUBSEQUENT ULTRASOUND EVENTUALLY SHOWED FLUID IN THE ABDOMEN, THIS FINDING WAS SPECULATED TO BE A RESULT OF THE RESUSCITATION EFFORTS. CONTINUED RESUSCITATION EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS SPECULATED TO HAVE BEEN PRECIPITATED BY THE ADMINISTRATION OF THE FENTANYL NARCOTIC. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
PHILIPS CPAP DEVICE | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING HUSBAND PASSED AWAY FROM CANCER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Dizziness; Dyspnea; Rash; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) WAS HOSPITALIZED AND EXPIRED ON (B)(6) 2025 APPROXIMATELY 26 MINUTES AFTER CCPD THERAPY BEGAN. FOLLOW-UP WITH THE PATIENT¿S INPATIENT PROGRAM MANAGER (IPM) REVEALED THE PATIENT PRESENTED TO THE EMERGENCY ROOM ON (B)(6) 2025 WITH COMPLAINTS OF DYSPNEA, DIZZINESS, AND A RASH. ALTHOUGH THE TIMELINE/SPECIFICS OF THE SERIOUS ADVERSE EVENTS ARE LARGELY UNKNOWN (PATIENT¿S ELECTRONIC MEDICAL RECORD CLOSED), THE IPM WAS ABLE TO RECALL THE PATIENT WAS DIAGNOSED WITH AN ALMOST COMPLETE ATRIOVENTRICULAR (AV) BLOCK (WHICH WOULD HAVE REQUIRED A PACEMAKER), AND A SEVERE CASE OF ADULT-ONSET HAND/FOOT/MOUTH DISEASE. THE IPM STATED THE PATIENT HAD A DO NOT RESUSCITATE (DNR) ORDER IN PLACE AND PASSED AWAY ON (B)(6) 2025. THE PATIENT WAS UNDERGOING CCPD AT THE TIME OF HER PASSING (26 MINUTES INTO TREATMENT), HOWEVER THE IPM WAS CONFIDENT THE CARDIAC ARREST OCCURRED DUE TO THE PATIENT¿S CARDIAC COMORBID CONDITIONS WHILE THE BODY WAS FIGHTING THE HAND/FOOT/MOUTH INFECTION. PER THE IPM, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINIC REVIEW: BASED ON THE AVAILABLE INFORMATION, THE PATIENT¿S LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM HAVING A CAUSAL OR CONTRIBUTORY ROLE IN THE ADVERSE EVENTS EXPERIENCED BY THE PATIENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A SERIOUS INJURY, PATIENT DEATH, OR OTHER SERIOUS ADVERSE EVENT(S) RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) OCCURRED. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | Breathing-related sleep disorder | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK FILTER, BREATHING ISSUE, AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THERE ARE NO REPORTABLE MALFUNCTION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PLUS | K091319 | Cancer | Degraded | (B)(4) - ALLEGATION OF PATIENT HARM. | IN THE PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN GENERAL INFORMATION, BOX B, BOX G AND BOX H. THE FOLLOWING CORRECTION HAS BEEN CORRECTED IN THIS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. GENERAL INFORMATION: 510K HAS BEEN CORRECTED. IN BOX B: DESCRIBE EVENT OR PROBLEM HAS BEEN CORRECTED. IN BOX G: 510K HAS BEEN CORRECTED. IN BOX H: TYPE OF REPORTED COMPLAINT HAS BEEN CORRECTED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY ISSUES AND SARCOMA CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT REPORTED THE PATIENT PASSED AWAY. IT COULD NOT BE CONFIRMED IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. ADDITIONAL DETAILS WERE OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT¿S SON REPORTED THAT THE PATIENT WAS FOUND UNRESPONSIVE ON (B)(6) 2025. THE SON WAS NOT AT THE PATIENT¿S HOME AT THE TIME OF THE EVENT. 911 WAS CALLED BY SOMEONE IN THE HOME. EMERGENCY MEDICAL SERVICES (EMS) RESPONDED TO THE HOME AND RESUSCITATIVE EFFORTS WERE ADMINISTERED, BUT NOT SUCCESSFUL. THE PATIENT¿S SON REPORTED THAT HE DID NOT ENTER THE ROOM WHERE THE PATIENT WAS FOUND. HE COULD NOT CONFIRM IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. THE DEATH WAS REPORTEDLY UNEXPECTED (NO HEALTH ISSUES REPORTED). NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR PD THERAPY PRIOR TO THE DEATH. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A POSSIBLE TEMPORAL RELATIONSHIP BETWEEN THE PERITONEAL DIALYSIS (PD) AND USE OF THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH PRECEDED BY UNRESPONSIVENESS. HOWEVER, IT IS NOT CONFIRMED IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE ADVERSE EVENT. THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. ALTHOUGH NO CAUSE OF DEATH WAS CONFIRMED, PATIENTS ON DIALYSIS HAVE A SUBSTANTIALLY HIGHER MULTIVARIABLE-ADJUSTED MORTALITY THAN PATIENTS NOT ON DIALYSIS WHO HAVE CANCER, DIABETES, OR CARDIOVASCULAR DISEASE. A LARGE NUMBER OF RISK FACTORS THAT ARE UNRELATED TO THE DIALYSIS PROCEDURE HAVE BEEN ASSOCIATED WITH DECREASED SURVIVAL AMONG PATIENTS ON DIALYSIS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
REMSTAR PRO C-FLEX+ | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer; Headache; Dizziness; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THE PATIENT HAS PASSED AWAY DUE TO LUNG CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | Cancer; Lymphoma | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND NON-HODGKIN'S LYMPHOMA (FOLLICULAR). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, UNSPECIFIED KIDNEY PROBLEM AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K091319 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | Asthma; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Bacterial Infection; Insufficient Information; Cancer; Lymphoma | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER NON-HODGKINS LYMPHOMA, BACTERIAL LUNG INFECTION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | Cancer; Insufficient Information; Unspecified Hepatic or Biliary Problem; Cardiac Arrest | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH, OTHER, LIVER AND LUNG CANCER AND HEART ATTACK. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER, LIVER (FATAL). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SERVICE THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED WHILE CONNECTED TO THEIR LIBERTY SELECT CYCLER. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REVEALED THAT THE PATIENT EXPIRED IN THE EARLY MORNING ON (B)(6) 2025 DUE TO A CARDIAC ARREST ATTRIBUTED TO A SIGNIFICANT CARDIAC DISEASE HISTORY. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO THEIR CYCLER AT THE TIME OF THE ARREST AND SUBSEQUENT DEATH. IT WAS EXPLAINED THAT THE PATIENT RECENTLY UNDERWENT AN EMERGENCY AORTIC VALVE REPLACEMENT (EXACT DATE NOT REPORTED) AND WAS DISCHARGED HOME WITH A HOLTER HEART MONITOR. THE OVERSEEING FACILITY NOTIFIED THE PATIENT¿S SPOUSE THAT THE PATIENT WAS PULSELESS, AND EMERGENCY MEDICAL SERVICES WERE ACTIVATED. THE PATIENT WAS NOT TRANSPORTED AND THEY WERE PRONOUNCED DECEASED IN THEIR HOME. IT WAS CONFIRMED THAT THE PATIENT¿S CARDIAC ARREST LEADING TO THEIR DEATH WAS UNRELATED TO PD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). THE PDRN STATED THE PATIENT'S CYCLER WOULD BE RETURNED TO FRESENIUS. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO THEIR DEATH. THE CAUSE OF THIS PATIENT¿S CARDIAC ARREST AND DEATH CAN BE ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE END-STAGE RENAL DISEASE (ESRD) POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOVASCULAR DISEASE AND RECENT INVASIVE CARDIAC SURGERY. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
DREAMSTATION AUTO BIPAP | K131982 | Cancer; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING THROAT CANCER, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY AND PATIENT HAS PASSED AWAY (DEATH). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | Cancer; Unspecified Kidney or Urinary Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER, DEATH AND GLIOBLASTOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Unspecified Hepatic or Biliary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, LIVER CANCER (FATAL) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | Pulmonary Dysfunction; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | Cancer; Insufficient Information; Pulmonary Embolism | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, OTHER (UNSPECIFIED EVENT), PROSTATE CANCER AND PULMONARY EMBOLISM. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Pulmonary Dysfunction; Unspecified Kidney or Urinary Problem; Unspecified Respiratory Problem | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE / TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, CUSTOMER COMPLAINT WAS NOT CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING PANCREATIC CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI XI | K131861 | Pulmonary Embolism | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED MALIGNANT HYSTERECTOMY SURGICAL PROCEDURE FOR ENDOMETRIAL CANCER, THE PATIENT CODED AND EXPIRED. THE SURGEON REPORTED THAT THE HYSTERECTOMY PORTION OF THE CASE WAS COMPLETED. DURING THE COLPOTOMY VAGINAL CUFF CLOSURE THE PATIENT CODED; THE OR TEAM EMERGENTLY REMOVED THE INSTRUMENTS, UNDOCKED THE SYSTEM AND PERFORMED CARDIOPULMONARY RESUSCITATION WITHOUT SUCCESS, AND DURING WHICH THERE WAS SOME UNSPECIFIED CHEST BLEEDING. THE SURGEON STATED HE BELIEVES THAT THE DA VINCI SYSTEM AND INSTRUMENTS HAD NOTHING TO DO WITH THE EVENT AS THERE WERE NO ISSUES REPORTED DURING THE HYSTERECTOMY PORTION; HE BELIEVES THAT THE PATIENT HAD AN UNDIAGNOSED CLOT, SUFFERED A PULMONARY EMBOLISM AND EXPIRED. NO ADDITIONAL DETAILS WERE PROVIDED. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, FENESTRATED FORCE BIPOLAR. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT SUDDENLY CODED DURING A MALIGNANT HYSTERECTOMY PROCEDURE. NO ISSUES OR COMPLAINTS WERE REPORTED WITH ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS UP TO THIS POINT. THEY EMERGENTLY UNDOCKED AND ATTEMPTED CPR BUT THE PATIENT DIED DESPITE THESE EFFORTS. THE CAUSE OF DEATH WAS ATTRIBUTED TO A PULMONARY EMBOLISM. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
ION | K182188 | Insufficient Information; Altered State of Consciousness | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNEVENTFUL ION ENDOLUMINAL LUNG BIOPSY PROCEDURE. THE PHYSICIAN REPORTED THAT THE PATIENT WAS STABLE THE ENTIRE TIME AND WAS EXTUBATED. THE BIOPSIED LESION WAS 5 CM IN SIZE, LOCATED IN THE LEFT UPPER LOBE FAR FROM THE PLEURA, AND WAS DETERMINED TO BE NECROTIC TISSUE. APPROXIMATELY 2.5 HOURS POST-PROCEDURE THE PHYSICIAN AND ANESTHESIOLOGIST WERE NOTIFIED THAT THE PATIENT WAS STILL DROWSY AND WAS NOT FOLLOWING COMMANDS; THE PATIENT WAS OTHERWISE HEMODYNAMICALLY STABLE. THE PATIENT WAS TRANSFERRED TO THE EMERGENCY DEPARTMENT FOR A STROKE WORK UP. THE PATIENT WAS INTUBATED AT SOME POINT OVERNIGHT WHILE IN THE ICU FOR DECLINING MENTAL STATUS AND EXPIRED AT AN UNSPECIFIED TIME LATER. THE CAUSE OF DEATH WAS UNKNOWN. AN AUTOPSY IS PENDING. PER THE PHYSICIAN, THE ION SYSTEM DID NOT EXHIBIT ANY ISSUES OR UNEXPECTED BEHAVIORS THAT CONTRIBUTED TO THE DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH A HISTORY OF ENDOMETRIAL CANCER UNDERWENT AN ION LUNG BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. POST EXTUBATION IN RECOVERY THE PATIENT¿S SENSORIUM REMAINED ALTERED 2.5 HOURS POST PROCEDURE PROMPTING WORK UP FOR A STROKE AND ADMISSION TO THE ICU. THE PATIENT WAS INTUBATED OVERNIGHT FOR WORSENING MENTAL STATUS BUT REMAINED OTHERWISE HEMODYNAMICALLY STABLE. THE PATIENT EVENTUALLY DIED AND AN AUTOPSY IS PENDING AND EFFORTS AT OBTAINING THE RESULTS ARE ONGOING. AT THIS JUNCTURE THE CAUSE OF DEATH IN THE SETTING OF AN ALTERED SENSORIUM POST PROCEDURE IS UNCLEAR BUT THERE IS NO EVIDENCE OF MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS, OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENTS WERE LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. SECTION B2: PATIENT DATE OF DEATH: INTUBATION WAS THE DAY AFTER THE PROCEDURE (05/01/2025); HOWEVER, THE DATE OF DEATH WAS NOT PROVIDED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Dyspnea; Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT¿S DAUGHTER REPORTED THAT AS THEY WERE ATTEMPTED TO DISCONNECT THE PATIENT FROM THE CYCLER AFTER THE EVENT OCCURRED, FLUID DID COME OUT AS SHE WAS REMOVING THE CASSETTE FROM CASSETTE DOOR. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS IN TREATMENT ON (B)(6) 2025 CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT HAD TO GET UP AND USE THE RESTROOM. THE PATIENT WAS IN THE RESTROOM WHEN SHE EXPRESSED TO HER SPOUSE THAT SHE COULD NOT BREATHE. THE PATIENT¿S SPOUSE CALLED 911. EMERGENCY MEDICAL SERVICES (EMS) RESPONDED. THE PATIENT HAD BECOME UNRESPONSIVE PRIOR TO EMS ARRIVAL. EMS ATTEMPTED RESUSCITATION, BUT IT WAS UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME. THE CAUSE OF THE PATIENT¿S DEATH IS UNKNOWN. THERE WERE NO ISSUES OR REPORTED ALARMS DURING THE TREATMENT PRIOR TO THE PATIENT¿S DEATH. THE FLUID LEAK REPORTED DURING THE DISCONNECTION OF THE PATIENT AND THE SUPPLIES OCCURRED BECAUSE OF THE PATIENT¿S DAUGHTER NOT KNOWING WHAT SHE WAS DOING. THE PATIENT¿S DAUGHTER WAS ATTEMPTING TO DISCONNECT THE CASSETTE, AND SHE CAUSED THE FLUID TO LEAK. THERE WAS NO LEAK INDICATED DURING THE PATIENT¿S TREATMENT. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS AND UTILIZATION OF THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF NOT BEING ABLE TO BREATH, BECOMING UNRESPONSIVE, AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE REPORTED FLUID LEAK WAS A RESULT OF AN ERROR THAT THE PATIENT¿S DAUGHTER MADE WHEN ATTEMPTING TO DISCONNECT THE SUPPLIES FROM THE CYCLER. THERE WERE NO REPORTED ALARMS OR ISSUES WITH THE TREATMENT PRIOR TO THE PATIENT EVENT. ALTHOUGH THE CAUSE OF DEATH WAS NOT CONFIRMED, SUDDEN CARDIAC DEATH (SCD) IS RESPONSIBLE FOR THE MAJORITY OF CARDIOVASCULAR-RELATED DEATHS IN PATIENTS WITH ESRD WITH STUDIES REPORTING THAT UP TO 25 % OF ALL DEATHS IN THIS HIGH-RISK POPULATION IS ATTRIBUTABLE TO SCD. BASED ON THE AVAILABLE INFORMATION AND NO EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
ALARIS SYSTEM | K211218 | Low Blood Pressure/ Hypotension; Cardiac Arrest | Inaccurate Delivery | INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST. | THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K211218 | Low Blood Pressure/ Hypotension; Cardiac Arrest | Inaccurate Delivery | INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST. | THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
PATIENT INFORMATION CENTER IX | K211900 | Insufficient Information | No Audible Alarm | IT WAS REPORTED NO ALARMS WERE GENERATED ON THE NURSE PHONES AND THE PATIENT PASSED AWAY. IT WAS INDICATED RED ALARMS WOULD BE DISPLAYED BRIEFLY ON THE NURSE PHONE BUT IT APPEARS THERE WAS NO AUDIBLE ALARM. | E1 REPORTING INSTITUTION PHONE#: (B)(6). IT WAS ALSO INDICATED THE DEVICE (ASCOM PHONE) WAS NOT USABLE AND A REPLACEMENT WAS REQUESTED. THE CUSTOMER ALLEGED THAT AN ASYSTOLE ALARM WAS MISSED. THE PATIENT UNDERWENT RESUSCITATION MEASURES, HOWEVER THE PATIENT PASSED AWAY. A PHILIPS PROJECT MANAGER AND APPLICATION SPECIALIST WERE CONTACTED. IT WAS NOTED THAT THE CONFIGURATION IS SET FOR DELAYED PHONE NOTIFICATION, HOWEVER THE ALLEGATION INDICATES THAT THE MESSAGE WAS RECEIVED (VERY BRIEFLY) AND DISAPPEARED. A PHILIPS REMOTE SERVICE ENGINEER (RSE) SPOKE WITH THE CUSTOMER AND NOTED THAT THE EVENT TOOK PLACE ON (B)(6) 2025 AT APPROXIMATELY 19:50 HOURS. THE ONSITE TECHNICIAN TRIGGERED A RED ALARM AND THIS WAS RECEIVED AT THE END DEVICE (MYCO 3 SMARTPHONE) WITH CAREASSIST 4.4.0. THE TEST WAS SUCCESSFUL. THE SOFTWARE WAS PROACTIVELY UPDATED TO VERSION 4.3.7. A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND THE PRELIMINARY INVESTIGATION AND FOUND THAT STARTING AT 19:55:07, AN ALARM WAS GENERATED AT THE PATIENT INFORMATION CENTER IX (PIC IX) INDICATING THAT THERE WAS NO MOBILE DEVICE AVAILABLE FOR ALERTING FOR THE AFFECTED BED LOCATION. FROM 19:55:15 AND 19:55:18, RED ALARMS WERE GENERATED ON THE MX40 MONITOR AND THE PICIX. THESE WERE ACKNOWLEDGED AT THE PICIX AT 19:55:28 AND 19:55:29, WHEREUPON THE ALARM WAS TERMINATED AT THE PICIX. ACCORDING TO THE CAREGIVER, THE NECESSARY ACTIONS WERE THEN IMMEDIATELY INITIATED FOR THE PATIENT. THE EARLY INVESTIGATION INDICATES THAT THE MOBILE PHONE WAS NOT IN USE. AS THIS IS CONTRADICTORY TO THE ORIGINAL ALLEGATION, QUALITY PERSONNEL ARE WORKING WITH THE FIELD ENGINEER TO CLARIFY THE EVENT. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A USER OF A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE REPORTS THE PATIENT WAS NOT USING THE DEVICE AT THE TIME OF PASSING. THE PATIENT HAS NOT USED THE DEVICE FOR A VERY LONG TIME. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE IS NO ALLEGATION THAT THE DEVICE CAUSED THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | Cancer; Appropriate Term / Code Not Available | Degraded | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED/RECERTIFIED DREAM STATION CPAP WITH AN ALLEGATION METASTATIC ESOPHAGEAL CARCINOMA. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER MAILING SENT OUT DUE TO A DEVICE RECALL. THE MANUFACTURER WAS INFORMED BY A FAMILY MEMBER THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Asystole; Hemorrhage/Blood Loss/Bleeding | Insufficient Information | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION B5 ADDITIONAL INFORMATION: THE MOBILIZATION OF THE TUMOR WAS ALMOST COMPLETED WHEN THE BLEEDING BEGAN. A SYNCHROSEAL INSTRUMENT AND UNSPECIFIED BIPOLAR INSTRUMENT WERE BEING USED DURING THIS TASK. THE CUSTOMER SAID THERE WAS NO MALFUNCTION OF A DA VINCI PRODUCT DURING THIS EVENT. THE BLEEDING WAS COMING FROM THE LIVER VEIN, DISTAL TO THE TUMOR. CENTRALLY, THE VEIN WAS ADHERED TO THE TUMOR. THE DECISION TO CONVERT TO OPEN WAS MORE DUE TO THE VEIN BEING ADHERED TO THE TUMOR IN THE CENTRAL PART THAN THE BLEEDING. THE RESECTION OF THE TUMOR AND HEMOSTASIS WERE BOTH COMPLETED WHILE OPEN. THE BLOOD SUPPLY TO THE LIVER AND INTESTINES WERE WELL MAINTAINED AT ALL TIMES. THE VENA CAVA WAS WELL FILLED AT ALL TIMES, SO THE RETURN FLOW TO THE HEART WAS MORE VIGOROUS THAN INSUFFICIENT. THE CUSTOMER STATED THAT THE ORIGIN OF THE ASYSTOLE REMAINS UNEXPLAINED. A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
REMSTAR PRO C-FLEX+ | K091319 | Insufficient Information | No Apparent Adverse Event | THE REPORTER, WHO IS THE RELATIVE OF THE PATIENT, REPORTED THE PATIENT EXPIRED. THE REPORTER STATED THE PATIENT PREVIOUSLY USED; HOWEVER, THE PATIENT WAS NOT USING THE DEVICE AT THE TIME OF DEATH. THE PATIENT WAS SWITCHED TO A DIFFERENT (UNSPECIFIED) DEVICE. THE REPORTER HAD TO MAKE ANOTHER CALL AND WAS UNABLE TO ANSWER ANY FURTHER QUESTIONS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ALARIS SYSTEM | K133532 | Low Blood Pressure/ Hypotension | Inaccurate Delivery | IT WAS REPORTED THAT THE CUSTOMER REQUESTED ALOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT BUT HARM WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE, THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR) NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | Low Blood Pressure/ Hypotension | Inaccurate Delivery | IT WAS REPORTED THAT THE CUSTOMER REQUESTED ALOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT BUT HARM WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE, THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR) NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH DUE TO ALZHEIMER'S. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded; Contamination | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTER HAS ALLEGED HER HUSBAND PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING INVESTIGATION, THE MANUFACTURER OBSERVED AN UNKNOWN LIGHT CONTAMINATION AT THE AIR INLET AND THE BOTTOM OF THE BLOWER BOX. LIGHT BLACK SPECKS IN THE BOTTOM ENCLOSURE. UNKNOWN LIGHT WHITE DUST-LIKE CONTAMINATION WAS ON TOP AND BOTTOM OF THE BLOWER MOTOR AND THE BLOWER MOTOR SEAL. MISSING AIR INLET FILTER. DURING INVESTIGATION THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER SUPPLY AND POWER CORD TO THE DEVICE AND VERIFIED AIRFLOW AND CONFIRMED THAT THE DEVICE POWERS ON PROVIDES THERAPY. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE COMPLAINT. NO SPECIFIC COMPLAINT GIVEN. THE MANUFACTURER OBSERVED DUST CONTAMINATION IN THE DEVICE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND IS ALSO REQUESTING FOR A RETURN LABEL SENT TO THE MAILING ADDRESS DOCUMENTED ABOVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION. THERE WAS 0 ERROR CODE FOUND. THE MANUFACTURER SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT HER HUSBAND PASSED AWAY AND THAT SHE NEEDED TO KNOW WHEN HE RECEIVED HIS REPLACEMENT DEVICE FOR BILLING PURPOSES FOR A LOANER DEVICE HE HAD USED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
ION | K182188 | Bradycardia; Pulmonary Embolism; Hemorrhage/Blood Loss/Bleeding | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE A PATIENT EXPIRED. THE BIOPSY WAS PERFORMED ON A METASTATIC MASS IN THE LEFT LUNG (PROBABLY LEFT UPPER LOBE). THE PROCEDURE WAS COMPLETED WITH NO MALFUNCTIONS OR ISSUES; HOWEVER, AT THE END OF THE PROCEDURE, WHEN THE PHYSICIAN WAS UNDOCKING THE ROBOT, BLEEDING WAS NOTICED FROM THE LUNG. THE PATIENT¿S HEART RATE THEN DROPPED TO 38BPM, ¿AND ESSENTIALLY THE PATIENT CRASHED AND SUBSEQUENTLY PASSED AWAY.¿ NO FURTHER DETAILS WERE PROVIDED. THE PHYSICIAN SPECULATED THAT THE PATIENT MAY HAVE EXPERIENCED A PULMONARY EMBOLUS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY FOR METASTATIC DISEASE TO THE LUNG. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. AS THE ION SYSTEM WAS UNDOCKED BLEEDING WAS NOTED ASSOCIATED WITH BRADYCARDIA INTO THE 30S. THE PATIENT DETERIORATED AND SUBSEQUENTLY DIED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENTS WERE LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DREAMSTATION AUTO BIPAP | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K153702 | Insufficient Information | Defective Alarm | IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN OBTAINING CLINICAL AUDIT LOGS FROM THE CENTRAL STATION FOLLOWING A PATIENT DEATH. THE CUSTOMER WANTED TO DETERMINE WHETHER THE DEVICE ALARMED APPROPRIATELY. THE CUSTOMER INDICATED THE PATIENT HAD ASPIRATED AT SOME POINT, WHICH LED TO DEATH. A LOG REVIEW WAS PERFORMED BY THE CUSTOMER, REVEALING THE ALARMS WERE GENERATED AS EXPECTED. THERE WERE MULTIPLE ALARMS WHICH WERE LATCHED BUT NOT ACKNOWLEDGED, WHICH MASKED LOWER-LEVEL ALARMS. THE LOG REVIEW PROVIDED ANSWERS TO THE CUSTOMER REQUEST. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | A PHILIPS PRODUCT SUPPORT ENGINEERS (PSE) REVIEWED THE LOGS AND DETERMINED THAT THEY REFLECT THAT THE ALARMS FOR THE BED IN QUESTION DURING THE TIME PERIOD OF INTEREST WERE NOTED. A PHILIPS CLINICAL APPLICATIONS SPECIALIST (CAS) WENT ON SITE CONFIRMED THAT THE UNIT IS WORKING AS EXPECTED, THE CLINICAL AUDIT LOGS WERE ATTACHED, AND ALARMS WERE PRESENT AS EXPECTED WITH LATCHED AND NON-LATCHED ALARMS. BASED ON THIS INFORMATION, THERE WAS NO MALFUNCTION OF THE DEVICE. IN ADDITION, THERE WAS NO ALLEGATION THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
ALARIS SYSTEM | K211218 | No Clinical Signs, Symptoms or Conditions | Excess Flow or Over-Infusion | IT WAS REPORTED THAT THE PUMP WAS INTENDED TO INFUSE AT A RATE OF 4 ML/HOUR. HOWEVER, MEDICATION BAG WAS FOUND TO BE EMPTY. FURTHER INFORMATION WAS RECEIVED THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT MORPHINE 100 ML WAS PROGRAMMED ON (B)(6) 2025 TO INFUSE AT 4 ML/HOUR AND DOUBLE VERIFIED. THE CLINICIAN NOTED AN AIR-IN-LINE ALARM AND EMPTY MEDICATION BAG APPROXIMATELY ONE (1) HOUR AND TWENTY (20) MINUTES LATER. THERE WAS NO REPORTED CHANGE IN THE PATIENT'S CONDITION, AND NO INTERVENTIONS WERE PROVIDED. HOWEVER, IT WAS REPORTED THE PATIENT PASSED AWAY ON (B)(6) 2025. CUSTOMER RISK MANAGEMENT STATED, "THE EVENT REPORTED WAS EQUIPMENT MALFUNCTION; HOWEVER, AT THIS POINT NOT FELT TO HAVE CONTRIBUTED TO THE DEATH." | BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE FAILURE INVESTIGATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | A RELATIVE OF THE PATIENT REPORTED THE PATIENT EXPIRED. THE REPORTER STATED THE PATIENT WAS NOT USING THE DEVICE AT THE TIME THE PATIENT PASSED AWAY. THE REPORTER STATED THE PATIENT WAS USING ANOTHER UNSPECIFIED DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
ALARIS SYSTEM | K133532 | Insufficient Information | Excess Flow or Over-Infusion | IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON (B)(6) 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | Insufficient Information | Excess Flow or Over-Infusion | IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON 27 MAY 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT PASSED AWAY ON THE CYCLER. THE CAUSE OF DEATH IS UNKNOWN. A RED BAG PICKUP FROM THE CLINIC PROCESS WAS REQUESTED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025. THE PATIENT REPORTEDLY AWOKE GASPING FOR AIR AND PASSED SHORTLY THEREAFTER. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
DAVINCI XI | K131861 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE PATIENT UNDERWENT A DA VINCI-ASSISTED ROUX-EN-Y GASTRIC BYPASS PROCEDURE. THE PATIENT SUBSEQUENTLY EXPIRED ON POST-OPERATIVE DAY 6. NO INFORMATION REGARDING THE EVENTS LEADING UP TO THE DEATH OR THE CAUSE OF DEATH WERE PROVIDED. THERE WAS NO REPORT OF ANY INTRAOPERATIVE COMPLICATIONS. THE SITE VICE PRESIDENT OF BARIATRICS REPORTED THAT NO SPECIFIC INFORMATION REGARDING THE POST-OPERATIVE COURSE WOULD BE AVAILABLE; HOWEVER, THERE WAS A POTENTIAL THAT THE PATIENT'S CAUSE OF DEATH COULD BE ATTRIBUTED TO A POST-OPERATIVE OVERDOSE. NO ADDITIONAL INFORMATION WAS PROVIDED. THERE WAS NO ALLEGATION THAT A DA VINCI SYSTEM, INSTRUMENT, OR ACCESSORY MALFUNCTION OCCURRED DURING THE PROCEDURE. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, PERMANENT CAUTERY HOOK, TIP-UP FENESTRATED GRASPER, CADIERE FORCEPS, AND MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. DHR REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. REVIEW OF THE SUREFORM 60 STAPLER LOGS FOUND USED IN THIS EVENT FOUND THE INSTRUMENT WAS INSTALLED ON THE SYSTEM 8 TIMES AND FIRED 7 WHITE RELOADS. ON INSTALL 1, A WHITE RELOAD WAS LOADED, HOWEVER NO CLAMPING OR FIRING WAS ATTEMPTED. ON INSTALL 2, THE FIRST CLAMP ATTEMPT STOPPED AT APPROXIMATELY 94% COMPLETION. THE NEXT TWO CLAMPS WERE SUCCESSFUL, AND THE FIRING WAS COMPLETED WITH 2-3 PAUSES FOR COMPRESSION. ON INSTALLS 3-5, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE COMPLETED WITH 1 PAUSE FOR COMPRESSION EACH. ON INSTALLS 6-8, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE COMPLETED WITH NO PAUSES FOR COMPRESSION EACH. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS, AND THERE WAS NO REPORT OF ANY ISSUES FROM THE OR TEAM REGARDING THE STAPLER INSTRUMENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT PATIENT IN THIS REPORT DIED 6 DAYS FOLLOWING A REPORTEDLY UNEVENTFUL GASTRIC BYPASS PROCEDURE. HOW THEY DIED AND WHAT EVENTS LED TO THEIR DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE PREVIOUS REPORT HAD THE INCORRECT HEALTH IMPACT CODE. THIS HAS BEEN UPDATED TO DEATH UNDER THE HEALTH IMPACT GRID.
RENTAL, DREAMSTATION ST30, GB | K102465 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | MANUFACTURER RECEIVED A CALL FROM THE CLINIC THAT THE PATIENT PASSED AWAY. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DAVINCI XI | K131861 | Hemorrhage/Blood Loss/Bleeding; Adhesion(s) | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RADICAL CYSTECTOMY AND URINARY DIVERSION PROCEDURE, THE PATIENT EXPIRED POST-OPERATIVELY DUE TO BLEEDING. THE PATIENT REPORTEDLY HAD SEVERE ADHESIONS, WAS UNDERGOING CHEMOTHERAPY, AND HAD FRAGILE BLOOD VESSELS. THE SURGEONS CONVERTED TO LAPAROTOMY, BUT COULD NOT IDENTIFY THE SOURCE OF THE BLEEDING. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION WERE MADE; NO RESPONSE WAS RECEIVED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 8MM 30 DEGREE ENDOSCOPE PLUS, 0 DEGREE ENDOSCOPE, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, LARGE SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND X2. A SITE HISTORY REVIEW FOUND THAT ALL MULTIPLE USE INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES WITH NO REPORTED ISSUES THAT WOULD BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A RADICAL CYSTECTOMY. BLEEDING IS SUGGESTED AS THE CAUSE OF DEATH AND POTENTIAL CONTRIBUTING FACTORS ARE SUGGESTED BUT HOW AND WHEN THE BLEEDING OCCURRED IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | ON (B)(6) 2025 A HOSPITAL NURSE REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT HAD PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT WAS AT THE END OF TREATMENT (STEP 8). THE NURSE DISCONNECTED THE PATIENT AND REALIZED THE PATIENT WAS NOT WAKING UP. THE NURSE TOOK THE PATIENT¿S VITAL SIGNS, AND HE SUBSEQUENTLY WENT INTO CARDIAC ARREST AND DIED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH THE DIALYSIS UNIT NURSE. THE PATIENT WAS AN IN-PATIENT REHABILITATION (REHAB) DIALYSIS PATIENT. THE PATIENT WAS IN REHAB DUE TO A LEFT BELOW KNEE AMPUTATION (BKA) WHICH WAS UNRELATED TO DIALYSIS. THE DATE OF THE ADMISSION TO THE REHAB WAS UNKNOWN. THE NURSE STATED THAT ON (B)(6) 2025, THE HOSPITAL NURSE WAS DOING ROUNDS WHEN THEY WENT TO DISCONNECT THIS PATIENT FROM THE LIBERTY SELECT CYCLER AT THE END OF TREATMENT. THE NURSE WENT TO TAKE THE PATIENT¿S VITALS AND FOUND THE PATIENT¿S PULSE TO BE THREADY. THE PATIENT SUBSEQUENTLY WENT INTO CARDIAC ARREST. THE NURSE COULD NOT CONFIRM WHAT RESUSCITATIVE EFFORTS WERE ADMINISTERED. THE EFFORTS WERE NOT SUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST. THE NURSE STATED THE DEATH WAS NOT RELATED TO PD THERAPY OR USE OF THE LIBERTY SELECT CYCLER OR OTHER FRESENIUS PRODUCT(S). NO FURTHER INFORMATION WAS PROVIDED. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NOT VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT. THERE WERE VISUAL INDICATION OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NOT BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. TEACH PUMPS TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
PATIENT INFORMATION CENTER IX | K183387 | Insufficient Information | Defective Alarm | IT WAS REPORTED THE ALARMS FOR APNEA, BRADYCARDIA, AND ASYSTOLE WERE NOT GENERATED. THE ICU PATIENT WAS FOUND PASSED AWAY IN THEIR BED AND RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL. TESTING OF THE DEVICE ONSITE REVEALED NO TECHNICAL ANOMALIES. IT WAS CLARIFIED NONE OF THE ALARMS, AUDIBLE OR VISUAL, WERE GENERATED AT THE CENTRAL STATION. THE CONDITION OF THE PATIENT PRIOR AND THE CAUSE OF DEATH WERE UNKNOWN, BUT IT WAS INDICATED THE PATIENT WAS ELDERLY. | A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE AND COLLECTED THE LOGS OF THE PIC IX AND LOGS AND CONFIGURATION OF THE MP5. THE FSE DID NOT FIND ANY TECHNICAL DEFECTS WHILE ONSITE. INITIAL REVIEW OF THE LOGS INDICATES MANUAL DEACTIVATION OF THE PULSE ALARMS. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. SECTION E REPORTING INSTITUTION PHONE #: (B)(6).
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | A NURSE REPORTED A PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT). DURING THE INTAKE PROCESS, IT WAS REPORTED THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN PATIENT EXPIRED. FOLLOW-UP WITH THE PATIENTS¿ HOME PROGRAM MANAGER (HPM) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. THE PATIENT REPORTEDLY EXPIRED WHILE SLEEPING AND WAS DISCOVERED BY THE PATIENT¿S SPOUSE. THE HPM REPORTED THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. | ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL ANALYSIS. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGN OF PHYSICAL DAMAGE.THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT.THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA.HERE WERE NOT BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE.SYSTEM AIR LEAK TEST PASSED.VALVE ACTUATION TEST PASSED.A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY ON THE BOTTOM COVER. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED.THERE WERE NOT DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S). UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DAVINCI XI | K131861 | Insufficient Information | Insufficient Information; Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED PROCEDURE, THE PATIENT LATER PRESENTED FOR UNSPECIFIED REASONS, UNDERWENT AN EXPLORATORY LAPAROTOMY, AND SUBSEQUENTLY EXPIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT NO RESPONSE HAS BEEN RECEIVED. | INSUFFICIENT INFORMATION IS AVAILABLE FOR PROCEDURE VERIFICATION; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN UNKNOWN ROBOTIC PROCEDURE. NO ADDITIONAL INFORMATION IS PROVIDED ABOUT THE PROCEDURE OR ANY INTRAOPERATIVE ISSUES. THE EVENTS THAT LED TO THEIR REOPERATION, THE FINDINGS OF THE REOPERATION, AND THE SUBSEQUENT ISSUES THAT LED TO THEIR DEATH AND CAUSE OF DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | Hemorrhage/Blood Loss/Bleeding | Insufficient Information | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LIVER RESECTION SURGICAL PROCEDURE, THE PATIENT EXPIRED DUE TO AN INJURY TO THE VENA CAVA, WHICH RESULTED IN A SIGNIFICANT AMOUNT OF BLOOD LOSS. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION HAVE BEEN MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MEGA SUTURECUT NEEDLE DRIVER, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. STAPLER LOG REVIEW SHOWS THE SUREFORM 60 STAPLER INSTRUMENT WAS INSTALLED ON THE SYSTEM TWICE AND SUCCESSFULLY FIRED 2 BLUE RELOADS. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE VESSEL SEALER EXTEND INSTRUMENT AND THE E-100 GENERATOR LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THE REPORT DIED AS THE RESULT OF BLEEDING FROM AN INJURY TO THE VENA CAVA DURING A LIVER RESECTION. HOW THIS EVENT OCCURRED OR DETAILS OF WHAT MAY HAVE LED TO THE INJURY ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED, BUT UNABLE TO GATHER ADDITIONAL INFORMATION RELATED DEVICE RETURN. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S DAUGHTER IN LAW, REQUESTING REPLACEMENT DEVICE SENT TO HER AS HER MOTHER-IN-LAW HAS PASSED AWAY. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS RECEIVING DIALYSIS AND EXPIRED WHILE IN TREATMENT. THE NURSE WAS ADVISED TO DISCONTINUE THE USE OF THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DAVINCI XI | K131861 | Appropriate Term / Code Not Available; Asystole | Insufficient Information | IT WAS REPORTED THAT A PATIENT EXPIRED DURING A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA SURGICAL PROCEDURE. THE SURGEON REPORTED THAT THE PATIENT HAD A HISTORY OF LUNG CANCER WITH RADIATION AND CHEMOTHERAPY TREATMENTS RESULTING IN DIFFICULT, SCARRED, AND WEAK TISSUE ANATOMY. WHILE DISSECTING TISSUE NEAR THE VAGUS NERVE, CONTROLLED BLEEDING OCCURRED AND PRESSURE WAS APPLIED. THE PATIENT THEN EXPERIENCED BRADYCARDIA FROM A VASOVAGAL RESPONSE THAT DETERIORATED TO ASYSTOLE. THE BLEEDING SITE WAS PACKED, THE ROBOT WAS EMERGENTLY UNDOCKED, AND THE PROCEDURE WAS CONVERTED TO OPEN. CHEST COMPRESSIONS WERE INITIATED. DESPITE RESUSCITATION EFFORTS THE PATIENT REMAINED IN ASYSTOLE; NO CARDIAC RHYTHM WAS RESTORED AND THE PATIENT EXPIRED. THE SURGEON STATED THAT NO MALFUNCTIONS OF ANY INTUITIVE PRODUCTS OCCURRED DURING THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM SHOW THE SYSTEM FUNCTIONED NORMALLY AND THERE WERE NO INDICATIONS RELATING TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, TWO CADIERE FORCEPS, PERMANENT CAUTERY HOOK, VESSEL SEALER EXTEND, SUCTION IRRIGATOR, 8MM 0 DEGREE ENDOSCOPE. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BRADYCARDIC EVENT LIKELY FROM A VAGAL RESPONSE DURING A HIATAL HERNIA REPAIR. THIS LED TO ASYSTOLE AND THE PATIENT EXPIRED DESPITE RESUSCITATIVE EFFORTS. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY SUDDENLY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOT REPORTED THAT THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF DEATH. ADDITIONAL INFORMATION REQUESTED BUT HAS NOT BEEN OBTAINED. | ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. CORRECTION PROVIDED IN G4. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR NO SHOWED SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K211900 | Insufficient Information | Use of Device Problem | IT WAS REPORTED THE USER DID NOT HEAR OR RESPOND TO THE ALARMS AND THE PATIENT PASSED WAY; HOWEVER, THERE WAS NO ALLEGATION OF ALARM FAILURE. IT WAS INDICATED THE USER DID NOT HEAR THE ECG LEADS OFF ALARMS AND DID NOT RESPOND TO THE ECG LEADS OFF YELLOW TECHNICAL INOP ALARMS. THE CUSTOMER REQUESTED ASSISTANCE OBTAINING THE UNFILTERED AUDIT LOGS, WHICH REVEALED THERE WERE TWO HOURS OF ECG LEADS OFF ALARMS WITH NO ALARM ACKNOWLEDGEMENT AND THE ALARMS WERE PAGED APPROPRIATELY. RESUSCITATION WAS ATTEMPTED BUT UNSUCCESSFUL AND THE PATIENT PASSED AWAY. THE NURSE MANAGER INDICATED THE ALARM VOLUME FOR YELLOW ALARMS IS SET LOWER THAN RED ALARMS AND THE STAFF IS NOT RESPONDING TO THE YELLOW ALARMS. THE CUSTOMER WILL BE REQUESTING THE ECG LEADS OFF TECHNICAL INOP BE CHANGED TO A RED ALARM VIA THE INTERNAL ALARM COUNCIL. | THE CLINICAL AUDIT LOGS WERE PULLED FOR INVESTIGATION. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | Insufficient Information | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Insufficient Information | Insufficient Information | IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED INGUINAL HERNIA REPAIR, THE PATIENT WAS DISCHARGED HOME. AN INTUITIVE CLINICAL TERRITORY ASSOCIATE WAS LATER INFORMED BY A HOSPITAL NURSE THAT THE PATIENT HAD PRESENTED TO THE EMERGENCY DEPARTMENT EIGHT HOURS LATER WITH UNSPECIFIED SYMPTOMS AND EXPIRED ON POST-OPERATIVE DAY 1. NO INFORMATION REGARDING THE CAUSE OF DEATH OR TREATMENT RENDERED WAS PROVIDED. THE ROBOTIC PROCEDURE WAS REPORTEDLY COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS OR DA VINCI SYSTEM ERRORS. ADDITIONAL INFORMATION WAS REQUESTED; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | REVIEW OF THE SYSTEM LOGS FOUND NO RELATED ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO RELATED ERRORS OCCURRED. THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MONOPOLAR CURVED SCISSORS, FORCE BIPOLAR, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN INGUINAL HERNIA REPAIR. ALTHOUGH NO INTRAOPERATIVE ISSUES OR COMPLICATIONS WERE REPORTED, THE REASON FOR THE RETURN TO THE EMERGENCY ROOM JUST 8 HOURS AFTER SURGERY DISCHARGE IS NOT PROVIDED. THE CAUSE OF DEATH IS ALSO NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION A2: PATIENT AGE IS ESTIMATED; REPORTED TO BE "EARLY TO MID 70'S."
RENTAL, DREAMSTATION ST30, GB | K102465 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE RENTAL DREAMSTATION ST30 DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED THAT THE DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Cardiac Arrest | Adverse Event Without Identified Device or Use Problem | IT WAS REPORTED BY THE PATIENT'S SPOUSE THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. THE PATIENT WAS DISCOVERED BY HER SPOUSE WHEN HE AWOKE IN THE MORNING. THE PDRN REPORTED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT TO THE MORGUE. THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST SECONDARY TO HER EXTENSIVE CARDIAC COMORBID CONDITIONS (E.G., ATRIAL FIBRILLATION, CORONARY ARTERY DISEASE, UNDERWENT MULTIPLE CARDIAC CATHETERIZATIONS WITH STENTING. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE PDRN ATTRIBUTED CAUSALITY TO THE PATIENTS¿ SIGNIFICANT CARDIAC COMORBID CONDITIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DS2ADV AUTO CPAP | K200480 | No Clinical Signs, Symptoms or Conditions | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER BECAME AWARE OF A DS2ADV AUTO CPAP STATING THAT THE PATIENT'S SIGNIFICANT OTHER ADVISED THAT THE PATIENT SADLY PASSED AWAY AND THAT THERE IS NO ALLEGATION THAT THE DEVICE HAS CONTRIBUTED TO THE DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | Cancer | Degraded | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | Insufficient Information; Vomiting | Adverse Event Without Identified Device or Use Problem | A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PD CLINIC. THE PATIENT¿S SPOUSE CONTACTED THE PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) ON (B)(6) 2025 AROUND 10:30AM TO REPORT THAT THE PATIENT HAD PASSED AWAY. THE PATIENT HAD BEEN IN TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER WHEN SHE BEGAN VOMITING. THE VOMIT WAS BLACK IN COLOR. THE SPOUSE DID AN EMERGENCY DISCONNECT FROM THE TREATMENT AND CALLED 911. WHEN EMERGENCY MEDICAL SERVICES (EMS) ARRIVED, THE PATIENT WAS NOT RESPONSIVE. EMS ATTEMPTED TO RESUSCITATE THE PATIENT BUT WERE UNSUCCESSFUL. THE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO REPORTED ISSUES WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT PRIOR TO THE EVENT. NO ADDITIONAL INFORMATION WAS AVAILABLE PERTAINING TO THE DEATH. | CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE PATIENT REPORTEDLY BEGAN VOMITING BLACK MATTER PRIOR TO BECOMING UNRESPONSIVE. ALTHOUGH THE CAUSE OF DEATH WAS NOT CONFIRMED, THIS IS AN INDICATION OF A GASTROINTESTINAL (GI) BLEED. IF THE PATIENT DID HAVE A SIGNIFICANT BLEED, SHE COULD HAVE GONE INTO HYPOVOLEMIC SHOCK CAUSING CARDIAC ARREST AND DEATH WHICH WOULD HAVE BEEN WHY THE PATIENT WAS UNRESPONSIVE WHEN PARAMEDICS ARRIVED. HOWEVER, WITHOUT ADDITIONAL INFORMATION, THAT CANNOT BE CONFIRMED. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO BIPAP | K131982 | Insufficient Information | Adverse Event Without Identified Device or Use Problem | THE MANUFACTURER RECEIVED INFORMATION A USER OF UNKNOWN DEVICE HAD PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | Sepsis; Bowel Perforation; Failure to Anastomose | Insufficient Information | IT WAS REPORTED THAT A DA VINCI-ASSISTED RADICAL CYSTECTOMY WITH INTRACORPOREAL ILEAL CONDUIT CREATION, BILATERAL URETERAL STENT PLACEMENT, AND BILATERAL EXTENDED PELVIC LYMPH NODE DISSECTION WAS PERFORMED. THERE WERE NO INTRAOPERATIVE COMPLICATIONS. THE INITIAL POST-OPERATIVE COURSE WAS UNCOMPLICATED AND THE PATIENT WAS DISCHARGED ON POST-OPERATIVE DAY (POD) 5. THE PATIENT ULTIMATELY EXPIRED ON POD 50 DUE TO SUBSEQUENT COMPLICATIONS. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT ON POD 8 WITH INTRA-ABDOMINAL FLUID COLLECTION AND SUBSEQUENTLY DEVELOPED SIGNS OF SEPTIC SHOCK AND ACIDOSIS AND WAS STABILIZED WITH VASOPRESSORS. DURING EMERGENT EXPLORATORY SURGERY 2 LITERS OF FRANK STOOL WAS FOUND IN THE ABDOMEN, AND A LARGE ABSCESS IN THE PELVIS WAS IRRIGATED AND SUCTIONED OUT. A 1 CM MID-ILEAL ANTI-MESENTERIC SMALL BOWEL PERFORATION WAS FOUND AND RESECTED. ADDITIONALLY, THE COMMON ENTEROTOMY AND DISTAL ILEAL-ILEAL ANASTOMOSIS STAPLE LINE WAS COMPLETELY DISRUPTED WITH FRANK LEAKAGE OF STOOL; THE AREA WAS RESECTED (20CM TOTAL) AND REPAIRED. THE PATIENT REMAINED CRITICALLY ILL AFTER THIS PROCEDURE. A SUBSEQUENT EXPLORATORY LAPAROTOMY ON POD 10 OCCURRED AND PER THE OPERATIVE NOTES, A RIGHT COLECTOMY, ABDOMINAL WASHOUT, ILEOCOLIC ANASTOMOSIS, SMALL-BOWEL TO SMALL-BOWEL ANASTOMOSIS, PLACEMENT OF A WOUND VAC AND FASCIAL CLOSURE WERE COMPLETED. EXAMINATION OF THE ILEAL CONDUIT FROM THE ORIGINAL ROBOTIC PROCEDURE BY UROLOGY FOUND NO SIGNS OF LEAKAGE; THE ANASTOMOSES AT THE CONDUIT ITSELF AS WELL AS THE URETERS APPEARED VIABLE WITHOUT SIGNS OF NECROSIS OR ISCHEMIA. NOTES AT POD 33 DOCUMENT A POST-PROCEDURE COURSE THAT INCLUDED ACUTE RENAL FAILURE REQUIRING DIALYSIS, ACUTE HYPOXIC RESPIRATORY FAILURE REQUIRING VENTILATION, URINE LEAK REQUIRING BILATERAL NEPHROSTOMY TUBE/DRAIN PLACEMENT, SEPTIC SHOCK REQUIRING ONGOING VASOPRESSORS, MULTIPLE ANTIBIOTICS RELATED TO ANASTOMOTIC LEAK, FUNGEMIA, AND LOWER ABDOMINAL MIDLINE WOUND DEHISCENCE WITH PURULENT DRAINAGE. NO ADDITIONAL INFORMATION WAS PROVIDED UP TO THE PROVIDED DEATH CERTIFICATE AT POD 50 WHICH LISTED CARDIAC ARREST, SEPSIS AND ABDOMINAL INFECTION AS CAUSE OF DEATH. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE EVENT. REVIEW FOUND THAT THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE OBSERVED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, VESSEL SEALER EXTEND, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, SUREFORM 60 STAPLER, LARGE NEEDLE DRIVER, LARGE SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. THE STAPLER INSTRUMENT LOGS FOR THIS PROCEDURE FOUND A SUREFORM 60 STAPLER INSTRUMENT WAS USED TO FIRE 12 RELOADS WITHOUT ANY ISSUES OBSERVED IN THE LOGS. THE VESSEL SEALER INSTRUMENT LOGS FOUND ONE CUT FAILURE ON THE INSTRUMENT'S 15TH CUT ATTEMPT. THE LOGS INDICATE THERE WERE TWO TIMES WHERE THE VESSEL SEALER EXTEND JAWS WERE OPEN TOO MUCH TO INITIATE A CUT OF THE TISSUE WITHIN THE JAWS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A PRIOR HISTORY OF PROSTATE CANCER AND PRIOR PROSTATECTOMY WHO WAS UNDERGOING A CYSTECTOMY FOR BLADDER CANCER WITH CREATION OF AN ILEAL CONDUIT. THE PROCEDURE REPORTEDLY WENT WELL AND NO INTRAOPERATIVE ISSUES WERE NOTED. POST-OPERATIVELY, THE PATIENT DEVELOPED A SMALL BOWEL LEAK SEVERAL DAYS AFTER THE ROBOTIC PROCEDURE FROM THE ANASTOMOSIS WHERE THE ILEAL CONDUIT WAS HARVESTED. THE REASON FOR THE LEAK AND POTENTIAL FACTORS WHICH MAY HAVE LED TO THE LEAK IS NOT PROVIDED. THE PATIENT UNDERWENT MULTIPLE ADDITIONAL PROCEDURES AS NOTED ABOVE FOLLOWING THE INITIAL COMPLICATION AND EVENTUALLY DIED SEVERAL WEEKS LATER. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT FROM THE INITIAL PROCEDURE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.