## Sheet1
Brand Name | K Number | Event Narrative | Manufacturer Narrative
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A REMSTAR PRO C-FLEX+ DEVICE. THERE WAS AN ALLEGATION OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR PRO C-FLEX+ DEVICE'S SOUND ABATEMENT FOAM. THERE WAS AN ALLEGATION OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. ADVERSE EVENT/PRODUCT PROBLEM (UPDATED FROM ADVERSE EVENT TO BOTH), OUTCOMES ATTRIBUTED TO AE, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID WERE UPDATED.
INJECTOMAT AGILIA BR | K121613 | THE FOLLOWING HAS BEEN REPORTED: HOSPITAL INFORMS THAT THE PUMP DID NOT WORK, CAUSING A DEATH IN THE NEONATAL UNIT OF THE (B)(6), AN INTERNAL ANALYSIS OF THE PUMPS HAS BEEN REQUESTED. HISTORICAL REPORT. MEDICATIONS: VASOPRESSIN TOTAL VALUE 9.6 ML, INFUSE: 0.2 ML/H IN 24 HOURS, INSTALLED AT 18:00 AND DOBUTAMINE TOTAL VALUE 4.8 ML, INFUSE 0.1 ML/H IN 24 HOURS, INSTALLED AT 08:00. AT 10:00 THE INFUSION WAS INCREASED TO 02.ML/H. IN TELEPHONE CONTACT WITH THE INITIAL NOTIFIER, SHE SAID THAT THE PUMPS WERE NOT WORKING AND THE PROGRAMMED VOLUMES HAD NOT BEEN ADMINISTERED, BUT THAT SHE WOULD CHECK MORE INFORMATION WITH THE DEPARTMENT RESPONSIBLE FOR THE PATIENT'S DIRECT CARE. THE PATIENT SUFFERED CARDIORESPIRATORY ARREST AND DIED. WE ARE IN CONTACT WITH THE CLIENT TO VERIFY MORE INFORMATION ABOUT THE REPORT. THE REPORT RECEIVED MENTIONS BOTH THE AGILIA SP MC BR, SERIAL NUMBER (B)(6), AND THE AGILIA INJECTOMAT SERINGA PUMP, SERIAL NUMBER (B)(6), HOWEVER IT DOES NOT CLARIFY WHICH DRUG WAS PROGRAMMED IN EACH PUMP. WHEN WE RECEIVE MORE INFORMATION, WE WILL SHARE IT IN THE COMMUNICATION. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EVENT EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.
AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | K150692 | THIS REPORT PERTAINS TO ONE OF TWO DEVICES USED DURING THE SAME PROCEDURE. REFER TO MANUFACTURER REPORT# 3005099803-2024-03613 FOR THE ASSOCIATED DEVICE INFORMATION. IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS WERE USED TRANSDUODENAL TO THE COMMON BILE DUCT TO TREAT METASTATIC PANCREATIC AND RENAL CANCER WITH AORTIC INVOLVEMENT DURING ENDOSCOPIC ULTRASOUND (EUS)-GUIDED BILIARY DRAINAGE (EUS-BD) AND ENDOSCOPIC ULTRASOUND (EUS)-GUIDED GALLBLADDER DRAINAGE (EUS-GBD) PROCEDURES PERFORMED ON (B)(6) 2024. DURING THE (EUS-BD) PROCEDURE, THE FIRST 8MM AXIOS STENT (SUBJECT OF MANUFACTURER REPORT # 3005099803-2024-03613) WAS ADVANCED THROUGH THE SCOPE. CANNULATION INTO THE COMMON BILE DUCT WAS SUCCESSFUL; HOWEVER, THE DELIVERY SYSTEM WAS NOT ABLE TO ADVANCE RETROGRADE OVER THE GUIDEWIRE (OTW). THE FIRST AXIOS STENT WAS THEN REMOVED, BUT THE STAINLESS-STEEL PLUG INSERT WAS DETACHED FROM THE HANDLE, LEAVING A STRAND OF EXPOSED COPPER WIRE PROTRUDING FROM THE PLUG ORIFICE. SUBSEQUENTLY, THE PHYSICIAN PROCEEDED TO PERFORM THE (EUS-GBD) PROCEDURE. DURING THE (EUS-GBD) PROCEDURE, THE SECOND 10MM AXIOS STENT (SUBJECT OF THIS REPORT) WAS SUCCESSFULLY IMPLANTED. HOWEVER, POST STENT PLACEMENT, THE PATIENT EXHIBITED LOW CARDIAC OUTPUT AND HAD CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS THEN PERFORMED FOR TEN MINUTES BEFORE A FEMORAL PULSE WAS DETECTED. THE PATIENT WAS TRANSFERRED TO ANGIOGRAPHY FOR IMAGING TO DETERMINE THE SOURCE OF THE BLOOD LOSS. A BLOOD TRANSFUSION WAS ADMINISTERED BEFORE AND DURING IMAGING WITH A PERFORATED ARTERIOLE OBSERVED UNDERNEATH THE DUODENAL FLANGE ON THE IMPLANTED 10MM AXIOS STENT. EMBOLIZATION WAS CONDUCTED AND HEMOSTASIS WAS SUCCESSFULLY ACHIEVED. THE PATIENT WAS TRANSFERRED TO ICU IN A STABLE CONDITION. | BLOCK H6: IMDRF PATIENT CODE E0511 CAPTURES THE REPORTABLE EVENT OF PERFORATION. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE EVENT OF MAJOR HEMORRHAGE. IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE EVENT OF CARDIAC ARREST. IMDRF PATIENT CODE E2401 CAPTURES THE REPORTABLE EVENT OF HYPOVOLEMIC SHOCK. IMDRF IMPACT CODE F2203 CAPTURES THE REPORTABLE EVENT OF FLUOROSCOPIC IMAGES WERE TAKEN. IMDRF IMPACT CODE F2302 CAPTURES THE REPORTABLE EVENT OF BLOOD TRANSFUSION WAS ADMINISTERED. IMDRF IMPACT CODE F2306 CAPTURES THE REPORTABLE EVENT OF CARDIOPULMONARY RESUSCITATION WAS PERFORMED. IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT (ICU) IMDRF IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF EMBOLISATION WAS CONDUCTED. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF THE PATIENT PASSED AWAY.
AGILIA SP MC BR | K121613 | THE FOLLOWING HAS BEEN REPORTED: HOSPITAL INFORMS THAT THE PUMP DID NOT WORK, CAUSING A DEATH IN THE NEONATAL UNIT OF THE VITÓRIA APART HOSPITAL, AN INTERNAL ANALYSIS OF THE PUMPS HAS BEEN REQUESTED. HISTORICAL REPORT. MEDICATIONS: VASOPRESSIN TOTAL VALUE 9.6 ML, INFUSE: 0.2 ML/H IN 24 HOURS, INSTALLED AT 18:00 AND DOBUTAMINE TOTAL VALUE 4.8 ML, INFUSE 0.1 ML/H IN 24 HOURS, INSTALLED AT 08:00. AT 10:00 THE INFUSION WAS INCREASED TO 02.ML/H. IN TELEPHONE CONTACT WITH THE INITIAL NOTIFIER, SHE SAID THAT THE PUMPS WERE NOT WORKING AND THE PROGRAMMED VOLUMES HAD NOT BEEN ADMINISTERED, BUT THAT SHE WOULD CHECK MORE INFORMATION WITH THE DEPARTMENT RESPONSIBLE FOR THE PATIENT'S DIRECT CARE. THE PATIENT SUFFERED CARDIORESPIRATORY ARREST AND DIED. WE ARE IN CONTACT WITH THE CLIENT TO VERIFY MORE INFORMATION ABOUT THE REPORT. THE REPORT RECEIVED MENTIONS BOTH THE AGILIA SP MC BR, SERIAL NUMBER (B)(6), AND THE AGILIA INJECTOMAT SERINGA PUMP, SERIAL NUMBER (B)(6), HOWEVER IT DOES NOT CLARIFY WHICH DRUG WAS PROGRAMMED IN EACH PUMP. WHEN WE RECEIVE MORE INFORMATION, WE WILL SHARE IT IN THE COMMUNICATION. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EVENT EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION BIPAP AVAPS30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION BIPAP AVAPS30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND GASTROESOPHAGEAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
EDWARDS ESHEATH PLUS INTRODUCER SHEATH | K200258 | AS REPORTED BY OUR EDWARDS LIFESCIENCES AFFILIATE IN GERMANY REGARDING A 26MM SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IMPLANT, DURING THE PROCEDURE, WHEN THE COMMANDER DELIVERY SYSTEM WITH CRIMPED VALVE WAS ADVANCED THROUGH THE ESHEATH, IT PERFORATED THE ESHEATH AND AORTA. THE PATIENT PASSED AWAY. | THE INVESTIGATION IS ONGOING.
MITRACLIP | K190167 | IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4. THE STEERABLE GUIDE CATHETER (SGC) WAS INSERTED AND ADVANCED INTO THE LEFT ATRIUM (LA). HOWEVER, WHILE REMOVING THE DILATOR, BUBBLES WERE OBSERVED IN THE LEFT VENTRICLE (LV) AND THE PATIENT HAD REDUCED MOBILITY IN THE VENTRICULAR WALL. MEDICATION WAS ADMINISTERED AND THE PATIENT WAS STABILIZED. THE PROCEDURE WAS CONTINUED AND TWO CLIPS WERE IMPLANTED, REDUCING MR TO A GRADE OF <1. HOWEVER, THE PATIENT DID NOT WAKE UP AFTER SEDATION AND DIED THE FOLLOWING MORNING DUE TO CEREBRAL EMBOLISM. | THE DEVICE IS EXPECTED TO BE RETURNED FOR EVALUATION. IT HAS NOT YET BEEN RECEIVED. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
CATH F5.2+ PIG 145° 110CM 6SH | K914007 | AS REPORTED, DUE TO A DEFECT IN THE UNKNOWN ANGIOGRAPHY PIGTAIL CATHETER (LEAKAGE FROM THE CATHETER), AIR LEAKED INTO THE CIRCULATION, LEADING TO AN AIR EMBOLISM. THE AIR EMBOLISM LED TO BRAIN DAMAGE WITH SUBSEQUENT CEREBRAL EDEMA. ULTIMATELY, THE PATIENT DIED NINE DAYS LATER. THE INCIDENT OCCURRED DURING A LEFT ATRIAL APPENDAGE (LAA) OCCLUDER PROCEDURE. WHILE HAVING THIS PROBLEM WITH AIR-EMBOLI, THEY REALIZED THAT THE DIAGNOSTIC CATHETER HAD A LEAKAGE. | ADDITIONAL INFORMATION IS PENDING AND WILL BE SUBMITTED WITHIN 30 DAYS UPON RECEIPT.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH DUE TO CANCER FOR USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED, AND MANUFACTURER COULD NOT CONFIRM PARTICLES IN AIR PATH DURING DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. H6 UPDATED/ CORRECTED.
ASTRAL 150 - AMER | K172875 | IT WAS REPORTED TO RESMED THAT A PATIENT WHO WAS VENTILATOR DEPENDENT AND USING AN ASTRAL 150 DEVICE EXPIRED ON JUNE 20, 2024. IT WAS REPORTED THAT A FEW DAYS BEFORE THE INCIDENT, THE PATIENT¿S EQUIPMENT PROVIDER HAD ADJUSTED THE ASTRAL DEVICE'S ALARMS. ON THE DAY OF THE INCIDENT, THE PATIENT WAS ATTENDED BY A HOME HEALTHCARE NURSE. THE PATIENT'S WIFE CLAIMED THAT THE NURSE FAILED TO PLUG THE DEVICE BACK INTO THE EXTERNAL POWER SOURCE AND THAT SHE NEVER HEARD ANY ALARMS. SHE RETURNED TO THE PATIENT SOMETIME AFTER THE NURSE HAD LEFT AND FOUND HIM DEAD. SHE ALSO FOUND THAT THE DEVICE WAS NOT WORKING. NO FURTHER INFORMATION HAS BEEN MADE AVAILABLE TO RESMED. | AN INVESTIGATION WILL BE PERFORMED ON ALL AVAILABLE INFORMATION BASED ON REPORTER¿S OBLIGATIONS UNDER APPLICABLE FDA REGULATIONS; HOWEVER, IT IS LIKELY THAT THIS MATTER MAY BECOME THE SUBJECT OF LITIGATION WHICH MAY LIMIT THE COMPANY¿S ABILITY TO DISCLOSE CONFIDENTIAL, PRIVILEGED ATTORNEY CLIENT COMMUNICATIONS AND WORK PRODUCT. INVESTIGATION METHODS, RESULTS AND CONCLUSIONS ARE NOT FINALIZED AT THIS STAGE PARTLY BECAUSE THE DEVICE IN QUESTION HAS NOT BEEN RETURNED FOR INSPECTION. RESMED HAS REQUESTED THE DEVICE BE RETURNED SO THAT AN INVESTIGATION CAN BE PERFORMED. RESMED REFERENCE#:(B)(4).
NIPRO SURDIAL DX HEMODIALYSIS SYSTEM | K182940 | PATIENT WAS ADMITTED TO THE HOSPITAL ON 06/21/2024 DUE TO LOW OXYGEN SATURATION AND GI ISSUES, DISCHARGED FROM THE HOSPITAL 07/02/202024. IT WAS HER FIRST TREATMENT IN THE CLINIC POST HOSPITALIZATION. TREATMENT STARTED 7:07AM WITH A BLOOD PRESSURE (BP) OF 110/49 AND IT STARTED DROPPING. 7:30 - BP 93/43 MM/HG: 8:00 - BP 95/52 MM/HG AND 9 AM - BP WAS 99/38 MM/HG AND THEY STOPPED THE ULTRAFILTRATION (UF). PATIENT STARTED VOMITING AND PASSED OUT, CPR WAS PERFORMED, FLUIDS WERE GIVEN AND PARAMEDICS WERE CALLED. THEY TRIED ENDING THE TREATMENT BUT WERE UNABLE TO UNTIL AFTER PATIENT WAS PRONOUNCED DEAD @ 10:24 AM. DIALYZER/BLOODLINES WERE SAVED. STAFF STATED THAT THERE WERE NO ALARMS TRIGGERED BEFORE OR DURING THE EVENT, EXCEPT FOR TREATMENT RELATED ALARMS PRIOR TO EVENT. CAUSE OF DEATH IS STILL UNKNOWN. THERE WERE NO ISSUES REPORTED WITH ANY OF THE NIPRO DEVICES USED. TREATMENT TIME - 3.5 HRS; BLOOD FLOW RATE (BFR) - 250; DIALYSATE FLOW RATE (DFR) - 600; A/B - 3K, 2.5; BICARB - 35, TEMP - 35.5; NA - 138; ACCESS: CVC; UF GOAL - 4L; UF REMOVED - 1.5 L; ESTIMATED DRY WEIGHT (EDW) - 74 KG; PRE WEIGHT - 77.5 KG MEDICATIONS GIVEN; MIRCERA - 60 MCG; VENOFER - 50 MG COMORBIDS: CANCER, CHRONIC OBSTRUCTION PULMONARY DISEASE (COPD) MACHINE S/N: 22EW0301. MACHINE PULLED OUT OF SERVICE, DIALYSATE SAMPLES WERE DRAWN. OTHER DEVICES USE: NIPRO BLOOD TUBING SET, CODE: BL+A430/V912, LOT# 23K29 NIPRO ELISIO-H DIALYZER, CODE: DD+ELISIO-17H, LOT# 24C28H
POWER-TRIALYSIS SHORT-TERM CURVED EXTENSION DIALYSIS | K133456 | IT WAS REPORTED BY THE CUSTOMER THAT PATIENT WITH ASPIRATION PNEUMONIA AND END-STAGE RENAL FAILURE. THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO A SHUNT CLOG, AND AN EMERGENCY PTA WAS PERFORMED, BUT THE SHUNT COULD NOT BE OPENED, AND A VASCAS WAS PLACED. ON THE 8TH DAY OF HOSPITALIZATION, THE PATIENT WAS SENT TO THE HOSPITAL FOR SOFT CELL INSERTION, BUT THE INSERTION WAS DIFFICULT, AND AS A RESULT, THE VASCAS WAS REINSERTED. ON THE 12TH DAY OF HOSPITALIZATION, THE PATIENT STARTED DIALYSIS AND IMPLEMENTATION WENT WELL FOR A WHILE. AFTER THAT, OXYGENATION DECREASED AND ASPIRATE AND OXYGEN ADMINISTRATION TEMPORARILY IMPROVED THE PATIENT'S CONDITION, BUT AFTER ABOUT 2 HOURS, A LARGE AMOUNT OF BUBBLES WERE OBSERVED IN THE CHAMBER ON THE DE-OXYGENATED SIDE. THERE WAS NO PROBLEM FOR A WHILE AFTER THE START OF DIALYSIS, BUT AFTER ABOUT 2 HOURS, A LARGE AMOUNT OF AIR BUBBLES WERE OBSERVED IN THE CHAMBER ON THE DEOXYGENATED SIDE. THE CAUSE OF THIS AIR CONTAMINATION IS UNKNOWN. DUE TO THE WISHES OF THE FAMILY, AUTOPSY AND AI HAVE NOT BEEN PERFORMED, AND THE CAUSE REMAINS UNKNOWN. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE DATE OF EVENT AND THE DATE THE PATIENT EXPIRED WERE NOT PROVIDED BY THE COMPLAINANT/REPORTER, THE DATE OF EVENT AND DATE OF DEATH REFLECTED IN THIS REPORT ARE THE DATE BD BECAME AWARE OF THE EVENT. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
NIPRO BLOOD TUBING SET W/TP | K182940 | PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2024 DUE TO LOW OXYGEN SATURATION AND GI ISSUES, DISCHARGED FROM THE HOSPITAL ON (B)(6) "2024". IT WAS HER FIRST TREATMENT IN THE CLINIC POST HOSPITALIZATION. TREATMENT STARTED 7:07AM WITH A BLOOD PRESSURE (BP) OF 110/49 AND IT STARTED DROPPING. 7:30 - BP 93/43 MM/HG: 8:00 - BP 95/52 MM/HG AND 9 AM - BP WAS 99/38 MM/HG AND THEY STOPPED THE ULTRAFILTRATION (UF). PATIENT STARTED VOMITING AND PASSED OUT, CPR WAS PERFORMED, FLUIDS WERE GIVEN AND PARAMEDICS WERE CALLED. THEY TRIED ENDING THE TREATMENT BUT WERE UNABLE TO UNTIL AFTER PATIENT WAS PRONOUNCED DEAD @ 10:24 AM. DIALYZER/BLOODLINES WERE SAVED. STAFF STATED THAT THERE WERE NO ALARMS TRIGGERED BEFORE OR DURING THE EVENT, EXCEPT FOR TREATMENT RELATED ALARMS PRIOR TO EVENT. CAUSE OF DEATH IS STILL UNKNOWN. THERE WERE NO ISSUES REPORTED WITH ANY OF THE NIPRO DEVICES USED. TREATMENT TIME - 3.5 HRS; BLOOD FLOW RATE (BFR) - 250; DIALYSATE FLOW RATE (DFR) - 600; A/B - 3K, 2.5; BICARB - 35, TEMP - 35.5; NA - 138; ACCESS: CVC; UF GOAL - 4L; UF REMOVED - 1.5 L; ESTIMATED DRY WEIGHT (EDW) - 74 KG; PRE WEIGHT - 77.5 KG MEDICATIONS GIVEN; MIRCERA - 60 MCG; VENOFER - 50 MG. COMORBIDS: CANCER, CHRONIC OBSTRUCTION PULMONARY DISEASE (COPD). MACHINE S/N: (B)(6) . MACHINE PULLED OUT OF SERVICE, DIALYSATE SAMPLES WERE DRAWN. OTHER DEVICES USE: NIPRO SURDIAL DX HEMODIALYSIS SYSTEM, CODE: MC+SDX01, SERIAL (B)(6) . NIPRO ELISIO-H DIALYZER, CODE: DD+ELISIO-17H, LOT# 24C28H. | 8/26/2024: CORRECTION MADE TO SECTION H5. PREVIOUSLY INCORRECTLY SELECTED "NO". CORRECTED TO "YES".
NIPRO ELISIO-H DIALYZER | K131935 | PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2024 DUE TO LOW OXYGEN SATURATION AND GI ISSUES, DISCHARGED FROM THE HOSPITAL (B)(6) 2024. IT WAS HER FIRST TREATMENT IN THE CLINIC POST HOSPITALIZATION. TREATMENT STARTED 7:07AM WITH A BLOOD PRESSURE (BP) OF 110/49 AND IT STARTED DROPPING. 7:30 - BP 93/43 MM/HG: 8:00 - BP 95/52 MM/HG AND 9 AM - BP WAS 99/38 MM/HG AND THEY STOPPED THE ULTRAFILTRATION (UF). PATIENT STARTED VOMITING AND PASSED OUT, CPR WAS PERFORMED, FLUIDS WERE GIVEN AND PARAMEDICS WERE CALLED. THEY TRIED ENDING THE TREATMENT BUT WERE UNABLE TO UNTIL AFTER PATIENT WAS PRONOUNCED DEAD @ 10:24 AM. DIALYZER/BLOODLINES WERE SAVED. STAFF STATED THAT THERE WERE NO ALARMS TRIGGERED BEFORE OR DURING THE EVENT, EXCEPT FOR TREATMENT RELATED ALARMS PRIOR TO EVENT. CAUSE OF DEATH IS STILL UNKNOWN. THERE WERE NO ISSUES REPORTED WITH ANY OF THE NIPRO DEVICES USED. TREATMENT TIME - 3.5 HRS; BLOOD FLOW RATE (BFR) - 250; DIALYSATE FLOW RATE (DFR) - 600; A/B - 3K, 2.5; BICARB - 35, TEMP - 35.5; NA - 138; ACCESS: CVC; UF GOAL - 4L; UF REMOVED - 1.5 L; ESTIMATED DRY WEIGHT (EDW) - 74 KG; PRE WEIGHT - 77.5 KG. MEDICATIONS GIVEN; MIRCERA - 60 MCG; VENOFER - 50 MG COMORBIDS: CANCER, CHRONIC OBSTRUCTION PULMONARY DISEASE (COPD) MACHINE S/N: (B)(6). MACHINE PULLED OUT OF SERVICE; DIALYSATE SAMPLES WERE DRAWN. OTHER DEVICES USE: NIPRO SURDIAL DX HEMODIALYSIS SYSTEM, CODE: MC+SDX01, SERIAL#(B)(6) NIPRO BLOOD TUBING SET W/TP, CODE: BL+A430/V912, LOT# 23K29.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED SURGICAL PROCEDURE, THE PATIENT EXPIRED. THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE WAS NOTIFIED OF THE PATIENT DEATH BY A SURGEON; NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS TO OBTAIN FURTHER INFORMATION WERE MADE BUT HAVE NOT BEEN SUCCESSFUL. | THE FOLLOWING INVESTIGATIONS COULD NOT BE PERFORMED DUE TO INSUFFICIENT INFORMATION PROVIDED (I.E. PROCEDURE/EVENT DATE, UNCONFIRMED SURGEON NAME, PROCEDURE TYPE): EVENT VERIFICATION, SYSTEM/INSTRUMENT LOG REVIEW FOR THE EVENT, PROCEDURE INFORMATION. THE USER FACILITY HAS TWO DAVINCI ROBOTIC SYSTEMS. THE INFORMATION IN SECTION D, SUSPECT MEDICAL DEVICE, IS POTENTIALLY THE DEVICE USED IN THE EVENT. THEIR OTHER DAVINCI SYSTEM IS MODEL# 380652-55, SERIAL (B)(6), UDI (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED ON BOTH OF THE DAVINCI SYSTEMS CURRENTLY IN USE AT THIS HOSPITAL WITH NO NON-CONFORMANCE'S IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SITE'S COMPLAINT HISTORY DOES NOT SHOW ANY REPORTED EVENTS RELATED TO THIS EVENT. A REVIEW OF THE EVENT WAS PERFORMED BY AN INTUITIVE SURGICAL, INC. (ISI) MEDICAL SAFETY OFFICER (MSO) WHO CONCLUDED THAT THE TYPE OF PROCEDURE AND THE EVENTS AROUND THIS PATIENT¿S DEATH WERE NOT REPORTED. NO ALLEGATIONS AGAINST ANY INTUITIVE PRODUCTS WERE MADE. INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL SYSTEM OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED SURGICAL PROCEDURE, THE PATIENT EXPIRED. THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE WAS NOTIFIED OF THE PATIENT DEATH BY A SURGEON; NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS TO OBTAIN FURTHER INFORMATION WERE MADE BUT HAVE NOT BEEN SUCCESSFUL. | THE FOLLOWING INVESTIGATIONS COULD NOT BE PERFORMED DUE TO INSUFFICIENT INFORMATION PROVIDED (I.E. PROCEDURE/EVENT DATE, UNCONFIRMED SURGEON NAME, PROCEDURE TYPE): EVENT VERIFICATION, SYSTEM/INSTRUMENT LOG REVIEW FOR THE EVENT, PROCEDURE INFORMATION. THE USER FACILITY HAS TWO DAVINCI ROBOTIC SYSTEMS. THE INFORMATION IN SECTION D, SUSPECT MEDICAL DEVICE, IS POTENTIALLY THE DEVICE USED IN THE EVENT. THEIR OTHER DAVINCI SYSTEM IS MODEL# 380652-55, SERIAL# (B)(6), UDI (B)(4). DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED ON BOTH OF THE DAVINCI SYSTEMS CURRENTLY IN USE AT THIS HOSPITAL WITH NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SITE'S COMPLAINT HISTORY DOES NOT SHOW ANY REPORTED EVENTS RELATED TO THIS EVENT. A REVIEW OF THE EVENT WAS PERFORMED BY AN INTUITIVE SURGICAL, INC. (ISI) MEDICAL SAFETY OFFICER (MSO) WHO CONCLUDED THAT THE TYPE OF PROCEDURE AND THE EVENTS AROUND THIS PATIENT¿S DEATH WERE NOT REPORTED. NO ALLEGATIONS AGAINST ANY INTUITIVE PRODUCTS WERE MADE. INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL SYSTEM OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS ZERO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT CORRECTED. SECTION H6 UPDATED IN THIS REPORT.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED EYE IRRITATION, NOSE IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED EYE IRRITATION, NOSE IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REACT CATHETER | K182097 | MEDTRONIC RECEIVED A REPORT THAT THE PATIENT WAS DISCHARGED FROM THROMBECTOMY AND ADMITTED TO THE HOSPITAL'S REHABILITATION UNIT FOR REHABILITATION AND DIED ON (B)(6) 2024 AFTER RECURRENT FEVER FROM SEVERE PNEUMONIA. BASELINE MODIFIED RANKIN SCALE (MRS) SCORE 0, NATIONAL INSTITUTES OF HEALTH STROKE SCALE (NIHSS) 16. THIS EVENT HAD AN ONSET DATE OF 2024-02-07. OTHER ACTIONS TAKEN INCLUDED INTRAVENOUS INFUSION OF ANTI-INFECTION AND ANTIFUNGAL MEDICATIONS; VENTILATION BY HUMIDIFIED RESPIRATOR. DIAGNOSTIC TESTS ON (B)(6) 2024 C-REACTIVE PROTEIN WITH RESULT 630.7UG/L, SIGNIFICANT; BODY TEMPERATURE RESULT 38 DEGREES CELSIUS, SIGNIFICANT; CALCITONINOGEN RESULT 2.14UG/ML, SIGNIFICANT. THIS EVENT OCCURRED POST-PROCEDURE AND WAS NOT A RECURRENT OR NEW STROKE. ADVERSE EVENT (AE) WAS NOT RELATED TO THE DISEASE UNDER STUDY AND HAD A CAUSAL RELATIONSHIP WITH AN UNDERLYING CONDITION OR DISEASE. AE DID NOT RESULT FROM A DEVICE DEFICIENCY. THE SITE ASSESSED THIS EVENT AS NOT RELATED TO THE STUDY DEVICE OR PROCEDURE. ON 2024-JUL-31 THE CLINICAL EVENTS COMMITTEE (CEC) ADJUDICATED THIS EVENT AS POSSIBLY RELATED TO THE STUDY PROCEDURE. THE PRE-PROCEDURE MTICI SCORE WAS 0, FINAL POST-PROCEDURE MTICI SCORE WAS 3. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
NXSTAGE SYSTEM ONE | K122051 | A REPORT WAS RECEIVED ON 11 JUL 2024 FROM THE HOME THERAPY NURSE (HTN) OF A 70 YEAR OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING PRIOR CARDIAC EVENTS, CONGESTIVE HEART FAILURE, ACUTE RESPIRATORY FAILURE AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPERIENCED SHORTNESS OF BREATH, CHEST PAIN AND BECAME UNRESPONSIVE DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 18 JUL 2024 FROM THE HTN WHO STATED THE PATIENT WAS TRANSFERRED TO HOSPITAL AND WAS PULSELESS AND APNEIC UPON ARRIVAL. RESUSCITATIVE EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT EXPIRED ON (B)(6) 2024. PER THE HOSPITAL RECORD PROVIDED BY THE HTN, THE CAUSE OF DEATH WAS CARDIOPULMONARY ARREST. | THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
BIPAP A40 PRO | K121623 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A PATIENT HAD SADLY PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING THAT A PATIENT HAD SADLY PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS EVALUATED BY THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY (PIL) AND UNABLE TO CONFIRM AN INTERRUPTION OF THERAPY OF THIS BIPAP A40 PRO. EXTERNAL VISUAL INSPECTION FOUND NO DEFECTS. PIL DOWNLOADED THE DEVICE ERROR LOG TO REVIEW THE EVENTS FROM THE LOG. THE INCIDENT WAS REPORTED ON 07/09/2024. THE PATIENT WAS REPORTED TO BE USING A DIFFERENT DEVICE AT THE TIME OF THE INCIDENT. THE ERROR LOG REVIEW DID FIND PATIENT RELATED ALARMS, AND NOT DEFECTIVE COMPONENT ALARMS. THIS DEVICE WAS POWERED OFF (B)(6) 2024 AND NOT POWERED UP AGAIN UNTIL (B)(6) 2025. AT PIL, THE UNIT RAN 30 MINUTES ON QUICKLUNG F7797, WITHOUT ALARM OR ERROR. NO ALLEGATIONS HAVE BEEN MADE AGAINST THIS DEVICE. THIS UNIT WILL BE SCRAPPED PER THE REQUIREMENTS SET FORTH IN WORK INSTRUCTION 8.4-1131 REV 9 AND DISPOSED OF ACCORDINGLY. PRIOR TO IDENTIFYING THE ISSUE DESCRIBED IN FSN 2023-CC-SRC-039, COMPLAINTS ASSOCIATED WITH THE ISSUE DID NOT MEET REQUIREMENTS FOR VIGILANCE REPORTING AS A REPORTABLE INCIDENT. AS PART OF THE HHE PROCESS, A RETROSPECTIVE REVIEW OF COMPLAINTS ASSOCIATED WITH THE ISSUE DESCRIBED IN FSN 2023-CC-SRC-039 WAS NECESSARY TO REASSESS REPORTABILITY. UPON THIS FURTHER ASSESSMENT, THESE COMPLAINTS ARE BEING REPORTED. THE BIPAP A40 PRO (RGBX3100S19) IS SUBSTANTIALLY SIMILAR TO THE BIPAP A40 AND WILL BE REPORTED IN THE UNITED STATES UNDER BIPAP A40, 510K NUMBER: K121623.
TEMPUS LS | K200849 | IT WAS REPORTED WHEN WORKING A CARDIAC ARREST EVENT, THE DEVICE MALFUNCTIONED AND STARTED READING HARDWARE FAILURE (DPM). THE DEVICE WAS IN CLINICAL USE AT THE TIME OF THE EVENT. A PATIENT DEATH WAS REPORTED.
VISISHEATH DILATOR SHEATH | K092378 | A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD DUE TO FRACTURE, VENTRICULAR TACHYCARDIA (VT) STATUS POST ICD LEAD PLACEMENT, AND THROMBOSED INNOMINATE VEIN. A SPECTRANETICS LLD EZ LEAD LOCKING DEVICE (LLD EZ) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION. MULTIPLE SPECTRANETICS DEVICES (14F GLIDELIGHT LASER SHEATH, 16F GLIDELIGHT, 13F TIGHTRAIL ROTATING DILATOR SHEATH, LARGE VISISHEATH DILATOR SHEATH) WERE USED DURING THE PROCEDURE. HOWEVER, WHILE THE 16F GLIDELIGHT WAS IN USE AND THE VISISHEATH WAS PUSHED OVER THE LEAD'S COIL, PROGRESS STALLED IN THE SUPERIOR VENA CAVA (SVC). THEN, THE VISISHEATH WAS USED AGGRESSIVELY IN AN ATTEMPT TO ADVANCE THE DEVICE DOWN THE VASCULATURE, BUT THE PATIENT'S BLOOD PRESSURE DROPPED DRAMATICALLY. RESCUE EFFORTS BEGAN IMMEDIATELY, INCLUDING RESCUE BALLOON, PERICARDIOCENTESIS, STERNOTOMY, AND BYPASS. AS THE INJURY LOCATION WAS BEING DETERMINED, IT WAS DISCOVERED THAT THE PATIENT'S RIGHT INTERNAL JUGULAR AND LEFT SUBCLAVIAN/INNOMINATE VEIN WERE COMPLETELY DISCONNECTED FROM THE SUPERIOR VENA CAVA (SVC). AFTER HOURS OF ATTEMPTING TO REPAIR THE INJURY AT THE SVC, RESCUE EFFORTS WERE UNSUCCESSFUL AND THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE LARGE VISISHEATH IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. | A2): PATIENT DATE OF BIRTH UNK. A4): PATIENT WEIGHT UNK. B7): OTHER RELEVANT HISTORY UNK. D4/H4: THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3): THE VISISHEATH WAS NOT RETURNED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6): ALTHOUGH GREAT VESSEL PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE VISISHEATH, THE VISISHEATH WAS USED AGGRESSIVELY TO ATTEMPT ADVANCEMENT DOWN THE VASCULATURE. THE INSTRUCTIONS FOR USE (IFU) STATES TO ALWAYS MAINTAIN ADEQUATE TENSION AND COAXIAL ALIGNMENT ON THE LEAD TO SUPPORT THE MANEUVERING OF THE DILATOR SHEATHS AND ANY ACCOMPANYING INNER SHEATHS TO PROPERLY GUIDE THEM WITHIN THE PATIENT ANATOMY. IN ADDITION, IT STATES IF THE DILATOR SHEATH FAILS TO PROGRESS AFTER INITIAL SUCCESS, OR IF ADVANCING THE SHEATH WAS DIFFICULT, REMOVE THE SHEATH TO INSPECT THE TIP. IF THE TIP IS DISTORTED OR FRAYED, EXCHANGE THE DAMAGED SHEATH FOR A NEW SHEATH BEFORE CONTINUING TREATMENT. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NAUSEA / VOMITING, LIVER DISEASE/TOXICITY, CANCER, DEATH AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
SUREFORM | K173721 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED TOTAL GASTRECTOMY FOR GASTRIC CANCER, THE PATIENT EXPERIENCED POST-OPERATIVE ANASTOMOTIC LEAKS THAT RESULTED IN SEPSIS, AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY (POD) 12. THE SURGEON STATED THERE WERE NO INTRA-OPERATIVE COMPLICATIONS DURING THE PROCEDURE. THE SUREFORM 60 STAPLER INSTRUMENT WITH BLUE AND GREEN RELOADS WAS USED DURING THE PROCEDURE. ON POD 1, THE PATIENT EXPERIENCED RESPIRATORY CHALLENGES AND BILATERAL PLEURAL EXUDATE REQUIRING CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) TREATMENT. ON POD 2, THE PATIENT SHOWED SIGNS OF SEPSIS AND A CT SCAN WITH INTRALUMINAL CONTRAST FOUND LEAKAGE FROM THE STAPLED ESOPHAGO-JEJUNO ANASTOMOSIS. IN THE FOLLOWING DAYS, THE PATIENT WAS ALSO DIAGNOSED WITH INTESTINAL LEAKAGE FROM THE DUODENAL AND ENTERO-ENTERO STAPLER ANASTOMOSES. POST-OPERATIVE TREATMENT INCLUDED RESPIRATORY SUPPORT, UNSPECIFIED "DRAINAGE", AND ANTIBIOTICS. NO INFORMATION WAS AVAILABLE REGARDING IF ANY ADDITIONAL SURGICAL PROCEDURES WERE PERFORMED. THE SURGEON STATED THAT ALL THREE SITES STAPLED WITH THE SUREFORM 60 STAPLER AND RELOADS ARE PRESUMED TO BE DEFECTIVE AS THE STAPLE LINES EXPERIENCED POST-OPERATIVE ANASTOMOTIC DEFECTS. THE CAUSE OF DEATH WAS REPORTED AS SEPSIS AND MULTIPLE ORGAN FAILURE. | THE SUREFORM 60 STAPLER INSTRUMENT AND BLUE AND GREEN RELOADS ARE NOT AVAILABLE FOR FAILURE ANALYSIS AS THEY WERE DISCARDED AT THE HOSPITAL. REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES RELATED TO THIS EVENT. THE SUREFORM 60 STAPLER LOGS DID NOT SHOW ANY RELEVANT ERRORS. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE, HARMONIC ACE, SYNCHROSEAL, SUREFORM STAPLER 60, TIP-UP FENESTRATED GRASPER, LARGE NEEDLE DRIVER, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, LARGE SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND ONE REPORT WAS RECEIVED FOR THE LARGE SUTURECUT NEEDLE DRIVER INSTRUMENT USED DURING THIS PROCEDURE STATING THE INSTRUMENT JAWS WOULD NOT OPEN WHILE IT WAS BEING USED TO SUTURE. THIS COMPLAINT IS NOT RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT WAS PERFORMED BY AN INTUITIVE SURGICAL, INC. (ISI) MEDICAL SAFETY OFFICER (MSO) WHO CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A TOTAL GASTRECTOMY FOR GASTRIC CANCER WITH MULTIPLE USES OF THE SUREFORM 60 STAPLER. ALTHOUGH NO INTRAOPERATIVE ISSUES WERE NOTED, THE STAPLE LINES EVENTUALLY LEAKED POSTOPERATIVELY WHICH LED TO SEPSIS AND ADDITIONAL COMPLICATIONS. HOW THE LEAKS WERE EVALUATED FOR POTENTIAL CAUSES OR CONTRIBUTING FACTORS AND HOW THE LEAKS WERE ADDRESSED WAS NOT REPORTED. THE PATIENT EVENTUALLY DIED. THERE ARE MULTIPLE POSSIBLE CONTRIBUTING OR CAUSATIVE FACTORS WHICH INDIVIDUALLY OR CUMULATIVELY COULD LEAD TO A STAPLE LINE FAILURE. THESE INCLUDE SURGICAL TECHNIQUE, TISSUE ISCHEMIA, AND PATIENT DISEASE. FOR EXAMPLE, THIS PATIENT HAD CLAUDICATION WHICH SUGGESTS VASCULAR DISEASE. THIS PATIENT ALSO HAD GASTRIC CANCER AND THE MARGINS MAY NOT HAVE BEEN CLEAR. THE FACT THAT MULTIPLE LEAKS OCCURRED DESPITE NO DEVICE OR ANASTOMOSIS ISSUES NOTED INTRAOPERATIVELY MAY SUGGEST A PATIENT RELATED SYSTEMIC ISSUE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
SUREFORM | K173721 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED TOTAL GASTRECTOMY FOR GASTRIC CANCER, THE PATIENT EXPERIENCED POST-OPERATIVE ANASTOMOTIC LEAKS THAT RESULTED IN SEPSIS, AND ULTIMATELY EXPIRED ON POST-OPERATIVE DAY (POD) 12. THE SURGEON STATED THERE WERE NO INTRA-OPERATIVE COMPLICATIONS DURING THE PROCEDURE. THE SUREFORM 60 STAPLER INSTRUMENT WITH BLUE AND GREEN RELOADS WAS USED DURING THE PROCEDURE. ON POD 1, THE PATIENT EXPERIENCED RESPIRATORY CHALLENGES AND BILATERAL PLEURAL EXUDATE REQUIRING CONTINUOUS POSITIVE AIRWAY PRESSURE (CPAP) TREATMENT. ON POD 2, THE PATIENT SHOWED SIGNS OF SEPSIS AND A CT SCAN WITH INTRALUMINAL CONTRAST FOUND LEAKAGE FROM THE STAPLED ESOPHAGO-JEJUNO ANASTOMOSIS. IN THE FOLLOWING DAYS, THE PATIENT WAS ALSO DIAGNOSED WITH INTESTINAL LEAKAGE FROM THE DUODENAL AND ENTERO-ENTERO STAPLER ANASTOMOSES. POST-OPERATIVE TREATMENT INCLUDED RESPIRATORY SUPPORT, UNSPECIFIED "DRAINAGE", AND ANTIBIOTICS. NO INFORMATION WAS AVAILABLE REGARDING IF ANY ADDITIONAL SURGICAL PROCEDURES WERE PERFORMED. THE SURGEON STATED THAT ALL THREE SITES STAPLED WITH THE SUREFORM 60 STAPLER AND RELOADS ARE PRESUMED TO BE DEFECTIVE AS THE STAPLE LINES EXPERIENCED POST-OPERATIVE ANASTOMOTIC DEFECTS. THE CAUSE OF DEATH WAS REPORTED AS SEPSIS AND MULTIPLE ORGAN FAILURE. | THE SUREFORM 60 STAPLER INSTRUMENT AND BLUE AND GREEN RELOADS ARE NOT AVAILABLE FOR FAILURE ANALYSIS AS THEY WERE DISCARDED AT THE HOSPITAL. REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCE'S RELATED TO THIS EVENT. THE SUREFORM 60 STAPLER LOGS DID NOT SHOW ANY RELEVANT ERRORS. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE, HARMONIC ACE, SYNCHROSEAL, SUREFORM STAPLER 60, TIP-UP FENESTRATED GRASPER, LARGE NEEDLE DRIVER, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, LARGE SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND ONE REPORT WAS RECEIVED FOR THE LARGE SUTURECUT NEEDLE DRIVER INSTRUMENT USED DURING THIS PROCEDURE STATING THE INSTRUMENT JAWS WOULD NOT OPEN WHILE IT WAS BEING USED TO SUTURE. THIS COMPLAINT IS NOT RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT WAS PERFORMED BY AN INTUITIVE SURGICAL, INC. (ISI) MEDICAL SAFETY OFFICER (MSO) WHO CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A TOTAL GASTRECTOMY FOR GASTRIC CANCER WITH MULTIPLE USES OF THE SUREFORM 60 STAPLER. ALTHOUGH NO INTRAOPERATIVE ISSUES WERE NOTED, THE STAPLE LINES EVENTUALLY LEAKED POSTOPERATIVELY WHICH LED TO SEPSIS AND ADDITIONAL COMPLICATIONS. HOW THE LEAKS WERE EVALUATED FOR POTENTIAL CAUSES OR CONTRIBUTING FACTORS AND HOW THE LEAKS WERE ADDRESSED WAS NOT REPORTED. THE PATIENT EVENTUALLY DIED. THERE ARE MULTIPLE POSSIBLE CONTRIBUTING OR CAUSATIVE FACTORS WHICH INDIVIDUALLY OR CUMULATIVELY COULD LEAD TO A STAPLE LINE FAILURE. THESE INCLUDE SURGICAL TECHNIQUE, TISSUE ISCHEMIA, AND PATIENT DISEASE. FOR EXAMPLE, THIS PATIENT HAD CLAUDICATION WHICH SUGGESTS VASCULAR DISEASE. THIS PATIENT ALSO HAD GASTRIC CANCER AND THE MARGINS MAY NOT HAVE BEEN CLEAR. THE FACT THAT MULTIPLE LEAKS OCCURRED DESPITE NO DEVICE OR ANASTOMOSIS ISSUES NOTED INTRAOPERATIVELY MAY SUGGEST A PATIENT RELATED SYSTEMIC ISSUE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
ION | K182188 | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING AND SUBSEQUENTLY EXPIRED. WHILE THE PHYSICIAN WAS PERFORMING A BIOPSY OF A LESION IN THE RIGHT MIDDLE LOBE, BLUSHING WAS VISIBLE ON 2-D FLUOROSCOPY IMAGING. THE ION VISION PROBE WAS REINSERTED INTO THE ION CATHETER AND A SMALL AMOUNT OF BLOOD WAS NOTED. BRONCHOALVEOLAR LAVAGE (BAL) WAS PERFORMED TO RETRIEVE THE BIOPSY SPECIMEN, AFTER WHICH APPROXIMATELY 20ML OF BLOOD WAS NOTED. THE ION SYSTEM WAS UNDOCKED AND A MANUAL THERAPEUTIC BRONCHOSCOPE WAS PLACED. BLOOD WAS OBSERVED IN THE ENDOTRACHEAL TUBE AND ¿A LOT¿ OF BLOOD WAS SUCTIONED OUT. THE PATIENT STOPPED BREATHING AND THE LUNGS COULD NOT BE OXYGENATED; ADDITIONAL BLOOD WAS SUCTIONED ALLOWING THE PATIENT TO SUCCESSFULLY BE PLACED ON VENTILATOR SUPPORT. AFTER STABILIZING FOR APPROXIMATELY 60 MINUTES, THE PATIENT CODED, REPORTEDLY DUE TO MANY BLOOD CLOTS THAT CREATED PRESSURE IN THE PATIENT'S LUNGS. RESUSCITATION WAS PERFORMED SUCCESSFULLY. THE PATIENT WAS TRANSFERRED TO THE ICU AND EXPIRED APPROXIMATELY 90 MINUTES LATER. | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING AND SUBSEQUENTLY EXPIRED. WHILE THE PHYSICIAN WAS PERFORMING A BIOPSY OF A LESION IN THE RIGHT MIDDLE LOBE, BLUSHING WAS VISIBLE ON 2-D FLUOROSCOPY IMAGING. THE ION VISION PROBE WAS REINSERTED INTO THE ION CATHETER AND A SMALL AMOUNT OF BLOOD WAS NOTED. BRONCHOALVEOLAR LAVAGE (BAL) WAS PERFORMED TO RETRIEVE THE BIOPSY SPECIMEN, AFTER WHICH APPROXIMATELY 20ML OF BLOOD WAS NOTED. THE ION SYSTEM WAS UNDOCKED AND A MANUAL THERAPEUTIC BRONCHOSCOPE WAS PLACED. BLOOD WAS OBSERVED IN THE ENDOTRACHEAL TUBE AND ¿A LOT¿ OF BLOOD WAS SUCTIONED OUT. THE PATIENT STOPPED BREATHING AND THE LUNGS COULD NOT BE OXYGENATED; ADDITIONAL BLOOD WAS SUCTIONED ALLOWING THE PATIENT TO SUCCESSFULLY BE PLACED ON VENTILATOR SUPPORT. AFTER STABILIZING FOR APPROXIMATELY 60 MINUTES, THE PATIENT CODED, REPORTEDLY DUE TO MANY BLOOD CLOTS THAT CREATED PRESSURE IN THE PATIENT'S LUNGS. RESUSCITATION WAS PERFORMED SUCCESSFULLY. THE PATIENT WAS TRANSFERRED TO THE ICU AND EXPIRED APPROXIMATELY 90 MINUTES LATER. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH . MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS TWENTY-FOUR ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
BIPAP AUTO | K043607 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAD PASSED AWAY, SO THE PATIENT'S CHILD WANTS THE NAME ON THE CHEQUE TO BE CHANGED TO HER NAME INSTEAD OF HER FATHER'S NAME. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
COBLATION HALO WAND | K192027 | IT WAS REPORTED THAT NINE DAYS AFTER A TONSILLECTOMY USING A COBLATION HALO WAND, THE PATIENT HAD A MEDICAL INCIDENT AND PASSED AWAY. THE CAUSE OR FURTHER INFORMATION IS UNKNOWN AT THIS TIME. | INTERNAL COMPLAINT REFERENCE (B)(4). H10 H3, H6: A DEVICE DEFICIENCY WAS NOT IDENTIFIED, AND THE ROOT CAUSE OF THE REPORTED EVENT COULD NOT BE DETERMINED DUE TO THE LIMITED CLINICAL INFORMATION PROVIDED. INSUFFICIENT PRODUCT IDENTIFICATION INFORMATION WAS PROVIDED AND THUS A MANUFACTURING RECORD REVIEW COULD NOT BE CONDUCTED. BASED ON THE INFORMATION AVAILABLE, THERE WERE NO INDICATIONS TO SUGGEST THAT THE PRODUCT DID NOT MEET MANUFACTURING SPECIFICATIONS UPON RELEASE FOR DISTRIBUTION. A COMPLAINT HISTORY REVIEW FOUND SIMILAR REPORTED EVENTS. A RISK MANAGEMENT REVIEW FOUND THAT THERE WERE NO INDICATIONS TO SUGGEST THE ANTICIPATED RISK IS NOT ADEQUATE. THE INSTRUCTIONS FOR USE WERE REVIEWED AND FOUND TO INCLUDE CONDITIONS OF OFF LABEL USE AND TECHNIQUE SPECIFICS, AS WELL AS PRECAUTIONS AND WARNINGS RELATED TO THE USE OF THE DEVICE. A CLINICAL REVIEW STATES THAT IT IS UNKNOWN WHAT TYPE OF ¿MEDICAL INCIDENT¿ THE PATIENT EXPERIENCED. AS OF THE DATE OF THIS MEDICAL INVESTIGATION, NO RELEVANT SUPPORTING CLINICAL INFORMATION HAS BEEN PROVIDED TO ASSIST WITH A CLINICAL INVESTIGATION SURROUNDING THE CIRCUMSTANCES OF THE PATIENT¿S DEATH. THEREFORE A THOROUGH MEDICAL ASSESSMENT COULD NOT BE PERFORMED; AND WE ARE UNABLE TO CONCLUDE SPECIFIC FACTORS KNOWN TO CONTRIBUTE TO THE ALLEGED REPORT. BASED ON THIS INVESTIGATION, THE NEED FOR CORRECTIVE ACTION IS NOT INDICATED.
EDWARDS ESHEATH+ INTRODUCER | K200258 | PER REPORT, THE PATIENT PASSED AWAY DUE TO A DISTAL AORTA PERFORATION. | A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED TO EDWARDS LIFESCIENCES FOR EVALUATION, HOWEVER, A DEVICE HISTORY RECORD (DHR) REVIEW WAS DONE AND DID NOT REVEAL ANY MANUFACTURING NONCONFORMANCE ISSUES THAT WOULD HAVE CONTRIBUTED TO THE EVENT. A LOT HISTORY REVIEW WAS PERFORMED AND REVEALED COMPLAINTS RELATING TO THE REPORTED EVENT. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY, UNDERSIZED VESSEL) LIKELY CONTRIBUTED TO THE RELATED EVENT. THERE WAS NO DEVICE OR LABELING PROBLEM WAS IDENTIFIED DURING THE EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE IFU, CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THE 3MENSIO IMAGERY WAS PROVIDED BY THE SITE AND REVEALED THE FOLLOWING: PATIENT'S RIGHT ACCESS VESSEL HAD PRESENCE OF TORTUOSITY AND CALCIFICATION. THE REPORTED EVENT OF INABILITY TO ADVANCE THROUGH SHEATH AND LINER PUNCTURED WERE UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF THE DEVICE/APPLICABLE IMAGERY. PER THE TRAINING MANUAL, 'PUSH FORCE CAN VARY DUE TO ANGLE OF ACCESS AND INSERTION, THV SIZE, VESSEL DIAMETER, TORTUOSITY AND DEGREE OF CALCIFICATION'. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. SIMILARLY, TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. IF HIGH INSERTION FORCE WAS USED TO OVERCOME ANY RESISTANCE EXPERIENCED NAVIGATING THE TORTUOUS ANATOMY, IT IS POSSIBLE TO EXACERBATE THE INTERACTION OF THE VALVE AND THE LINER AT SUBOPTIMAL ANGLES, CONTRIBUTING TO THE REPORTED LINER PUNCTURE. THE PROCEDURAL TRAINING MANUAL STATES, 'IF INSERTION FORCE IS HIGH, CONSIDER SLIGHTLY PULLING BACK THE SHEATH WHILE ADVANCING THE THV/DELIVERY SYSTEM 1-2 CM'. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY) AND PROCEDURAL FACTORS (HIGH INSERTION FORCE) MAY HAVE CONTRIBUTED TO THE COMPLAINT EVENT. HOWEVER, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PNEUMONIA, HYPERCARBIA, RESPIRATORY ARREST, RESPIRATORY TRACT IRRITATION AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
INTELLIVUE MX450 PATIENT MONITOR | K130849 | IT WAS REPORTED THE CUSTOMER CALLED TO REQUEST ASSISTANCE GETTING LOGS FROM THE BEDSIDE MONITOR. THE CUSTOMER HUNG UP PRIOR TO PROVIDING MORE INFORMATION AND ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE UNSUCCESSFUL. THE KEY MARKET HAS CONFIRMED THE DEVICE WAS NOT RELATED TO THE PATIENT EXPIRATION. | BOX B : UPDATED INFORMATION, CHANGED OUTCOMES ATTRIBUTED TO AE, FROM SERIOUS INJURY TO DEATH. BOX B: UPDATED INFORMATION: CHANGED EVENT DATE, FROM 07/12/2024 TO 07/08/2024.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED. ON THE PREVIOUSLY SUBMITTED REPORTED, THE WRONG TYPE OF REPORTED COMPLAINT WAS SELECTED IN SECTION H. IT IS CORRECTED ON THIS REPORT.
SENSATION PLUS 8FR. 50CC IAB & ACCESSORIES (W/PRESSURE TUBES, NO STYLET) | K133074 | IT WAS REPORTED THAT DURING INTRA-AORTIC BALLOON (IAB) THERAPY, THE IAB WOULD NOT FULLY INFLATE ON A CS300 INTRA-AORTIC BALLOON PUMP (IABP). THE FACILITY SWAPPED OUT THE CS300 FOR A CARDIOSAVE IABP, AND TRIED TO USE ANOTHER IAB, BUT THE BALLOON DID NOT INFLATE. THEY WERE NEVER ABLE TO INITIATE THERAPY. THE PATIENT ULTIMATELY PASSED AWAY. THIS REPORT IS FOR THE 1ST IAB. A SEPARATE REPORT WILL BE SUBMITTED FOR THE 2ND IAB. THIS REPORT IS FOR THE IAB. TWO SEPARATE REPORTS FOR THE CS300 IABP & CARDIOSAVE IABP WERE SUBMITTED UNDER MFG REPORT NUMBER 2249723-2024-01188 & MFG REPORT NUMBER 2249723-2024-01205. | UDI # IS INCO UDI RELATED DATA QUALITY UPDATES ONLY. PROVIDING UPDATED DEVICE IDENTIFICATION INFORMATION IN ALIGNMENT WITH GUDID (D1, D2, D3, D4, G4, H4, H5). UDI # IS INCOMPLETE AS THE CATALOG#, LOT #, MANUFACTURE DATE, EXP. DATE WERE NOT PROVIDED. THIS INFORMATION WAS REQUESTED FROM THE CUSTOMER; HOWEVER, DESPITE OUR BEST EFFORTS, NO INFORMATION HAS BEEN RECEIVED. IF ANY PERTINENT INFORMATION IS RECEIVED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. REFERENCE COMPLAINT #(B)(4).
EMBOTRAP III 6.5 MM X 45 MM | K211338 | IT WAS REPORTED, VIA A PERSONAL INTERACTION, THAT AN EMBOTRAP III 6.5 MM X 45 MM (ET309645/ 24C158AV) AND AN UNKNOWN CEREGLIDE71 INTERMEDIATE CATHETER (PRODUCT CODE/LOT # UNKNOWN) WERE USED FOR A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT AN OCCLUSION AT THE M1 SEGMENT OF THE LEFT MIDDLE CEREBRAL ARTERY (MCA), WHICH WAS A CARDIOGENIC CEREBRAL EMBOLISM. DURING THE PROCEDURE, THE USER REPORTED WITHDRAWAL DIFFICULTY OF THE EMBOTRAP III INTO THE INTERMEDIATE CATHETER. RECANALIZATION WAS ACHIEVED BY 2 PASSES. NO BLEEDING WAS OBSERVED ON POST-RETRIEVAL IMAGING; HOWEVER, POST-PROCEDURAL MRI IMAGING SHOWED A SMALL AMOUNT OF HEMORRHAGE NEAR BASAL GANGLIA. MEDICATION WAS PROVIDED FOR MANAGING BLOOD PRESSURE. THE NEXT DAY, MRI FOUND ENLARGEMENT OF HEMORRHAGE, AND THE PATIENT SUBSEQUENTLY DIED IN A FEW DAYS. THE EVENT WAS REPORTED AS SUCH, ¿THE PROCEDURE WAS A MECHANICAL THROMBECTOMY OF LEFT MIDDLE CEREBRAL ARTERY M1 OCCLUSION FOR CARDIOGENIC CEREBRAL EMBOLISM. STENOSIS WAS OBSERVED PRIOR TO C3 CURVE. THE EMBOTRAP III AND RED72 WERE ATTEMPTED TO ADVANCE TO THE LESION, BUT THE RED72 COULD NOT BE ADVANCED, SO REPLACED WITH THE CEREGLIDE71. THE EMBOTRAP III WAS MANAGED TO ADVANCE NEAR C3 STENOSED LESION (COULD NOT ADVANCE ANY FURTHER THAN C3), THE STENT WAS DEPLOYED AND RETRIEVED THE THROMBUS. ALTHOUGH THERE WAS RESISTANCE DURING RETRIEVAL, THE STENT WAS RETRIEVED AS IT WAS BECAUSE THE MICROCATHETER HAD BEEN REMOVED USING SAVE METHOD AND COULD NOT RE-SHEATH THE STENT WITH THE MICROCATHETER. RECANALIZATION WAS ACHIEVED BY 2 PASSES. NO BLEEDING WAS OBSERVED ON POST-RETRIEVAL IMAGING; HOWEVER, POST-PROCEDURAL MRI SHOWED A SMALL AMOUNT OF HEMORRHAGE NEAR BASAL GANGLIA. MEDICATION WAS PROVIDED FOR MANAGING BLOOD PRESSURE. THE NEXT DAY, MRI FOUND ENLARGEMENT OF HEMORRHAGE, AND THE PATIENT SUBSEQUENTLY DIED IN A FEW DAYS. THE PATIENT DIED A FEW DAYS AFTER THE PROCEDURE. PROGRESS: THE PATIENT DIED A FEW DAYS AFTER THE PROCEDURE DUE TO BLEEDING FROM STENT WITHDRAWAL OR HEMORRHAGIC INFARCTION (IT WAS CONSIDERED BLOOD FLOW RETURNED AND BLOOD VESSEL COULD NOT WITHSTAND BLOOD PRESSURE, AND THE BLEEDING AREA WAS ENLARGED). THE PHYSICIAN ASSESSED THE EVENT NON-SERIOUS (MODERATE/MINOR), THOUGH THE DEATH DUE TO POSTOPERATIVE HEMORRHAGE WAS SERIOUS. THE PHYSICIAN DETERMINED THAT THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE EVENT AND THE PRODUCT. THE PHYSICIAN¿S COMMENT ON THE CAUSAL RELATIONSHIP BETWEEN THE EVENT AND THE PRODUCT: HE FELT RESISTANCE WHEN RETRIEVING THE STENT DUE TO THE STENOSIS OF C3. AT THAT TIME, THE MICROCATHETER HAD BEEN REMOVED USING SAVE METHOD EVEN THOUGH HE KNEW THAT RE-SHEATH OF THE STENT WAS OPTIMAL, AND ALSO THE ASPIRATION CATHETER COULD NOT CROSS THE STENOSIS AT C3, SO HE RETRIEVED THE STENT AS IT WAS DESPITE FEELING RESISTANCE. FINALLY, RECANALIZATION WAS ACHIEVED BY 2 PASSES AND THE THROMBUS COULD BE RETRIEVED, BUT THE BLOOD FLOW RETURNED, AND THE VESSEL COULD NOT WITHSTAND THE BLOOD PRESSURE, WHICH MOST LIKELY CAUSED A HEMORRHAGIC INFARCTION. HE BELIEVED THE ABOVE WAS THE MOST LIKELY CAUSE, SINCE THE POST-RETRIEVAL IMAGING AND POST-PROCEDURAL MRI DID NOT SHOW MASSIVE BLEEDING. THE DEVICES WERE USED ACCORDING TO IFU. MEDICAL HISTORY/TREATMENT HISTORY/OTHER DISEASES CURRENTLY BEING TREATED: HYPERTROPHIC CARDIOMYOPATHY, HYPERTENSION, ATRIAL FIBRILLATION. PAST SURGICAL HISTORY: TWO SURGERIES AT THE SAME HOSPITAL FOR CHRONIC SUBDURAL HEMATOMA. ALLERGIES, SMOKING, ETC.: UNKNOWN. CONTINUOUS FLUSH WAS UNKNOWN. OTHER CONCOMITANT DEVICE WAS TRAK MICROCATHETER.¿ NO FURTHER INFORMATION WAS MADE AVAILABLE AT THE TIME. | PRODUCT COMPLAINT # (B)(4). INFORMATION REGARDING PATIENT WEIGHT, HEIGHT, RACE, AND ETHNICITY WAS NOT REPORTED. SECTION B2 ¿ DATE OF DEATH: THE DATE OF DEATH WAS NOT REPORTED. SECTION E1. INITIAL REPORTER PHONE: (B)(6). THE DEVICE WAS DISCARDED; THEREFORE, NO FURTHER INVESTIGATION CAN BE PERFORMED. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH LOT 24C158AV PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESS THAT CAN BE RELATED TO THE REPORTED COMPLAINT. DIFFICULTY WITHDRAWING THE EMBOTRAP III DEVICE INTO THE INTERMEDIATE/GUIDE CATHETER AFTER DEPLOYMENT CAN RESULT IN THE UNINTENDED LOSS OF ACCESS TO THE OCCLUDED TREATMENT SITE REQUIRING RE-ACCESS WITH INCREASED RISK FOR VASOSPASM, EMBOLI, VESSEL DAMAGE, ISCHEMIA, OR INFARCT, AND/OR THE NEED FOR ADDITIONAL INTERVENTION. THE TREATING PHYSICIAN FELT THE RESISTANCE WAS DUE TO THE STENOSIS OF THE VESSEL. IT IS IMPORTANT TO NOTE, THE INSTRUCTIONS FOR USE (IFU) FOR THE EMBOTRAP III STATES USERS SHOULD NOT WITHDRAW THE DEVICE AGAINST SIGNIFICANT RESISTANCE; SHOULD THERE BE SIGNIFICANT RESISTANCE UPON WITHDRAWAL, THE USER SHOULD ASSESS THE CAUSE OF RESISTANCE USING FLUOROSCOPY AND, IF NECESSARY, ADVANCE THE MICROCATHETER OVER THE DEVICE TO RE-SHEATH OR PARTIALLY RE-SHEATH TO AID WITHDRAWAL. IN THIS CASE, THE EVENT RESULTED IN A ¿SMALL AMOUNT OF HEMORRHAGE NEAR BASAL GANGLIA,¿ WHICH WAS WORSENED BY HYPERTENSION/HEMORRHAGIC TRANSFORMATION AND THE PATIENT ULTIMATELY EXPIRED A FEW DAYS LATER. THE PHYSICIAN DETERMINED THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE EVENT AND THE PRODUCT. THEREFORE, THIS EVENT DOES MEET US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED TO IDENTIFY STATISTICAL SIGNALS FOR CONSIDERATION OF FURTHER CORRECTION ACTION. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. MISSING INFORMATION FROM THIS REPORT IS IDENTIFIED AS BLANK; THIS INFORMATION WAS NOT PROVIDED IN THE REPORTED EVENT OR AVAILABLE AT THE TIME OF REPORT SUBMISSION. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803 (AND/OR PART 4, AS APPLICABLE). THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY CERENOVUS, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, CERENOVUS, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K150709 | IT WAS REPORTED THAT THROMBOSIS AND DEATH OCCURRED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT ACCESS SOLUTION KIT WAS SELECTED FOR USE. PRE-PROCEDURE LEFT ATRIAL APPENDAGE (LAA) THROMBUS CHECK WAS PERFORMED AND CLEARED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE); PASS ASSESSMENT FOR SEAL WAS COMPLETE ALL AROUND THE DEVICE, NO PERI-DEVICE LEAK (PDL) NOTED. HEPARIN WAS ADMINISTERED TO PATIENT PRIOR TO TRANSSEPTAL CROSSING. FIVE TO SIX MINUTES BETWEEN HEPARIN BEING GIVEN AND FIRST ACT DRAWN; ACT RESULT 265. THE TRANSSEPTAL CROSSING WAS SAFELY PERFORMED, WITH NO ISSUES REPORTED. A NON-BOSTON SCIENTIFIC PIGTAIL CATHETER WAS THEN INSERTED INTO THE WATCHMAN SHEATH AND WAS USED TO NAVIGATE FROM LEFT SUPERIOR PULMONARY VEIN (LSPV) INTO LAA. CONTRAST INJECTION(S) WERE PERFORMED TO ASSESS LAA MORPHOLOGY AND SIZING FOR CLOSURE DEVICE. THE PIGTAIL WAS WITHDRAWN FROM THE LAA THROUGH THE SHEATH AND THROMBUS WAS NOTED ON THE DISTAL END OF THE PIGTAIL, PROXIMAL TO THE CURL. THE WATCHMAN SHEATH WAS ASPIRATED PRIOR TO INSERTING 24MM WATCHMAN DEVICE. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) WAS PERFORMED AS A SWEEP OF THE LAA AND AORTIC VALVE TO ASSESS FOR THROMBUS. THERE WAS A POSSIBLE THROMBUS DENOTED AT THE AORTIC VALVE. THE PHYSICIAN SUSPECTED THAT IT WAS POSSIBLE THAT PORTION(S) OF THE THROMBUS WERE SHEARED OFF OF THE PIGTAIL INTO THE LEFT ATRIUM WHEN PULLED INTO SHEATH. THE PHYSICIAN CONTINUED THE PROCEDURE ADVANCING THE 24MM WATCHMAN DEVICE INTO THE LAA AND DEPLOYED ONCE, NO RECAPTURE. POSITION, ANCHOR, SIZE AND SEAL (PASS) CRITERIA WAS CHECKED AND MET REQUIREMENTS. THE WATCHMAN CLOSURE DEVICE WAS RELEASED, AND THE WATCHMAN SHEATH WITHDRAWN FROM BODY. THE PHYSICIAN USED PERCLOSE TO CLOSE THE GROIN ACCESS, AND ANESTHESIA BEGAN POST-PROCEDURE PROCESS OF REVERSING AND WAKING PATIENT UP. THE PATIENT WAS SENT TO THE POST-ANESTHESIA CARE UNIT (PACU) FOR RECOVERY. WITHIN ONE HOUR POST PROCEDURE THE PATIENT UNDERWENT SEVER BRADYCARDIA AND CODED IN THE PACU AND UNKNOWN INTERVENTION WAS PERFORMED BUT THE PATIENT DIED. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). THE VERSACROSS RF WIRE WAS YET INSIDE THE PATIENT BODY WHEN THE THROMBUS WAS NOTED. NO VERSACROSS DEVICE MALFUNCTIONS WERE REPORTED. IT IS UNSURE IF THE PATIENT HAD A STROKE OR CARDIAC ARREST. IN THE PHYSICIAN OPINION, THE VERSACROSS DEVICE DID NOT IMPACT DIRECTLY TO THE PATIENT COMPLICATION OR DEATH, BUT IT IS SUSPECTED A THROMBOEMBOLIC EVENT AS THE CAUSE OF DEATH. THE CAUSE OF DEATH IS POSSIBLY STROKE OR CARDIAC ARREST. | CORRECTION TO THE INITIAL MDR: THE FIELDS B2 (OUTCOMES ATTRIB TO ADV EVENT), F10 AND H6 (IMPACT CODES (3)) WERE UPDATED, THUS THIS SUPPLEMENTAL MDR IS BEING FILED. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON A REPLACEMENT CPAP DEVICE ISSUED TO A PATIENT THAT IS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
SENSATION PLUS 8FR. 50CC IAB & ACCESSORIES (W/PRESSURE TUBES, NO STYLET) | K133074 | IT WAS REPORTED THAT DURING INTRA-AORTIC BALLOON (IAB) THERAPY, THE IAB WOULD NOT FULLY INFLATE ON A CS300 INTRA-AORTIC BALLOON PUMP (IABP). THE FACILITY SWAPPED OUT THE CS300 FOR A CARDIOSAVE IABP, AND TRIED TO USE ANOTHER IAB, BUT THE BALLOON DID NOT INFLATE. THEY WERE NEVER ABLE TO INITIATE THERAPY. THE PATIENT ULTIMATELY PASSED AWAY. THIS REPORT IS FOR THE 2ND IAB. A SEPARATE REPORT HAS BEEN SUBMITTED FOR THE 1ST IAB UNDER MFG REPORT NUMBER 2248146-2024-00475. THIS REPORT IS FOR THE IAB. TWO SEPARATE REPORTS FOR THE CS300 IABP & CARDIOSAVE IABP WERE SUBMITTED UNDER MFG REPORT NUMBER 2249723-2024-01188 & MFG REPORT NUMBER 2249723-2024-01205. | **UDI RELATED DATA QUALITY UPDATES ONLY** PROVIDING UPDATED DEVICE IDENTIFICATION INFORMATION IN ALIGNMENT WITH GUDID (D1, D2, D3, D4, G4, H4, H5). UDI # IS INCOMPLETE AS THE CATALOG#, LOT #, MANUFACTURE DATE, EXP. DATE WERE NOT PROVIDED. THIS INFORMATION WAS REQUESTED FROM THE CUSTOMER; HOWEVER, DESPITE OUR BEST EFFORTS, NO INFORMATION HAS BEEN RECEIVED. IF ANY PERTINENT INFORMATION IS RECEIVED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. COMPLAINT RECORD ID # (B)(4).
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REP DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEXX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | UPON REVIEW OF THIS COMPLAINT, THERE WAS NO SERIOUS INJURY ALLEGED. IN ADDITION, THE MALFUNCTION WILL NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY IF THE EVENT WERE TO REOCCUR. THIS COMPLAINT DOES NOT MEET THE DEFINITION OF A REPORTABLE EVENT.
GIRAFFE OMNIBED | K101788 | THE HOSPITAL REPORTED A PATIENT THAT USED A GIRAFFE OMNIBED DEVELOPED A SKIN INFECTION. THE SUSCEPTIBLE PATIENT WAS EXPOSED TO A FUNGUS (LICHTHEIMIA RAMOSA) FROM THE ZYGOMYCETE FAMILY WHICH WAS THE SOURCE OF INFECTION. REPORTEDLY, THE PATIENT INFECTION WAS TREATED WITH AMPHOTERICIN B. IT WAS REPORTED THE PATIENT DIED. GE HEALTHCARE WILL SUBMIT A FOLLOW-UP REPORT WHEN THE INVESTIGATION HAS BEEN COMPLETED. | THE END USER CONFIRMED THAT EXTERNAL TESTING OF THE INFECTION IDENTIFIED THE SAME STRAIN (LICHTHEIMIA RAMOSA FROM THE ZYGOMYCETE FAMILY ) AS THEIR IN-HOUSE TESTING. THE EXTERNAL TESTS ALSO SHOW THAT THE GE HEALTHCARE GIRAFFE OMNIBED RESULTS WERE ALL NEGATIVE AND THUS UNRELATED TO THE ZYGOMYCETE INFECTION. BASED ON THE INFORMATION AVAILABLE AT THIS TIME, GE HEALTHCARES (GEHC) REVIEW OF THE EVENT DETERMINED THAT THE INFECTION WAS NOT CAUSED OR RELATED TO THE USE OF A GEHC DEVICE. THE GIRAFFE OMNI BED DID NOT CAUSE OR CONTRIBUTE TO THE INFECTION.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DREAMSTATION AUTO BIPAP THAT THE PATIENT ALLEGING RESPIRATORY DISTRESS, RESPIRATORY FAILURE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGES RESPIRATORY DISTRESS, RESPIRATORY FAILURE AND DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING THE INVESTIGATION, THE MANUFACTURER OBSERVED THE CONTROL KNOB WAS MISSING. UNKNOWN BLACK CONTAMINATION, INCONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM, IN THE AIR INLET OF THE BLOWER BOX AND ON THE UI (USER INTERFACE) PANEL. UNKNOWN BROWN RESIDUE ON THE SD (SECURE DIGITAL) CARD FLIP DOOR. PHILIPS INVESTIGATION LABORATORY WAS NOT ABLE TO GET CARE ORCHESTRATOR INFORMATION FROM THE DEVICE. POTENTIAL DRIED LIQUID SPOTS OR MINERAL DEPOSITS ON THE OUTSIDE AND INSIDE OF THE BLOWER MOTOR, INDICATING POTENTIAL LIQUID INGRESS. BLACK CONTAMINATION, CONSISTENT WITH KERATIN, AT THE AIR OUTLET AND RUBBER GROMMET RETAINMENT SLOT OF THE BLOWER BOX. HAIR-LIKE FIBER ON THE BOTTOM ENCLOSURE. DUST-LIKE CONTAMINATION THROUGHOUT THE DEVICE. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. 0 ERRORS WERE LOGGED. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW PHILIPS INVESTIGATION LABORATORY USED A FITT (FOAM INTEGRITY TEST TOOL) AND THE ASSOCIATED PROCEDURE AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE MANUFACTURER CONCLUDES THAT THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGESTS THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE AND WAS UNABLE TO ADDRESS THE PATIENT'S SYMPTOMS. IN BOX D: DEVICE AVAILABLE FOR EVAL? AND DATE RETURNED TO MFG HAS BEEN UPDATED. IN BOX H: DEVICE EVAL. BY MFG?, (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER AND LUNG DISEASE. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
INVACARE® PLATINUM® OXYGEN CONCENTRATOR | K020386 | THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THE FOLLOWING PATIENT EVENT. THIS PATIENT EVENT HAD BEEN REPORTED TO INVACARE BY A DISTRIBUTOR IN THE UK. ¿THE INFORMATION WE HAVE AT THE TIME OF THIS REPORT IS VERY LIMITED. WE HAVE BEEN INFORMED THAT AN END USER IS DECEASED AND WE HAVE BEEN INFORMED BY (B)(6) POLICE THAT THEY HAVE A CONCENTRATOR THAT THEY HAVE REMOVED FROM A DECEASED PATIENT'S HOME ON THE ISLE OF WIGHT AND WOULD LIKE IT TESTED PLEASE. WE HAVE ALSO RECEIVED A REQUEST FOR INVACARE TO PROVIDE A TEST OF THE PLATINUM 9 SERIAL NO:(B)(6) AND HOMEFILL COMPRESSOR SERIAL NO: (B)(6). THIS IS TO BE VIDEOED. THE POLICE HAVE NOT PROVIDED FURTHER INFORMATION ON THE CIRCUMSTANCES SURROUNDING THE DEATH. WE HAVE NO INDICATION THAT THE INCIDENT HAS BEEN REPORTED TO THE COMPETENT AUTHORITIES.¿ INVACARE ALSO ADVISED VENTEC THAT THE PATIENT¿S PRESCRIPTION WAS 2 LPM 16-20HPD. THE EXACT DATE OF THE PATIENT¿S DEATH WAS NOT REPORTED. AS A RESULT, SECTION B2, DATE OF DEATH, WAS LEFT BLANK. NO FURTHER DETAILS ABOUT THE PATIENT OR THE EVENT WERE PROVIDED. | H6: THE PREVIOUS DEVICE MANUFACTURER, INVACARE, INVESTIGATED THE REPORTED EVENT. INVACARE WAS INFORMED BY THE HOME HEALTHCARE SERVICE PROVIDER, DOLBY VIVISOL, THAT THE PATIENT SUFFERED FROM COPD. SECTION B7 HAS BEEN UPDATED ACCORDINGLY. INVACARE ALSO RECEIVED THE FOLLOWING ADDITIONAL INFORMATION FROM THE POLICE: "TO CONFIRM THE CONCENTRATOR WAS SEIZED FROM THE HOME ADDRESS OF PATIENT QA12132 AFTER CARERS FOUND THE HOME OWNER AND USER OF THE CONCENTRATOR SADLY PASSED AWAY. THE CONCENTRATOR WAS FOUND WITHIN THE SPARE BEDROOM OF THE PATIENTS FIRST FLOOR FLAT CLOSEST TO THE WINDOW BUT NOT TOO CLOSE TO THE WALL. THE ROOM WAS AVERAGE TEMPERATURE AND UNCLUTTERED. THE CONCENTRATOR WAS PLUGGED IN WITH THE SWITCHES ON HOWEVER WAS SWITCHED OFF ON TOP AND LOWER SWITCHES. NO ALARMS WERE SOUNDING HOWEVER OF NOTE I UNDERSTAND THAT PRIOR TO THE PATIENTS PASSING A PHONE CALL WAS MADE BY HER FRIEND ASKING FOR ASSISTANCE WITH THE CONCENTRATOR (I DO NOT KNOW IF THE ALARM WAS SOUNDING AT THIS STAGE). I BELIEVE THERE IS A RECORDING OF THIS PHONE CALL WHICH WAS SENT TO US BY [REDACTED] WHICH MAY ASSIST YOU. THERE WAS NO CYLINDER ON THE CONCENTRATOR. TO CONFIRM THE CONCENTRATOR WAS REMOVED FROM THE SAME ROOM IT WAS FOUND IN AND EXHIBITED AS 2 ITEMS WRAPPED IN PLASTIC SHEETING. I UNPLUGGED THE MACHINE AND REMOVED THE PLASTIC TUBING FROM IT (OF NOTE ON REMOVING THE TUBING A SCREW BECAME LOOSE WHICH I SECURED WITH AN EVIDENCE BAG). NO OTHER AREAS WERE HANDLED ON THE MACHINE. THE TUBING RAN FROM THE CONCENTRATOR IN THE SPARE ROOM ALONG THE HALLWAY AND THROUGH INTO THE PATIENT'S BEDROOM AND WAS FOUND ON THE FLOOR." ADDITIONAL INFORMATION WAS ALSO RECEIVED FROM DOLBY VIVISOL: "THE PATIENT RELATIVE REPORTED ON A CALL THE PATIENT SAYING THEY COULD NOT FEEL THE OXYGEN THROUGH THE TUBING. VIVISOL ADVISED THEM TO CALL 111 BUT THEY REFUSED. AN URGENT JOB WAS RAISED AND OUR TECHNICIAN CAME OUT AND CHECKED OVER THE TUBING AND CONCENTRATOR. HE CONFIRMED HE COMPLETED CHECKS ON THE MACHINE CHECKING THE PRESSURE, CONCENTRATION AND ALL ALARMS WERE WORKING .NO ISSUES WERE FOUND. HE THEN CHECKED ALL TUBING TO ENSURE NO SPLITS IN THE TUBING, AGAIN NO ISSUES WERE FOUND. WE HAVE HAD THE TUBING BACK FROM THE POLICE AND COMPLETED TESTS ON IT. NO NON CONFORMANCES IDENTIFIED WITH IT." INVACARE'S INVESTIGATION RESULTS ARE AS FOLLOWS: ¿THE AFFECTED PLATINUM 9 OXYGEN CONCENTRATOR (SN: (B)(6)) AND THE HOMEFILL COMPRESSOR (SN: (B)(6)) WERE RETURNED AND INVESTIGATED AT THE INVACARE FACILITY IN PENCOED, UK. IT WAS NOT POSSIBLE TO TEST THE HOMEFILL COMPRESSOR DUE TO A BROKEN CONNECTION TUBE. THE AFFECTED PLATINUM 9 WAS INSPECTED AND FUNCTIONALLY TESTED. ON INITIAL INSPECTION THE HOMEFILL CURLY CONNECTOR WAS PLACED ON TOP OF THE CONCENTRATOR WITH ONE END BROKEN INSIDE A POLICE EVIDENCE BAG, ON CLOSER INSPECTION IT WAS VISIBLE THAT ONE END OF THE CURLY CONNECTOR WAS BROKEN INSIDE THE HOMEFILL OUTLET PORT. SOME SCRATCHES ON THE OUTSIDE SHROUD BUT NO OTHER DAMAGE WAS VISIBLE. THE UNIT ARRIVED WITH POWER BUTTON SWITCHED ON, NO FAULT CODE FOUND ON THE PCB. THE SHROUD FILTERS WERE FOUND TO BE CLEAN, INTERNAL USED BUT NO ISSUE. THE SNAPPED HOMEFILL PORT WAS REMOVED FROM THE OUTLET THEN THE UNIT WAS SWITCHED ON. THE UNIT FLOW METER WAS SET AT 3LPM ON ARRIVAL AND THE FIRST TEST WAS CONDUCTED AT THIS LEVEL. THE OXYGEN CONCENTRATOR WAS FUNCTIONALLY TESTED AT FLOW RATES FROM 2 TO 9 LPM FOR SEVERAL HOURS ON EACH FLOW RATE. THE OXYGEN CONCENTRATION WAS WITHIN THE SPECIFICATIONS ON ALL TESTED FLOW RATES AND NO FAULTS WERE FOUND. FURTHERMORE, ADDITIONAL TEST WERE PERFORMED ON THE CONCENTRATOR. SNAPPED HOMEFILL PORT WAS RE-INSERTED TO THE OUTLET PORT AND TESTED AT 7.5LPM, THE REASON THIS LEVEL WAS TO DEMONSTRATE THAT THE UNIT HAS A MAXIMUM PRODUCTION CAPACITY OF 9LPM BUT BY HAVING AN OPEN HOMEFILL PORT THIS REQUIRES 2LPM AND THE OUTLET REQUIRES 7.5LPM THE UNIT IS BEING ASKED TO SUPPLY 9.5 LPM. DURING THIS TEST THE UNIT WASN'T ABLE TO PRODUCE THE REQUIRED LEVEL OF OXYGEN PURITY AND WENT INTO A YELLOW (LOW 02) ALARM (APPROX. 35-40MIN). LOW FLOW TEST: CONDUCTED WITH THE SNAPPED HOMEFILL PORT INSERTED INTO THE OUTLET PORT, UNIT SET AT 5LPM AND RESTRICTED THE FLOW FROM THE USER OUTLET (FINGER OVER THE PORT) THE ALARM DIDN'T SOUND DUE TO THE 2LPM LEAVING THE UNIT THROUGH THE HOMEFILL OUTLET PORT THIS PREVENTED ANY PRESSURE BUILD UP. CONDUCTED WITHOUT THE SNAPPED HOMEFILL PORT, UNIT SET AT 5LPM AND RESTRICTED THE FLOW FROM THE USER OUTLET (FINGER OVER THE PORT) THE ALARM SOUNDED AFTER 20SEC TO INDICATE LOW FLOW FROM THE UNIT. NO POWER TEST: UNIT SWITCHED ON AND RUN AT 5LPM, THE MAIN CABLE REMOVED FROM THE UNIT AND A CONTINUOUS ALARM WAS SOUNDED. THE INVESTIGATION SHOWED THAT THE PLATINUM 9 UNIT WORKS CORRECTLY AS INTENDED PER USER MANUAL. HAVING THE DAMAGED COMPONENT FROM THE HOMEFILL CONNECTOR INSERTED IN THE OUTLET PORT OF THE CONCENTRATOR, IT COULD CREATE A SCENARIO WHEREBY THE LOW FLOW WOULD NOT GIVE AN ALARM GIVEN THE RIGHT CIRCUMSTANCES (FLOW METER SET TO 5LPM AND CANNULA/TUBING KINKED OR BLOCKED). IT WAS STATED THAT THERE WAS "NO CYLINDER ON THE CONCENTRATOR" AND IT IS UNKNOWN WHETHER THE HOMEFILL CONNECTOR WAS INSERTED IN THE OUTLET PORT AT THE TIME OF THE INCIDENT. IT IS STATED IN THE USER MANUAL OF THE PLATINUM 9 THAT WHEN THE HOMEFILL COMPRESSOR IS NOT IN USE "THE PLUG PROVIDED WITH THE CONCENTRATOR SHOULD BE INSERTED INTO THE OUTLET FITTING". IF THE CONCENTRATOR WAS CONNECTED TO THE HOMEFILL COMPRESSOR WITH THE BROKEN HOMEFILL CONNECTOR, THE YELLOW INDICATOR LIGHT ON THE HOMEFILL COMPRESSOR WOULD HAVE BEEN SIGNALING THAT IT ENCOUNTERED A PROBLEM. THE YELLOW INDICATOR LIGHT IS DESCRIBED AS FOLLOWS IN THE TROUBLESHOOTING SECTION OF THE HOMEFILL COMPRESSOR: ¿PROBLEM: O2 BELOW NORMAL (YELLOW) LIGHT STAYS ON SOLUTION 1. ENSURE THE CONCENTRATOR HAS WARMED UP FOR AT LEAST 20 MINUTES. IF IT HAS NOT BEEN 20 MINUTES, TURN THE COMPRESSOR OFF UNTIL THE CONCENTRATOR WARMS UP. IF THE YELLOW LIGHT STILL REMAINS ON PROCEED TO STEP 2. 2. INSPECT THE CONNECTION BETWEEN THE COMPRESSOR AND THE CONCENTRATOR FOR DAMAGE OR PINCHING. IF DAMAGED, REPLACE. IF PINCHED, STRAIGHTEN TUBING. TURN THE COMPRESSOR OFF (O) FOR AT LEAST 10 SECONDS. TURN THE COMPRESSOR ON ( I ). IF THE YELLOW LIGHT STILL REMAINS ON AFTER 10 MINUTES, PROCEED TO STEP 3. 3. THE CONCENTRATOR MAY NEED SERVICE, CONTACT YOUR HOME HEALTH CARE PROVIDER OR INVACARE." THE FOLLOWING IS STATED IN THE USER MANUAL: "WARNING! RISK OF INJURY OR DAMAGE TO PREVENT INJURY OR DAMAGE DURING USE: ¿ IF YOU FEEL ILL OR UNCOMFORTABLE, OR IF YOU ARE UNABLE TO FEEL THE OXYGEN FLOW, CONSULT YOUR EQUIPMENT PROVIDER AND/OR YOUR PHYSICIAN IMMEDIATELY." WE RECEIVED THE INFORMATION THAT THE PATIENT HAD A PRESCRIBED FLOW RATE OF 2 LPM FOR 16-20 HOURS PER DAY, WHICH IS A QUITE LOW FLOW RATE. FURTHERMORE, WE WERE INFORMED THAT THE PATIENT SUFFERED FROM COPD. THE CAUSE OF DEATH OF THE PATIENT REMAINS UNKNOWN. THE PLATINUM 9 OXYGEN CONCENTRATOR IS INTENDED FOR INDIVIDUAL USE BY PATIENTS WITH RESPIRATORY DISORDERS WHO REQUIRE SUPPLEMENTAL OXYGEN. THE DEVICE IS TO BE USED AS AN OXYGEN SUPPLEMENT ONLY AND IS NOT INTENDED TO BE LIFE-SUPPORTING OR LIFE-SUSTAINING. THIS IS ALSO STATED IN THE USER MANUAL. THE FOLLOWING SAFETY-RELEVANT INFORMATION CAN BE FOUND IN THE USER MANUAL OF THE PLATINUM 9: "1.2 INTENDED USE: DANGER! RISK OF INJURY OR DEATH. THIS PRODUCT IS TO BE USED AS AN OXYGEN SUPPLEMENT AND IS NOT INTENDED TO BE LIFE-SUPPORTING OR LIFE-SUSTAINING. ONLY USE THIS PRODUCT IF THE PATIENT IS CAPABLE OF SPONTANEOUS BREATH, ABLE TO INHALE AND EXHALE WITHOUT THE USE OF A MACHINE. ¿ DO NOT USE IN PARALLEL OR SERIES WITH OTHER OXYGEN CONCENTRATORS OR OXYGEN THERAPY DEVICES. [...] DANGER! RISK OF INJURY OR DEATH. WHILE INVACARE STRIVES TO PRODUCE THE BEST OXYGEN CONCENTRATOR IN THE MARKET TODAY, THIS OXYGEN CONCENTRATOR CAN FAIL TO PRODUCE OXYGEN DUE TO POWER FAILURE OR DEVICE MALFUNCTION. ¿ ALWAYS HAVE A BACKUP SOURCE OF OXYGEN READILY AVAILABLE" "3 COMPONENTS: 3.1 COMPONENT IDENTIFICATION: THE HOMEFILL OUTLET FITTING I IS TO BE USED ONLY FOR FILLING OXYGEN CYLINDERS WITH THE HOMEFILL HOME OXYGEN COMPRESSOR. THE OUTLET FITTING DOES NOT AFFECT CONCENTRATOR PERFORMANCE. REFER TO THE HOMEFILL SYSTEM USER MANUAL FOR CONNECTION AND OPERATING INSTRUCTIONS. WHEN NOT IN USE, THE PLUG PROVIDED WITH THE CONCENTRATOR SHOULD BE INSERTED INTO THE OUTLET FITTING. FOR MORE INFORMATION ABOUT THE HOMEFILL SYSTEM, CONTACT YOUR INVACARE DEALER." "6.3.2 CONNECTING/POSITIONING THE NASAL CANNULA CAUTION! RISK OF MINOR INJURY OR DISCOMFORT: TO ENSURE PROPER OXYGEN DELIVERY: ¿ ENSURE THE CANNULA PRONGS ARE POSITIONED PROPERLY IN YOUR NOSE. THIS IS CRITICAL TO THE EFFECTIVENESS OF THE OXYGEN THERAPY. ¿ DO NOT USE TUBING/CANNULA LENGTH EXCEEDING 15 M (50 FT). ¿ USE CRUSH-PROOF OXYGEN TUBING. ¿ CHECK FOR GAS FLOW AT THE OUTLET OF THE CANNULA. [...] 6.3.4 FLOWRATE: WARNING! RISK OF MINOR INJURY OR DISCOMFORT: IT IS VERY IMPORTANT TO SELECT THE PRESCRIBED OXYGEN FLOW SETTING. THIS WILL ENSURE YOU WILL RECEIVE THE THERAPEUTIC AMOUNT OF OXYGEN ACCORDING TO YOUR MEDICAL CONDITION. DO NOT INCREASE OR DECREASE THE L/MIN FLOW SETTING UNLESS A CHANGE HAS BEEN PRESCRIBED BY YOUR PHYSICIAN OR THERAPIST. THE THERAPEUTIC EFFECTIVENESS OF THE PRESCRIBED OXYGEN FLOW SETTING SHOULD BE PERIODICALLY REASSESSED. USE ONLY THE TUBING AND ACCESSORIES THAT WERE USED TO DETERMINE THE PRESCRIBED OXYGEN FLOW SETTING. [...] 2. IF THE FLOWRATE ON THE FLOWMETER EVER FALLS BELOW 0.5 L/MIN FOR MORE THAN ABOUT ONE MINUTE, THE POTENTIAL OBSTRUCTION ALERT WILL BE TRIGGERED. THIS IS A RAPID BEEPING OF THE AUDIBLE ALARM. CHECK YOUR TUBING OR ACCESSORIES FOR BLOCKED OR KINKED TUBING OR A DEFECTIVE HUMIDIFIER BOTTLE. AFTER RATED FLOW IS RESTORED TO BETWEEN 0.5 L/MIN AND 0.75 L/MIN, THE POTENTIAL OBSTRUCTION ALERT WILL TURN OFF. A POTENTIAL OBSTRUCTION ALERT INDICATES A CONDITION THAT MAY BE ASSOCIATED WITH A PARTIAL OR COMPLETE OBSTRUCTION OF OXYGEN OUTPUT. THE USE OF SOME ACCESSORIES AND THE HOMEFILL COMPRESSOR WILL DEACTIVATE THE POTENTIAL OBSTRUCTION ALERT." THE FOLLOWING FURTHER SAFETY-RELEVANT INFORMATION CAN BE FOUND IN THE USER MANUAL OF THE HOMEFILL: "2 SAFETY: 2.1 GENERAL GUIDELINES. [...] WARNING! RISK OF DEATH, INJURY OR DAMAGE: IF IMPROPERLY POSITIONED AND SECURED, THE POWER CORD AND INTERCONNECT HOSE CAN CAUSE INJURY DUE TO TRIPPING, FALLING, AND STRANGULATION. PRODUCT DAMAGE MAY ALSO OCCUR. PROPERLY STORE AND POSITION ELECTRICAL CORDS AND/OR TUBING TO PREVENT TRIPPING AND STRANGULATION HAZARDS. AVOID POSITIONING POWER CORDS ACROSS AREAS OF HIGH TRAFFIC. BEFORE MOVING OR REPOSITIONING EITHER THE COMPRESSOR OR CONCENTRATOR, ALWAYS DISCONNECT THE AC POWER CORDS AND THE INTERCONNECT HOSE." "4.3 CONNECTING THE CYLINDER TO THE COMPRESSOR. [...] 1. PRIOR TO EACH USE, INSPECT PRODUCT FOR VISIBLE DAMAGE. REFER TO THE EXTERNAL EXAMINATION SECTION. DO NOT USE IF ANY DAMAGE IS FOUND. IF FOR ANY REASON, ANY LABEL BECOMES ILLEGIBLE OR LOST, CONTACT YOUR EQUIPMENT PROVIDER."" THE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED.
HOMEFILL® OXYGEN SYSTEM | K021685 | THE DEVICE¿S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THE FOLLOWING PATIENT EVENT. THIS PATIENT EVENT HAD BEEN REPORTED TO INVACARE BY A DISTRIBUTOR IN THE UK. ¿THE INFORMATION WE HAVE AT THE TIME OF THIS REPORT IS VERY LIMITED. WE HAVE BEEN INFORMED THAT AN END USER IS DECEASED AND WE HAVE BEEN INFORMED BY (B)(6) POLICE THAT THEY HAVE A CONCENTRATOR THAT THEY HAVE REMOVED FROM A DECEASED PATIENT'S HOME ON THE ISLE OF WIGHT AND WOULD LIKE IT TESTED PLEASE. WE HAVE ALSO RECEIVED A REQUEST FOR INVACARE TO PROVIDE A TEST OF THE PLATINUM 9 SERIAL NO: (B)(6) AND HOMEFILL COMPRESSOR SERIAL NO: (B)(6). THIS IS TO BE VIDEOED. THE POLICE HAVE NOT PROVIDED FURTHER INFORMATION ON THE CIRCUMSTANCES SURROUNDING THE DEATH. WE HAVE NO INDICATION THAT THE INCIDENT HAS BEEN REPORTED TO THE COMPETENT AUTHORITIES.¿ INVACARE ALSO ADVISED VENTEC THAT THE PATIENT¿S PRESCRIPTION WAS 2 LPM 16-20HPD. THE EXACT DATE OF THE PATIENT¿S DEATH WAS NOT REPORTED. AS A RESULT, SECTION B2, DATE OF DEATH, WAS LEFT BLANK. NO FURTHER DETAILS ABOUT THE PATIENT OR THE EVENT WERE PROVIDED. | H6: THE PREVIOUS DEVICE MANUFACTURER, INVACARE, INVESTIGATED THE REPORTED EVENT. INVACARE WAS INFORMED BY THE HOME HEALTHCARE SERVICE PROVIDER, DOLBY VIVISOL, THAT THE PATIENT SUFFERED FROM COPD. SECTION B7 HAS BEEN UPDATED ACCORDINGLY. INVACARE ALSO RECEIVED THE FOLLOWING ADDITIONAL INFORMATION FROM THE POLICE: "TO CONFIRM THE CONCENTRATOR WAS SEIZED FROM THE HOME ADDRESS OF PATIENT (B)(6) AFTER CARERS FOUND THE HOME OWNER AND USER OF THE CONCENTRATOR SADLY PASSED AWAY. THE CONCENTRATOR WAS FOUND WITHIN THE SPARE BEDROOM OF THE PATIENTS FIRST FLOOR FLAT CLOSEST TO THE WINDOW BUT NOT TOO CLOSE TO THE WALL. THE ROOM WAS AVERAGE TEMPERATURE AND UNCLUTTERED. THE CONCENTRATOR WAS PLUGGED IN WITH THE SWITCHES ON HOWEVER WAS SWITCHED OFF ON TOP AND LOWER SWITCHES. NO ALARMS WERE SOUNDING HOWEVER OF NOTE I UNDERSTAND THAT PRIOR TO THE PATIENTS PASSING A PHONE CALL WAS MADE BY HER FRIEND ASKING FOR ASSISTANCE WITH THE CONCENTRATOR (I DO NOT KNOW IF THE ALARM WAS SOUNDING AT THIS STAGE). I BELIEVE THERE IS A RECORDING OF THIS PHONE CALL WHICH WAS SENT TO US BY [REDACTED] WHICH MAY ASSIST YOU. THERE WAS NO CYLINDER ON THE CONCENTRATOR. TO CONFIRM THE CONCENTRATOR WAS REMOVED FROM THE SAME ROOM IT WAS FOUND IN AND EXHIBITED AS 2 ITEMS WRAPPED IN PLASTIC SHEETING. I UNPLUGGED THE MACHINE AND REMOVED THE PLASTIC TUBING FROM IT (OF NOTE ON REMOVING THE TUBING A SCREW BECAME LOOSE WHICH I SECURED WITH AN EVIDENCE BAG). NO OTHER AREAS WERE HANDLED ON THE MACHINE. THE TUBING RAN FROM THE CONCENTRATOR IN THE SPARE ROOM ALONG THE HALLWAY AND THROUGH INTO THE PATIENTS BEDROOM AND WAS FOUND ON THE FLOOR." ADDITIONAL INFORMATION WAS ALSO RECEIVED FROM DOLBY VIVISOL: "THE PATIENT RELATIVE REPORTED ON A CALL THE PATIENT SAYING THEY COULD NOT FEEL THE OXYGEN THROUGH THE TUBING. VIVISOL ADVISED THEM TO CALL 111 BUT THEY REFUSED. AN URGENT JOB WAS RAISED AND OUR TECHNICIAN CAME OUT AND CHECKED OVER THE TUBING AND CONCENTRATOR. HE CONFIRMED HE COMPLETED CHECKS ON THE MACHINE CHECKING THE PRESSURE, CONCENTRATION AND ALL ALARMS WERE WORKING .NO ISSUES WERE FOUND. HE THEN CHECKED ALL TUBING TO ENSURE NO SPLITS IN THE TUBING, AGAIN NO ISSUES WERE FOUND. WE HAVE HAD THE TUBING BACK FROM THE POLICE AND COMPLETED TESTS ON IT. NO NON CONFORMANCES IDENTIFIED WITH IT." INVACARE'S INVESTIGATION RESULTS ARE AS FOLLOWS: ¿THE AFFECTED PLATINUM 9 OXYGEN CONCENTRATOR (SN: (B)(6)) AND THE HOMEFILL COMPRESSOR (SN: (B)(6)) WERE RETURNED AND INVESTIGATED AT THE INVACARE FACILITY IN PENCOED, UK. IT WAS NOT POSSIBLE TO TEST THE HOMEFILL COMPRESSOR DUE TO A BROKEN CONNECTION TUBE. THE AFFECTED PLATINUM 9 WAS INSPECTED AND FUNCTIONALLY TESTED. ON INITIAL INSPECTION THE HOMEFILL CURLY CONNECTOR WAS PLACED ON TOP OF THE CONCENTRATOR WITH ONE END BROKEN INSIDE A POLICE EVIDENCE BAG, ON CLOSER INSPECTION IT WAS VISIBLE THAT ONE END OF THE CURLY CONNECTOR WAS BROKEN INSIDE THE HOMEFILL OUTLET PORT. SOME SCRATCHES ON THE OUTSIDE SHROUD BUT NO OTHER DAMAGE WAS VISIBLE. THE UNIT ARRIVED WITH POWER BUTTON SWITCHED ON, NO FAULT CODE FOUND ON THE PCB. THE SHROUD FILTERS WERE FOUND TO BE CLEAN, INTERNAL USED BUT NO ISSUE. THE SNAPPED HOMEFILL PORT WAS REMOVED FROM THE OUTLET THEN THE UNIT WAS SWITCHED ON. THE UNIT FLOW METER WAS SET AT 3LPM ON ARRIVAL AND THE FIRST TEST WAS CONDUCTED AT THIS LEVEL. THE OXYGEN CONCENTRATOR WAS FUNCTIONALLY TESTED AT FLOW RATES FROM 2 TO 9 LPM FOR SEVERAL HOURS ON EACH FLOW RATE. THE OXYGEN CONCENTRATION WAS WITHIN THE SPECIFICATIONS ON ALL TESTED FLOW RATES AND NO FAULTS WERE FOUND. FURTHERMORE, ADDITIONAL TEST WERE PERFORMED ON THE CONCENTRATOR. SNAPPED HOMEFILL PORT WAS RE-INSERTED TO THE OUTLET PORT AND TESTED AT 7.5LPM, THE REASON THIS LEVEL WAS TO DEMONSTRATE THAT THE UNIT HAS A MAXIMUM PRODUCTION CAPACITY OF 9LPM BUT BY HAVING AN OPEN HOMEFILL PORT THIS REQUIRES 2LPM AND THE OUTLET REQUIRES 7.5LPM THE UNIT IS BEING ASKED TO SUPPLY 9.5 LPM. DURING THIS TEST THE UNIT WASN'T ABLE TO PRODUCE THE REQUIRED LEVEL OF OXYGEN PURITY AND WENT INTO A YELLOW (LOW 02) ALARM (APPROX. 35-40MIN). LOW FLOW TEST: CONDUCTED WITH THE SNAPPED HOMEFILL PORT INSERTED INTO THE OUTLET PORT, UNIT SET AT 5LPM AND RESTRICTED THE FLOW FROM THE USER OUTLET (FINGER OVER THE PORT) THE ALARM DIDN'T SOUND DUE TO THE 2LPM LEAVING THE UNIT THROUGH THE HOMEFILL OUTLET PORT THIS PREVENTED ANY PRESSURE BUILD UP. CONDUCTED WITHOUT THE SNAPPED HOMEFILL PORT, UNIT SET AT 5LPM AND RESTRICTED THE FLOW FROM THE USER OUTLET (FINGER OVER THE PORT) THE ALARM SOUNDED AFTER 20SEC TO INDICATE LOW FLOW FROM THE UNIT. NO POWER TEST: UNIT SWITCHED ON AND RUN AT 5LPM, THE MAIN CABLE REMOVED FROM THE UNIT AND A CONTINUOUS ALARM WAS SOUNDED. THE INVESTIGATION SHOWED THAT THE PLATINUM 9 UNIT WORKS CORRECTLY AS INTENDED PER USER MANUAL. HAVING THE DAMAGED COMPONENT FROM THE HOMEFILL CONNECTOR INSERTED IN THE OUTLET PORT OF THE CONCENTRATOR, IT COULD CREATE A SCENARIO WHEREBY THE LOW FLOW WOULD NOT GIVE AN ALARM GIVEN THE RIGHT CIRCUMSTANCES (FLOW METER SET TO 5LPM AND CANNULA/TUBING KINKED OR BLOCKED). IT WAS STATED THAT THERE WAS "NO CYLINDER ON THE CONCENTRATOR" AND IT IS UNKNOWN WHETHER THE HOMEFILL CONNECTOR WAS INSERTED IN THE OUTLET PORT AT THE TIME OF THE INCIDENT. IT IS STATED IN THE USER MANUAL OF THE PLATINUM 9 THAT WHEN THE HOMEFILL COMPRESSOR IS NOT IN USE "THE PLUG PROVIDED WITH THE CONCENTRATOR SHOULD BE INSERTED INTO THE OUTLET FITTING". IF THE CONCENTRATOR WAS CONNECTED TO THE HOMEFILL COMPRESSOR WITH THE BROKEN HOMEFILL CONNECTOR, THE YELLOW INDICATOR LIGHT ON THE HOMEFILL COMPRESSOR WOULD HAVE BEEN SIGNALING THAT IT ENCOUNTERED A PROBLEM. THE YELLOW INDICATOR LIGHT IS DESCRIBED AS FOLLOWS IN THE TROUBLESHOOTING SECTION OF THE HOMEFILL COMPRESSOR: ¿PROBLEM: O2 BELOW NORMAL (YELLOW) LIGHT STAYS ON SOLUTION 1. ENSURE THE CONCENTRATOR HAS WARMED UP FOR AT LEAST 20 MINUTES. IF IT HAS NOT BEEN 20 MINUTES, TURN THE COMPRESSOR OFF UNTIL THE CONCENTRATOR WARMS UP. IF THE YELLOW LIGHT STILL REMAINS ON PROCEED TO STEP 2. 2. INSPECT THE CONNECTION BETWEEN THE COMPRESSOR AND THE CONCENTRATOR FOR DAMAGE OR PINCHING. IF DAMAGED, REPLACE. IF PINCHED, STRAIGHTEN TUBING. TURN THE COMPRESSOR OFF (O) FOR AT LEAST 10 SECONDS. TURN THE COMPRESSOR ON ( I ). IF THE YELLOW LIGHT STILL REMAINS ON AFTER 10 MINUTES, PROCEED TO STEP 3. 3. THE CONCENTRATOR MAY NEED SERVICE, CONTACT YOUR HOME HEALTH CARE PROVIDER OR INVACARE." THE FOLLOWING IS STATED IN THE USER MANUAL: "WARNING! RISK OF INJURY OR DAMAGE TO PREVENT INJURY OR DAMAGE DURING USE: ¿ IF YOU FEEL ILL OR UNCOMFORTABLE, OR IF YOU ARE UNABLE TO FEEL THE OXYGEN FLOW, CONSULT YOUR EQUIPMENT PROVIDER AND/OR YOUR PHYSICIAN IMMEDIATELY." WE RECEIVED THE INFORMATION THAT THE PATIENT HAD A PRESCRIBED FLOW RATE OF 2 LPM FOR 16-20 HOURS PER DAY, WHICH IS A QUITE LOW FLOW RATE. FURTHERMORE, WE WERE INFORMED THAT THE PATIENT SUFFERED FROM COPD. THE CAUSE OF DEATH OF THE PATIENT REMAINS UNKNOWN. THE PLATINUM 9 OXYGEN CONCENTRATOR IS INTENDED FOR INDIVIDUAL USE BY PATIENTS WITH RESPIRATORY DISORDERS WHO REQUIRE SUPPLEMENTAL OXYGEN. THE DEVICE IS TO BE USED AS AN OXYGEN SUPPLEMENT ONLY AND IS NOT INTENDED TO BE LIFE-SUPPORTING OR LIFE-SUSTAINING. THIS IS ALSO STATED IN THE USER MANUAL. THE FOLLOWING SAFETY-RELEVANT INFORMATION CAN BE FOUND IN THE USER MANUAL OF THE PLATINUM 9: "1.2 INTENDED USE. DANGER! RISK OF INJURY OR DEATH: THIS PRODUCT IS TO BE USED AS AN OXYGEN SUPPLEMENT AND IS NOT INTENDED TO BE LIFE-SUPPORTING OR LIFE-SUSTAINING. ONLY USE THIS PRODUCT IF THE PATIENT IS CAPABLE OF SPONTANEOUS BREATH, ABLE TO INHALE AND EXHALE WITHOUT THE USE OF A MACHINE. ¿ DO NOT USE IN PARALLEL OR SERIES WITH OTHER OXYGEN CONCENTRATORS OR OXYGEN THERAPY DEVICES. [...] DANGER! RISK OF INJURY OR DEATH: WHILE INVACARE STRIVES TO PRODUCE THE BEST OXYGEN CONCENTRATOR IN THE MARKET TODAY, THIS OXYGEN CONCENTRATOR CAN FAIL TO PRODUCE OXYGEN DUE TO POWER FAILURE OR DEVICE MALFUNCTION. ¿ ALWAYS HAVE A BACKUP SOURCE OF OXYGEN READILY AVAILABLE" "3 COMPONENTS: 3.1 COMPONENT IDENTIFICATION: THE HOMEFILL OUTLET FITTING I IS TO BE USED ONLY FOR FILLING OXYGEN CYLINDERS WITH THE HOMEFILL HOME OXYGEN COMPRESSOR. THE OUTLET FITTING DOES NOT AFFECT CONCENTRATOR PERFORMANCE. REFER TO THE HOMEFILL SYSTEM USER MANUAL FOR CONNECTION AND OPERATING INSTRUCTIONS. WHEN NOT IN USE, THE PLUG PROVIDED WITH THE CONCENTRATOR SHOULD BE INSERTED INTO THE OUTLET FITTING. FOR MORE INFORMATION ABOUT THE HOMEFILL SYSTEM, CONTACT YOUR INVACARE DEALER." "6.3.2 CONNECTING/POSITIONING THE NASAL CANNULA: CAUTION! RISK OF MINOR INJURY OR DISCOMFORT: TO ENSURE PROPER OXYGEN DELIVERY: ¿ ENSURE THE CANNULA PRONGS ARE POSITIONED PROPERLY IN YOUR NOSE. THIS IS CRITICAL TO THE EFFECTIVENESS OF THE OXYGEN THERAPY. ¿ DO NOT USE TUBING/CANNULA LENGTH EXCEEDING 15 M (50 FT). ¿ USE CRUSH-PROOF OXYGEN TUBING. ¿ CHECK FOR GAS FLOW AT THE OUTLET OF THE CANNULA. [...] 6.3.4 FLOWRATE: WARNING! RISK OF MINOR INJURY OR DISCOMFORT: IT IS VERY IMPORTANT TO SELECT THE PRESCRIBED OXYGEN FLOW SETTING. THIS WILL ENSURE YOU WILL RECEIVE THE THERAPEUTIC AMOUNT OF OXYGEN ACCORDING TO YOUR MEDICAL CONDITION. DO NOT INCREASE OR DECREASE THE L/MIN FLOW SETTING UNLESS A CHANGE HAS BEEN PRESCRIBED BY YOUR PHYSICIAN OR THERAPIST. THE THERAPEUTIC EFFECTIVENESS OF THE PRESCRIBED OXYGEN FLOW SETTING SHOULD BE PERIODICALLY REASSESSED. USE ONLY THE TUBING AND ACCESSORIES THAT WERE USED TO DETERMINE THE PRESCRIBED OXYGEN FLOW SETTING. [...] 2. IF THE FLOWRATE ON THE FLOWMETER EVER FALLS BELOW 0.5 L/MIN FOR MORE THAN ABOUT ONE MINUTE, THE POTENTIAL OBSTRUCTION ALERT WILL BE TRIGGERED. THIS IS A RAPID BEEPING OF THE AUDIBLE ALARM. CHECK YOUR TUBING OR ACCESSORIES FOR BLOCKED OR KINKED TUBING OR A DEFECTIVE HUMIDIFIER BOTTLE. AFTER RATED FLOW IS RESTORED TO BETWEEN 0.5 L/MIN AND 0.75 L/MIN, THE POTENTIAL OBSTRUCTION ALERT WILL TURN OFF. A POTENTIAL OBSTRUCTION ALERT INDICATES A CONDITION THAT MAY BE ASSOCIATED WITH A PARTIAL OR COMPLETE OBSTRUCTION OF OXYGEN OUTPUT. THE USE OF SOME ACCESSORIES AND THE HOMEFILL COMPRESSOR WILL DEACTIVATE THE POTENTIAL OBSTRUCTION ALERT." THE FOLLOWING FURTHER SAFETY-RELEVANT INFORMATION CAN BE FOUND IN THE USER MANUAL OF THE HOMEFILL: "2 SAFETY, 2.1 GENERAL GUIDELINES: [...] WARNING! RISK OF DEATH, INJURY OR DAMAGE: IF IMPROPERLY POSITIONED AND SECURED, THE POWER CORD AND INTERCONNECT HOSE CAN CAUSE INJURY DUE TO TRIPPING, FALLING, AND STRANGULATION. PRODUCT DAMAGE MAY ALSO OCCUR. PROPERLY STORE AND POSITION ELECTRICAL CORDS AND/OR TUBING TO PREVENT TRIPPING AND STRANGULATION HAZARDS. AVOID POSITIONING POWER CORDS ACROSS AREAS OF HIGH TRAFFIC. BEFORE MOVING OR REPOSITIONING EITHER THE COMPRESSOR OR CONCENTRATOR, ALWAYS DISCONNECT THE AC POWER CORDS AND THE INTERCONNECT HOSE." "4.3 CONNECTING THE CYLINDER TO THE COMPRESSOR. [...] 1. PRIOR TO EACH USE, INSPECT PRODUCT FOR VISIBLE DAMAGE. REFER TO THE EXTERNAL EXAMINATION SECTION. DO NOT USE IF ANY DAMAGE IS FOUND. IF FOR ANY REASON, ANY LABEL BECOMES ILLEGIBLE OR LOST, CONTACT YOUR EQUIPMENT PROVIDER."" THE CAUSE FOR THE REPORTED ISSUE COULD NOT BE DETERMINED.
SHILEY | K090352 | ACCORDING TO THE REPORTER, DURING USE AND WHEN TESTING THE PRESSURE OF THE ENDOTRACHEAL TUBE, THE PRESSURE WAS INSUFFICIENT, INDICATING A LEAK. THE AIR IN THE CUFF WAS REPLENISHED, AND THE PRESSURE WAS CORRECTED. AFTER THE PATIENT¿S TUBE BECAME DISLODGED, A TEST SHOWED THAT THE TRACHEAL TUBE CUFF CONTAINED ONLY 7 C.C. OF AIR. AN HOUR AND 25 MINUTES LATER, THE PATIENT EXPERIENCED CONFUSION, INCOHERENT SPEECH, AND SHOUTING. FOLLOW-UP ARTERIAL BLOOD GAS ANALYSIS SHOWED: PH 7.260, PCO2 50.70, PO2 75.30, HCO3 22.20, BE -4.80, O2 SATURATION 90.40%. THE ON-DUTY PHYSICIAN WAS INFORMED AND, AFTER EVALUATION, A NEW TRACHEAL TUBE WAS PLACED. INFORMATION WAS RECEIVED THAT THE PATIENT DIED AND THE CAUSE OF DEATH WAS PNEUMONIA WITH ACUTE RESPIRATORY FAILURE. | ADDITIONAL INFORMATION: G3, H3, H6 H3 EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. THE DEVICE WAS NOT RETURNED, BUT A PHOTO WAS AVAILABLE FOR EVALUATION. PICTURE EVALUATION SHOWED AN ENDOTRACHEAL TUBE, THE CUFF APPEARED TO BE DEFLATED. IT WAS REPORTED THAT DURING USE, THE ENDOTRACHEAL TUBE'S CUFF HAD AN AIR LEAK AFTER 12 DAYS. THE REPORTED ISSUE WAS CONFIRMED. THE MOST LIKELY CAUSE COULD NOT BE ESTABLISHED FROM THE INFORMATION AVAILABLE. THE MANUFACTURING RECORDS FOR EACH DEVICE ARE THOROUGHLY REVIEWED PRIOR TO RELEASE TO ENSURE THAT IT MEETS ALL MEDTRONIC QUALITY SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO LUNG DISEASE AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO LUNG DISEASE AND CANCER. PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED WITH WRONG TYPE OF REPORTED COMPLAINT AS SERIOUS INJURY. IN THIS REPORT, THE TYPE OF REPORTED COMPLAINT HAS BEEN UPDATED CORRECTLY AS DEATH. SECTION TYPE OF REPORTED COMPLAINT IN BOX H HAS BEEN UPDATED/ CORRECTED.
PRLNE BLU 36IN 4-0 D/A SH | K133356 | IT WAS REPORTED THAT A PATIENT UNDERWENT AN ELECTIVE CARDIAC PROCEDURE ON (B)(6) 2024 AND SUTURE WAS USED TO CLOSE THE INCISION. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. SIX HOURS POST-OP, THE PATIENT¿S INCISION REOPENED AND THE PATIENT EXSANGUINATED INSTANTANEOUSLY, RESULTING IN IMMEDIATE PATIENT DEATH. ADDITIONAL INFORMATION WAS REQUESTED. | PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: NAME OF INDEX SURGICAL PROCEDURE? THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? MITRAL STENOSIS AND INSUFFICIENCY. WHAT WAS THE INITIAL APPROACH FOR THE INDEX SURGICAL PROCEDURE? (OPEN, LAPAROSCOPIC OR OTHER)? OPEN. WHAT INSTRUMENTS WERE USED IN THIS PROCEDURE TO HANDLE THE SUTURE? CASTRO NEEDLE DRIVER. ON WHAT TISSUE WAS THE SUTURE USED/LOCATION OF SUTURE PLACEMENT FOR EACH SUTURE? LEFT ATRIUM. WHAT TISSUE DEHISCED? LEFT ATRIUM. WHAT WAS THE TISSUE CONDITION (NORMAL, THIN, CALCIFIED, FRAGILE, DISEASED)? NORMAL. HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? CONTINUOUS. HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? 5 KNOTS. HOW WAS THE DEHISCENCE MANAGED? PATIENT WAS OPENED EMERGENTLY AT THE BEDSIDE IN THE POSTOP ICU WHAT WAS THE SOURCE AND TRIGGERING EVENT OF BLEEDING? UNKNOWN. SUTURE WAS FOUND BROKEN WHAT WAS THE VOLUME OF BLOOD LOSS? MASSIVE 6+L. HOW WAS THE BLEEDING TREATED? OPENED THE PATIENT AND BLOOD GIVEN WAS ANY MEDICAL/SURGICAL INTERVENTION PERFORMED POST-OP WHEN THE INCISION RE-OPENED TO TRY AND STOP THE BLEEDING? IF YES, PLEASE DESCRIBE. ALL THE BLOOD HAD EXITED THE PATIENT. THIS WAS A MASSIVE BLOOD LOSS WERE THERE ANY PATIENT STRESS FACTORS THAT PRECIPITATED THE EVENT OF THE INCISION REOPENING AND BLEEDING? NO DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? NOT IN THIS CASE OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? NONE WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? SUTURE FRACTURE HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? AUTOPSY NOT NEEDED SINCE WE OPENED THE PATIENT WOULD THE PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? YES.
PRLNE BLU 36IN 4-0 D/A SH | K133356 | IT WAS REPORTED THAT A PATIENT UNDERWENT AN ELECTIVE CARDIAC PROCEDURE ON (B)(6) 2024 AND SUTURE WAS USED TO CLOSE THE INCISION. THE PROCEDURE WAS SUCCESSFULLY COMPLETED. SIX HOURS POST-OP, THE PATIENT¿S INCISION REOPENED AND THE PATIENT EXSANGUINATED INSTANTANEOUSLY, RESULTING IN IMMEDIATE PATIENT DEATH. ADDITIONAL INFORMATION WAS REQUESTED. | PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ADDITIONAL INFORMATION WAS REQUESTED, AND THE FOLLOWING WAS OBTAINED: NAME OF INDEX SURGICAL PROCEDURE? THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? MITRAL STENOSIS AND INSUFFICIENCY WHAT WAS THE INITIAL APPROACH FOR THE INDEX SURGICAL PROCEDURE? (OPEN, LAPAROSCOPIC OR OTHER)? OPEN WHAT INSTRUMENTS WERE USED IN THIS PROCEDURE TO HANDLE THE SUTURE? CASTRO NEEDLE DRIVER ON WHAT TISSUE WAS THE SUTURE USED/LOCATION OF SUTURE PLACEMENT FOR EACH SUTURE? LEFT ATRIUM WHAT TISSUE DEHISCED? LEFT ATRIUM WHAT WAS THE TISSUE CONDITION (NORMAL, THIN, CALCIFIED, FRAGILE, DISEASED)? NORMAL HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? CONTINUOUS HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? 5 KNOTS HOW WAS THE DEHISCENCE MANAGED? PATIENT WAS OPENED EMERGENTLY AT THE BEDSIDE IN THE POSTOP ICU WHAT WAS THE SOURCE AND TRIGGERING EVENT OF BLEEDING? UNKNOWN. SUTURE WAS FOUND BROKEN WHAT WAS THE VOLUME OF BLOOD LOSS? MASSIVE 6+L HOW WAS THE BLEEDING TREATED? OPENED THE PATIENT AND BLOOD GIVEN WAS ANY MEDICAL/SURGICAL INTERVENTION PERFORMED POST-OP WHEN THE INCISION RE-OPENED TO TRY AND STOP THE BLEEDING? IF YES, PLEASE DESCRIBE. ALL THE BLOOD HAD EXITED THE PATIENT. THIS WAS A MASSIVE BLOOD LOSS WERE THERE ANY PATIENT STRESS FACTORS THAT PRECIPITATED THE EVENT OF THE INCISION REOPENING AND BLEEDING? NO DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? NOT IN THIS CASE OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? NONE WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT?SUTURE FRACTURE HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? AUTOPSY NOT NEEDED SINCE WE OPENED THE PATIENT WOULD THE PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? YES.
SOLITAIRE FR | K113455 | ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE DEATH WAS NOT THOUGHT TO BE RELATED TO THE MEDTRONIC DEVICE OR THERAPY. THE PATIENT¿S CONDITION WAS UNKNOWN AFTER DISCHARGE AND COULD NOT BE CONTACTED, AND THE OCCLUDED PART COULD NOT BE UNBLOCKED. THE CLOT COULD NOT BE REMOVED FROM THE BLOOD VESSEL. THE INSTRUMENT COULD NOT ENTER, SO THE CHARACTERISTICS OF THE CLOT WERE UNKNOWN. THE CAUSE OF THE SOLITAIRE BREAK/SEPARATION WERE NOT DETERMINED. THE INITIAL ESTIMATE WAS THAT IT MAY HAVE BEEN THAT THE BLOOD VESSELS WERE EXTREMELY TORTUOUS, RESULTING IN MANUAL DETACHMENT OF THE DETACHMENT POINT. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REMSTAR PLUS | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH AND GASTROESOPHAGEAL CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.
EVERFLO | K061261 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A EVERFLO. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION REGARDING A EVERFLO. THE PATIENT WAS PASSED AWAY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
EDWARDS ESHEATH PLUS INTRODUCER SHEATH | K200258 | AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 29 MM SAPIEN 3 ULTRA RESILIA VALVE IN THE AORTIC POSITION, THE 16F ESHEATH+ WAS INSERTED AT A STEEP ANGLE, BUT WITHOUT INCIDENT ON THE PATIENT'S LEFT SIDE. AS THE IMPLANTING CARDIOLOGIST WAS INSERTING THE COMMANDER INTO THE ESHEATH+, HE FELT RESISTANCE BUT NOTED IT WAS NOT MORE THAN USUAL, BUT THEN IT FELT LIKE THE COMMANDER WAS STUCK. IMAGING WAS PANNED DOWN TO THE GROIN AREA, AND THE TEAM SAW THE DISTAL END OF THE ESHEATH+ WITH THE TIP OF THE COMMANDER OUT OF IT BUT DID NOT VISUALIZE LOWER TO SEE THE VALVE. HE PANNED BACK UP TO THE HEART TO VISUALIZE THE WIRE ACROSS THE AORTIC VALVE AND RESUMED PUSHING THE COMMANDER INTO THE ESHEATH. AT THIS POINT HE FELT INCREASED RESISTANCE AND PANNED DOWN TO THE GROIN TO VISUALIZE THE VALVE. IT WAS REALIZED THEN THAT THE VALVE AND COMMANDER HAD PUSHED THROUGH THE SIDE WALL OF THE ESHEATH+. THE PHYSICIAN WAS UNABLE TO ADVANCE THE VALVE FORWARD OR PULL IT BACK INTO THE ESHEATH+. HE ATTEMPTED TO PULL BOTH THE COMMANDER AND ESHEATH BACK TOGETHER AT THE SAME TIME AND MET RESISTANCE. THE PATIENTS PRESSURED DROPPED AND VASCULAR WAS CALLED. A CUTDOWN WAS PERFORMED ON THE LEFT LEG AND THE ESHEATH+ AND COMMANDER WERE REMOVED WITH A PIECE OF TISSUE ASSUMED TO BE THE PATIENTS ILIAC ATTACHED. THE ILIAC WAS TORN WHEN TRYING TO REMOVE THE ESHEATH+ AND COMMANDER. THE PATIENT EXPIRED. | INVESTIGATION IS ONGOING.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | ON 30/JUL/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2024 WHILE UNDERGOING CCPD THERAPY. DURING INTAKE THE PATIENT¿S PD REGISTERED NURSE (PDRN) REPORTED THE PATIENT¿S EXPIRATION WAS UNRELATED TO PD THERAPY. NO ADDITIONAL INFORMATION PROVIDED DURING INTAKE. ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. | PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY DUE TO CANCER, SO HIS WIFE WILL NEED A RETURN LABEL TO RETURN THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH, PANCREATIC, ESOPHAGEAL. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
SAFARI2¿ | K151244 | EVENT DESCRIPTION: A TRANSCATHETER AORTIC VALVE IMPLANTATION WAS PERFORMED AND A 25 MM NAVITOR TRANSCATHETER BIOPROSTHETIC VALVE SYSTEM WAS ATTEMPTED. DURING THE IMPLANTATION OF THE NAVITOR, THE BLOOD PRESSURE DECREASED AND AN ECHOCARDIOGRAM WAS PERFORMED, WHICH REVEALED A PERICARDIAL EFFUSION. THE PHYSICIAN DETERMINED THAT THE USE OF THE SAFARI GUIDEWIRE HAD CAUSED LEFT VENTRICULAR PERFORATION, LEADING TO CARDIAC TAMPONADE. PERICARDIAL DRAINAGE WAS PERFORMED, AND THE PATIENT IMMEDIATELY UNDERWENT OPEN-CHEST SURGERY. HEMOSTASIS WAS ATTEMPTED AT THE PERFORATION SITE BY SUTURING, BUT AS THIS WAS NOT ACHIEVED, LEFT VENTRICULAR RECONSTRUCTION WAS PERFORMED. HOWEVER, AS THE BLEEDING WAS STILL NOT COMPLETELY STOPPED, THE PATIENT WAS RETURNED TO THE INTENSIVE CARE UNIT (ICU) WITH SOME DEGREE OF HEMOSTASIS ACHIEVED. ON (B)(6) 2024, THE PATIENT DIED. REGARDING THE CAUSE OF DEATH, THE PHYSICIAN STATED THAT THE CAUSE OF DEATH WAS HEMORRHAGE DUE TO PERFORATION OF THE LEFT VENTRICULAR APEX BY A GUIDEWIRE. | PRODUCT WAS DISCARDED; THEREFORE, NO VISUAL OR PHYSICAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE (DFU) WARNINGS SECTION: ·MONITOR WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP. ·WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH-RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. ·NEVER ADVANCE THE GUIDEWIRE AGAINST RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. ·THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED IN THE VENTRICLE. ·THE CURVE OF THE SAFARI2 GUIDEWIRE SHOULD BE CONSTRAINED WITH IN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 7.ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8.CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9.MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED TO DATE, IT APPEARS THAT CLINICAL AND/OR PROCEDURAL FACTORS HAVE CONTRIBUTED TO THE EVENT. ALTHOUGH THE EXACT CAUSE OF THE PERFORATION HAS NOT BEEN DETERMINED, CARDIAC PERFORATION, TAMPONADE, PERICARDIAL EFFUSION, AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS AND ARE LISTED IN THE SAFARI2 INSTRUCTIONS FOR USE (IFU). THE SECTIONS LEFT BLANK OR UNKNOWN ON THIS FORM COULD NOT BE COMPLETED DUE TO MISSING INFORMATION. AN ATTEMPT TO GATHER FURTHER DETAILS TO OBTAIN THE INFORMATION WAS MADE; HOWEVER, PER THE DISTRIBUTOR, GOOD FAITH EFFORTS HAVE BEEN PERFORMED AND AT THE TIME OF THIS REPORT NO ADDITIONAL INFORMATION HAS BEEN RECEIVED REGARDING THIS EVENT. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 09/11/2024, AND SECTION H10 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID WERE UPDATED IN THIS REPORT.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY, BUT THE PATIENT DID NOT PASS AWAY DUE TO CPAP DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT THE PATIENT PASSED AWAY, BUT THE PATIENT DID NOT PASS AWAY DUE TO CPAP DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION GO AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, AND ANAPLASTIC ASTROCYTOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY DUE TO NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY DUE TO NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY AND CANCER. PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED WITH WRONG TYPE OF REPORTED COMPLAINT AS SERIOUS INJURY. IN THIS REPORT, THE TYPE OF REPORTED COMPLAINT HAS BEEN UPDATED CORRECTLY AS DEATH. SECTION TYPE OF REPORTED COMPLAINT IN BOX H HAS BEEN UPDATED/ CORRECTED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG, COPD. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2024, A PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM USING GORE® DRYSEAL FLEX INTRODUCER SHEATH (DSF), TREO DEVICE (TERUMO, PROXIMAL 28 MM, DISTAL 14 MM, LENGTH 100 MM). DURING THE TREATMENT, ATTEMPT WAS MADE TO INSERT A 18 FR DSF FROM THE LEFT FEMORAL ARTERY, HOWEVER IT WAS NOT ABLE TO BE ADVANCED DUE TO HIGH CALCIFICATION OF THE ACCESS VESSEL. A 14 FR DSF WAS INSERTED FROM THE RIGHT FEMORAL ARTERY. THE TREO WAS DEPLOYED BUT COULD NOT BE IMPLANTED PERPENDICULAR TO THE VESSEL. A PROXIMAL TYPE I ENDOLEAK FROM THE TREO WAS OBSERVED. FOR TREATMENT, AN ENDURANT (AORTA EXTENSION, 28 MM) WAS IMPLANTED. THE PROXIMAL TYPE I ENDOLEAK WAS REDUCED BUT REMAINED. ISCHEMIA OF THE RIGHT LEG WAS CONFIRMED AND DISSECTION WAS OBSERVED. THE VESSEL WAS DISSECTED AND ANASTOMOSIS WAS ATTEMPTED. DURING THE DISSECTION ANASTOMOSIS, A THROMBUS WAS FOUND IN THE SURROUNDING AREA, THEREFORE A THROMBECTOMY WAS ALSO PERFORMED. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2024, ISCHEMIA (DETAIL UNKNOWN) WAS OBSERVED. A BARE STENT WAS IMPLANTED FROM THE DISTAL LEG OF THE TREO TO THE FEMORAL ARTERY. THE PATIENT WAS STABLE WITH NO ISCHEMIA. ON (B)(6) 2024, THE PATIENT PASSED AWAY OF INTESTINAL RUPTURE DUE TO ISCHEMIC ENTERITIS. COMMENTS FROM THE MEDICAL INSTITUTION WERE AS FOLLOWS: THE PATIENT HAD CALCIFICATION OF THE ACCESS VESSEL AND A TORTUOSITY AT THE SITE OF IMPLANTATION, THEREFORE, WE WERE AWARE OF THE RISK BEFORE THE PROCEDURE. IT WAS THOUGHT THAT A BLOOD CLOT FLEW INTO THE SMA (SUPERIOR MESENTERIC ARTERY) DUE TO RUBBING OF THE VESSEL WHEN THE DEVICE WAS BEING INSERTED OR REMOVED. | H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICES NOT AVAILABLE FOR RETURN. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: ACCORDING TO THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: VASCULAR TRAUMA (I.E., DISSECTION, RUPTURE, PERFORATION, TEAR, ETC.), EMBOLIZATION (MICRO OR MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA, DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2024, A PATIENT UNDERWENT ENDOVASCULAR TREATMENT OF AN ABDOMINAL AORTIC ANEURYSM USING GORE® DRYSEAL FLEX INTRODUCER SHEATH (DSF), TREO DEVICE (TERUMO, PROXIMAL 28 MM, DISTAL 14 MM, LENGTH 100 MM). DURING THE TREATMENT, ATTEMPT WAS MADE TO INSERT A 18 FR DSF FROM THE LEFT FEMORAL ARTERY, HOWEVER IT WAS NOT ABLE TO BE ADVANCED DUE TO HIGH CALCIFICATION OF THE ACCESS VESSEL. A 14 FR DSF WAS INSERTED FROM THE RIGHT FEMORAL ARTERY. THE TREO WAS DEPLOYED BUT COULD NOT BE IMPLANTED PERPENDICULAR TO THE VESSEL. A PROXIMAL TYPE I ENDOLEAK FROM THE TREO WAS OBSERVED. FOR TREATMENT, AN ENDURANT (AORTA EXTENSION, 28 MM) WAS IMPLANTED. THE PROXIMAL TYPE I ENDOLEAK WAS REDUCED BUT REMAINED. ISCHEMIA OF THE RIGHT LEG WAS CONFIRMED AND DISSECTION WAS OBSERVED. THE VESSEL WAS DISSECTED AND ANASTOMOSIS WAS ATTEMPTED. DURING THE DISSECTION ANASTOMOSIS, A THROMBUS WAS FOUND IN THE SURROUNDING AREA, THEREFORE A THROMBECTOMY WAS ALSO PERFORMED. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2024, ISCHEMIA (DETAIL UNKNOWN) WAS OBSERVED. A BARE STENT WAS IMPLANTED FROM THE DISTAL LEG OF THE TREO TO THE FEMORAL ARTERY. THE PATIENT WAS STABLE WITH NO ISCHEMIA. ON (B)(6) 2024, THE PATIENT PASSED AWAY OF INTESTINAL RUPTURE DUE TO ISCHEMIC ENTERITIS. COMMENTS FROM THE MEDICAL INSTITUTION WERE AS FOLLOWS: THE PATIENT HAD CALCIFICATION OF THE ACCESS VESSEL AND A TORTUOSITY AT THE SITE OF IMPLANTATION, THEREFORE, WE WERE AWARE OF THE RISK BEFORE THE PROCEDURE. IT WAS THOUGHT THAT A BLOOD CLOT FLEW INTO THE SMA (SUPERIOR MESENTERIC ARTERY) DUE TO RUBBING OF THE VESSEL WHEN THE DEVICE WAS BEING INSERTED OR REMOVED. | H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICES NOT AVAILABLE FOR RETURN. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: ACCORDING TO THE GORE® EXCLUDER® AAA ENDOPROSTHESIS INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: EMBOLIZATION (MICRO OR MACRO) WITH TRANSIENT OR PERMANENT ISCHEMIA, DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
MESH ¿ VENTRALEX | K132441 | ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF AN UNSPECIFIED BARD/DAVOL VENTRALEX HERNIA PATCH ON (B)(6) 2007. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST THE VENTRALEX HERNIA PATCH. ATTORNEY ALLEGES THAT THE PATIENT HAD SUBSEQUENT SURGICAL INTERVENTION DUE TO THE HERNIA MESH DEVICE AND WRONGFUL DEATH OF THE PATIENT. IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. | NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING SUBSEQUENT SURGICAL INTERVENTION AND DEATH. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : NOT RETURNED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTOA-FLEX DEVICE'S SOUND ABATEMENT FOAM THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH. THE PATIENT ALSO ALLEGED KIDNEY DISEASE/TOXICITY, CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. | REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
SEPRAMESH IP | K063739 | ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF AN UNSPECIFIED BARD/DAVOL SEPRAMESH IP ON (B)(6) 2016. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST THE SEPRAMESH IP. ATTORNEY ALLEGES THAT THE PATIENT SUSTAINED PERSONAL INJURY DUE TO THE HERNIA MESH DEVICE AND WRONGFUL DEATH OF THE PATIENT. IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. | NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT : NOT RETURNED.
TRILOGY 100 | K083526 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE REPORTER ALLEGED OF HAVING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH FROM A TRILOGY 100 DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE REPORTER ALLEGED OF HAVING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH FROM A TRILOGY 100 DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS EVALUATED BY THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY (PIL) AND FOUND THE DEVICE FUNCTIONED AS DESIGNED AND OPERATED WITHOUT ISSUE OR ERROR CODES. THE VENTILATOR PASSED THE POSTTEST¿S APPLICABLE TEST STEPS ON THE TRILOGY MULTIFUNCTION TEST STATION. NO PARTICLES WERE FOUND ON THE INSIDE OF THE AIR FLOW PATH ON THE DEVICE. PIL HAS DETERMINED THAT THE FLOW SENSOR (MT5) ON THE SENSOR BOARD IS OUT-OF-TOLERANCE AT HIGHER FLOWS AND IS INDICATIVE TO CONTAMINATION OF THE FLOW SENSOR. THIS FAILURE HAS BEEN INVESTIGATED AND IDENTIFIED, REFERRING TO THE ENGINEERING ANALYSIS ON SENSOR FAILURES IN (B)(4). THE OUT-OF-TOLERANCE HIGHER POSITIVE FLOWS OF THE FLOW SENSOR WOULD NOT BE THE CAUSE OF ANY OF THE ALLEGATIONS IN THE COMPLAINTS.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF ASTHMA, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, REDUCED CARDIOPULMONARY RESERVE. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
OMNIPOD 5 POD | K203768 | AN EMAIL WAS RECEIVED FROM AN UNCONFIRMED CUSTOMER ALLEGING, " FAMILY FRIEND FEW WEEKS AGO THAT HAD DIED OF AN OVERDOSE OF INSULIN FROM OMNIPOD." THREE ATTEMPTS TO FOLLOW UP WITH THE CUSTOMER WERE COMPLETED BUT NO NEW INFORMATION BECAME AVAILABLE. IF NEW INFORMATION BECOMES AVAILABLE WE WILL REOPEN AND SUPPLEMENT THIS FILE. | ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED PATIENT'S PASSING. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED.
DREAMSTATION BIPAP AUTOSV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, LUNG DISEASE AND CANCER. THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA AND REDUCED CARDIOPULMONARY RESERVE, INFLAMMATORY RESPONSE, HYPERSENSITIVITY, DIZZINESS/HEADACHE, NAUSEA/VOMITING AND EYE, SKIN, NOSE AND RESPIRATORY TRACT IRRITATION. THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE HARM REPORTED. BASED ON AVAILABLE INFORMATION, AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE HAS BEEN NO COMMUNICATION FROM THE CUSTOMER ON THE RETURN OF THE DEVICE. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, ASTHMA (NEW OR WORSENING). THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A PARA-ESOPHAGEAL HERNIA REPAIR PROCEDURE, THE PATIENT'S VITAL SIGNS STARTED TO DROP REQUIRING CARDIOPULMONARY RESUSCITATION; THE PATIENT EXPIRED. THE PROCEDURE WAS ALMOST COMPLETED WHEN IT WAS OBSERVED THAT THE PATIENT WAS HYPOTENSIVE AND HAD AN ERRATIC HEART RATE WITH AN UNSPECIFIED CARDIAC ARRHYTHMIA. WHILE UNDOCKING THE DA VINCI ROBOT TO CONVERT TO AN EMERGENT OPEN PROCEDURE, THE PATIENT "FLAT LINED" AND CPR WAS PERFORMED FOR APPROXIMATELY 30 MINUTES WITHOUT SUCCESS. THE PATIENT'S FAMILY DECLINED AN AUTOPSY AND INFORMED THE SURGEON THAT THE PATIENT'S UNCLE HAD DIED DURING A SURGERY IN THE SAME WAY YEARS PRIOR. ALTHOUGH NO TESTING WAS DONE, THE SURGEON STATED THAT THE PATIENT DEATH WAS POTENTIALLY EITHER A MYOCARDIAL INFARCTION/CARDIAC EVENT OR A MASSIVE PULMONARY EMBOLISM. THE SURGEON STATED NO DA VINCI PRODUCTS CONTRIBUTED TO THE EVENT. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, TIP-UP FENESTRATED GRASPER, LARGE NEEDLE DRIVER, CADIERE FORCEPS, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED DURING A PARA-ESOPHAGEAL HERNIA PROCEDURE. THERE WERE NO INTRAOPERATIVE ISSUES UP UNTIL THE TIME OF THE EVENT AND NO INJURIES WERE REPORTED. THERE WERE NO ALLEGATIONS OF ANY ISSUES OR PROBLEMS WITH ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. THE CAUSE OF THE DEATH IS UNKNOWN. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
SYSTEM ONE HTD HUMID ASSY ,DOM | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATIONRESPIRATORY TRACT IRRITATIONDIZZINESS AND/OR HEADACHENAUSEA / VOMITINGDEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION THAT AN END USER DEVELOPED DEATH WHILE USING AN DS2ADV AUTO CPAP. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY BECAME AWARE OF AN ALLEGATION THAT AN END USER DEVELOPED DEATH WHILE USING AN DS2ADV AUTO CPAP. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IN PREVIOUS REPORT, MANUFACTURER MENTIONED BOX B (ADVERSE EVENT/PRODUCT PROBLEM) INCORRECTLY, WHICH HAS BEEN CORRECTED IN THIS REPORT. THE DEVICE IS NOT A PART OF RECALL. IN THIS REPORT, BOX B, H HAS BEEN UPDATED OR CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING EVALUATION OF THE DEVICE AT A THIRD-PARTY SERVICE CENTER, THE DEVICE LOGS WERE DOWNLOADED AND NO ERRORS CODE WERE FOUND. THERE WAS NO EVIDENCE OF FOAM PARTICLES DURING THE TESTING OF THE DEVICE. NO OTHER DEVICE PROBLEMS WERE FOUND. THE DEVICE WAS SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING IRRITATION (EYE, NOSE, SKIN, RESPIRATORY TRACT), DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA/VOMITING, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, LUNG DISEASE (UNSPECIFIED). THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA / VOMITINGINFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED OF EYE IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA/VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, RESPIRATORY TRACT IRRITATION AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
SENSATION PLUS 8FR. 50CC IAB & ACCESSORIES (W/PRESSURE TUBES, NO STYLET) | K133074 | IT WAS REPORTED THAT DURING THERAPY, AN INTRA-AORTIC BALLOON (IAB) PERFORATION AND LEAK OCCURRED. THE PUMP WAS PLACED IN STANDBY PER PROTOCOL AND THERAPY WAS DISCONTINUED. IT WAS NOTED THAT PRIOR TO THE LEAK TWO "LOW GAS" ALARMS HAD OCCURRED BUT NO BLOOD WAS VISIBLE IN THE TUBING AT THAT TIME. THE PATIENT HAD ALREADY BEEN VERY SICK AND ULTIMATELY EXPIRED APPROXIMATELY 6 HOURS AFTER DISCONTINUATION OF IAB THERAPY. THE IAB COULD NOT BE REMOVED DUE TO COAGULOPATHY. BLOOD PRODUCTS ORDERED. THE HOSPITAL CHOSE NOT TO REPLACE RUPTURED IAB DUE TO FAILING AND DETERIORATING CONDITION OF THE PATIENT. | OCCUPATION - MSN, RN, ACNS ¿ BC, APRN, CCRN - CMC / CLINICAL NURSE SPECIALIST. ADDITIONAL INITIAL REPORTERS - (B)(6), ICU RNS & (B)(6). THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER SO WE ARE UNABLE TO COMPLETE AN EVALUATION. IF PROVIDED WE WILL SEND A SUPPLEMENTAL REPORT WITH OUR ADDITIONAL FINDINGS. COMPLAINT RECORD ID # (B)(4). H3 OTHER TEXT : DEVICE NOT RETURNED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED FUNGAL INFECTION IN LUNGS. THE PATIENT DIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
AZURION | K200917 | IT WAS REPORTED TO PHILIPS THAT AT THE START OF AN EMBOLISM CASE, THE AZURION¿S C-ARM MOVEMENT WAS RESTRICTED, AND IT COULD NOT FULLY ROTATE TO THE DOCTOR'S SIDE OF THE TABLE. A MOBILE C-ARM WAS BROUGHT INTO THE ROOM TO PROCEED WITH THE CASE. IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS. | PHILIPS HAS INVESTIGATED THIS COMPLAINT. ACCORDING TO THE ADDITIONAL INFORMATION COLLECTED, IT WAS CONFIRMED BY HOSPITAL STAFF THAT THE PATIENT'S DEATH WAS NOT RELATED TO THE PHILIPS SYSTEM. A PHILIPS FIELD SERVICE ENGINEER (FSE) INSPECTED THE SYSTEM ON-SITE AND CONFIRMED THE SYSTEM WAS WORKING AS INTENDED. THE ON-SITE INVESTIGATION CONCLUDED THAT THE FRONTAL STAND WAS MOVED OUT OF ITS WORKING AREA FOR MOTORIZED MOVEMENTS WHICH CAUSED THE REPORTED ISSUE. PER THE SYSTEM¿S INSTRUCTIONS FOR USE (IFU): ¿IF YOU MOVE THE STAND OUT OF ITS WORKING AREA FOR MOTORIZED MOVEMENTS, A GUIDANCE MESSAGE IS DISPLAYED IN THE STATUS AREA INDICATING THE STAND IS OUT OF RANGE. MOVE THE STAND BACK TO THE WORKING AREA MANUALLY TO CONTINUE WITH MOTORIZED MOVEMENT.¿ A REVIEW OF THE LOG FILES CONFIRMED THAT THE SYSTEM PROVIDED THE APPROPRIATE USER GUIDANCE AS PRESCRIBED IN THE IFU. THE FSE PROVIDED THE CUSTOMER WITH ADDITIONAL TRAINING ON THE PHILIPS SYSTEM MOVEMENTS. THE SYSTEM WAS HANDED OVER TO THE CUSTOMER IN GOOD WORKING ORDER. AT THE TIME THIS COMPLAINT WAS RECEIVED, PHILIPS DID NOT HAVE ENOUGH INFORMATION TO EXCLUDE THAT A PRODUCT MALFUNCTION OCCURRED AND CONTRIBUTED TO THE REPORTED PATIENT OUTCOME. SINCE THAT TIME, NEW INFORMATION HAS BEEN ACQUIRED, LEADING TO THE DETERMINATION THAT THE PHILIPS SYSTEM WAS WORKING AS INTENDED, AND THAT THE PATIENT¿S DEATH WAS UNRELATED TO THE PHILIPS SYSTEM. AS SUCH, THIS COMPLAINT HAS BEEN RE-EVALUATED AS NOT REPORTABLE HEALTH IMPACT CODE WAS CORRECTED.
FORTILINK-C IBF SYSTEM | K192718 | A 4-LEVEL ACDF PROCEDURE WAS COMPLETED SUCCESSFULLY AND DURING THE FLIP OF THE PATIENT A CARDIAC ARREST EVENT HAPPENED AND THE PATIENT EXPIRED. IT WAS CONFIRMED THAT THIS WAS NOT A PRODUCT RELATED EVENT. | THE ANTERIOR/ACDF WAS COMPLETED WITH NO COMPLICATIONS. AS THE SUPERFICIAL WOULD WAS BEING CLOSED THE PATIENT CODED. THE PATIENT WAS REVIVED. THE POSTERIOR PORTION OF THE PROCEDURE NEEDED TO BE COMPLETED. IT WAS BELIEVED THAT THE PATIENT THREW A PULMONARY EMBOLUS WHILE IN THE PROCESS OF BEING TURNED OVER FROM ANTERIOR TO POSTERIOR, AND THE PATIENT WAS NOT ABLE TO BE REVIVED. THERE IS NO ALLEGATION AGAINST THE DEVICE.
COMPANION 5 | K121167 | AS REPORTED: PATIENT FELL ASLEEP WHILE SMOKING WITH OXYGEN ON, IGNITING OXYGEN. PATIENT SUSTAINED GRADE THREE INHALATION BURNS AND ADDITIONAL PARTIAL THICKNESS BURNS TO HANDS AND FEET. SOOT NOTED THROUGH THE MOUTH AND NOSE, AND BILATERAL CORNEAL ABRASIONS. PATIENT TAKEN TO TRANSFERRED TO ORLANDO REGIONAL MEDICAL CENTER BURN UNIT. PATIENT IS CURRENTLY INTUBATED WITH VENTILATOR SUPPORT. UPDATED 17 JULY: PATIENT DEATH DATE (B)(6) 2024. | PURSUANT TO TITLE 21 - FOOD AND DRUGS, CHAPTER I - FOOD AND DRUG ADMINISTRATION DEPARTMENT OF HEALTH AND HUMAN SERVICES, SUBCHAPTER H -0 MEDICAL DEVICE, PART 803 - MEDICAL DEVICE REPORTING, SUBPART A - GENERAL PROVISIONS, SECTION 803.16, NEITHER THIS REPORT NOR ANY INFORMATION SUBMITTED HEREIN CONSTITUTES AN ADMISSION BY CAIRE INC. THAT THE DEVICE STATED IN THIS REPORT, CAIRE INC., OR CAIRE INC.'S EMPLOYEES, CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT STATED HEREIN. THE INFORMATION PROVIDED REPRESENTS THE BEST INFORMATION AVAILABLE TO CAIRE. THE DEVICE WILL NOT BE AVAILABLE FOR ENGINEERING EVALUATION AND THE AVAILALBE EVIDENCE INDICATES CONFIRMED USE ERROR. THE END USER REPORTED FALLING ASLEEP WHILE SMOKING A CIGARETTE AND WHILE USING THE DEVICE. NO SMOKING WARNINGS ON THE EQUIPMENT AND IN THE IFU WERE REVIEWED AND DEEMED ADEQUATE IN THE COURSE OF THE INVESTIGATION. "NO SMOKING" SYMBOLS (REG# P002) ARE AFFIXED TO THE FACE OF THE UNITS. THROUGHOUT THE UNIT IFU, WARNINGS STATE NOT TO ALLOW SMOKING OR OPEN FLAME NEAR THE DEVICE AND TO KEEP ALL FLAMMABLE MATERIALS AWAY FROM THE EQUIPMENT. ADDITIONAL WARNINGS STATE THAT SMOKING WHILE WEARING AN OXYGEN CANNULA CAN CAUSE FACIAL BURNS AND POSSIBLY RESULT IN DEATH.
FORTILINK-C IBF SYSTEM | K192718 | A 4-LEVEL ACDF PROCEDURE WAS COMPLETED SUCCESSFULLY AND DURING THE FLIP OF THE PATIENT A CARDIAC ARREST EVENT HAPPENED AND THE PATIENT EXPIRED. IT WAS CONFIRMED THAT THIS WAS NOT A PRODUCT RELATED EVENT. | THIS FOLLOW UP IS TO CORRECT THE UDI NUMBER THAT WAS INCORRECTLY ENTERED ON THE INITIAL REPORT. THE ANTERIOR/ACDF WAS COMPLETED WITH NO COMPLICATIONS. AS THE SUPERFICIAL WOULD WAS BEING CLOSED THE PATIENT CODED. THE PATIENT WAS REVIVED. THE POSTERIOR PORTION OF THE PROCEDURE NEEDED TO BE COMPLETED. IT WAS BELIEVED THAT THE PATIENT THREW A PULMONARY EMBOLUS WHILE IN THE PROCESS OF BEING TURNED OVER FROM ANTERIOR TO POSTERIOR, AND THE PATIENT WAS NOT ABLE TO BE REVIVED. THERE IS NO ALLEGATION AGAINST THE DEVICE.
DJO SURGICAL | K143242 | IT WAS REPORTED THAT THE PATIENT PASSED AWAY ON (B)(6) 2023 FROM "EXSANGUINATION DUE TO LACERATION OF THE LEFT POPLITEAL ARTERY, DUE TO MECHANICAL FAILURE OF LEFT PROSTHETIC KNEE JOINT." | ADDITIONAL REPORTING ON THIS EVENT WILL BE PROVIDED AS A SUPPLEMENTAL REPORT TO THIS DOCUMENT AS SOON AS IT BECOMES AVAILABLE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING INFLAMMATORY RESPONSE, CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), LUNG DISEASE. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
AGILIA VP MC BR | K121613 | THE FOLLOWING HAS BEEN REPORTED: ASSESS DEVICE OPERATION SUCH AS PROGRAMMING, TOTAL VOLUME CALCULATION, ALARMS, KVO AND KVO ALARM. ON (B)(6) 2024 AT 10:00 A.M. PATIENT IN ICU USING NORADRENALINE INFUSING 30ML/H FOR AN INDEFINITE PERIOD AND 100ML SODIUM CHLORIDE BAG. THE VOLUME OF THE ALARM WAS NOT HEARD BY THE TEAM - AT THE END OF THE DRUG INFUSION THE PUMP DID NOT EMIT AN AUDIBLE ALARM. THE PATIENT DIED. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EVENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION. | N/A
OMNIPOD 5 CONTROLLER | K203768 | IT WAS REPORTED THAT THE PATIENT PASSED AWAY DUE TO A HYPOGLYCEMIC EVENT AND WAS EXPERIENCING FLUCTUATIONS IN BLOOD GLUCOSE (BG) LEVELS, AS WELL AS, HAVING TECHNICAL ISSUES WITH THE CONTINUOUS GLUCOSE MONITOR (CGM). REPORTEDLY, THE HEALTHCARE PROFESSIONAL (HCP) HAS POSSESSION OF THE OMNIPOD 5 CONTROLLER- ATTEMPTS ARE BEING MADE TO RETRIEVE THE DEVICE. THE SERIAL NUMBER OF THE LAST DEVICE TO COMMUNICATE WITH THE CLOUD WAS IN MANUAL MODE AND DELIVERING THE PROGRAMMED BASAL RATE. NO BG VALUES WERE BEING RECEIVED FROM A SENSOR AND THE DEVICE WAS GETTING THE MISSING SENSOR VALUES ALERT. THERE'S A GAP IN THE DATA BETWEEN (B)(6) 2024 AND (B)(6) 2024 THAT WOULD INDICATE THE CONTROLLER WAS OFF FOR THIS PERIOD. LAST DATA POINT WAS ON (B)(6) 2024. AT THIS TIME, THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. | ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO CONFIRM THE REPORTED THE REPORTED PATIENT'S DEATH. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. NO ANDROID LOG FILES WERE UPLOADED TO THE CLOUD SYSTEM. REVIEW OF AVAILABLE CLOUD DATA SHOWED NO ABNORMALITIES. SYSTEM WAS USED IN MANUAL MODE DELIVERING USER-PROGRAMMED BASAL PROFILE FOR THE DAYS LEADING UP TO DATE OF OCCURRENCE. THE SYSTEM WAS NOT RECEIVING BGS FROM A SENSOR FOR THIS PERIOD AND THE USER WAS MADE AWARE THROUGH ALERTS. CONTROLLER WAS OFF FOR A PERIOD STARTING ON (B)(6) 2024 AND ENDING ON (B)(6) 2024. LAST DATA POINT WAS GENERATED ON (B)(6) 2024. CLOUD DATA SHOWS CONTROLLER IN USE WAS OF SN (B)(6). CONTROLLER SN ALSO REGISTERED TO USER (B)(6) WAS NOT USED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY DUE TO INFLAMMATORY RESPONSE, CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMST ST30 H | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMST ST30 DEVICE PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K063377 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT TWO MONTHS AND ONE DAY LATER POST PORT PLACEMENT PROCEDURE FOR CHEMOTHERAPY TREATMENT FOR PANCREATIC CANCER, THE PATIENT ALLEGEDLY DEVELOPED WITH BACTERIAL AND FUNGAL INFECTION. IT WAS FURTHER REPORTED THAT PATIENT HAD PAIN OVER THE PORT SITE AND DEVELOPED WITH SEPSIS. REPORTEDLY, THE PATIENT WAS TREATED WITH ANTIBIOTICS FOR INFECTION AND EVENTUALLY EXPIRED. THE CAUSE OF DEATH WAS NOT PROVIDED. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H10: D4 (EXPIRY DATE: 02/2025). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
NELLCOR | K120773 | ACCORDING TO THE REPORTER, DURING USE, THE DEVICE'S SPO2 VALUE DECREASED AS THE PATIENT'S CONDITION SUDDENLY CHANGED, BUT THE NURSE NOTICED IT TOO LATE. IT WAS UNKNOWN WHETHER THE ALARM SOUNDED OR NOT. FURTHERMORE, IT WAS UNKNOWN WHETHER OR NOT THE ALARM VOLUME WAS CHANGED FROM THE DEFAULT SETTINGS OF THE HOSPITAL. THE PATIENT DIED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
TRILOGY O2 VENTILATOR, JAPAN | K083526 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A PATIENT HAS DIED WHILE USING A VENTILATOR ON (B)(6) 2024. THE PATIENT WAS USING TRILOGY O2 PLUS. THE COMPLAINT INITIALLY STATES "HE REQUESTS THAT WE RETRIEVE THE AED. HE ALSO ASKED ME TO EXTRACT DEVICE DATA. WHEN I CHECKED THE DETAILS, I CONFIRMED THAT A PATIENT HAD DIED WHILE THE EQUIPMENT WAS IN USE AT 16:40. FURUKAWA, THE PERSON IN CHARGE AT THE HOSPITAL, WILL VISIT THE WARD. I CHECKED THE EVENT HISTORY FOR THE DEVICE. WHEN THE CASE OCCURS, I CONFIRMED THAT THE OFF CIRCUIT (1,489 SECONDS) AND THE MINUTE VENTILATION VOLUME LOWER LIMIT (1,476 SECONDS) ALARMS OCCURRED. AT THE SAME TIME, I CHECKED THE SD CARD DATA AND CONFIRMED FLUCTUATIONS IN THE WAVEFORM DATA, ETC. AT THE TIME THE PROBLEM OCCURRED. THE DIRECTION FOR FUTURE INVESTIGATION AND CONFIRMATION HAD NOT BEEN DECIDED, SO I WILL VISIT AGAIN THE NEXT WEEK I EXPLAINED THAT HE WOULD COLLECT THE DEVICE AND INSPECT IT, AND I COMPLETED THE RESPONSE." ADDITIONALLY, THE COMPLAINT STATES THAT ON (B)(6) AT 11:53 A CALL FROM A HEAD NURSE OF THE HOSPITAL WARD WAS RECEIVED. THEY ASKED US TO RETRIEVE THE UNIT AS THE PATIENT WHO HAD BEEN USING IT HAD DIED. THEY ALSO REQUESTED FOR DATA EXTRACTION. THE SALES REPRESENTATIVE CONFIRMED WITH THEM THAT THE PATIENT HAD DIED WHILE USING THE UNIT AT APPROXIMATELY 16:40. THE SALES REP VISITED THE HOSPITAL WARD. HE CHECKED THE EVENT LOG AND CONFIRMED THAT CIRCUIT DISCONNECT ALARM HAD BEEN OCCURRING FOR 1,489 SECONDS AND LOW MINUTE VENTILATION ALARM HAD BEEN OCCURRING FOR 1,476 SECONDS AT THE TIME OF THE EVENT. HE ALSO CHECKED THE SD CARD DATA AND CONFIRMED THAT THERE HAD BEEN FLUCTUATIONS IN THE WAVEFORM DATA AND SUCH DURING THE TIME OF THE EVENT. SINCE THEY HAVE NOT DECIDED THE DIRECTION FOR FUTURE INVESTIGATION, THE REPRESENTATIVE WILL VISIT THEM AGAIN NEXT WEEK AND RETRIEVE THE UNIT FOR INVESTIGATION. THEY ARE ASKING FOR THE DETAILED DATA FOR THE DAY OF THE EVENT AND FOR THE DETAILED DATA AND EVENT LOG AT THE TIME OF THE EVENT. PLEASE NOTE THAT THE DEVICE TIME IS AROUND 10 MINUTES BEHIND THE ACTUAL TIME. RESUSCITATION MEASURES WERE PERFORMED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING THAT A PATIENT HAS DIED WHILE USING A VENTILATOR ON (B)(6) 2024. THE PATIENT WAS USING TRILOGY O2 PLUS. THE COMPLAINT INITIALLY STATES "HE REQUESTS THAT WE RETRIEVE THE AED. HE ALSO ASKED ME TO EXTRACT DEVICE DATA. WHEN I CHECKED THE DETAILS, I CONFIRMED THAT A PATIENT HAD DIED WHILE THE EQUIPMENT WAS IN USE AT 16:40. FURUKAWA, THE PERSON IN CHARGE AT THE HOSPITAL, WILL VISIT THE WARD. I CHECKED THE EVENT HISTORY FOR THE DEVICE. WHEN THE CASE OCCURS, I CONFIRMED THAT THE OFF CIRCUIT (1,489 SECONDS) AND THE MINUTE VENTILATION VOLUME LOWER LIMIT (1,476 SECONDS) ALARMS OCCURRED. AT THE SAME TIME, I CHECKED THE SD CARD DATA AND CONFIRMED FLUCTUATIONS IN THE WAVEFORM DATA, ETC. AT THE TIME THE PROBLEM OCCURRED. THE DIRECTION FOR FUTURE INVESTIGATION AND CONFIRMATION HAD NOT BEEN DECIDED, SO I WILL VISIT AGAIN THE NEXT WEEK I EXPLAINED THAT HE WOULD COLLECT THE DEVICE AND INSPECT IT, AND I COMPLETED THE RESPONSE." ADDITIONALLY, THE COMPLAINT STATES THAT ON (B)(6) 2024 AT 11:53 A CALL FROM A HEAD NURSE OF THE HOSPITAL WARD WAS RECEIVED. THEY ASKED US TO RETRIEVE THE UNIT AS THE PATIENT WHO HAD BEEN USING IT HAD DIED. THEY ALSO REQUESTED FOR DATA EXTRACTION. THE SALES REPRESENTATIVE CONFIRMED WITH THEM THAT THE PATIENT HAD DIED WHILE USING THE UNIT AT APPROXIMATELY 16:40. THE SALES REP VISITED THE HOSPITAL WARD. HE CHECKED THE EVENT LOG AND CONFIRMED THAT CIRCUIT DISCONNECT ALARM HAD BEEN OCCURRING FOR 1,489 SECONDS AND LOW MINUTE VENTILATION ALARM HAD BEEN OCCURRING FOR 1,476 SECONDS AT THE TIME OF THE EVENT. HE ALSO CHECKED THE SD CARD DATA AND CONFIRMED THAT THERE HAD BEEN FLUCTUATIONS IN THE WAVEFORM DATA AND SUCH DURING THE TIME OF THE EVENT. SINCE THEY HAVE NOT DECIDED THE DIRECTION FOR FUTURE INVESTIGATION, THE REPRESENTATIVE WILL VISIT THEM AGAIN NEXT WEEK AND RETRIEVE THE UNIT FOR INVESTIGATION. THEY ARE ASKING FOR THE DETAILED DATA FOR THE DAY OF THE EVENT AND FOR THE DETAILED DATA AND EVENT LOG AT THE TIME OF THE EVENT. PLEASE NOTE THAT THE DEVICE TIME IS AROUND 10 MINUTES BEHIND THE ACTUAL TIME. RESUSCITATION MEASURES WERE PERFORMED. THE CLINICAL ASSESSMENT STATES THAT THE DEVICE MAY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT VIA POSSIBLE FAILURE, MALFUNCTION, IMPROPER OR INADEQUATE DESIGN, MANUFACTURE, LABELING, AND/OR USER ERROR PER 21 CFR 803.3. BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY AN UNINTENTIONAL USER ERROR ¿ FAILURE TO RESPOND TO CIRCUIT DISCONNECT AND LOW MINUTE VENTILATION ALARMS. DEVICE EVALUATIONS AND LOG ATTACHMENTS ARE CURRENTLY PENDING. THEREFORE, THE ALLEGED DEATH WHILE ON THE DEVICE IS ASSESSED AS POSSIBLE RELATED. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, THERE WERE NO ABNORMALITIES CONFIRMED DURING AN OPERATIONAL CHECK. WHEN THE DEVICE WAS MADE TO OPERATE WITH THE CIRCUIT REMOVED, THE 'CIRCUIT DISCONNECT' AND 'LOW MINUTE VENTILATION' ALARMS WERE CONFIRMED TO SOUND PROPERLY. THE DEVICE'S ERROR LOG CONFIRMED THE RECORDS INDICATING THAT THE 'CIRCUIT DISCONNECT' HAD SOUNDED FOR 1,489 SECONDS AND THE 'LOW MINUTE VENTILATION' SOUNDED FOR 1,476 SECONDS. THE SERVICE TECHNICIAN CONFIRMED THAT THERE WAS NO ERROR INDICATING A DEVICE FAILURE. ADDITIONALLY, WHEN A REPAIR TEST WAS PERFORMED AS A VERIFICATION RELATING TO A MEASUREMENT OF FLOW AT 0 L/MIN, IT HAD RESULTED IN A FAIL. IT IS CONFIRMED THAT ALL THE OTHER REPAIR TESTS WERE A PASS. THE WAVEFORM DATA SHOWS THAT THE FLOW VALUE WHEN THE 'CIRCUIT DISCONNECT' OCCURRED WAS APPROXIMATELY 155 L/MIN WITH THE INSPIRATORY PRESSURE SETTING OF 15 CMH2O AND APPROXIMATELY 83 L/MIN WITH THE PEEP SETTING OF 5 CMH2O. THEREFORE, IT IS CONSIDERED THAT IT WAS NOT AFFECTED BY THE ITEMS THAT FAILED THE TEST. IT IS CONSIDERED THAT THE CIRCUIT WAS DISCONNECTED DURING THE TIME WHEN THE PATIENT WAS USING THE DEVICE, CAUSING THE 'CIRCUIT DISCONNECT' AND 'LOW MINUTE VENTILATION' ALARMS TO SOUND. THE SENSOR PRINTED CIRCUIT ASSEMBLY (PCA) BOARD NEEDS TO BE REPLACED TO ADDRESS THE TEST FAILURE. SINCE THE LEASE TERM IS ABOUT TO END, THE REPAIR WAS CANCELLED AND THE DEVICE WILL BE SCRAPPED. THE SENSOR PCA BOARD WAS SENT TO THE PRODUCT INVESTIGATION LAB (PIL) FOR FURTHER TESTING. PIL WAS ABLE TO DUPLICATE THE FAILURE OF THE SENSOR BOARD. THE ROOT CAUSE OF THE FAILURE IS INDICATIVE TO CONTAMINATION OF THE AIR FLOW SENSOR (MT5). PIL INSTALLED THE SUSPECT BOARD INTO THE TRILOGY TEST UNIT AND TESTED THE SENSOR BOARD ON THE TRILOGY TEST STATION. THE DEVICE FAILED THE ZERO, POSITIVE, AND NEGATIVE FLOW VERIFICATION STEPS. PIL DETERMINED THAT THE MT5 IS THE CAUSE OF THE FAILURE AND THIS FAILURE HAS BEEN INVESTIGATED AND IDENTIFIED. NO FURTHER INVESTIGATION IS REQUIRED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE PATIENT ALSO ALLEGED EYE IRRITATION, RESPIRATORY TRACT IRRITATION, REDUCED CARDIOPULMONARY RESERVE, CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED CANCER, STAGE 4 ESOPHAGEAL CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE AND CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
VIVO 45 LS | K193586 | ON 07/16/2024, BREAS RECEIVED THE FOLLOWING INFORMATION FROM DISTRIBUTOR (B)(4). "PT PASSED ON (B)(6) 2024 POSSIBLY IN LATE AFTERNOON. (B)(4) WASN'T INFORMED UNTIL 7/16 AND RECEIVED MINIMAL DETAILS. THE INCIDENT MAY HAVE HAPPENED WHILE THE PATIENT WAS IN TRANSPORT FROM A PPEC FACILITY THAT WATCHES CHILDREN. A COUSIN OF THE FAMILY WORKS ON A REFERRAL BASIS FOR A PERSONAL INJURY LAW FIRM AND MENTIONED THE SITUATION THE LAW FIRM. THE FAMILY ASKED FOR THE DATA, AND WHEN THEY WERE TOLD THAT (B)(4) WOULD NEED TO DOWN-LOAD THE DATA AND THEN GIVE THEM A COPY, THE ATTORNEYS TOOK POSSESSION OF THE DEVICE. (B)(4) INFORMED THE LAWYERS THAT THE DEVICE IS NOT THEIRS TO TAKE. THE LAWYERS ASKED FOR A PURCHASE OR RENTAL PRICE FOR THE DEVICE SO THAT THEY COULD RETAIN CUSTODY OF THE DEVICE. (B)(4) SUPPLIED THEM WITH A NUMBER THE LAWYERS THEN SAID THAT THEY WILL LOOK FOR "SOMEONE WHO CAN DOWNLOAD THE DATA". (B)(4) INFORMED THEM THAT WAS NOT A GOOD IDEA. THEY INFORMED THEM THAT AN UNQUALIFIED PERSON COULD POTENTIALLY DELETE ALL THE DATA WITH A PUSH OF THE WRONG BUTTON.". IT IS ALLEGED THAT THE PATIENT WAS CONNECTED TO A VIVO 45LS DEVICE, PART NUMBER: 230216 AND SERIAL NUMBER: (B)(6), AT THE TIME OF THE ADVERSE EVENT, BUT THIS IS NOT SUPPORTED BY THE LOG FILES OBTAINED FROM THE DEVICE (SEE ADDITIONAL COMMENTS). THERE HAS BEEN NO INDICATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE EVENT AT THIS TIME. | BREAS MEDICAL INC. DOES NOT HAVE ACCESS TO THE PHYSICAL DEVICE AT THIS TIME; HOWEVER, LOG FILES FROM THE DEVICE WERE OBTAINED WHICH WERE ANALYZED BY BREAS R&D. LOG FILES INDICATE THAT THE DEVICE, SN (B)(6), WAS NOT IN USE AT THE TIME THE EVENT WAS REPORTED TO HAVE OCCURRED ON JULY 15, 2024. THE LOG FILES SHOW DATA FROM (B)(6) 2024. TREATMENT DATA ENDS ON (B)(6) 2024. DATA FROM THE (B)(6) WAS RELATED TO DEVICE STATUS ONLY WITH THERE NO TREATMENT OCCURRING ON THOSE DAYS. THE LAST DATA IN THE LOG FILES IS ON (B)(6) AT THE TIME THE DOWNLOAD WAS DONE BY A BREAS EMPLOYEE. AFTER ANALYZING THIS DATA, THERE IS NO INDICATION OF CORRUPTION TO THE FILE. TREATMENT SESSIONS ARE RECORDED AS WELL AS ALARMS, ALL OF WHICH ARE CONSISTENT WITH DEVICE SPECIFICATIONS.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN PRODUCT UDI AND PATIENT OUTCOME CODE GRID WHICH WERE CORRECTED IN THIS REPORT. DESCRIPTION EVENT OR PROBLEM HAS BEEN CORRECTED AS THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER, STOMACH, SPLEEN, INTESTINE, CANCER (UNSPECIFIED) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, LUNG, KIDNEY, LIVER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
BIPAP AUTOSV ADV | K090539 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN PREVIOUS REPORT DESCRIBE EVENT OR PROBLEM WAS CAPTURED INCORRECTLY AND IT HAS BEEN CORRECTED IN THIS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED DEATH, PLEURAL EFFUSION, E COLI PNEUMONIA, ACUTE RESPIRATORY FAILURE, ACUTE BLOOD LOSS ANEMIA, BROKEN ARM, HOSPITALIZATION AND NURSING HOME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. IN BOX B: DATE OF DEATH CAPTURED. IN BOX D: DEVICE SERVICED BY 3RDP? AND DEVICE AVAILABLE FOR EVAL? HAS BEEN UPDATED. IN BOX G: DATE RECEIVED BY MFG, REPORT TYPE AND IN BOX H, IF FOLLOW-UP, WHAT TYPE?, DEVICE PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID, REMEDIAL ACTION INITIATED AND RECALL (Z) NUMBER HAS BEEN UPDATED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, DEATH NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
SENTRANT INTRODUCER SHEATH | K123990 | A SENTRANT SHEATH (00181873) WAS USED AS A CONDUIT DURING THE IMPLANT OF A TRANSCATHETER BIOPROSTHETIC VALVE. IT WAS REPORTED DURING THE PROCEDURE, THE DELIVERY CATHETER SYSTEM (DCS) FAILED TO RECAPTURE THE VALVE. THE DCS BECAME BENT, EXPOSING METAL BELIEVED TO BE THE CAPSULE STRUCTURE AND LEAVING 10 TO 15% OF THE VALVE WAS EXPOSED ALMOST TO NODE 3, MAKING CAPTURE AND RELEASE OF THE VALVE IMPOSSIBLE AND PREVENTING ITS EXTRACTION. A SENTRANT SHEATH WAS USED DURING THE PROCEDURE BLOCKED THE DCS FROM EXITING THE PATIENT'S BODY. SURGICAL INTERVENTION WAS REQUIRED TO REMOVE THE VALVE. THE VALVE WAS WITHDRAWN WITHIN THE DCS AND INTRODUCER SHEATH, AND A DISSECTION WAS OBSERVED IN THE LEFT FEMORAL ARTERY. THE (CONFIDA) GUIDEWIRE WAS SUSPECTED TO HAVE PERFORATED THE LEFT VENTRICLE, RESULTING IN THE SUBSEQUENT DEATH OF THE PATIENT. PER THE PHYISICAN THE CAUSE OF DEATH WAS DUE TO THE DELIVERY SYSTEM FAILURE DURING THE PROCEDURE WHICH PREVENTED THE RECAPTURE OF THE VALVE, NECESSITATING EMERGENCY SURGERY, WHICH ENDED WITH A PERFORATION TO THE PATIENTS LEFT VENTRICLE. IT WAS NOTED THAT THE VALVE DID NOT CAUSE OR CONTRIBUTE TO THE INJURY. THE DCS CAUSED DAMAGE TO THE INSERT AND EXPOSURE OF MATERIALS, WHILE THE SENTRANT SHEATH CONTRIBUTED TO THE INJURY BY PREVENTING THE EXIT OF THE DCS. | B5: ADDITIONAL INFORMATION RECEIVED: IT WAS REPORTED THAT THE DCS AND VALVE WERE INTRODUCED INTO THE PATIENT USING A SENTRANT SHEATH. IT WAS CLARIFIED THAT THE DELIVERY SYSTEM KINKED , WHICH PREVENTED PASSAGE THROUGH THE SENTRANT SHEATH, CAUSING A JAM AND PREVENTING THE EXIT OF THE RELEASE DEVICE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
TUBING SET | K112360 | COMPLAINT ID# (B)(4). | THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT THE DIALYSIS LOCK WITH VALVE WAS ACCESSED BRIEFLY, AND WHEN THE CLINICIANS REMOVED THE TUBING (THE EMERGENCY PRIMING LINE THAT COMES IN THE HLS SET), THE SPRING APPARENTLY FAILED AND DID NOT SEAL OFF THE PORT. A FAIR AMOUNT OF BLOOD WAS LOST BY THE TIME THEY GOT THE CAP PUT BACK ONTO THE PORT. THE ORIGINAL CAP THAT CAME ON THE PORT WAS USED. INITIALLY, THAT STOPPED THE BLEEDING, BUT ABOUT AN HOUR LATER, THAT PORT STARTED SPRAYING BLOOD. THE SPECIALIST USED HIS FINGER TO STOP THE BLOOD FROM SPRAYING OUT, AND ANOTHER EMERGENCY LINE WAS GRABBED TO PUT ON THE PORT. THE LINE WAS CLAMPED, AND THIS STOPPED THE BLOOD FROM COMING OUT. UPON INSPECTION OF THE CAP PLACED ON THE PORT, IT WAS NOTICED THAT THERE WAS A BIG OPEN HOLE IN THE CENTER. THE PRE-PRIMED BACK UP CIRCUIT WAS LOOKED AT. IT HAD THE SAME CAP, BUT IT WAS REALIZED THAT IT IS COVERED BY A WHITE ¿STICKER.¿ ACCORDING TO THE CLINICIAN, IT WAS JUST NOTED THAT THE CAPS HAVE HOLES IN THEM THAT ARE THEN COVERED BY THE WHITE STICKER GIVEN THAT IT¿S THE FIRST TIME THIS PROBLEM OCCURS. THE SPRING FAILED IN THE PORT. THE CUSTOMER SAID THAT THE DON¿T USE THIS PORT OFTEN. THE CIRCUIT WAS 28 HOURS OLD. THIS INCIDENT CAUSED THE PATIENT TO LOSE A SIGNIFICANT AMOUNT OF BLOOD AND IT WAS NEEDING A TRANSFUSING OF 2 UNITS OF BLOOD. ON 2024-08-05, THE INFORMATION WAS RECEIVED THAT THE USER DID NOT SWAP THE HLS SET WHEN THE FAILURE OCCURRED. INSTEAD, THE CUSTOMER FIXED THE FAILURE AND CONTINUED TO USE. THE HLS SET WAS ONLY DISCONNECTED WHEN THE PATIENT COMPLETED THE THERAPY. ON 2024-08-13 THE INFORMATION WAS RECEIVED THAT THE CUSTOMER PUT A QUICK PRIME LINE ON THE PORT AND CLAMPED THE QUICK PRIME LINE. HE TIED IT IN A COUPLE KNOTS TOO. THE PATIENT EXPIRED ON (B)(6) 2024. THE PATIENT EXPIRED DUE TO HIS ILLNESS NOT DUE TO THE CIRCUIT CONFIRMED BY THE GETINGE SERVICE AND SALES UNIT (REFER TO COMMUNICATION GRID). THE AFFECTED PRODUCT WAS INVESTIGATED BY THE GETINGE LABORATORY ON 2024-12-13 WITH FOLLOWING CONCLUSION. DURING TESTING A LEAKAGE AT THE DIALYSIS LOCK WAS DETECTED. DURING VISUAL INSPECTION IT COULD BE SEEN DAMAGES ON THE DIALYSIS CONNECTOR AND THE VALVES ON THE INNER SIDE OF THE CONNECTOR. THUS THE REPORTED FAILURE " LEAKAGE ON DIALYSIS LOCK" COULD BE CONFIRMED. THE DAMAGES COULD BE RELATED TO EXCESSIVE EXTERNAL FORCES WHICH LEAD TO THE LEAKAGE. AS FAR AS THE CUSTOMER REPORTED THE DIALYSIS CONNECTOR WAS NOT SHOWING ANY MALFUNCTION DURING PRIMING NOR LATER DURING USE. WHEN DISCONNECTING THE EMERGENCY PRIMING LINE BY THE CUSTOMER IT LEAD TO THE LEAKAGE. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON (B)(6) 2024 WITH FOLLOWING CONCLUSION:"THE COMPLAINT NARRATIVE DESCRIBES AN INCIDENT IN TEXAS, USA, ON (B)(6) 2024. THE AFFECTED PATIENT, 34 YEARS OLD, 193 CM, 74 KG, WAS PLACED ON VV ECMO ON (B)(6) 2024 DUE TO ARDS, RESPIRATORY FAILURE, AND SEPTIC SHOCK, WITH CO-MORBIDITIES INCLUDING PULMONARY TUBERCULOSIS, INFLUENZA, AND RIGHT VENTRICULAR DYSFUNCTION. THE DIALYSIS LOCK WAS BRIEFLY ACCESSED FOR DIALYSIS, AND WHEN THE EMERGENCY PRIMING LINE WAS REMOVED, THE SPRING IN THE PORT MALFUNCTIONED AND FAILED TO SEAL IT OFF, CAUSING SIGNIFICANT BLOOD LOSS. THE ORIGINAL CAP WAS REATTACHED, INITIALLY STOPPING THE BLEEDING, BUT ABOUT AN HOUR LATER, THE PORT BEGAN TO LEAK BLOOD UNDER PRESSURE. THE BLEEDING WAS CONTROLLED BY APPLYING EXTERNAL PRESSURE, AND ANOTHER EMERGENCY LINE WAS ATTACHED AND CLAMPED TO STOP THE BLOOD FLOW. UPON INSPECTION OF THE CAP ON THE PORT, A LARGE HOLE WAS OBSERVED IN THE CENTER. THE PRE-PRIMED BACKUP CIRCUIT, WHICH FEATURED A SIMILAR CAP, WAS EXAMINED, AND FOUND TO HAVE A HOLE COVERED BY A WHITE STICKER. THE CLINICIAN NOTED THAT THE CAPS ARE INTENTIONALLY DESIGNED WITH HOLES THAT ARE TYPICALLY SEALED BY THE STICKER. FOLLOWING THE SPRING FAILURE IN THE PORT, THE CUSTOMER QUESTIONED THE RATIONALE BEHIND THE CAP'S DESIGN, AS IT RELIES ON THE SPRING TO FUNCTION WITHOUT FAILURE. THE CIRCUIT, 28 HOURS OLD, HAD BEEN REPLACED THE DAY PRIOR. THIS INCIDENT LED TO SIGNIFICANT BLOOD LOSS FOR THE PATIENT, WHO REQUIRED A TRANSFUSION OF TWO UNITS OF BLOOD. THE PATIENT STABILIZED AFTER THE TRANSFUSION, AND THE HLS SET WAS REPLACED DURING THE TREATMENT. ON (B)(6) 2024, THE PATIENT UNFORTUNATELY SUCCUMBED DUE TO EXISTING MORBIDITIES. THROUGH LABORATORY INVESTIGATIONS OF THE AFFECTED HLS SET AND THE SUBSEQUENT LCE REPORT, A LEAK AT THE DIALYSIS LOCK WAS CONFIRMED. ADDITIONALLY, SEVERAL SIGNIFICANT DAMAGES TO THE DIALYSIS CONNECTOR - SPECIFICALLY THE CLOSURE PIN, DIALYSIS CONNECTOR, AND THE BLOOD OUTLET CAP WERE IDENTIFIED IN THE FORM OF CRACKS. ACCORDING TO THE LCE REPORT, SUCH DAMAGES ¿[...] CAN ONLY OCCUR THROUGH SUBSTANTIAL EXTERNAL FORCE (E.G., CLAMPING, ETC.) [...]¿ WHICH LIKELY CONTRIBUTED TO DAMAGE IN THE FRONT AREA OF THE CLOSURE PIN AND POTENTIALLY LED TO THE MISALIGNMENT OF THE CLOSURE VALVE. AN INSPECTION OF ALL INDIVIDUAL VALVE COMPONENTS CONFIRMED THAT ALL PARTS WERE CORRECTLY POSITIONED, MAKING INCORRECT ASSEMBLY AS A ROOT CAUSE LESS LIKELY. FURTHERMORE, THE LCE REPORT INDICATES THAT THE VALVE MISALIGNMENT CAUSED BY THE DAMAGE TO THE CLOSURE PIN WAS ¿HIGHLY LIKELY¿ TO HAVE RESULTED IN THE LEAKAGE DESCRIBED BY THE CUSTOMER AFTER THE DIALYSIS LINE WAS REMOVED. THEREFORE. IT CANNOT BE DEFINITIVELY DETERMINED IF THE SUBSTANTIAL DAMAGES OCCURRED DURING THE INITIAL CONNECTION OR DURING SUBSEQUENT SEALING ATTEMPTS. BASED ON THE COMPLETED QUESTIONNAIRE, SPECIFICALLY QUESTION 3 "WAS LEAKING OBSERVED FROM THE DIALYSIS PORT AT THE TIME OF PRIMING?", THE AFFECTED CUSTOMER RESPONDED AS FOLLOWS: "NO, THERE WAS NO LEAKING BEFORE OR WHEN WE INITIALLY PLACED THE EMERGENCY LINE ONTO THE DIALYSIS PORT. THE LEAKING STARTED WHEN WE REMOVED THE EMERGENCY LINE FROM THE PORT". IN THIS REGARD, IT BECOMES EVIDENT THAT NO DAMAGE TO THE DIALYSIS CONNECTOR WAS PRESENT DURING THE PRIMING PROCEDURE OR PRIOR TO THE CONNECTION OF THE EMERGENCY PRIMING LINE. CONSEQUENTLY, DUE TO THE LACK OF FURTHER INFORMATION REGARDING THE INCIDENT ITSELF, IT CANNOT BE FURTHER CLARIFIED WHETHER THE SUBSTANTIAL DAMAGES OCCURRED DURING THE INITIAL CONNECTION OR DURING SUBSEQUENT SEALING ATTEMPTS. FURTHERMORE, THE INVESTIGATIONS IN THE LCE REPORT REVEALED THAT THE FAILURE PATTERN, INCLUDING COMPARABLE DAMAGES TO THE CLOSURE PIN, HAD ALREADY BEEN IDENTIFIED AND OCCURRED MULTIPLE TIMES. IN THIS REGARD, THE CLOSURE PIN AND THE AFFECTED BATCH OF THE COMPONENT WERE EXAMINED MORE CLOSELY. THE ROOT CAUSES AND ASSOCIATED RISKS IN THE FIELD WERE ASSESSED IN A FIELD SAFETY CORRECTIVE ACTION (FSCA) DECISION, AND ACTIONS TO ADDRESS THE ROOT CAUSE WERE DEFINED IN AN ENGINEERING CHANGE REQUEST (ECR). HOWEVER, NO ABNORMALITIES WERE FOUND IN THE BATCH NUMBER OF THE AFFECTED HLS SET DURING THIS INVESTIGATION. GIVEN THE CURRENT FINDINGS, A MANUFACTURING OR PRODUCTION-RELATED CAUSE IS UNLIKELY. IF ISSUES WERE PRESENT IN THE INJECTION MOLDING PROCESS, THEY WOULD HAVE BEEN MORE SYSTEMATIC AND RECURRENT. VISUAL AND DIMENSIONAL INSPECTIONS CONDUCTED DURING INCOMING GOODS CHECKS AND PRODUCTION WOULD PREVENT THE INSTALLATION OF DEFECTIVE PARTS. IN CONCLUSION, THE PRIMARY CAUSE OF THE LEAKAGE IS CONSIDERED TO BE THE MISALIGNMENT OF THE VALVE CAUSED BY DEFORMATIONS OF THE CLOSURE PIN, ALONG WITH ADDITIONAL DAMAGE TO THE DIALYSIS CONNECTOR. THIS LED TO AN INCOMPLETE SEAL BETWEEN THE O-RING AND THE BLOOD OUTLET CAP. HOWEVER, NO CLEAR ROOT CAUSES COULD BE IDENTIFIED FOR THE DAMAGES OBSERVED IN THE DIALYSIS CONNECTOR ASSEMBLY, INCLUDING THE DEFORMATIONS OF THE CLOSURE PIN. FINALLY, IT IS DIFFICULT TO DETERMINE IF THE REPORTED EVENT HAS A DIRECT INFLUENCE ON THE OUTCOME OF THE PATIENT OR IF THE OUTCOME OF THE PATIENT WAS A CONFLUENCE OF CONFOUNDING FACTORS AS THE PATIENT WAS SUFFERING FROM SEVERE CONDITIONS (ARDS, RESPIRATORY FAILURE, AND SEPTIC SHOCK) AND NUMEROUS COMORBIDITIES (PULMONARY TUBERCULOSIS, INFLUENZA, AND RIGHT VENTRICULAR DYSFUNCTION). HOWEVER, THE INCIDENT IN QUESTION, INCLUDING FRANK LOSS OF BLOOD VOLUME WHICH NECESSITATED ADMINISTRATION OF BANKED BLOOD PRODUCTS, MAY HAVE CONTRIBUTED TO THE PATIENT'S OVERALL DETERIORATION. HOWEVER, WITHOUT MORE DETAILS REGARDING THE EVENTS SURROUNDING THE COMPLIANT, IT IS CHALLENGING TO DEFINITIVELY ASSIGN THE REPORTED EVENT TO EITHER A DIMINUTION IN PERFORMANCE OR TO A MALPERFORMANCE OF THE PRODUCT. " AS STATED IN THE INSTRUCTIONS FOR USE HLS SET ADVANCED 5.0 / 7.0, HIT SET ADVANCED 5.0 / 7.0 - IN CHAPTER PREPARATION AND INSTALLATION ¿PERFORM A CAREFUL VISUAL INSPECTION OF THE DEVICE BEFORE USE. IN PARTICULAR, ENSURE THERE IS NO DAMAGE TO THE MATERIAL, CRACKS, BURRS OR FISSURES.¿ FURTHERMORE, IT IS STATED IN CHAPTER PRIMING THE SYSTEM ¿CHECK THE DEVICE FOR LEAKS DURING PRIMING. DO NOT USE THE DEVICE IF THERE ARE ANY LEAKS.¿ IN THE CHAPTER USING DIALYSIS LOCK, IT IS INDICATED THAT ¿IF A TUBE IS CONNECTED TO THE DIALYSIS LOCK WITH VALVE DURING ONGOING PERFUSION, AIR MAY ENTER THE OXYGENATOR. THIS CAN LEAD TO AN AIR EMBOLISM IN THE PATIENT. THEREFORE, SCREW THE TUBE TO THE DIALYSIS LOCK WITH A SWIFT TURN-AND-PRESS MOVEMENT. MAKE SURE THAT THE TUBE IS NOT CLAMPED AND THE AIR IN THE TUBE IS DISPLACED BY THE BLOOD FLOWING OUT, TO PREVENT A PULSATING BACKFLOW OF AIR INTO THE OXYGENATOR. MAKE SURE THAT THE PRESSURE IN THE OXYGENATOR IS HIGHER THAN THE AMBIENT PRESSURE¿. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON (B)(6) 2024. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "LEAKAGE DIALYSIS LOCK" COULD BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING AND NOT NEEDING IT ANYMORE. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
SIGNIA | K191070 | ACCORDING TO THE REPORTER, DURING A PROCEDURE TO REMOVE A RIGHT UPPER LOBE DUE TO LUNG CANCER, THE PATIENT'S SATURATION LEVELS AND PRESSURE KEPT DROPPING. PATENT WAS CONVERTED FROM A ROBOTIC SURGERY TO AN OPEN SURGERY TO GAIN MORE CONTROL FOR ANESTHESIA TO INFLATE THE LUNGS. THE SURGEON FIRED A SMALL-DIAMETER RELOAD (PARTIALLY FIRED) ON THE PULMONARY ARTERY AND HAD SOME OOZING (THERE WAS BLOOD LOSS). TO CONTROL THE BLEEDING, A BLOOD TRANSFUSION AND A CELL SAVER WERE MADE. SUTURING TO REPAIR THE TEAR IN THE VESSEL WOULD HAVE BEEN THE RESOLUTION, BUT THE PATIENT COULD NOT GET STABLE. THE PATIENT THEN BEGAN TO CRASH FURTHER. A CARDIAC SURGEON WAS CALLED IN, AND RESUSCITATION EFFORTS COMMENCED. THE PATIENT DIED DUE TO CARDIAC ARREST. | ADDITIONAL INFORMATION: B5, G3, H6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
SIGNIA | K160176 | ACCORDING TO THE REPORTER, DURING A PROCEDURE TO REMOVE A RIGHT UPPER LOBE DUE TO LUNG CANCER, THE PATIENT'S SATURATION LEVELS AND PRESSURE KEPT DROPPING. PATENT WAS CONVERTED FROM A ROBOTIC SURGERY TO AN OPEN SURGERY TO GAIN MORE CONTROL FOR ANESTHESIA TO INFLATE THE LUNGS. THE SURGEON FIRED A SMALL-DIAMETER RELOAD (PARTIALLY FIRED) ON THE PULMONARY ARTERY AND HAD SOME OOZING (THERE WAS BLOOD LOSS). TO CONTROL THE BLEEDING, A BLOOD TRANSFUSION AND A CELL SAVER WERE MADE. SUTURING TO REPAIR THE TEAR IN THE VESSEL WOULD HAVE BEEN THE RESOLUTION, BUT THE PATIENT COULD NOT GET STABLE. THE PATIENT THEN BEGAN TO CRASH FURTHER. A CARDIAC SURGEON WAS CALLED IN, AND RESUSCITATION EFFORTS COMMENCED. THE PATIENT DIED DUE TO CARDIAC ARREST. | ADDITIONAL INFORMATION: B5, G3, H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL-DREAMSTATION AUTO CPAP PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
BIPAP A40 | K121623 | THIS COMPLAINT HAS BEEN REVIEWED FOR THE INFORMATION THE MANUFACTURER RECEIVED THROUGH A REPRESENTATIVE OF THE CUSTOMER STATING "WE HAVE BEEN ADVISED BY THE CLINIC THAT THE PATIENT HAS SADLY PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH". COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION THE MANUFACTURER RECEIVED THROUGH A REPRESENTATIVE OF THE CUSTOMER STATING "WE HAVE BEEN ADVISED BY THE CLINIC THAT THE PATIENT HAS SADLY PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH." COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE WAS EVALUATED BY THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY (PIL). "THE PRODUCT INVESTIGATION LAB (PIL) IS UNABLE TO CONFIRM AN INTERRUPTION OF THERAPY OF THIS BIPAP A40 PRO. EXTERNAL VISUAL INSPECTION FOUND NO DEFECTS. PIL DOWNLOADED THE DEVICE ERROR LOG TO REVIEW THE EVENTS FROM THE LOG. THE INCIDENT WAS REPORTED ON (B)(6) 2024. THE ERROR LOG REVIEW DID FIND PATIENT RELATED ALARMS, AND NOT DEFECTIVE COMPONENT ALARMS. THIS DEVICE WAS POWERED OFF (B)(6) 2024 AND NOT POWERED UP AGAIN UNTIL (B)(6) 2025. AT THE PIL, THE UNIT RAN 25 MINUTES ON QUICKLUNG F7797, WITHOUT ALARM OR ERROR. NO ALLEGATIONS HAVE BEEN MADE AGAINST THIS DEVICE. THIS UNIT WILL BE SCRAPPED PER THE REQUIREMENTS SET FORTH IN WORK INSTRUCTION 8.4-1131 REV 9 AND DISPOSED OF ACCORDINGLY".
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING REDUCED CARDIOPULMONARY RESERVE. IN ADDITION, THE PATIENT ALSO ALLEGED EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), LUNG DISEASE. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER, WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE, DUE TO A PATIENT PASSING. AND NOT NEEDING IT ANYMORE. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION, THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED, FOR AN ALLEGATION OF A PATIENT DEATH ONLY, DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED, WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION, NO FAULTS WERE FOUND. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS RETURNED TO STOCK.
REACT CATHETER | K182097 | ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE UNDERLYING DISEASE 'STEMI' WAS TREATED FIRST AND THE CEREBROVASCULAR PROCEDURE WAS DELAYED. THEREFORE, THIS ADVERSE EVENT (AE) WAS JUDGED TO BE UNRELATED TO THIS STUDY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
SOLITAIRE X | K203358 | ADDITIONAL INFORMATION RECEIVED REPORTED THAT THE UNDERLYING DISEASE 'STEMI' WAS TREATED FIRST AND THE CEREBROVASCULAR PROCEDURE WAS DELAYED. THEREFORE, THIS ADVERSE EVENT (AE) WAS JUDGED TO BE UNRELATED TO THIS STUDY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
6VT-D | K223832 | THE CUSTOMER HAS CONTACTED GE HEALTHCARE TO INVESTIGATE DEVICE CONTRIBUTION TO AND/OR IMPACT UPON A PATIENT DEATH. CUSTOMER REPORTS THAT DURING THE EVENING AFTER AN ATRIAL FIBRILLATION ABLATION PROCEDURE, THE PATIENT WENT THROUGH A CARDIAC DECOMPENSATION, AND THE DOCTORS DIAGNOSED A STOMACH PERFORATION AND ESOPHAGUS THICKENING. ONCE THE DIAGNOSIS WAS ESTABLISHED, THE PATIENT WAS TRANSFERRED IN ICU AND INTUBATED, UNDER MECHANICAL VENTILATION, ANTIBIOTIC THERAPY, AND ENDOSCOPIC TREATMENT (PLACEMENT OF A CLIP ON AN ACTIVE BLEEDING AREA). THE SPECIALISTS THEN DECIDED ON A CONSERVATIVE APPROACH FOR THE ESOPHAGEAL PERFORATION, DIRECTING HEALING THROUGH A NASO-JEJUNAL PROBE. SURGEONS DID NOT NOTICE ANY ISSUE WITH THE TEE PROBE DURING THIS SPECIFIC PROCEDURE. CUSTOMER REPORTS THAT SEVEN DAYS AFTER THE PROCEDURE THE PATIENT DIED. THE CUSTOMER DOES NOT FAULT THE DEVICE, AND THEY ASK GE HEALTHCARE TO DEMONSTRATE THAT NEITHER PROBE NOR ECHOGRAPH COULD HAVE CAUSED THE ISSUE. GE HEALTHCARE HAS INVESTIGATION IS IN PROGRESS. | GE HEALTHCARE HAS INVESTIGATED THE POSSIBLE CONTRIBUTION OF A 6VT-D DIAGNOSTIC ULTRASOUND PROBE WITH VIVID S70 SYSTEM TO PATIENT INJURY AND DEATH. THE CUSTOMER PROVIDED TO GE HEALTHCARE THE 6VT-D PROBE AND THE VIVID S70 SYSTEM FOR EVALUATION. THE PROBE WAS EVALUATED FOR SAFETY AND EFFECTIVENESS, AND NO DEFECTS WERE FOUND WITH THE PROBE. THE CONSOLE LOGFILES WERE ASSESSED FOR OVERHEATING DURING THE PATIENT EXAM, AND NO INSTANCES OF THE PROBE OVERHEATING DURING THE PATIENT EXAM WERE FOUND WITHIN LOGFILES. THEREFORE, GE HEALTHCARE HAS CONCLUDED NEITHER THE PROBE NOR THE CONSOLE MALFUNCTIONED DURING THE REPORTED INCIDENT, AND NO FURTHER ACTION IS REQUIRED.
ZOOM REPERFUSION CATHETER | K210996 | A 67-YEAR-OLD FEMALE PATIENT WAS TREATED FOR AN OCCLUSION IN THE LEFT M2 SEGMENT. THE PATIENT WAS REPORTED TO HAVE ATHEROSCLEROTIC DISEASE WITH AN ACUTE BEND IN CERVICAL INTERNAL CAROTID ARTERY (ICA) AND HAIRPIN BEND BEYOND CAROTID BULB. ACCESS WAS OBTAINED WITH AN ARISTOTLE GUIDEWIRE AND A ZOOM 88 TO THE DISTAL PETROUS. THE PHYSICIAN INDICATED THAT ADVANCING THE ZOOM 88 OVER THE STIFF GUIDEWIRE WAS "TRICKY". DURING THE FIRST PASS, A ZOOM 55 CATHETER WAS ADVANCED THROUGH THE ZOOM 88 TO THE FACE OF THE CLOT. THE TREATING PHYSICIAN ENCOUNTERED SIGNIFICANT RESISTANCE WHILE ADVANCING THE ZOOM 55 THROUGH THE ZOOM 88. ASPIRATION WAS APPLIED TO THE ZOOM 55, AND THE CLOT WAS REMOVED FROM THE M2 SEGMENT, ACHIEVING A TICI 3 SCORE. THE ZOOM 88 WAS RETRACTED INTO THE COMMON CAROTID WITH THE ZOOM 55 INSIDE IT. THERE WAS RESISTANCE RETRACTING THE ZOOM 55 THROUGH THE ZOOM 88. AFTER THE ZOOM 55 WAS REMOVED FROM THE PATIENT, IT WAS NOTICED THAT THE DISTAL 10MM OF THE CATHETER WAS BROKEN. THE PHYSICIAN ALSO NOTED THERE WAS NO CLOT OBSERVED IN THE FILTER (ZOOM POD). THE PHYSICIAN BELIEVED THAT THE DISTAL SEGMENT OF THE ZOOM 55 AND CLOT WERE STILL INSIDE THE ZOOM 88. A CONTRAST RUN WAS THEN PERFORMED. IMAGING SHOWED THAT THE 10MM DISTAL PORTION OF THE ZOOM 55 WAS IN THE CERVICAL CAROTID. IT WAS BELIEVED THAT THE CONTRAST INJECTION FLUSHED THE ZOOM 55 DISTAL PORTION OUT OF THE ZOOM 88 AND INTO THE CERVICAL CAROTID. THE PHYSICIAN ATTEMPTED TO REMOVE THE ZOOM 55 THROUGH ZOOM 88 WITH ASPIRATION BUT WAS UNSUCCESSFUL. ANOTHER CONTRAST RUN WAS PERFORMED AND SHOWED THAT A CLOT WAS BEGINNING TO FORM AROUND THE BROKEN PIECE THAT HAD MOVED FROM THE CERVICAL CAROTID INTO THE ICA. ATTEMPTS WERE MADE TO CAPTURE THE BROKEN PIECE USING 2MM AND 4MM SNARES RESULTING IN IT BEING PUSHED FURTHER INTO THE SIPHON. ANTICOAGULANTS/ANTIPLATELETS DRUGS WERE ALSO ADMINISTERED, BUT CLOT BURDEN REMAINED. ON THE FINAL RETRIEVAL ATTEMPT, THE PHYSICIAN USED A STENT RETRIEVER UNSUCCESSFULLY. THE SAME ARISTOTLE GUIDE WIRE WAS THEN USED TO ATTEMPT TO PASS THROUGH THE CLOT, BUT THE TIP OF IT WAS NOTED TO HAVE BROKEN OFF DURING THE RETRIEVAL ATTEMPT. AFTER ALL ATTEMPTS TO CAPTURE THE CLOT AND BROKEN PIECES OF DEVICES, OCCLUSION WAS NOTED BETWEEN THE ICA AND M1. THE PATIENT PASSED AWAY WITHIN 24 HOURS AFTER THE PROCEDURE AND ACCORDING TO THE PHYSICIAN, THE CAUSE OF DEATH WAS ATTRIBUTED TO THE BROKEN DISTAL PORTION OF THE ZOOM 55 RESULTING IN RE-OCCLUSION. | THE ZOOM 55 HAS NOT BEEN RETURNED FOR INVESTIGATION AS IT IS BEING RETAINED BY THE HOSPITAL RISK MANAGEMENT DEPARTMENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY ISSUES PERTAINING TO DESIGN, MANUFACTURING, OR QUALITY. AS THE DEVICE WAS NOT RETURNED THE EXACT ROOT CAUSE OF THE SHAFT BREAK COULD NOT BE DETERMINED. PER USER FACILITY REPORT (MW5157834) FILED BY THE HOSPITAL, "AN ARISTOTLE 24 GUIDEWIRE USED TO TRY TO OBTAIN THE BROKEN TIP ALSO FRACTURED AND CONTRIBUTED TO THE OCCLUSION." IT IS THEREFORE UNKNOWN WHETHER THE CAUSE OF THE RE-OCCLUSION IS ATTRIBUTABLE TO THE ZOOM 55, THE GUIDEWIRE TIP, OR A COMBINATION OF BOTH.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT DECEASED DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA / VOMITING, COPD. THE PATIENT PASSED AWAY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER,DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO EYE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY AND LUNG DISEASE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
MX40 1.4 GHZ SMART HOPPING | K113125 | IT WAS REPORTED THE STAFF DID NOT HEAR THE ALARM FOR ASYSTOLE AT THE TIME OF DEATH ON A PACED DO NOT RESUSCITATE (DNR) PATIENT. THE FACILITY BIOMEDICAL ENGINEER INDICATED PACER DETECTION WAS TURNED ON AT THE TIME OF THE EVENT. AN INITIAL REVIEW OF THE CLINICAL AUDIT LOG BY REMOTE SERVICE REVEALED THE PATIENT HAD A PROGRESSION OF ARRHYTHMIA ALARMS WHICH ULTIMATELY DETERIORATED TO ASYSTOLE. THE HOSPITAL CSU CLINICAL SUPERVISOR FURTHER CLARIFIED THE ALLEGATION TO POINT TO A DISCREPANCY BETWEEN THE PHILIPS DEFIBRILLATOR MONITOR AND THE CENTRAL MONITORING SYSTEM. THE CODE BLUE LOG STARTS WITH ASYSTOLE DOCUMENTED AT 22:10, AS CONFIRMED BY THE PHILIPS DEFIBRILLATOR MONITOR. THE CENTRAL MONITORING SYSTEM DID NOT DETECT THE LETHAL ARRHYTHMIA AND DID NOT ALARM UNTIL 22:13. THE PHILPS FIELD SERVICE ENGINEER AND THE HOSPITAL BIOMEDICAL ENGINEER ARE CURRENTLY INVESTIGATING WHETHER THE DEFIBRILLATOR'S TIME IS SYNCED WITH THE PHILIPS CENTRAL MONITORING SYSTEM. THE INVESTIGATION IS IN PROGRESS. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
HEART LUNG MACHINE | K112360 | COMPLAINT ID# (B)(4). | IT WAS INITIALLY REPORTED THAT THE CUSTOMER¿S TEAM HAD PRIMED AN HLS SET AND HAD IT ATTACHED TO A CARDIOHELP ON SATURDAY, (B)(6). THE TREATMENT WAS INITIATED WITH THIS PRIMED HLS SET AND CARDIOHELP SOMETIME ON TUESDAY, (B)(6). THE PATIENT WAS DOING WELL ON THE CARDIOHELP. IT WAS NOTICED THAT THE INTERNAL PRESSURE (PINT) OF THE HLS SET WAS REALLY HIGH (READING 698 MMHG PRESSURE). THE ARTERIAL PRESSURE (PART) WAS FAIRLY NORMAL AT 95 MMHG PRESSURE. THIS HIGH PINT MINUS THE PART IS WHAT GIVES A HIGH DELTA PRESSURE, WHICH WAS SHOWING AT 601MMHG. THE CUSTOMER INFORMED THAT THE INTERNAL SENSOR PORT ON THE HLS SET APPEARED TO HAVE SALT-LIKE CRYSTALS AROUND SOME OF THE METAL PINS INSIDE THE PORT. SHE ALSO SAID THAT THE END OF THE INTERNAL SENSOR CABLE FOR DISPOSABLES OF THE CARDIOHELP LOOKS SOMEWHAT CORRODED, MOST LIKELY FROM WHERE IT HAS BEEN IN CONTACT WITH THIS SALT-LIKE SUBSTANCE FOR THE PAST 11 DAYS. THE CUSTOMER USES PLASMALYTE FLUID AS THEIR PRIMING FLUID WHEN THEY PRIME THEIR HLS SETS PRIOR TO USING AS THERAPY ON A PATIENT. IT IS THE CUSTOMER¿S BEST ASSUMPTION THAT SOME OF THIS PLASMALYTE FLUID DRIPPED OR WAS SPILLED INTO THIS INTERNAL SENSOR PORT OF THE HLS SET AND THEN THE SENSOR CABLE TO PLUGGED INTO THIS SET FOR A NUMBER OF DAYS. THE CUSTOMER BELIEVES THIS COULD HAVE CAUSED THE FORMATION OF THESE SALT-LIKE CRYSTALS AND THIS MATERIAL COULD VERY LIKELY CAUSE THE CORROSION ON THE END OF THE INTERNAL SENSOR CABLE. THE AFFECTED CARDIOHELP DEVICE WILL BE INVESTIGATED IN COMPLAINT# (B)(4). ON (B)(6) 2024 THE INFORMATION WAS RECEIVED THAT THE PATIENT EXPIRED DUE TO A SEPSIS A FEW DAYS AFTER HE WAS WEANED OFF AND DECANNULATED. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
ARISTOTLE 24 | K231954 | ON 22 JUL 2024, SCIENTIA VASCULAR WAS NOTIFIED THAT AN A24-200-002 GUIDEWIRE FRACTURED DURING A NEUROVASCULAR PROCEDURE, ATTEMPTS TO GATHER MORE INFORMATION BETWEEN THE SCIENTIA VASCULAR SALES REPRESENTATIVE AND THE HOSPITAL OCCUR, AND NO ADDITIONAL INFORMATION WAS PROVIDED. ON 26 JUL 2024, DETAILS OF THE EVENT WERE PROVIDED TO SCIENTIA VASCULAR. THE EVENT WAS DESCRIBED AS FOLLOWS: "DURING MECHANICAL THROMBECTOMY FOR ACUTE STROKE, THE DISTAL TIP OF THE REPERFUSION CATHETER TIP SEPARATED. AFTER SUCCESSFUL RECANALIZATION, IT APPEARED THAT THE CATHETER TIP WAS RETAINED IN THE THROMBOSIS AND ULTIMATELY MIGRATED; RESULTED IN M1 OCCLUSION. AN ARISTOTLE 24 GUIDEWIRE WAS USED TO TRY TO OBTAIN THE BROKEN CATHETER TIP, THE GUIDEWIRE ALSO FRACTURED AND CONTRIBUTED TO THE OCCLUSION. OCCLUSION COULD NOT BE CLEARED. MANY ATTEMPTS TO RETRIEVE CATHETER TIP AND GUIDEWIRE WITH A VARIETY OF STRATEGIES WERE UNSUCCESSFUL." THE COMPLAINT GUIDEWIRE WAS NOT RETURNED TO SCIENTIA VASCULAR FOR EVALUATION, DESPITE SALES REPRESENTATIVE ATTEMPTS TO COLLECT THE GUIDEWIRE. | FROM THE REVIEW OF THE MANUFACTURING LOT HISTORY RECORD, THERE IS NO INDICATION THAT THE BREAK OCCURRED DUE TO A MANUFACTURING DEFECT, AND IT CONFIRMS THE LOT MET THE RELEASING CRITERIA. THE COMPLAINT DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION. FOR THESE REASONS, NO RESULTS WERE OBTAINED.
HEART LUNG MACHINE | K133598 | COMPLAINT ID (B)(4). | A GETINGE FIELD SERVICE TECHNICIAN (FST) FINISHED THE INVESTIGATION ON 2024-11-06. THE HLS CABLE HAS CORROSION IN THE FEMALE CONNECTOR AND COULD NOT BE PLUGGED INTO THE PORT. THE HLS CABLE WAS REPLACED. THE DEVICE WAS RUNNING FOR ONE MONTH WITHOUT ANY FAILURES. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON 2024-08-20 FOR THE PERIOD OF 2018-04-06 TO 2024-07-26. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2018-04-06. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE OF A RENTAL DREAMSTATION AUTO CPAP DEVICE. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
CARDIOSAVE HYBRID, TYPE B PLUG | K112372 | N/A. | UPDATED DATA: B4, G3, G6, H2, H10, H11. CORRECTED DATA: H1, H2, H6(CLINICAL CODE).
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH, LIVER CANCER. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2024 WHILE UNDERGOING CCPD THERAPY. DURING INTAKE IT WAS REPORTED THE PATIENT¿S EXPIRATION WAS UNRELATED TO PD THERAPY. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME ON (B)(6) 2024 WHILE UNDERGOING CCPD THERAPY. THE PATIENTS¿ SPOUSE REPORTEDLY DISCOVERED HER HUSBAND HAD PASSED IN HIS SLEEP AFTER SHE AROSE ¿DURING THE MIDDLE OF THE NIGHT¿ ON (B)(6) 2024. DESPITE THE PRIMARY AND SECONDARY CAUSE OF DEATH BEING UNKNOWN, THE PDRN REPORTED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO PD THERAPY. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT¿S EXPIRATION OCCURRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS CURRENTLY UNKNOWN; THEREFORE, CAUSALITY CANNOT BE FIRMLY ESTABLISHED. HOWEVER, DURING FOLLOW-UP THE PDRN STATED THE PATIENT¿S DEATH WAS UNRELATED TO PD THERAPY, NOR WAS THERE ANY REPORT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY OCCURRED. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
| K181166 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A TRILOGY EVO VENTILATOR LOST POWER AND SWITCHED OFF FOR AN UNKNOWN PERIOD OF TIME WITHOUT EMITTING AN AUDIBLE ALARM ON A VENTILATOR DEPENDENT TRACHEOSTOMY PATIENT. THE VENTILATOR WAS SET TO AC-PC MODE, AVAPS ON. THE VENTILATOR WAS REPORTED TO BE PLUGGED IN TO A POWER SOURCE AT THE TIME OF THE INCIDENT. THERE WAS NO ONE PRESENT DURING THE TIME OF THE REPORTED OCCURRENCE. THE PATIENT'S FAMILY MEMBER LATER OBSERVED THAT THE VENTILATOR WAS NOT OPERATING AND INSTALLED A NEW EXTERNAL BATTERY WHICH RESTORED THE DEVICE TO NORMAL OPERATION. THE PATIENT WAS REPORTED TO SUFFER CARDIAC AND RESPIRATORY ARREST AND WAS ADMITTED TO THE CRITICAL CARE UNIT. THE PATIENT IS REPORTED TO BE OFF SEDATION AND WAKING AT THIS TIME. | RESULTS OF LAB INVESTIGATION 11 FEB 2025: THE PRODUCT INVESTIGATION LAB (PIL) RECEIVED A TRILOGY EVO GREAT BRITAIN DEVICE WITH SERIAL NUMBER (B)(6) WITH THE ALLEGATION OF THE VENTILATOR LOSING POWER AND SWITCHED OFF ON THE PATIENT. THE PIL VISUALLY INSPECTED THE TRILOGY EVO FOR VISUAL DEFECTS AND FOUND NONE. NOTE, THE DEVICE WAS RETURNED WITH BOTH FILTERS INSTALLED. THE PIL RETRIEVED AND REVIEWED THE EVENT LOG FROM THE DEVICE AND PER THE EVENT LOG, THE DEVICE WAS SHUT DOWN ON (B)(6) 2024 AT 9:34 AND THEN AC POWER WAS DISCONNECTED AT 12:31 ON (B)(6) 2024. THE DEVICE WAS THEN POWERED BACK ON AT 19:06 AND THE DEVICE SHUT DOWN AND ALARMED FOR POWER FAIL. AC POWER WAS THEN CONNECTED TO THE DEVICE AT 00:26 ON (B)(6) 2024 AND THE DEVICE WAS POWERED BACK ON. AT 01:06 ON (B)(6) 2024, THE DETACHABLE BATTERY WAS REMOVED, AND ANOTHER DETACHABLE BATTERY WAS INSTALLED. AC POWER WAS THEN DISCONNECTED AT 01:08 AND THEN RECONNECTED AT 01:17 ON (B)(6) 2024. NOTE: ALL DATES AND TIMES RE IN UTC TIME. THE PIL NOTED THAT THE SOFTWARE VERSION ON THE DEVICE IS 1/06/2010. NOTE, THE SOFTWARE WAS UPGRADED FROM 1.06.10 AFTER THE DATE OF THE INCIDENT. THE PIL NOTED THAT ACCORDING TO THE EVENT LOG, THE DEVICE SETTINGS WERE CHANGED AFTER THE INCIDENT AND THE DEVICE SETTINGS WHEN RECEIVED AT THE PIL WERE AS FOLLOWS: MODE: S/T, CIRCUIT TYPE: PASSIVE, CIRCUIT SIZE: ADULT, ACTIVE HUMIDIFICATION: OFF, EPAP: 5 CMH2O, IPAP: 15 CMH2O, INSP. TIME: 1.5 S, BREATH RATE: 15 BPM, TRIGGER TYPE: AUTO-TRAK, TRIGGER SENS.: AUTO, RISE TIME: 2. THE PIL THEN STARTED THERAPY UTILIZING THE EXISTING DEVICE SETTINGS WITH A TEST LUNG. THE PIL RAN THERAPY FOR APPROXIMATELY 16.5 HOURS AND THE DEVICE OPERATED WITHOUT ISSUE. THE PIL THEN RETRIEVED AND REVIEWED THE EVENT LOG FROM THE DEVICE AND FOUND NOTHING OF NOTE. THE DEVICE OPERATED ON AC POWER WITHOUT ISSUE. THE PIL THEN TESTED THE DEVICE ON THE PIL TRILOGY EVO MULTIFUNCTIONAL TEST STATION AT PRETEST FOR AUDIO ALARM VERIFICATION AND POWER FAIL VERIFICATION AND THEN PERFORMED A POST TEST OF THE DEVICE AND PASSED ALL TESTING. THE PIL WAS UNABLE TO CONFIRM A FAILURE OF THE TRILOGY EVO DEVICE. THE PIL CONCLUDED THAT THE TRILOGY EVO GREAT BRITAIN DEVICE OPERATES AS DESIGNED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, EYE IRRITATION, SKIN IRRITATION. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVED INFORMATION ALLEGING CANCER, EYE IRRITATION, SKIN IRRITATION. THERE WAS NO MEDICAL INTERVENTION REQUIRED BY THE PATIENT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN BOX B5 HAS BEEN CORRECTED TO REFLECT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, EYE IRRITATION, SKIN IRRITATION. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, DIZZINESS, HEADACHE, AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K200480 | THE MANUFACTURER RECEIVED INFORMATION PATIENT PASSED AWAY. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT CANCER. THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW, BLOWER CONTAMINATION WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER SERVICE CENTER. THERE WAS 8 ERROR CODES FOUND. THE MANUFACTURER SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT PASSED FINAL TEST.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LIVER DISEASE/TOXICITY, CANCER, LIVER CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED ASTHMA, RESPIRATORY TRACT IRRITATION, LUNG DISEASE, SQUAMOUS CELL CANCER OF LEFT LUNG, DEATH, COPD. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA, VOMITING, LIVER DISEASE/TOXICITY, CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A BIPAP PRO BIFLEX. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. IT WAS ALLEGED THAT THE PATIENT WAS DEAD. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA (NEW OR WORSENING). THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
MILLENNIUM M600 WITHOUT OPI DOMESTIC | K043006 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY DUE TO RESPIRATORY TRACT IRRITATION, CANCER, NASAL CANCER, COVID/SEVERE SEPSIS, PNEUMONIA AND RESPIRATORY FAILURE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY AND PNEUMONIA. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY; CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
S5 HEART-LUNG-MACHINE | K071318 | LIVANOVA DEUTSCHLAND RECEIVED A REPORT FROM AEMPS OF A DEATH CASE IN WHICH A S5 SYSTEM WAS USED. CUSTOMER HAS NOT NOTIFIED DIRECTLY LIVANOVA OF THE EVENT WHEN IT OCCURRED, AS HE DID NOT CONSIDER THE S5 DEVICE TO BE INVOLVED. | THROUGH FOLLOW-UP COMMUNICATION WITH THE CUSTOMER, LIVANOVA LEARNED THAT THE PATIENT WAS A YOUNG BOY AND DURING THE INTERVENTION THERE WAS A MASSIVE ENTRY OF AIR THAT CAUSED THE DEATH OF THE PATIENT. S5 WORKED PROPERLY AND NO ALARMS WENT OFF. CUSTOMER IS IN CONTACT WITH THE COMMERCIAL COMPANY THAT PROVIDED HIM WITH THE NORMOTHERMIA MODULES AS WELL AS THE CORRESPONDING CONSUMABLES. FROM THE REVIEW OF THE DATABASE, NO FURTHER EVENT WAS REPORTED ON THIS CONSOLE, THAT WAS MANUFACTURED IN 2016. A REVIEW OF THE DHR COULD NOT IDENTIFY ANY DEVIATIONS OR NONCONFORMITIES RELEVANT TO THE ISSUE. TECHNICAL SERVICE INTERVENTIONS AND PREVENTIVE MAINTENANCE ACTIVITIES WERE VERIFIED AND CONFIRMED TO HAVE BEEN CARRIED OUT ON TIME AND WITH NO ISSUE. BASED ON ALL INFORMATION COLLECTED, NO ADVERSE EVENT RELATED TO LIVANOVA MACHINE WAS CONFIRMED NOR ANY CORRELATION TO PATIENT'S DEATH WAS COMPLAINED BY CUSTOMER. THEREFORE, THERE IS NO RELATIONSHIP BETWEEN THE REPORTED ADVERSE EVENT AND LIVANOVA DEVICE. NO OTHER ACTION IS DEEMED NECESSARY. THE RISK IS IN THE ACCEPTABLE REGION. LIVANOVA MAINTAINS AND DOCUMENTS PERIODIC CUSTOMER EVENTS MONITORING PROCESS TO EVALUATE ACTIONS FOR PRODUCTS IMPROVEMENT. LIVANOVA WILL KEEP MONITORING THE MARKET.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | THE PROGRAM MANAGER FOR THE ACUTE CLINIC MADE NOTIFICATION THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE ACUTE CLINIC¿S LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PROGRAM MANAGER. THE PATIENT WAS INITIALLY IN TREATMENT ON HIS HOME LIBERTY SELECT CYCLER ON (B)(6) 2024 WHEN HE EXPERIENCED APHASIA AND STROKE. THE PATIENT WAS ADMITTED TO THE HOSPITAL. NO ADDITIONAL INFORMATION PERTAINING TO THAT EVENT WAS AVAILABLE. DURING THE HOSPITALIZATION, ON (B)(6) 2024. THE ON-CALL DIALYSIS NURSE FROM THE ACUTE CLINIC WAS CALLED BY HOSPITAL STAFF TO RESPOND TO THE PATIENT¿S ROOM. THE PATIENT NEEDED TO BE DISCONNECTED FROM TREATMENT ON THE LIBERTY SELECT CYCLER TO GO FOR A COMPUTED TOMOGRAPHY (CT) SCAN. WHEN THE NURSE ARRIVED AT THE PATIENT¿S ROOM, THE PATIENT WAS CODING. THE HOSPITAL STAFF WAS ACTIVELY RUNNING THE CODE. THEY WERE UNSUCCESSFUL IN RESUSCITATING THE PATIENT AND THE PATIENT WAS PRONOUNCED DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST LIKELY SECONDARY TO ASPIRATION PNEUMONIA VERSUS WORSENING STROKE AND QUESTIONABLE PERITONITIS. IT IS UNKNOWN WHAT CYCLE OR PHASE OF TREATMENT THE PATIENT WAS IN AT THE TIME OF THE EVENT. THE NURSE DID NOT MAKE NOTE PRIOR TO BREAKING DOWN THE CYCLER FROM THE TREATMENT. IT IS UNKNOWN IF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY WERE RELATED TO THE PATIENT DEATH, BUT IT IS NOT LIKELY PER THE PROGRAM MANAGER. THIS PATIENT HAD COMORBID CONDITIONS OF THIS PATIENT HAD COMORBID CONDITIONS OF CORONARY ARTERY DISEASE (CAD), CONGESTIVE HEART FAILURE (CHF), HYPERTENSION (HTN), DEPRESSION, ALCOHOL ABUSE, STROKE, LEFT FRONTAL INFARCT, AND HOSPITAL ACQUIRED ASPIRATION PNEUMONIA. A FOLLOW-UP WITH THE PATIENT¿S CLINIC YIELDED NO ADDITIONAL INFORMATION AS THEY WERE NOT INFORMED OF THE PATIENT¿S ADMITTING DIAGNOSIS OR THE CIRCUMSTANCES RELATED TO THE PATIENT DEATH. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS AND UTILIZATION OF THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF APHASIA AND STOKE WITH HOSPITALIZATION AND SUBSEQUENT CARDIAC ARREST AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE INITIAL HOSPITALIZATION FOR STROKE AND THE SUBSEQUENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE PATIENT WAS INITIALLY HOSPITALIZED FOR A STROKE AND SUBSEQUENTLY EXPIRED AFTER EXPERIENCING CARDIAC ARREST. THE PATIENT HAD COMORBID CONDITIONS OF HYPERTENSION, CORONARY ARTERY DISEASE (CAD), AND CONGESTIVE HEART FAILURE (CHF). CARDIOVASCULAR DISEASE IS THE LEADING CAUSE OF DEATH IN PATIENTS WITH KIDNEY FAILURE WITH THE RISK OF CARDIOVASCULAR DISEASES, INCLUDING STROKE IN PATIENTS WITH END STAGE RENAL DISEASE (ESRD) BEING 5-30 TIMES GREATER THAN THE GENERAL PUBLIC. HYPERTENSION CONTINUES TO BE THE MAJOR MODIFIABLE RISK FACTOR FOR BOTH ISCHEMIC AND HEMORRHAGIC STROKE. THE INFORMATION MADE MENTION OF POSSIBLE PERITONITIS; HOWEVER, THERE WAS NO DOCUMENTATION TO CONFIRM THAT THE PATIENT WAS EVER DIAGNOSED WITH PERITONITIS DURING THE HOSPITALIZATION. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S INITIAL HOSPITALIZATION FOR APHASIA AND STROKE AND SUBSEQUENT CARDIAC ARREST AND DEATH.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO KIDNEY DISEASE/TOXICITY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | K233318 | IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT AN AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM WAS IMPLANTED TRANSGASTRIC TO THE JEJUNUM TO TREAT A GASTRIC OUTLET OBSTRUCTION DURING AN ENDOSCOPIC ULTRASOUND-GUIDED GASTROJEJUNOSTOMY (EUS-GJ) PROCEDURE PERFORMED ON (B)(6) 2024. DURING THE PROCEDURE, INITIAL PUNCTURE WAS NOT SUCCESSFUL. HOWEVER, THE PHYSICIAN DID NOT SEE CATHETER IN THE JEJUNAL LUMEN SO THE DEVICE WAS PULLED BACK INTO THE SCOPE, AND RE-PUNCTURED IN A DIFFERENT LOCATION WITH SUCCESSFUL PLACEMENT. POST STENT PLACEMENT, PATIENT EXPERIENCED ABDOMINAL PAIN. A COMPUTED TOMOGRAPHY (CT) SCAN WAS PERFORMED, AND SHOWED NO SIGNS OF PERFORATION OR AIR IN THE ABDOMEN. PATIENT WAS LATER ADMITTED TO THE INTENSIVE CARE UNIT (ICU) FOR HYPOTENSION., AND THE PATIENT PASSED AWAY. IN THE PHYSICIAN'S ASSESSMENT, THE CAUSE OF DEATH WAS PULMONARY EMBOLISM. ADDITIONAL INFORMATION ON SEPTEMBER 17, 2024. IT WAS REPORTED THAT THE AUTOPSY OF THE PATIENT SHOWED A LARGE HEMOPERITONEUM. AS PER DR. (B)(6), THIS IS CONSISTENT WITH HIGH LIKELIHOOD OF INTRAPERITONEAL BLEEDING RELATED TO THE PROCEDURE THAT LED TO PATIENTS DEATH. | NOTE: BLOCKS B5 AND H6 (PATIENT CODES) HAVE BEEN UPDATED FOLLOWING A MEDICAL REVIEW CONFIRMING THAT THE AUTOPSY SHOWED A LARGE HEMOPERITONEUM NOT AN EMBOLISM. BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE SUSPECT DEVICE LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. BLOCK H6: IMDRF PATIENT CODE E1002 CAPTURES THE REPORTABLE EVENT OF ABDOMINAL PAIN. IMDRF PATIENT CODE E2321 CAPTURES THE REPORTABLE EVENT OF HYPOTENSION. IMDRF PATIENT CODE E2311 CAPTURES THE REPORTABLE EVENT OF DISCOMFORT. IMDRF PATIENT CODE E0506 CAPTURES THE REPORTABLE HEMOPERITONEUM. IMDRF IMPACT CODE F2202 CAPTURES THE REPORTABLE EVENT OF CRE DILATION POST STENT PLACEMENT. IMDRF IMPACT CODE F2203 CAPTURES THE REPORTABLE EVENT OF CT SCAN SHOWED NO SIGN OF PERFORATION. IMDRF IMPACT CODE F0801 CAPTURES THE REPORTABLE EVENT OF PATIENT WAS LATER ADMITTED TO ICU. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT PASSED AWAY. IMDRF IMPACT CODE F12 CAPTURES THE REPORTABLE EVENT OF SERIOUS INJURY.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, LUNG DISEASE AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER AND KIDNEY DISEASE/TOXICITY. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION KIDNEY DISEASE, CANCER AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
HIGH FLOW INSUFFLATION UNIT | K122180 | IT WAS REPORTED THAT DURING A LAPAROSCOPIC UROLOGICAL SURGERY USING THE HIGH FLOW INSUFFLATION UNIT, THE PNEUMOPERITONEUM PRESSURE FLUCTUATED SIGNIFICANTLY WITHOUT ANY CLEAR CAUSE. THE DEVICE ALARM SOUNDED, AND THE PRESSURE ROSE. THEREAFTER, THE PATIENT WENT INTO CARDIAC ARREST AND IS HOSPITALIZED IN THE INTENSIVE CARE UNIT. THE PROCEDURE COULD NOT BE COMPLETED. AFTER THE MALFUNCTION, THE PRODUCT WAS NOT REPLACED. THE PHYSICIAN¿S OPINIONATED THAT DURING THE PROCEDURE (LAPAROSCOPIC TOTAL NEPHRECTOMY) IT WAS DISCOVERED THAT THE PATIENT DEVELOPED A PULMONARY EMBOLISM, CAUSING THE PNEUMOPERITONEUM PRESSURE TO BECOME UNSTABLE AND THEREFORE INCREASE PRESSURE ON THE DIAPHRAGM FROM THE LUNG SIDE. THEREFORE, THE PHYSICIAN BELIEVED THAT THE CAUSE WAS NOT THE DEVICE MALFUNCTION AND THAT NO SPECIAL INSPECTION WAS NECESSARY. DEVICE INSPECTION WAS CARRIED OUT USING A PNEUMOPERITONEUM CHECKER. THERE WAS NOTHING WRONG WITH THE ACTUAL PRODUCT, AND THE NUMBERS WERE WITHIN THE APPROPRIATE RANGE. ON FURTHER FOLLOW-UP THE CUSTOMER INDICATED THAT THE ABDOMINAL PRESSURE BECAME UNSTABLE ABOUT 50 MINUTES INTO THE PROCEDURE, ("DOCTOR DIDN'T DETECT LEAKAGE") THE PNEUMOPERITONEUM PRESSURE FLUCTUATED GREATLY (4MMHG TO 12MMHG) AND THE PATIENT'S BLOOD PRESSURE DROPPED AT THE SAME TIME. THE SURGERY WAS ABORTED IN PROGRESS. SHE DIED THE DAY AFTER CARDIAC ARREST ((B)(6) 2024). IT WAS BELIEVED THAT THE PATIENT HAD CONCURRENT AORTIC VALVE STENOSIS, WHICH MAY HAVE CAUSED THE ISSUE. THE PATIENT WAS SENT TO THE ICU DUE TO CARDIAC ARREST BUT WAS CONFIRMED DEAD THE NEXT DAY. IT IS UNKNOWN WHICH DRUGS WERE INFUSING AT THE TIME OF THE PROCEDURE. THE CUSTOMER THINKS THE CAUSE OF DEATH WAS A COMPLICATION OF AORTIC VALVE STENOSIS AND/OR PULMONARY EMBOLISM. NO SPECIFIC KNOWN INTERVENTIONS AND MEDICATION INFORMATION IS MADE AVAILABLE AND ONLY INFORMATION AVAILABLE WAS THAT THE PATIENT WAS HOSPITALIZED IN ICU. DEVICE SETTINGS NOTED WERE: PRESSURE: 10MMHG. RELIEF MODE: OFF. ALARM DELAY: 0 SEC. FLOW RATE: LOW 0.1, NORMAL 9, HIGH 24 (L/MIN). THE FUNCTIONAL MODE DEVICE OUTPUT WAS BETWEEN 4MMHG T0 12 MMHG. THE DEVICE WAS NOT INSPECTED BEFORE USE. THE OVERPRESSURE ALARM SOUNDED. | E1: COMPLETE ESTABLISHMENT NAME: (B)(6) HOSPITAL. THE DEVICE EVALUATION IS ONGOING. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
TELEHEALTH CENTER PRODUCT LICENSE | K211046 | THE CUSTOMER REPORTING A DELAY IN PROVIDING TREATMENT TO A PATIENT IN THE SICU IN CARDIAC ARREST, REPORTEDLY ASSOCIATED WITH AN ISSUE WHERE THE CAREGILITY EQUIPMENT WAS NOT ALLOWING MORE THAN ONE PERSON ON A CAMERA AT A TIME. THE BEDSIDE RN¿S WERE UNENGAGED AND APPARENTLY UNTRAINED IN CPR AND THE PATIENT WAS IN CARDIAC ARREST FOR AT LEAST 90 SECONDS. THE RN WAS ON CAMERA AND IMPLORING THE BEDSIDE TO GET HELP AND BEGIN CPR. THE PHYSICIAN WAS AWARE, BUT GIVEN THE LIMITATIONS OF CAREGILITY, UNABLE TO OPERATE THE CAMERA TO BEGIN CARE. IN THE MIDST OF THIS, ASSISTANT NURSE MANAGER HAD TO CALL THE MICU DIRECTOR TO ASK FOR BEDSIDE ASSISTANCE, WHICH WAS PROVIDED. INITIAL GFE RESPONSE FROM THE CUSTOMER STATES THAT, THEY ARE UNCLEAR WHETHER ANY ALARMS GENERATED BUT ¿OFTEN THOUGH CARDIAC ARREST DOES NOT TRIGGER ALARMS.¿ THE CUSTOMER WAS UNAWARE OF THE PATIENT OUTCOME BUT STATED THAT THEY WERE CRITICALLY ILL, IN CARDIAC ARREST AND THEY FEEL THAT THE DELAY IN CARE WOULD BE A CONTRIBUTING FACTOR TO A NEGATIVE OUTCOME, IF SUCH AN OUTCOME DID OCCUR. OF NOTE, IN TERMS OF THE HOSPITAL WORKFLOW AND STAFFING RATIOS, GFE RESPONSE INDICATES THE ¿RN¿S WORK A 1:50 RATIO¿ AND FINDING/GATHERING THE APPROPRIATE SUPPORT TO PERFORM CPR ¿TOOK UP SUBSTANTIAL TIME FOR BOTH AN RN AND THE ASSISTANT NURSE MANAGER. FOLLOW-UP GFE CONFIRMS THAT THERE WAS DEATH INVOLVED. NO FURTHER INFORMATION IS AVAILABLE. | SECTIONS D4, G4 AND H3 ARE UPDATED IN THIS REPORT. IN SECTION G4, K012171 WAS MENTIONED AS 510K NUMBER INSTEAD OF K211046. UPDATED AS K211046. IN SECTION D4 (B)(4) WAS MENTIONED INSTEAD OF (B)(4). UPDATED AS (B)(4). IN FIRST FOLLOW UP, SECTION H3, ¿DEVICE EVALUATED BY MANUFACTURER?¿ WAS LEFT BLANK. IT IS UPDATED AS YES. DURING THE INTERNAL REVIEW OF THE RECORD, ABOVE CHANGES WERE NOTICED. HENCE DATA IS CORRECTED AND SUBMITTED FOR UPDATES.
AERO TRACHEOBRONCHIAL STENT | K082284 | THE ACCOUNT ALLEGES THAT DURING A ROUTINE AERO TRACHEOBRONCHIAL POLYURETHANE STENT FOLLOW-UP APPOINTMENT FOR A FEMALE CANCER PATIENT CURRENTLY RECEIVING PALLIATIVE CARE DUE TO CANCER. THE AERO STENT WAS PLACED FOR A CANCER-RELATED STRICTURE APPROXIMATELY TWO WEEKS AGO WITHOUT ISSUE. THE PATIENT RECEIVED ONE TREATMENT OF RADIATION THERAPY POST-STENT IMPLANT ON AN UNKNOWN DATE. DURING THE FOLLOW-UP APPOINTMENT, THE PHYSICIAN NOTED SOME NARROWING AND WITH MINOR GRANULATED TISSUE [CONNECTIVE TISSUE] PRESENT AT THE AERO STENT SITE. THE INTERVENTIONAL PULMONOLOGIST RECOMMENDED REMOVING THE EXISTING AERO STENT AND UPSIZING IT TO A LARGER STENT FOR THE PATIENT. DURING THE STENT REMOVAL PROCEDURE, A MERIT PULMONARY BALLOON DILATATION CATHETER WAS USED TO DILATE THE INNER DIAMETER OF THE STENT UNDER FLUOROSCOPIC GUIDANCE AND AN ATTEMPT TO REMOVE THE STENT WITH A PAIR OF RAPTOR FORCEPS WAS INITIATED SUCCESSFULLY. THE ACTUAL STENT REMOVAL PROCESS WAS UNREMARKABLE. SOON AFTER THE STENT WAS REMOVED BLEEDING WAS NOTED AT THE STENT SITE. SUCTION WAS ATTEMPTED BUT UNSUCCESSFUL IN CONTROLLING THE HEMORRHAGING. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE AND A CODE BLUE WAS INITIATED. LIFE-SAVING PROTOCOLS WERE ATTEMPTED FOR THIS PATIENT WITH NO SUCCESS. THE PATIENT EXPIRED EARLY AFTERNOON [APPROX 2 PM 8-22-2024] IN THE PROCEDURE ROOM. THE PHYSICIAN STATES THE MERIT BALLOON AND AERO STENT SYSTEM DID NOT MALFUNCTION DURING THIS PROCEDURE, BOTH MEDICAL DEVICES OPERATED AS INTENDED. NO DEVICE RETURNING AT THIS TIME. | THE SUSPECT DEVICE IS NOT EXPECTED TO RETURN FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD AND COMPLAINT DATABASE COULD NOT BE PERFORMED AS A LOT NUMBER WAS NOT PROVIDED.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF KIDNEY DISEASE/TOXICITY. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | PREVIOUSLY REPORTED BY THE MANUFACTURER: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. NEW INFORMATION: CORRECTION TO BOX B, DATE OF DEATH.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. A GOOD FAITH EFFORT ATTEMPT WAS MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT THE CUSTOMER DECLINED TO RESPOND TO THE QUESTIONS AND TERMINATED THE CALL. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
TREK | K123279 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, ARRHYTHMIAS, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. | THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORD FOR THIS PRODUCT COULD NOT BE REVIEWED AND A COMPLAINT REVIEW COULD NOT BE PERFORMED BECAUSE THE PRODUCT WAS NOT RETURNED FOR EVALUATION AND THE PART AND LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECTS OF DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), TREK OTW AND MINI TREK OTW, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, A CONCLUSIVE CAUSE FOR THE REPORTED COMPLAINTS COULD NOT BE DETERMINED. ADDITIONALLY, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS LISTED IN THE REPORTED ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE ¿POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS¿ THE NC TRAVELER RX DEVICE IS CURRENTLY NOT COMMERCIALLY AVAILABLE IN THE US; HOWEVER, IT IS SIMILAR TO A DEVICE SOLD IN THE US. B2: DATE OF DEATH ESTIMATED AS (B)(6) 2024. B3: DATE OF EVENT ESTIMATED AS (B)(6) 2024. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
MINI TREK¿ | K123279 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, ARRHYTHMIAS, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. | THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORD FOR THIS PRODUCT COULD NOT BE REVIEWED AND A COMPLAINT REVIEW COULD NOT BE PERFORMED BECAUSE THE PRODUCT WAS NOT RETURNED FOR EVALUATION AND THE PART AND LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECTS OF DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), MINI TREK II OTW, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, A CONCLUSIVE CAUSE FOR THE REPORTED COMPLAINTS COULD NOT BE DETERMINED. ADDITIONALLY, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS LISTED IN THE REPORTED ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE ¿POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS.¿ THE NC TRAVELER RX DEVICE IS CURRENTLY NOT COMMERCIALLY AVAILABLE IN THE US; HOWEVER, IT IS SIMILAR TO A DEVICE SOLD IN THE US. B2: DATE OF DEATH ESTIMATED AS (B)(6) 2024. B3: DATE OF EVENT ESTIMATED AS 3/1/2024. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
NC TRAVELER CORONARY DILATATION CATHETER | K110134 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, ARRHYTHMIAS, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. | THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORD FOR THIS PRODUCT COULD NOT BE REVIEWED AND A COMPLAINT REVIEW COULD NOT BE PERFORMED BECAUSE THE PRODUCT WAS NOT RETURNED FOR EVALUATION AND THE PART AND LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECTS OF DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), NC TRAVELER, RX, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, A CONCLUSIVE CAUSE FOR THE REPORTED COMPLAINTS COULD NOT BE DETERMINED. ADDITIONALLY, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS LISTED IN THE REPORTED ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE ¿POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS¿ THE NC TRAVELER RX DEVICE IS CURRENTLY NOT COMMERCIALLY AVAILABLE IN THE US; HOWEVER, IT IS SIMILAR TO A DEVICE SOLD IN THE US. B2: DATE OF DEATH ESTIMATED AS (B)(6) 2024. B3: DATE OF EVENT ESTIMATED AS 3/1/2024. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
TRAVELER RX CORONARY DILATATION CATHETER | K103110 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, ARRHYTHMIAS, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS AND DEVICES REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS". | THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORD FOR THIS PRODUCT COULD NOT BE REVIEWED AND A COMPLAINT REVIEW COULD NOT BE PERFORMED BECAUSE THE PRODUCT WAS NOT RETURNED FOR EVALUATION AND THE PART AND LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECTS OF DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), TRAVELER RX, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, A CONCLUSIVE CAUSE FOR THE REPORTED COMPLAINTS COULD NOT BE DETERMINED. ADDITIONALLY, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS LISTED IN THE REPORTED ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE ¿POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS¿. THE NC TRAVELER RX DEVICE IS CURRENTLY NOT COMMERCIALLY AVAILABLE IN THE US; HOWEVER, IT IS SIMILAR TO A DEVICE SOLD IN THE US. B2: DATE OF DEATH ESTIMATED AS (B)(6) 2024. B3: DATE OF EVENT ESTIMATED AS (B)(6) 2024. D4: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
TREK | K103110 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, THROMBUS, HYPOTENSION, ARRHYTHMIAS, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. | THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORD FOR THIS PRODUCT COULD NOT BE REVIEWED AND A COMPLAINT REVIEW COULD NOT BE PERFORMED BECAUSE THE PRODUCT WAS NOT RETURNED FOR EVALUATION AND THE PART AND LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECT DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), TREK RX AND MINI TREK RX, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, A CONCLUSIVE CAUSE FOR THE REPORTED COMPLAINTS COULD NOT BE DETERMINED. ADDITIONALLY, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECTS, AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS AND DEVICES REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE ¿POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS¿. B2 - DATE OF DEATH: ESTIMATED AS (B)(6) 2024. B3 - DATE OF EVENT: ESTIMATED AS (B)(6) 2024. D4 - PRIMARY UDI NUMBER: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
MINI TREK | K103110 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, THROMBUS, HYPOTENSION, ARRHYTHMIAS, STENOSIS, ANGINA, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. | THE DEVICE WAS NOT RETURNED FOR EVALUATION. THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEW WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE PRODUCTION RECORD FOR THIS PRODUCT COULD NOT BE REVIEWED AND A COMPLAINT REVIEW COULD NOT BE PERFORMED BECAUSE THE PRODUCT WAS NOT RETURNED FOR EVALUATION AND THE PART AND LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), TREK RX AND MINI TREK RX, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, A CONCLUSIVE CAUSE FOR THE REPORTED COMPLAINTS COULD NOT BE DETERMINED. ADDITIONALLY, A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT(S), AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS AND DEVICES REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS". B2 - DATE OF DEATH : ESTIMATED AS (B)(6) 2024. B3 - DATE OF EVENT: ESTIMATED AS 3/1/2024. D4 - PRIMARY UDI NUMBER: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
NC TREK | K113464 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, ARRHYTHMIAS, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) DATABASE WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. A REVIEW OF PRODUCTION RECORDS AND THE COMPLAINT HANDLING DATABASE WAS NOT PERFORMED BECAUSE THE PART/LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECTS OF DEATH IS LISTED IN THE NC TREK OTW INSTRUCTION FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS AND DEVICES REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS" B2 - DATE OF DEATH : ESTIMATED AS (B)(6) 2024. B3 - DATE OF EVENT: ESTIMATED AS (B)(6) 2024. D4 - PRIMARY UDI NUMBER: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
NC TREK | K110134 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE OCCLUSION, ACUTE MYOCARDIAL INFARCTION (AMI), EMBOLISM, VESSEL DISSECTION, VESSEL PERFORATION, THROMBUS, ARRHYTHMIAS, ANGINA, AND DEATH. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, AND DEFLATION OF THE CDCS. OFF-LABEL USE WAS REPORTED FOR USE IN CHRONIC TOTAL OCCLUSION (CTO) LESIONS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) DATABASE WAS PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. A REVIEW OF PRODUCTION RECORDS AND THE COMPLAINT HANDLING DATABASE WAS NOT PERFORMED BECAUSE THE PART/LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE NC TREK INSTRUCTION FOR USE AS A KNOWN PATIENT EFFECT OF CORONARY PROCEDURES. A CONCLUSIVE CAUSE FOR THE REPORTED DEATH CANNOT BE DETERMINED. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING; THEREFORE, NO PRODUCT-RELATED CORRECTIVE ACTION WILL BE IMPLEMENTED IN THIS CASE. THE ADDITIONAL PATIENT EFFECTS, DEVICE MALFUNCTIONS AND DEVICES REPORTED IN THE PMCF ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS" B2 - DATE OF DEATH : ESTIMATED AS (B)(6) 2024. B3 - DATE OF EVENT: ESTIMATED AS (B)(6) 2024. D4 - PRIMARY UDI NUMBER: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY; CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING KIDNEY DISEASE/TOXICITY. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ILET BIONIC PANCREAS | K231485 | ON 7/30/24 A BETA BIONICS CLINICAL DIABETES SPECIALIST (CDS) BECAME AWARE THAT AN ILET USER HAD PASSED AWAY. THE DATE OF DEATH WAS NOT INITIALLY REPORTED BUT IS ESTIMATED TO HAVE OCCURRED ON (B)(6) 2024. AT THE TIME OF THIS REPORT, ATTEMPTS BY BETA BIONICS TO DETERMINE IF THE USER WAS ON THE ILET AT THE TIME OF THE EVENT AS WELL AS CAUSE OF DEATH HAVE BEEN UNSUCCESSFUL. THE CDS DID REPORT THE USER WAS 65 YEARS OLD AND HAD AND WAS RECEIVING TREATMENT FOR ADVANCED BREAST CANCER WITH LIVER AND PANCREAS METASTASES. | NO PRODUCT WAS RETURNED FOR EVALUATION. DATA UPLOADS FROM THE ILET WERE NOT COMPLETED AFTER (B)(6) 2024, PRIOR TO THE PRESUMED EVENT DATE; CGM GLUCOSE DATA WERE THEREFORE NOT AVAILABLE TO CONFIRM THE EVENT, AND ILET PERFORMANCE DURING THE EVENT COULD NOT BE EVALUATED. NO PRODUCT PERFORMANCE ISSUES CAN BE IDENTIFIED. IF THE PRODUCT IS RECEIVED AT A LATER DATE, OR THE ILET DATA IS UPLOADED, THE COMPLAINT WILL BE REOPENED AND INVESTIGATED ACCORDINGLY. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. REVIEW OF THE AVAILABLE DEVICE DATA SHOWED THE ILET WAS OPERATING AS INTENDED AT THAT TIME. WITHOUT MORE SPECIFIC INFORMATION INCLUDING THE EVENT DATE, CAUSE OF DEATH, AND IF THE USER WAS WEARING THE ILET AT THE TIME, AND UP TO DATE DEVICE DATA, THE PERFORMANCE OF THE DEVICE DURING THIS EVENT CANNOT BE EVALUATED AND THE CAUSE OF THE EVENT CANNOT BE DETERMINED. IT IS LIKELY THE USER'S DEATH WAS RELATED TO THEIR CANCER DIAGNOSIS.
BIPAP AUTO BIFLEX | K091319 | PATIENT'S DAUGHTER CALLED IN REGARDING HER FATHER'S DEVICES THAT SHE WANTS TO RETURN TO PHILIPS BECAUSE HER FATHER PASSED AWAY LAST 2022. CALLER REQUESTED A RETURN SHIPPING LABEL THROUGH EMAIL ADDRESS. CALLER WANTS TO REGISTER THE SYSTEM ONE DEVICE TO RECALL FOR FINANCIAL REIMBURSEMENT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH WHILE USING THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT EXPIRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT EXPIRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
NC TREK NEO¿ | K220634 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT PROACTIVELY COLLECTED AND EVALUATED DATA FROM SEVERAL SOURCES, SUCH AS IN-HOSPITAL OUTCOMES AND PHYSICIAN SURVEYS, TO CONFIRM THE SAFETY AND PERFORMANCE OF ABBOTT'S NC TREK NEO CORONARY DILATATION CATHETERS (CDCS). KEY SAFETY OUTCOMES COLLECTED WERE PERFORATION, DISSECTION, EMBOLISM, ARRHYTHMIAS, MI, OCCLUSION, THROMBUS, AND DEATH. THE WEIGHTED AVERAGE OF THE SAFETY COMPOSITE RANGED FROM 0% TO (B)(4)%. DEVICE PERFORMANCE WAS DEFINED AS SUCCESSFUL DELIVERY, INFLATION, DEFLATION, AND WITHDRAWAL OF THE CDCS. THERE WERE ZERO INSTANCES OF BALLOON DEFLATION ISSUES FOR THIS DEVICE. THERE WAS OFF-LABEL USAGE AS A SMALL PERCENTAGE OF THIS DEVICE WAS NOT USED TO TREAT CORONARY VESSELS. DETAILS ARE LISTED IN THE ATTACHED REPORT, TITLED POST-MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CORONARY DILATATION CATHETERS. PLEASE SEE THE ATTACHED POST MARKET CLINICAL FOLLOW UP EVALUATION REPORT FOR SPECIFIC INFORMATION. | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. B2 - DATE OF DEATH: ESTIMATED AS (B)(6) 2024. B3 - DATE OF EVENT: ESTIMATED AS 03/01/2024. D4 - PRIMARY UDI NUMBER: THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL DEVICE REFERENCED IN B5 IS FILED UNDER A SEPARATE MEDWATCH REPORT NUMBER. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS LISTED IN THE REPORT ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. LITERATURE ATTACHMENT: ARTICLE TITLE "POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT NC TREK NEO CORONARY DILATATION CATHETER".
HEMODIALYSIS BLOODLINES | K080807 | AS REPORTED BY THE USER FACILITY: CUSTOMER REPROTS THAT ON (B)(6), A PATIENT WAS HOOKED UP TO THE DIALOG FOR DIALYSIS AND 15 MINUTES INTO TREATMENT THE PATIENTS VENOUS BLOODLINE CAME DETACHED FROM HIS CVC. THE PATIENT EXSANGUINATED AND CODED. THE PATIENT WAS BRIEFLY RESCUSUTATED AND TRANSPORTED TO THE HOSPITAL, WHICH WAS ABOUT 1/4 MILE AWAY FROM THE DIALYSIS CENTER, WHERE HE CODED A SECOND TIME AND EXPIRED. PATIENT WAS 70 YEAR OLD MALE. HE WAS NOT NEW TO DIALYSIS, BUT WAS ALSO NOT A LONG-TERM PATIENT. HIS ACCESS WAS A CVC THE FOLLOWING EMAIL RESPONSES WERE RECEIVED BY THE CUSTOMER ON 07AUGUST2024: 1. HOW WAS THE INCIDENT DISCOVERED? DISCOVERED BY PCT WORKING WITH PATIENT. 2. WHAT HOSPITAL WAS THE PATIENT TRANSPORTED TO? (B)(6) HOSPITAL. 3. WHAT MEDICAL INTERVENTIONS WERE NEEDED AT THE DIALYSIS CENTER AND WHAT WAS NEEDED LATER AT THE HOSPITAL? CPR< CLAMPING BLOOD LINES / INTUBATION AND BLOOD ADMINISTRATION. 4. DID THE MACHINE ALARM AT THE TIME OF THE EVENT? YES - LOGS SHOWED A VENOUS ALARM WAS SILENCED AND MANUAL BLOOD PRESSURE WAS STARTED. 5. CAN ADDITIONAL INFORMATION BE PROVIDED ON THE BLOODLINE DETACHMENT. SPECIFICALLY, WHERE ON THE BLOODLINE DID THE DETACHMENT OCCUR? CONNECTION BETWEEN CVC VENOUS PORT AND BLOODLINES NOT TIGHT. 6. WAS THERE ANY VISUAL DAMAGE OR ISSUE IDENTIFIED WITH THE BLOODLINE. NO. 7. IS THE BLOODLINE BATCH AVAILABLE? IF SO, CAN IT BE PROVIDED? YES - 00619221B4. 8. WAS AN AUTOPSY PERFORMED? IF SO, CAN IT BE PROVIDED? YES, DO NOT HAVE A COPY. 9. WAS THERE AN OFFICIAL CAUSE OF DEATH? YES, DO NOT HAVE ON FILE. THE FOLLWING EMAIL RESPONSE WAS RECEIVED BY THE CUSTOMER ON 08AUGUST2024: YES, THE LETTER B WAS MISINTERPRETED. 0061922184 IS THE PROPER NUMBER. | THIS REPORT HAS BEEN IDENTIFIED AS B. BRAUN MEDICAL INTERNAL REPORT NUMBER (B)(4). THE INVESTIGATION IS ONGOING AT THIS TIME. A FOLLOW-UP WILL BE SUBMITTED WHEN THE INVESTIGATION RESULTS BECOME AVAILABLE. NOTE: ORIGINAL REPORT WAS SENT UNDER 2521402-2024-00086. AFTER RECEIVING THE CORRECT LOT NUMBER FROM THE CUSTOMER VIA EMAIL ON 08AUG2024, IT WAS DETERMINED THAT THE INCORRECT MANUFACTURING REPORT # WAS USED. THIS REPORT CREATED BASED ON THE NEW INFORMATION.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND THE PATIENT EXPIRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION 2 ADVANCE AUTO CPAP | K200480 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE DREAMSTATION 2 ADVANCE AUTO CPAP HAD STATES PATIENT HAS PASSED AWAY SO SHIPPING LABEL IS REQUIRED TO RETURN THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, AND THE PATIENT EXPIRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE TECHNICIAN FOUND NO VISUAL EVIDENCE OF FOAM PARTICLES. THERE WAS NO MALFUNCTION NOTED. THERE WAS EVIDENCE OF ENVIRONMENTAL CONTAMINATION. NO COMPONENTS WERE REPLACED. THE DEVICE WAS SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION BIPAP AUTOSV | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED. BOX H WAS UPDATED IN THIS REPORTED.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
AIRSEAL | K190303 | THIS COMPLAINT WAS CREATED FROM AN EMAIL RECEIVED BY OUR LEGAL DEPARTMENT. THE DEVICE AS-IFS1, AIRSEAL IFS, 110V, WAS USED DURING A ROBOTIC SALPINGO-OOPHORECTOMY ON (B)(6) 2024. THE REPORTER STATED ¿IN THE COURSE OF INSERTING AN INTUITIVE SURGICAL EXTRA-LONG TROCAR, SHE NICKED THE VENA CAVA ON INSERTION, APPARENTLY RESULTING IN THE PATIENT¿S DEATH BY EXSANGUINATION. THE DOCTOR HAS ASSERTED THAT THE IFS SYSTEM WAS AT FAULT FOR INSUFFICIENTLY INSUFFLATING THE ABDOMINAL CAVITY. IT SHOULD BE NOTED THAT IN THE TWO DAYS AFTER THE SURGERY, THE MEDICAL CENTER CONTINUED TO USE THE SUBJECT DEVICE ON MULTIPLE SURGERIES, WHERE IT PERFORMED SAFELY AND WITHOUT INCIDENT.¿. FURTHER ASSESSMENT WAS REQUESTED BUT TO DATE NO FURTHER INFORMATION IS AVAILABLE. THIS REPORT IS BEING RAISED ON THE REPORTED DEATH OF A PATIENT. | THE REPORTED DEVICE HAS BEEN RETURNED TO CONMED; HOWEVER, THE EVALUATION OF THE DEVICE HAS NOT BEEN COMPLETED. A SUPPLEMENTAL AND FINAL REPORT WILL BE FILED FOLLOWING THE COMPLETION OF THE DEVICE EVALUATION AND COMPLAINT INVESTIGATION. WE WILL CONTINUE TO MONITOR FOR TRENDS THROUGH THE COMPLAINT SYSTEM TO ASSURE PATIENT SAFETY.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NAUSEA/VOMITING, LIVER DISEASE/TOXICITY, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
V-18 CONTROL WIRE | K033742 | IT WAS REPORTED THAT THE PATIENT DIED. THE PATIENT UNDERWENT TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR). VASCULAR ACCESS WAS OBTAINED VIA THE RIGHT FEMORAL ACCESS WHICH WAS REVEALED TO BE BORDERLINE AND MODERATELY CALCIFIED. THE LESION ANATOMY WAS LOCATED IN THE FEMORAL AORTIC-ILIAC COMPLEX WITH MODERATE STENOSIS, AND MODERATE CALCIFICATION ON THE LEAFLETS. IT WAS ALSO NOTED THAT THERE WAS MODERATE TORTUOSITY IN THE ILIO-FEMORAL ARTERY, CALCIFIED AORTIC BIFURCATION, MULTIPLE LOCATIONS OF CALCIFICATION AND ATHEROSCLEROTIC PLAQUE THROUGHOUT THE ABDOMINAL AORTA. FOLLOWING PRE-DILATATION WITH A NON- BOSTON SCIENTIFIC (BSC) BALLOON, A 300CM, 8CM V-18 CONTROL WIRE WAS CONTRALATERALLY ADVANCED AND POSITIONED IN THE RIGHT SUPERFICIAL FEMORAL ARTERY (SFA). A 14F ISLEEVE INTRODUCER SET WAS POSITIONED SUCCESSFULLY AND WITHOUT COMPLICATION, AND THE PROSTHESIS WAS SUCCESSFULLY IMPLANTED. HOWEVER, DURING WITHDRAWAL, THE ACURATE NEO2 DELIVERY SYSTEM WAS DIFFICULT TO REMOVE. THE PHYSICIAN WAS ADVISED TO VISUALIZE THE REMOVAL OF THE DELIVERY SYSTEM VIA THE ISLEEVE, BUT PREFERRED TO VISUALIZE THE SAFARI GUIDEWIRE INSIDE THE LEFT VENTRICLE. WHEN REPORTING THE DIFFICULTY IN REMOVING THE DELIVERY SYSTEM, THE IMAGE WAS MOVED TO THE FEMORAL AORTIC-ILIAC COMPLEX, OBSERVING A KINK ON THE V18 GUIDEWIRE OUTSIDE THE INTRODUCER, WHICH MAY HAVE INJURED THE ARTERIAL WALL. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, AND A RUPTURE OF THE ILIAC ARTERY WAS SEEN, WITH BLEEDING INTO THE RETROPERITONEUM. CONTRALATERAL ELEVATION OF THE MUSTANG VASCULAR BALLOON WAS PERFORMED TO CONTAIN BLEEDING. THE SENTINEL WAS REMOVED THROUGH THE BRACHIALIS WHERE THE CONTRAST INJECTIONS WERE PERFORMED. THE ISLEEVE WAS REMOVED ALONG WITH THE DELIVERY SYSTEM AS A SINGLE UNIT. ANOTHER BALLOON WAS PLACED IN THE COMMON ILIAC UP TO THE BIFURCATION AND THEN AN 8.0 AND 6.0 NON-BSC STENTS WERE PLACED, BUT THE PATIENT CONTINUED TO BLEED. AFTER THE BALLOON WAS INFLATED AT THE BLEEDING SITE, THE PATIENT WAS SENT TO EXPLORATORY LAPAROTOMY TO REPAIR THE ARTERIAL VESSEL, BUT THE PATIENT DIED. | DEVICE EVAL BY MFR: THE DEVICE WAS NOT RETURNED. HOWEVER, A PHOTO WAS PROVIDED FROM THE HEALTHCARE FACILITY. IT IS NOT POSSIBLE TO OBSERVE THE V-18 CONTROL WIRE. IT IS A NON-BOSTON SCIENTIFIC (BSC) DEVICE WHAT IS SHOWN IN THE PHOTOS.
ZOOM REPERFUSION CATHETER | K211476 | A 36-YEAR-OLD FEMALE PATIENT WAS TREATED FOR AN OCCLUSION AT THE LEFT M2 SEGMENT. ACCESS WAS OBTAINED USING A GUIDEWIRE AND A ZOOM 88. A ZOOM 71 AND ZOOM 35 WERE ADVANCED THROUGH ZOOM 88 TO THE CLOT, AND ASPIRATION WAS APPLIED. THE PHYSICIAN COMPLETED FOUR PASSES USING THE SAME CATHETER COMBINATION WITHOUT REPORTING ANY ISSUES. DURING THE PROCEDURE, IMAGING SHOWED EXTRAVASATION, AND THE PHYSICIAN STATED A VESSEL DISSECTION OCCURRED. ACCORDING TO THE PHYSICIAN IT WAS DIFFICULT TO IDENTIFY THE EXACT ANATOMICAL LOCATION OF THE DISSECTION. THE PHYSICIAN ENDED THE CASE WITHOUT PERFORMING ANY MEDICAL INTERVENTION. THE ZOOM 71 CATHETER AND THE GUIDEWIRE WERE REPORTED TO BE PRESENT AT THE LOCATION OF THE VESSEL DISSECTION. THE PHYSICIAN STATED THAT THE ZOOM 71 DID NOT MALFUNCTION. THE PATIENT PASSED AWAY TWO DAYS LATER POST-PROCEDURE AND ACCORDING TO THE PHYSICIAN, THE CAUSE OF DEATH WAS ATTRIBUTED TO THE EXTRAVASATION. | THE DEVICE WAS DISCARDED AFTER USE AND THEREFORE NOT AVAILABLE FOR RETURN AND INVESTIGATION. THERE WERE NO DEVICE DEFICIENCIES REPORTED RELATED TO THE ZOOM 71. BASED ON THE INFORMATION PROVIDED AND WITH LIMITED CASE IMAGES, THE EXACT CAUSE OF THE VESSEL DISSECTION IS UNKNOWN. THE RELATIONSHIP WITH THE ZOOM 71 DEVICE TO THE VESSEL DISSECTION COULD NOT BE ESTABLISHED. IT IS NOT CLEAR IF THE DISSECTION WAS CAUSED BY THE ZOOM 71, THE GUIDEWIRE, OR OTHER REASONS. THE MANUFACTURING RECORDS FOR THE ZOOM 71 WERE REVIEWED AND DEMONSTRATED THAT THE PRODUCT MET ALL THE DESIGN AND MANUFACTURING SPECIFICATIONS.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION NOSE IRRITATION, DEATH, CANCER. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
SAFARI2¿ | K151244 | EVENT DESCRIPTION: ACURATE NEO2 IMPLANT. PATIENT HAD SMALL LEFT VENTRICLE AND SEVERE LEFT VENTRICLE HYPERTROPHY WITH SEVERE ECCENTRIC CALCIFICATION OF THE AORTIC ANNULUS. PATIENT FRAIL PRIOR TO PROCEDURE. COMORBIDITIES: MYELOID LEUKEMIA WITH ALLOGENEIC BONE MARROW TRANSPLANTATION (2014), REMISSION. XS SAFARI2 GUIDEWIRE SELECTED FOR USE DUE TO THE CHALLENGING LEFT VENTRICLE ANATOMY, WHICH THEY ARE NOT SO USED TO USING IN THEIR PRACTICE. LEFT FEMORAL ARTERY FOR DELIVERY SYSTEM, RIGHT RADIAL ARTERY FOR PIGTAIL, RIGHT FEMORAL VEIN FOR PACING ON THE WIRE. PREDILATATION WITH NUCLEUS 22X40MM UNDER PACING. SIZE MEDIUM ACURATE NEO2 VALVE PREPPED AND LOADED ONTO ACURATE NEO2 TF DELIVERY SYSTEM WITHOUT ISSUES. THE SIZE MEDIUM ACURATE NEO2 VALVE WAS IN A HIGH POSITION AND THERE WAS SUBOPTIMAL GUIDEWIRE POSITIONING WITH NO SPACE FOR THE NOSE CONE OF THE ACURATE NEO2 TF DELIVERY SYSTEM TO TRAVEL DURING DEPLOYMENT, AND THE SIZE MEDIUM ACURATE NEO2 VALVE WAS FULLY RELEASED IN A HIGH LOCATION WITH MODERATE AORTIC REGURGITATION. POST-DIL NOT PERFORMED DUE TO THE VALVE POSITION. THE LEFT FEMORAL ARTERY WAS CLOSED WITH ONE PROSTYLE AND 8F ANGIOSEAL (NO BLEEDING AFTER SUTURES). FOUR HOURS POST-PROCEDURE, RETROPERITONEAL BLEEDING WAS IDENTIFIED VIA CT SCAN; HOWEVER, THE BLEEDING WAS NOT SEEN INVASIVELY AFTER THIS CT SCAN. THE PATIENT THEN BECAME UNSTABLE. AN ECHO REVEALED TAMPONADE WHICH WAS DRAINED PERCUTANEOUSLY. AFTER 15 MINUTES, THE DRAIN WAS NOT FUNCTIONING AND THE PATIENT DETERIORATED. SURGICAL INTERVENTION WAS INITIATED. THE CHEST WAS OPENED AND A PERFORATION WAS IDENTIFIED IN THE LATERAL LEFT VENTRICLE WALL. THE PERFORATION WAS REPAIRED WITH SUTURES. IN DISCUSSION WITH THE PHYSICIAN, WE NOTICED THE LEVEL OF PUSH OF THE XS SAFARI WIRE IS INCREASED BECAUSE OF CHALLENGES IN MAINTAINING THE VALVE POSITION STABLE DURING RELEASE. THIS COULD BE THE ROOT CAUSE TO THE LEFT VENTRICLE TRAUMA THAT LEAD TO THE TAMPONADE. AS OF (B)(6) 2024 PATIENT REMAINS HOSPITALIZED DUE TO POST-PROCEDURE INFECTION. THE PHYSICIAN NOTED THAT THE PATIENT'S PRE-EXISTING IMMUNO-DEPRESSED CONDITION IS PLAYING A ROLE IN THE CHALLENGES OF RECOVERED FROM THE PROCEDURE. PATIENT ADMITTED TO HOSPITAL BEYOND THE STANDARD OF CARE: YES. PATIENT DISCHARGED FROM HOSPITAL?: UNKNOWN. PATIENT OUTCOME: UNKNOWN. PROCEDURE DATE-UNKNOWN (NOTE: JUNE 21, 2024 WAS DOCUMENTED AS THE DATE OF THE EVENT ON THIS FORM, AS THIS WAS THE DATE THE DISTRIBUTOR WAS FIRST NOTIFIED OF THE EVENT). THE DISTRIBUTOR WAS FIRST NOTIFIED OF THIS EVENT 21-JUNE-2024, HOWEVER WAS NOT MADE AWARE THAT A SAFARI2 WAS USED UNTIL 13-AUGUST-2024. EXACT PROCEDURE DATE HAS BEEN REQUESTED, BUT IS UNKNOWN AT THIS TIME. LOT NUMBER HAS BEEN REQUESTED, BUT IS UNKNOWN AT THIS TIME. ADDITIONAL INFORMATION RECEIVED FROM THE DISTRIBUTOR 22AUG2024: WE DID RECEIVE MEDIA BUT IT DID NOT SHOW OR CONFIRM THE CAUSE OR TIMING OF THE PERFORATION/TRAUMA. MEDIA SUMMARY: A REVIEW OF THE PROVIDED MEDIA WAS CONDUCTED. THERE ARE 10 SERIES OF ANGIOGRAPHIC MEDIA. CONTRAST SHOWS THE 3-CUSP VIEW. PRE-DILATION IS SHOWN. THE PRE-DILATION BALLOON BURSTS DURING INFLATION. MEDIA SHOWS THE GUIDEWIRE IN THE VENTRICLE NOT IN THE OPTIMAL POSITION WITH THE LOOP FACING UP. CONTRAST SHOWS INITIAL POSITIONING. STEP 1A (UPPER CROWNS RELEASED) AND STEP 1B (STABILIZATION ARCHES RELEASED) ARE NOT SHOWN. DURING FINAL DEPLOYMENT THE GUIDEWIRE VERY HIGH (NOT DEEP IN THE VENTRICLE), WHICH DOES NOT LEAVE ROOM FOR THE CAPSULE TO MOVE FORWARD. CONTRAST SHOWS THE FULLY DEPLOYED VALVE IN POSITION WITH A LEAK VISIBLE. CONTRAST INJECTION IN THE VENTRICLE TO CHECK FOR DAMAGE. CONTRAST SHOWS BLOOD FLOW THROUGH THE FEMORAL ARTERY. | PRODUCT WAS DISCARDED BY THE HEALTHCARE FACILITY; THEREFORE, NO PHYSICAL OR VISUAL ANALYSIS COULD BE COMPLETED. THE LACK OF LOT TRACEABILITY PREVENTS PERFORMING A DEVICE HISTORY RECORDS REVIEW FOR ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. A REVIEW OF THE MANUFACTURING PROCESSES INDICATES THE PRODUCTION OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE (DFU) WARNINGS SECTION: THE SAFARI2 GUIDEWIRE SHOULD BE USED ONLY BY PHYSICIANS TRAINED IN THE INTRODUCTION AND PLACEMENT OF INTERVENTIONAL DEVICES INCLUDING THOSE USED WITHIN TRANSCATHETER AORTIC VALVE PROCEDURES. CAREFULLY READ ALL INSTRUCTIONS PRIOR TO USE. OBSERVE ALL WARNINGS AND PRECAUTIONS. FAILURE TO DO SO MAY RESULT IN COMPLICATIONS. THE MONITOR WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH-RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. NEVER ADVANCE THE GUIDEWIRE AGAINST RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED IN THE VENTRICLE. THE CURVE OF THE SAFARI2 GUIDEWIRE SHOULD BE CONSTRAINED WITH IN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 1. ENSURE A CATHETER IS APPROPRIATELY PLACED IN THE VENTRICLE OR OTHER INTENDED TREATMENT AREA. 2. CAREFULLY REMOVE THE SAFARI2 GUIDEWIRE, FROM THE HOOP BY GRASPING THE PROXIMAL END OF THE J-STRAIGHTENER SEPARATING THE STRAIGHTENER FROM THE HOOP. 3. AFTER INSPECTION OF THE SAFARI2 GUIDEWIRE, ADVANCE THE J-STRAIGHTENER OVER THE DISTAL SECTION OF THE SAFARI2 WIRE TO STRAIGHTEN THE CURVE. 4. INSERT THE SAFARI2 GUIDEWIRE INTO THE HUB OF THE CATHETER WITH THE AID OF THE J- STRAIGHTENER. 5. CAREFULLY ADVANCE THE DISTAL TIP OF THE SAFARI2 GUIDEWIRE INTO THE LUMEN OF THE CATHETER. 6. REMOVE THE SAFARI2 GUIDEWIRE J- STRAIGHTENER BY WITHDRAWING IT OVER THE LENGTH OF THE SAFARI2 GUIDEWIRE. 7. ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8. CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9. MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. 10. TO REMOVE THE GUIDEWIRE FROM THE TREATMENT AREA: A. A CATHETER MUST BE ADVANCED INTO THE TREATMENT AREA OVER THE SAFARI2 GUIDEWIRE PRIOR TO WITHDRAWING THE WIRE. B. THE SAFARI2 GUIDEWIRE IS PULLED BACK COMPLETELY INTO THE CATHETER. C. THE SAFARI2 GUIDEWIRE MAY THEN BE WITHDRAWN FROM THE CATHETER. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT CLINICAL AND/OR PROCEDURAL FACTORS HAVE CONTRIBUTED TO THE EVENT. CARDIAC PERFORATION, TAMPONADE, BLEEDING, AND ADDITIONAL SURGICAL PROCEDURE ARE KNOWN POTENTIAL ADVERSE EVENTS AND ARE LISTED IN THE SAFARI2 DEVICE INSTRUCTIONS FOR USE (IFU). THE EXACT PROCEDURE DATE IS UNKNOWN AT THE TIME OF THIS REPORT. THE BEST ESTIMATE OF THE EVENT WAS DOCUMENTED AS JUNE 21, 2024, WHICH WAS THE DATE THE DISTRIBUTOR WAS FIRST NOTIFIED OF THE EVENT. THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED; THEREFORE, SECTION D COULD NOT BE COMPLETED, INCLUDING THE UDI NUMBER. THE SECTIONS LEFT BLANK OR UNKNOWN ON THIS FORM COULD NOT BE COMPLETED DUE TO MISSING INFORMATION. AN ATTEMPT TO GATHER FURTHER DETAILS TO OBTAIN THE INFORMATION WAS MADE; HOWEVER, AT THE TIME OF THIS REPORT NO FURTHER INFORMATION HAS BEEN RECEIVED REGARDING THIS EVENT. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM.THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, REDUCED CARDIOPULMONARY RESERVE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION BIPAP AUTOSV | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
FILMARRAY® PNEUMONIA PANEL PLUS | K181324 | SUMMARY: DISCORDANT CTX-M RESULTS BETWEEN THE FILMARRAY PNEUMONIA PANEL PLUS AND CONVENTIONAL CULTURE WERE MENTIONED IN A PUBLICATION BY BAY ET AL. TITLED "PERFORMANCE AND IMPACT OF RAPID MULTIPLEX PCR ON DIAGNOSIS AND TREATMENT OF VENTILATED HOSPITAL-ACQUIRED PNEUMONIA IN PATIENTS WITH EXTENDED-SPECTRUM BETA-LACTAMASE-PRODUCING ENTEROBACTERALES RECTAL CARRIAGE." A PATIENT WITH A DISCORDANT CTX-M RESULT RECEIVED MEROPENEM PLUS AMIKACIN AS EMPIRICAL ANTIBIOTIC THERAPY AFTER THE FILMARRAY PNEUMONIA PANEL PLUS RESULT. THE OUTCOME OF THE PATIENT WAS NOTED AS UNFAVORABLE, AND THEY PASSED AWAY WITHIN 48 HOURS DUE TO MULTIORGAN FAILURE. IT IS UNKNOWN IF THE FILMARRAY PNEUMONIA PANEL PLUS RESULT CONTRIBUTED TO THE PATIENT'S DEATH. BIOFIRE IS SUBMITTING THIS REPORT OUT OF AN ABUNDANCE OF CAUTION. BIOFIRE HAS REQUESTED FURTHER INFORMATION FROM THE AUTHOR AND IS CURRENTLY INVESTIGATING THIS EVENT. NO REMEDIAL ACTION, CORRECTIVE ACTION, PREVENTIVE ACTION, OR FIELD SAFETY CORRECTIVE ACTION (FSCA) HAS BEEN DEEMED NECESSARY AT THIS TIME. PLEASE NOTE THAT THE PATIENT'S ENDOTRACHEAL SAMPLE THAT WAS TESTED ON THE FILMARRAY PNEUMONIA PANEL PLUS HAD BEEN PRE-TREATED WITH A MUCOLYTIC SOLUTION. TREATING A FILMARRAY PNEUMONIA PANEL PLUS SAMPLE WITH A MUCOLYTIC SOLUTION IS CONSIDERED AN OFF-LABEL USE OF THE PANEL. | INVESTIGATION: DISCORDANT CTX-M RESULTS BETWEEN THE FILMARRAY PNEUMONIA PANEL PLUS AND CONVENTIONAL CULTURE WERE MENTIONED IN A PUBLICATION BY BAY ET AL. TITLED "PERFORMANCE AND IMPACT OF RAPID MULTIPLEX PCR ON DIAGNOSIS AND TREATMENT OF VENTILATED HOSPITAL-ACQUIRED PNEUMONIA IN PATIENTS WITH EXTENDED-SPECTRUM BETA-LACTAMASE-PRODUCING ENTEROBACTERALES RECTAL CARRIAGE." A PATIENT WITH A DISCORDANT CTX-M RESULT RECEIVED MEROPENEM PLUS AMIKACIN AS EMPIRICAL ANTIBIOTIC THERAPY AFTER THE FILMARRAY PNEUMONIA PANEL PLUS RESULT. THE OUTCOME OF THE PATIENT WAS NOTED AS UNFAVORABLE, AND THEY PASSED AWAY WITHIN 48 HOURS DUE TO MULTIORGAN FAILURE. IT IS UNKNOWN IF THE FILMARRAY PNEUMONIA PANEL PLUS RESULT CONTRIBUTED TO THE PATIENT'S DEATH. PLEASE NOTE THAT THE PATIENT'S ENDOTRACHEAL SAMPLE THAT WAS TESTED ON THE FILMARRAY PNEUMONIA PANEL PLUS HAD BEEN PRE-TREATED WITH A MUCOLYTIC SOLUTION. TREATING A FILMARRAY PNEUMONIA PANEL PLUS SAMPLE WITH A MUCOLYTIC SOLUTION IS CONSIDERED AN OFF-LABEL USE OF THE PANEL. BIOFIRE'S INVESTIGATION INTO THIS EVENT IS ONGOING AND FURTHER INFORMATION HAS BEEN REQUESTED FROM THE AUTHOR. THE FULL INVESTIGATION AND ASSOCIATED CONCLUSIONS WILL BE PROVIDED IN THE SUPPLEMENTAL REPORT. NO REMEDIAL ACTION, CORRECTIVE ACTION, PREVENTIVE ACTION, OR FSCA HAS BEEN DEEMED NECESSARY AT THIS TIME. REFERENCE: BAY, PIERRE, ET AL. "PERFORMANCE AND IMPACT OF RAPID MULTIPLEX PCR ON DIAGNOSIS AND TREATMENT OF VENTILATED HOSPITAL-ACQUIRED PNEUMONIA IN PATIENTS WITH EXTENDED-SPECTRUM BETA-LACTAMASE-PRODUCING ENTEROBACTERALES RECTAL CARRIAGE." ANNALS OF INTENSIVE CARE, VOL. 14, NO. 1, 29 JULY 2024, HTTPS://DOI.ORG/10.1186/S13613-024-01348-5. CONCLUSION: N/A FOR INITIAL REPORT.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED NOSE IRRITATION, SKIN IRRITATION, ASTHMA, LUNG DISEASE. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION INFLAMMATORY RESPONSE DEATH OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), LUNG DISEASE, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), LUNG DISEASE, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. SECONDARY FINDINGS WERE OBSERVED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, CANCER. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED AN UNKNOWN DUST CONTAMINANT ON THE FLIP DOOR, PCA, TOP ENCLOSURE, REAR PANEL, UI PANEL, BOTTOM ENCLOSURE, INSIDE ISO PORT, BLOWER, AND BLOWER SEAL. MANUFACTURE OBSERVED BLACK HAIRLIKE CONTAMINANT ON THE FLIP DOOR, TOP ENCLOSURE, UI PANEL, REAR PANEL, BOTTOM ENCLOSURE, AND BLOWER. EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS NOT OBSERVED. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER TO THE DEVICE AND VERIFIED AIRFLOW. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION AND THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE. THE MANUFACTURER OBSERVED DUST CONTAMINATION IN THE DEVICE AND ARE CONSISTENT WITH BEING FROM AN EXTERNAL SOURCE. (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID ARE UPDATED
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED WAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
NXSTAGE SYSTEM ONE | K122051 | A REPORT WAS RECEIVED ON (B)(6) 2024 FROM A HOME THERAPY NURSE (HTN) REGARDING A 55-YEAR-OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING DIABETES AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 07 AUG 2024 FROM THE HOME THERAPY NURSE (HTN) WHO STATED THE PATIENT STOOD UP AFTER APPROXIMATELY THREE HOURS INTO THERAPY AND COLLAPSED WITH NO VERBAL COMMENT, DISLODGING HIS CENTRAL VENOUS CATHETER. PER THE HTN, A PULSE WAS PALPABLE, AND EMERGENCY SERVICES (EMS) WERE CALLED, WITH THE PATIENT EXPIRING BY THE TIME EMS ARRIVED. ALTHOUGH REQUESTED, NO CAUSE OF DEATH WAS PROVIDED. | THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
ASTRAL 150 DBL KIT FRA | K172875 | IT WAS REPORTED TO RESMED THAT A PATIENT RECEIVING VENTILATION FROM AN ASTRAL 150 DEVICE EXPIRED. IT WAS REPORTED THAT THE PATIENT WAS FOUND UNCONSCIOUS IN BED WITH THE MASK ON AND THE DEVICE WORKING. IT WAS ALLEGED THAT THE ASTRAL DEVICE DID NOT ALARM. | THE DEVICE HAS BEEN RETURNED TO RESMED FOR AN ENGINEERING INVESTIGATION. THE INVESTIGATION METHODS, RESULTS AND CONCLUSION ARE NOT FINALIZED AT THIS STAGE. IF MORE INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THE ASTRAL 100/150 USER GUIDE PROVIDE THE FOLLOWING WARNINGS: ¿FOR VENTILATOR-DEPENDENT PATIENTS, ALWAYS HAVE ALTERNATE VENTILATION EQUIPMENT AVAILABLE, SUCH AS A BACK-UP VENTILATOR, MANUAL RESUSCITATOR OR SIMILAR DEVICE. FAILURE TO DO SO MAY RESULT IN PATIENT INJURY OR DEATH. VENTILATOR-DEPENDENT PATIENTS SHOULD BE CONTINUOUSLY MONITORED BY QUALIFIED PERSONNEL OR ADEQUATELY TRAINED CARERS. THESE PERSONNEL AND CARERS MUST BE CAPABLE OF TAKING THE NECESSARY CORRECTIVE ACTION IN THE EVENT OF A VENTILATOR ALARM OR MALFUNCTION.¿ RESMED REFERENCE # PR 3136830
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 68-91 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 4 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 5.67. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 2.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 4 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 78-86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE AND 3 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 4 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 1.25. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(6). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(6). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, CONCLUSION CODE GRID HAS BEEN UPDATED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 18 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 61-97 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALES AND 15 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 18 MAJOR VASCULAR COMPLICATION DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 11.94. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 1 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 80 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER (B)(4) AND 1 MAJOR VASCULAR COMPLICATION DEATH EVENT FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | CORRECTED DATA: LEGACY COMPLAINT ID FROM "BLANK" TO (B)(4).
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 VALVE. THIS REPORT SUMMARIZES 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 75-86 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 7 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.44. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 56-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 5 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 3.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 6 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 71-82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 3 FEMALES AND 3 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 1.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENT FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 7 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 72-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALES AND 3 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 7 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.86. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 20 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 56-(B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 10 FEMALES AND 10 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 20 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 2.55. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 21 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 56-89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 11 FEMALES AND 10 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 21 MAJOR VASCULAR COMPLICATION DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 66.62. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENT FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - RETROPERITONEAL DEATH EVENT FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 77 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - RETROPERITONEAL DEATH EVENT FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4).THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENT FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 77-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA VALVE. THIS REPORT SUMMARIZES 3 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS IS 77-87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 3 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 3 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 64-84 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALES AND 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 3 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 1.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(6). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(6). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - RETROPERITONEAL AT ACCESS SITER DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 71 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - RETROPERITONEAL AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 2 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 55-74 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 2 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 20 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 61-95 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 9 FEMALES AND 11 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 20 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 7.8. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4).ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENTS ARE 61-84 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALES AND 5 MALES. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 9 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 0.89. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 67 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 VASCULAR COMPLICATION - MAJOR DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 3.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 85 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 6.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 3 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 67 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 3.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4).ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE VASCULAR INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED THROUGH THE PLANNED ACCESS SITE. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE VESSELS THAT ARE NOT AMENABLE TO PERCUTANEOUS VALVE DELIVERY OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | MITRAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THIS EVENT IS 67 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE. | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND 1 BLEEDING - ACCESS SITE DEATH EVENT FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE "TIME TO EVENT" (TTE, IN DAYS) FOR THIS EVENT WAS 2.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST BLEEDING REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, INFLAMMATORY RESPONSE, LUNG DISEASE, DEATH, LUNG DAMAGE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED LIVER DISEASE/TOXICITY; CANCER; DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF NOSE IRRITATION, RESPIRATORY TRACK IRRITATION, CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION HUMIDIFIER, GB-RENT | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING AND NOT NEEDING IT ANYMORE. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K211900 | THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS FIELD SERVICE PERSONNEL AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON A PATIENT INFORMATION CENTER IX SYSTEM INDICATING THAT THE DEVICE FAILED TO ALARM. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. THE PATIENT WAS NOTED TO HAVE DIED. HOWEVER, IT WAS CONFIRMED BY A CUSTOMER NURSE THAT THE DEVICE DID NOT CONTRIBUTE TO THE DEATH. | A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT TO THE CUSTOMER SITE, AND TESTED THE DEVICE. THE DEVICE PASSED TESTING. ALARM LOGS WERE PROVIDED AND WERE REVIEWED BY A PHILIPS PRODUCT SUPPORT ENGINEER (PSE). THE PSE REVIEWED THE LOGS FOR A REQUESTED TIME FROM (B)(6) 2024 16:23:47 TO (B)(6) 2024 16:23:13 FOR BED LABEL TELB142. THE PSE FOUND THAT THE DEVICE HAD ALARMED FOR MULTIPLE ALARM TYPES STARTING AT 16:42:02 ON (B)(6) 2024 THROUGH 14:40:19 ON (B)(6) 2024; ALARMS INCLUDE INOPERATIVE ALARMS, SUCH AS "CANNOT ANALYZE ECG", YELLOW LEVEL ALARMS, SUCH AS "IRREGULAR HEART RATE", AND RED LEVEL ALARMS, SUCH AS "VENTRICULAR TACHYCARDIA (VTACH)", "VENTRICULAR FIBRILLATION (VENT FIB)", AND "EXTREME TACHYCARDIA (XTACHY)". BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, PHILIPS WAS UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
SURPASS EVOLVE 4.0MM X 17MM - PMAS | K170024 | IT WAS REPORTED THAT DURING FLOW DIVERSION OF LEFT POSTERIOR COMMUNICATING (PCOM) ANEURYSM, THE SUBJECT FLOW DIVERTER HAD A DIFFICULTY OPENING AT THE DISTAL SEGMENT. PHYSICIAN EXPLAINED A STENOSIS IN THE AREA OF CONCERN CAUSED BY A PREVIOUSLY PLACED NON-STRYKER FLOW DIVERTER DEVICE COULD HAVE RESULTED IN SUBJECT STENT OPENING DIFFICULTY. THE OPERATOR ATTEMPTED ANGIOPLASTY, BUT THE SUBJECT FLOW DIVERTER STILL DID NOT OPEN AND FLOW BEYOND SUBJECT STENT WAS LIMITED WHICH MAY HAVE CAUSED A STROKE. THE FLOW DIVERTER ONLY OPENED AFTER A THIRD ANGIOPLASTY WITH MORE PRESSURE. HOWEVER, THE VESSEL RUPTURED CAUSING CONTRAST EXTRAVASATION. AS A RESULT, THE PATIENT WAS DEEMED AS MEETING "BRAIN DEATH" CRITERIA AS THERE WAS NO DISTAL PERFUSION BEYOND THE INTERNAL CAROTID ARTERY (ICA). NO FURTHER INFORMATION IS AVAILABLE. | D4 GTIN - CORRECTED. G4 PMA/510K - CORRECTED. DUE TO THE AUTOMATED MANUFACTURING EXECUTION SYSTEM (MES) SYSTEM THERE ARE CONTROLS IN THE MANUFACTURING PROCESS TO ENSURE THE PRODUCT MET SPECIFICATIONS UPON RELEASE. VISUAL AND FUNCTIONAL INSPECTIONS WERE NOT PERFORMED DUE TO THE DEVICE WAS NOT RETURNED. THE REPORTED EVENT IS COVERED IN THE DEVICE DIRECTIONS FOR USE (DFU). AS WELL, THE RISK OF THE REPORTED EVENT IS DOCUMENTED IN THE RISK DOCUMENTATION AND THERE ARE CURRENT CONTROLS TO MITIGATE THE RISK OF THE AS REPORTED EVENT. THE REPORTED COMPLAINT COULD NOT BE CONFIRMED, AND IT COULD NOT BE DEFINITIVELY DETERMINED IF THE DEVICE FAILED TO MEET SPECIFICATIONS BECAUSE THE PRODUCT WAS NOT RETURNED. ADDITIONAL INFORMATION PROVIDED BY THE CUSTOMER INDICATED THE FLOW DIVERTER STENT DID NOT OPEN DESPITE AGGRESSIVE ANGIOPLASTY WITH THE BALLOON CENTER AROUND THE AREA OF STENOSIS FOR THE FIRST ANGIOPLASTY AND WITH THE TIP OF THE BALLOON ACROSS THE STENOSIS FOR THE SECOND ANGIOPLASTY. AFTER BALLOON ANGIOPLASTY X2 THE AREA OF STENOSIS IS AGAIN DOWN TO 1 MM AND THERE IS DELAYED PASSAGE OF CONTRAST BUT NO EVIDENCE FOR RUPTURE. REGULAR BALLOON ANGIOPLASTY WAS UNABLE TO OPEN THE FLOW DIVERTER EVEN THOUGH THE BALLOON WAS INFLATED TO 2 MM IN DIAMETER. THE FLOW DIVERTER ONLY OPENED AFTER A THIRD ANGIOPLASTY WITH MORE PRESSURE, BUT THIS ATTEMPT ALSO RUPTURED THE ARTERY AND RESULTED IN THE DEATH OF THE PATIENT. THE FLOW DIVERTER IS NOW OPEN AT THE PREVIOUS AREA OF STENOSIS. ALL ANGIOPLASTIES WERE CARRIED OUT WITH THE BALLOON INSIDE THE FLOW DIVERTER. SO, THE DEVICE COULD ONLY BE OPENED WITH EXTREME PRESSURES". THE PATIENT RECEIVED DUAL ANTI-PLATELET AGENTS PRIOR TO THE PROCEDURE, THE DEVICE WAS PREPARED AS PER DFU, A MICROCATHETER WAS USED TO ADVANCE THE FLOW DIVERTER, CONTINUOUS FLUSH WAS MAINTAINED THROUGHOUT THE PROCEDURE, THE PATIENT'S ANATOMY WAS SOMEWHAT TORTUOUS, THE PERCENTAGE STENOSIS PRESENT WAS 51-75%. ACCORDING TO THE PHYSICIAN, THE ADVERSE EVENT WAS RELATED TO ANGIOPLASTY-ANGIOPLASTY INSIDE EVOLVE AND CAUSED THE RUPTURE. THE EVOLVE WAS STILL NOT OPEN AND FLOW BEYOND DEVICE WAS LIMITED WHICH MAY HAVE CAUSED A STROKE. IT IS MOST LIKELY THAT ANATOMICAL OR PROCEDURAL FACTORS ENCOUNTERED DURING THE PROCEDURE CONTRIBUTED TO THE REPORTED EVENT. WHILE THERE ARE A NUMBER OF POTENTIAL CAUSES FOR THE REPORTED ISSUE, BECAUSE REVIEW AND ANALYSIS OF AVAILABLE INFORMATION FAILED TO IDENTIFY A DEFINITIVE CAUSE, A CAUSE OF UNDETERMINABLE WAS ASSIGNED TO THE REPORTED EVENTS ¿STENT FAILED/UNABLE TO OPEN¿ AND ¿STENT IMPROPERLY DEPLOYED'. AN ASSIGNABLE CAUSE OF ANTICIPATED PROCEDURAL COMPLICATION WILL BE ASSIGNED TO THE REPORTED EVENTS: ¿PATIENT DEATH¿, ¿PATIENT VESSEL DISSECTION¿ AND ¿PATIENT STROKE¿ AS A PRODUCT RELATED ROOT CAUSE DOES NOT APPLY AND THE ISSUE IS DUE TO A KNOWN PHYSIOLOGICAL EFFECT OF THE PROCEDURE AND/OR PATIENT CONDITION NOTED WITH THE DIRECTIONS FOR USE, PRODUCT LABELING AND/OR RISK DOCUMENTATION FILES.
NXSTAGE SYSTEM ONE | K122051 | A REPORT WAS RECEIVED ON 13 AUG 2024 FROM THE HOME THERAPY NURSE (HTN) OF A 70 YEAR OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING DIABETES AND END STAGE RENAL DISEASE, APPROVED FOR PERFORMING SOLO HOME HEMODIALYSIS, WHO STATED THE PATIENT WAS FOUND EXPIRED WITH THE VENOUS NEEDLE DISLODGED AND SURROUNDED BY AN UNSPECIFIED AMOUNT OF BLOOD DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 14 AUG 2024 FROM THE HTN WHO STATED THE PATIENT WAS FOUND BY HIS WIFE WHO CALLED EMERGENCY MEDICAL SERVICES (EMS) AND ATTEMPTED UNSUCCESSFUL CARDIOPULMONARY RESUSCITATION. THE PATIENT WAS PRONOUNCED DECEASED UPON EMS ARRIVAL. | THE CYCLER WAS RECEIVED FOR EVALUATION AND SUCCESSFULLY PASSED TESTING. EVALUATION OF THE AVAILABLE LOG FILES WAS UNREMARKABLE WITH NO PRODUCT DEFICIENCY IDENTIFIED. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
INTELLIVUE MX800 PATIENT MONITOR | K150310 | IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN REVIEWING THE CLINICAL AUDIT LOGS AFTER A PATIENT PASSED AWAY. THE NURSE INDICATED THEY DID NOT HEAR AN ALARM AND THE UNIT MANAGER WAS TRYING TO DETERMINE IF ALARMS WERE ACKNOWLEDGED OR PAUSED AT THE TIME OF THE PATIENT INCIDENT. IT WAS INDICATED THERE WAS NO DELAY AND THE CUSTOMER SIMPLY WANTED TO RESEARCH ALARMS AS THE CUSTOMER WAS INVESTIGATING A STAFF ISSUE. | THE PHILIPS REMOTE CLINICAL SUPPORT (RCS) CONTACTED THE CUSTOMER. IT WAS INDICATED THE X3 WAS DOCKED WITH THE MX800 HOST MONITOR, WHICH WAS NETWORKED WITH THE PIC IX. BOTH MONITORS WERE TESTED BY THE HOSPITAL AND FOUND TO BE WORKING AS EXPECTED. A REVIEW OF THE CLINICAL AUDIT LOGS REVEALED ALARMS WERE GENERATED OVER THE COURSE OF THE 5 HOUR EVENT TIMEFRAME; HOWEVER, THE USER DID NOT ACKNOWLEDGE OR PAUSE ALARMS DURING THAT TIME. THE CUSTOMER INDICATED THERE WAS NEVER ANY ALLEGATION OF DEVICE MALFUNCTION AND THE LOGS WERE REQUESTED DUE TO AN INTERNAL NURSING INVESTIGATION ONLY. THERE WAS NO DELAY IN CARE, THOUGH THE PATIENT WENT TO ICU AND LATER PASSED AWAY. AS THERE WAS NO MALFUNCTION AND NO DELAY IN CARE, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED PATIENT DEATH; HOWEVER, THE CAUSE REMAINS UNKNOWN.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION BIPAP PRO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF NOSE IRRITATION, RESPIRATORY TRACK IRRITATION, CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA/VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. PER THE REPORTER, THE CUSTOMER WAS FOUND UNRESPONSIVE WITH A BLOOD GLUCOSE LEVEL OF 1,500 MG/DL. ADDITIONALLY, THE CUSTOMER WENT TO THE INTENSIVE CARE UNIT (ICU). IT IS UNKNOWN IF THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF THE EVENT. AT THE TIME OF THIS REPORT, PUMP DATA IS NOT AVAILABLE FOR REVIEW. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO ADDITIONAL INFORMATION COULD BE OBTAINED. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | K201214 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. PER THE REPORTER, THE CUSTOMER PASSED AWAY DUE TO SEVERE HYPOGLYCEMIA AND SEIZURES. IT IS UNKNOWN IF THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF THE EVENT. AT THE TIME OF THIS REPORT, PUMP DATA IS NOT AVAILABLE FOR REVIEW. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO ADDITIONAL INFORMATION COULD BE OBTAINED. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED NOSE IRRITATION, SKIN IRRITATION, KIDNEY DISEASE/TOXICITY, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
AMNISURE ROM TEST | K081767 | A CUSTOMER REPORTED A FETAL DEATH TO ANSM AFTER AN UNEXPECTED NEGATIVE AMNISURE RESULT. MOTHER REPORTED LEAKAGE FOR SEVERAL DAYS PRIOR TO TESTING. | ADDITIONAL INFORMATION WAS PROVIDED BY USER. THEY DID NOT USE ANY OTHER ASSESSMENTS IN CONJUNCTION WITH THE AMNISURE TEST FOR DIAGNOSIS OF ROM. A SIGNIFICANT AMOUNT OF BLOOD WAS PRESENT IN THE SECOND AMNISURE TEST SAMPLE. WHEN THERE IS A SIGNIFICANT PRESENCE OF BLOOD ON THE SWAB, THE TEST CAN MALFUNCTION AND IS NOT RECOMMEND. OUR INVESTIGATION CONCLUDED THAT THE KIT LOT IS FUNCTIONING AS EXPECTED, AND NO CORRECTIVE ACTIONS ARE NEEDED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NON-HODGKIN'S LYMPHOMA AND CANCER. THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NON-HODGKIN'S LYMPHOMA AND CANCER. THERE WAS A REPORT OF PATIENT DEATH. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING SERVICING OF THE DEVICE, THE UNIT PASSED FINAL TESTS. THE DEVICE HAS BEEN REPAIRED. SECTION G2 AND H6 HAVE BEEN UPDATED ACCORDINGLY.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE ARE NO ALLEGATIONS THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN EVALUATED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WERE NOT OBSERVED. THERE WAS TWO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT CORRECTED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE ARE NO ALLEGATIONS THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE EVALUATION HAS NOT BEGUN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. SECTION-H HAS BEEN UPDATED.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | ON (B)(6) 2024, THIS PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR A THORACIC AORTIC ANEURYSM USING TWO VALIANT THORACIC STENT GRAFTS (MEDTRONIC JAPAN CO., LTD). A GORE® DRYSEAL FLEX INTRODUCER SHEATH (SHEATH) WAS USED DURING THE PROCEDURE. IT WAS REPORTED THAT PATIENT¿S RIGHT ACCESS VESSEL (FROM THE RIGHT COMMON ILIAC ARTERY (CIA) JUST BELOW THE TERMINAL AORTA THROUGH EXTERNAL ILIAC ARTERY (EIA) TO COMMON FEMORAL ARTERY (CFA)) WHERE THE SHEATH HAD BEEN INSERTED DURING THE PROCEDURE, BECAME OCCLUDED CAUSING PERFUSION INJURY TO THE PATIENT. AS A TREATMENT, A THROMBECTOMY AND A FEMORAL-FEMORAL CROSSOVER BYPASS SURGERY WAS PERFORMED ON (B)(6) 2024, HOWEVER, THE PATIENT WENT INTO CARDIOPULMONARY ARREST DUE TO THE REPERFUSION INJURY. AFTER FAILING TO CONTROL THE BLOOD PRESSURE, THE PATIENT EXPIRED ON (B)(6) 2024. ACCORDING TO THE OPERATING PHYSICIAN, THE PATIENT¿S ACCESS VESSELS ON THE RIGHT SIDE HAD UNFAVORABLE ANATOMY. THE RIGHT CIA AND THE RIGHT EIA MEASURED 8.2MM. THE RIGHT CIA AND THE RIGHT CFA WERE HIGHLY CALCIFIED. MILD STENOSIS WERE NOTED AT TERMINAL AORTA. THE PHYSICIAN SUSPECTED THE PATIENT WAS EXPOSED TO ISCHEMIC CONDITIONS FOR TOO LONG DUE TO THE USE OF THE LARGE-SIZED SHEATH DURING THE PROCEDURE, CAUSING THE PERFUSION INJURY THAT LEAD TO THE PATIENT'S DEATH. THE PHYSICIAN MADE NO ALLEGATION AGAINST THE PERFORMANCE OF THE SHEATH, AND HE DOES NOT THINK THE EVENT IS ATTRIBUTED TO THE SHEATH MALFUNCTION. | H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICE WAS DISCARDED AT FACILITY. RETURN NOT POSSIBLE. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTION FOR USE (IFU), THE POTENTIAL ADVERSE EVENTS WITH THE USE OF DEVICE MAY INCLUDE BUT ARE NOT LIMITED TO: TRANSIENT OR PERMANENT ISCHEMIA, DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
POWERPORT ISP M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | K063377 | IT WAS REPORTED THAT SOMETIME POST A PORT PLACEMENT, THE PORT ALLEGEDLY FELL OUT OF THE PATIENT AND HAD ALLEGEDLY LEAKED APPROXIMATELY 3OZ OR MORE OF CHEMOTHERAPY MEDICATION CAUSING CHEMICAL BURNS TO THE PATIENT¿S COLLAR BONE AREA WHERE PORT WAS PLACED. IT WAS FURTHER REPORTED THAT THE PATIENT WAS TREATED FOR THE ALLEGED BURNS, BUT IT IS UNKNOWN WHAT TYPE OF MEDICATION WAS USED FOR TREATMENT. REPORTEDLY, THE PORT WAS REMOVED AND REPLACED WITH A NEW PORT, AND THE PATIENT FINISHED TREATMENT WITH AN INTRAVENOUS CATHETER. EVIDENTLY, THE PATIENT GOT EXPIRED FROM CANCER DISEASE AND THE RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND THE DEVICE USED WAS NOT PROVIDED. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED FAILURE AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE FLUID LEAK AND DISCONNECTION COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3, H6 (METHOD). H11: D4 (UNIQUE IDENTIFIER (UDI) #). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
MICRA | K132030 | IT WAS REPORTED, THAT DURING THE IMPLANT PROCEDURE OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). DURING THE DEVICE POSITIONING PROCEDURE (NEVER RELEASED), THE PATIENT REPORTS CHEST PAIN. PERICARDIAL EFFUSION, CARDIAC TAMPONADE AND CARDIAC PERFORATION WERE OBSERVED. PERICARDIOCENTESIS AND RESUSCITATION PROCEDURES WERE STARTED IMMEDIATELY, BUT THE PATIENT DEVELOPED ELECTROMECHANICAL DISSOCIATION. RESULTING IN CARDIAC ARREST. PATIENT CONSEQUENTLY DIED. | CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, NOSE IRRITATION, DIZZINESS, HEADACHE, ASTHMA, LUNG DISEASE, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, DIZZINESS, HEADACHE, CANCER, NAUSEA, VOMITING, AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PERFIX PLUG LIGHT | K092032 | PER INFORMATION PROVIDED: (B)(6) 2020 - PATIENT WAS DIAGNOSED WITH LARGE INTRASCROTAL RIGHT INGUINAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF PERFIX PLUG LIGHT. PER OPERATIVE NOTES, ¿HERNIA WAS DISSECTED UP TO THE FLOOR OF THE INGUINAL CANAL AND THEN REDUCED ON MASS BACK INTO THE ABDOMEN. AN EXTRA-LARGE PERFIX PLUG WAS THEN PLACED INTO THE DEFECT AND TACKED IN SEPARATE PLACES ALONG CONJOINED TENDON AND THE SHELVING EDGE OF THE INGUINAL LIGAMENT. THE IN USING THE ONLAY PATCH AN OPENING WAS MADE TO ACCOMMODATE THE CORD STRUCTURES AND IT WAS PLACED AROUND THE CORD.¿ (B)(6) 2020 - PATIENT WAS DIAGNOSED WITH BLADDER OUTLET OBSTRUCTION SECONDARY TO BPH, ACUTE URINARY RETENTION THEREBY RESECTION OF TRANSURETHRAL PROSTATE. PER OPERATIVE NOTES, ¿THE POSTERIOR URETHRA WAS NOTED TO HAVE A GROSS OBSTRUCTION.¿ ATTORNEY ALLEGES THAT THE PATIENT HAD ADHESIONS, BOWEL OBSTRUCTION AND PERFORATION, HERNIA RECURRENCE, FISTULAE, INFECTION, PAIN & SUFFERING. IT WAS ALSO ALLEGED THAT THE PATIENT HAD SEPSIS, INFECTION, ABDOMINAL PAIN, RIGHT GROIN PAIN, URINARY RETENTION, TRANSURETHRAL RESECTION OF THE PROSTATE, PROSTATE HYPERTROPHY, URINARY TRACT SYMPTOMS, DIVERTICULITIS, ENCEPHALITIS, COLON AND INTESTINAL ULCERS, WEIGHT LOSS AND OTHERS. ADDENDUM PER LEGAL CLAIM: ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF AN BARD/DAVOL PERFIX LIGHT PLUG ON (B)(6) 2020. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST PERFIX LIGHT PLUG. ATTORNEY ALLEGES THAT THE PATIENT HAD SUBSEQUENT SURGICAL INTERVENTION DUE TO THE HERNIA MESH DEVICE AND WRONGFUL DEATH OF THE PATIENT. IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. | BASED ON THE INFORMATION PROVIDED, NO CONCLUSIONS CAN BE MADE. PER MEDICAL RECORDS REVIEW, ABOUT 3 WEEKS POST IMPLANT OF PERFIX PLUG LIGHT, PATIENT WAS DIAGNOSED WITH OBSTRUCTION AND ACUTE URINARY RETENTION THEREBY UNDERWENT HERNIA REPAIR. ADDENDUM: THIS IS AN ADDENDUM TO THE INITIAL EMDR SUBMITTED (MDR NUMBER 1213643-2024-094587). THIS SUPPLEMENTAL EMDR IS BEING SUBMITTED TO REPORT THE ALLEGATION OF PATIENT DEATH. NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING DEATH AND SUBSEQUENT SURGICAL INTERVENTION. NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. A REVIEW OF THE MANUFACTURING RECORDS WAS PERFORMED AND FOUND THAT THE LOT WAS MANUFACTURED TO SPECIFICATION. UPDATED FIELDS: B2 (OUTCOMES ATTRIBUTED TO ADV EV), B4, B5, G3, G6, H1, H2, H4 (MANUFACTURING DATE), H6, H10. NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, LUNG DISEASE AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ENDOWRIST | K173337 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | THE VIDEO RECEIVED OF THIS EVENT WAS REVIEWED BY A CLINICAL DEVELOPMENT ENGINEER (CDE) WHO PROVIDED THE FOLLOWING INFORMATION: AT AROUND 2:47 MINUTES INTO THE VIDEO: WHEN THE SURGEON IS SEALING THE VESSEL, A USER INTERFACE (UI) MESSAGE APPEARS, INFORMING THE SURGEON TO ENSURE THERE IS AN APPROPRIATE AMOUNT OF TISSUE IN THE JAWS AND THAT THE SEALING CYCLE WAS INCOMPLETE. THE USER MANUAL RECOMMENDS RE-GRASPING AN APPROPRIATE AMOUNT OF TISSUE BEFORE REAPPLYING ENERGY WHEN THIS ERROR OCCURS. FROM THE VIDEO, IT APPEARS THE SURGEON COULD HAVE SKELETONIZED THE VESSEL MORE THOROUGHLY. THERE IS STILL A SIGNIFICANT AMOUNT OF SURROUNDING FAT TISSUE, WHICH COULD HAVE LED TO THE VESSEL SIZE EXCEEDING THE VESSEL SEALER EXTEND'S (VSE) INDICATED DIAMETER OF 7MM.
ION | K182188 | IT WAS REPORTED THAT AFTER AN UNEVENTFUL ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING IN THE RIGHT LUNG AND WENT INTO CARDIAC ARREST. THE INTUITIVE ENDOLUMINAL CLINICAL SALES REPRESENTATIVE WAS INFORMED THAT THE ION PROCEDURE WAS COMPLETED AS PLANNED AND THE ION SYSTEM WAS UNDOCKED. THE PHYSICIAN INSERTED A THERAPEUTIC BRONCHOSCOPE PER NORMAL PRACTICE AND, PER A COLLEAGUE PHYSICIAN THAT WAS NOT AT THE PROCEDURE, A LARGE CLOT WAS DISCOVERED AT THE END OF THE ENDOTRACHEAL TUBE, WHICH RESULTED IN DECREASED VISUALIZATION. THE THERAPEUTIC SCOPE WAS ADVANCED INTO THE RIGHT LUNG AND IT APPEARED TO BE FILLED WITH BLOOD. THE PHYSICIAN PLACED AN ENDOBRONCHIAL BLOCKER IN THE RIGHT LUNG AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS REVIVED AND TRANSFERRED TO THE INTENSIVE CARE UNIT INTUBATED, WHERE PARALYTIC AGENTS WERE ADMINISTERED UNTIL A NEUROLOGICAL TEST COULD BE CONDUCTED. THE PATIENT EXPIRED AN UNSPECIFIED TIME LATER. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN THAT PERFORMED THE PROCEDURE HAVE BEEN SUCCESSFUL. | THE ION BIOPSY PROCEDURE WAS REPORTED TO HAVE BEEN COMPLETED AND THE SYSTEM WAS UNDOCKED WITH NO INDICATION OF ANY PRODUCT ISSUES. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED THAT WOULD BE RELEVANT TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE ION PROCEDURE FOUND THERE WERE NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THE EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT OF UNKNOWN CO-MORBIDITIES UNDERWENT AN ION ENDOLUMINAL BIOPSY OF AN UNKNOWN TARGET. THE BIOPSY WAS COMPLETED WITHOUT ISSUE AND THE ION CATHETER WAS REMOVED, THE SYSTEM WAS UNDOCKED AND THE PROCEDURE WAS TRANSITIONED TO MANUAL BRONCHOSCOPY. A CLOT WAS NOTED AT THE END OF THE ENDOTRACHEAL TUBE WITH MANUAL BRONCHOSCOPY. BLOOD WAS ALSO NOTED IN THE RIGHT LUNG FOR WHICH AN ENDOBRONCHIAL BLOCKER WAS PLACED. THE PATIENT THEN SUFFERED A CARDIAC ARREST, WAS SUCCESSFULLY RESUSCITATED AND TRANSFERRED TO THE ICU FOR ONGOING MANAGEMENT AND ULTIMATELY DIED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. NO FURTHER DATA IS AVAILABLE DESPITE MULTIPLE ATTEMPTS. BASED ON THE AVAILABLE DATA THE EVENTS ARE PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. SECTION B2 DATE OF DEATH IS ESTIMATED.
ION | K182188 | IT WAS REPORTED THAT FOLLOWING AN UNEVENTFUL ION ASSISTED TRANSBRONCHIAL LUNG BIOPSY PROCEDURE THE PATIENT EXPIRED. THE ION BIOPSY PROCEDURE AND SUBSEQUENT EBUS PROCEDURES WERE COMPLETED WITHOUT ISSUE. A POST-PROCEDURAL SCAN DID NOT IDENTIFY A PNEUMOTHORAX AND THE PATIENT WAS EXTUBATED AND MOVED TO RECOVERY IN STABLE CONDITION. THE PHYSICIAN REPORTED THAT AFTER 16 MINUTES THE PATIENT BEGAN EXPERIENCING DIFFICULTY BREATHING AND WENT INTO RESPIRATORY ARREST, LEADING TO A SUSPECTED BRADYCARDIC EVENT (HEART RATE 40 BPM) AND SUBSEQUENTLY CODED. CARDIOPULMONARY RESUSCITATION WAS PERFORMED FOR APPROXIMATELY 20 MINUTES WITHOUT REVIVAL. THE PHYSICIAN SUSPECTED A MASSIVE PULMONARY EMBOLISM BECAUSE OF THE SEQUENCE OF RESPIRATORY DISTRESS/ARREST LEADING TO BRADYCARDIA AND THEN CARDIAC ARREST. PER THE PHYSICIAN, THE ION SYSTEM DID NOT CAUSE OR CONTRIBUTE THE DEATH. THE PATIENT HAD A MEDICAL HISTORY OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND EMPHYSEMA. | A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE ION SYSTEM FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH COPD AND EMPHYSEMA UNDERWENT AN ION ENDOLUMINAL BIOPSY AND EBUS. THE PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS AND IMAGING REVEALED NO PNEUMOTHORAX. THE PATIENT WAS EXTUBATED AND TRANSPORTED TO RECOVERY. IN RECOVERY THE PATIENT REPORTED DIFFICULTY BREATHING FOLLOWED BY CARDIOPULMONARY ARREST. EFFORTS AT RESUSCITATION FOR 20 MINUTES WERE UNSUCCESSFUL AND THE PATIENT DIED. BASED ON THE AVAILABLE DATA A CARDIOVASCULAR EVENT INCLUDING PULMONARY EMBOLISM IS THE MOST LIKELY TRIGGERING EVENT. THERE WERE NO MALFUNCTIONS OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. THE EVENT WAS LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF (B)(4) WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). --OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPOR TED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | PLEASE DISREGARD THE ABOVE TW (B)(4) 2518422-2024-55662, IT IS DONE BY MISTAKE. THE FIRST INITIAL REPORT TW (B)(4) 2518422-2024-49835 IS THE CORRECT ONE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
NITREX | K040345 | IT WAS REPORTED, THAT A PATIENT WAS UNDERGOING A HEART FAILURE MONITORING DEVICE IMPLANT PROCEDURE. DURING PROCEDURE, THE PHYSICIAN GAINED ACCESS AND PROCEEDED TO USE A MEDTRONIC 0.18 NITREX GUIDEWIRE. THE PHYSICIAN ADVANCED THE WIRE TO THE PULMONARY ARTERY. THE PHYSICIAN USED A NON-MEDTRONIC CATHETER OVER THE WIRE. THE PATIENT COUGHED A FEW TIMES, BUT THE PHYSICIAN PROCEEDED WITH THE RIGHT HEART CATHETERIZATION. IT WAS THEN, THAT THE PATIENT EXPERIENCED HEMOPTYSIS. THE PATIENT'S STATUS WAS DETERIORATING, SO THE CODE TEAM WAS CALLED IN TO TRY TO RESUSCITATE THE PATIENT. THE PHYSICIAN DID A THORACOTOMY TO TRY TO STOP THE BLEEDING. THE PATIENT DIED IN THE CATH LAB. THE NON-MEDTRONIC HEART FAILURE MONITORING DEVICE WAS NOT OPENED OR INSERTED INTO THE BODY. NO SPECIFIC CAUSE OF DEATH WAS LISTED IN THE PATIENT¿S CHART. THE PHYSICIAN BELIEVED, THAT THERE MUST HAVE BEEN A PERFORATION, BUT ONE COULD NOT BE IDENTIFIED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK, BECAUSE THE INFORMATION IS CURRENTLY, UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING) LUNG DISEASE, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
SIMPLEX P WITH TOBRAMYCIN 1 PACK | K014199 | BHA WAS PERFORMED FOR RIGHT FEMORAL NECK FRACTURE. TWO MINUTES AFTER FILLING THE INTRATHECAL CAVITY WITH CEMENT AND INSERTING THE STEM, THE OXYGEN SATURATION LEVEL BEGAN TO GRADUALLY DECREASE, FOLLOWED BY A DROP IN BLOOD PRESSURE AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS CONFIRMED DEAD ON THE SAME DAY. | CORRECTED DATA: COMPLAINT HISTORY REVIEW (H11). AN EVENT REGARDING BONE CEMENT IMPLANTATION SYNDROME (BCIS) INVOLVING SIMPLEX CEMENT MIX WAS REPORTED. THE EVENT WAS NOT CONFIRMED. METHOD & RESULTS: PRODUCT EVALUATION AND RESULTS: MATERIAL ANALYSIS, VISUAL, FUNCTIONAL AND DIMENSIONAL INSPECTIONS COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED. CLINICIAN REVIEW: NO MEDICAL RECORDS WERE RECEIVED FOR REVIEW WITH A CLINICAL CONSULTANT. PRODUCT HISTORY REVIEW: REVIEW OF THE DEVICE HISTORY RECORDS INDICATE TEN-PACKS WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW: THERE HAVE BEEN 1 OTHER SIMILAR EVENTS FOR THE LOT REFERENCED, WHICH RELATES TO THE SAME PATIENT/DEVICE THEREFORE NO FURTHER TRENDING IS REQUIRED. CONCLUSIONS: IT WAS REPORTED THAT TWO MINUTES AFTER FILLING THE INTRATHECAL CAVITY WITH CEMENT AND INSERTING THE STEM, THE OXYGEN SATURATION LEVEL BEGAN TO GRADUALLY DECREASE, FOLLOWED BY A DROP IN BLOOD PRESSURE AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS CONFIRMED DEAD ON THE SAME DAY. THE EXACT CAUSE OF THE EVENT COULD NOT BE DETERMINED BECAUSE INSUFFICIENT INFORMATION WAS PROVIDED. FURTHER INFORMATION SUCH AS THE OPERATIVE REPORT AS WELL AS PATIENT HISTORY AND FOLLOW-UP NOTES ARE NEEDED TO COMPLETE THE INVESTIGATION FOR DETERMINING ROOT CAUSE. NO FURTHER INVESTIGATION FOR THIS EVENT IS POSSIBLE AT THIS TIME. IF ADDITIONAL INFORMATION BECOMES AVAILABLE TO INDICATE FURTHER EVALUATION IS WARRANTED, THIS RECORD WILL BE REOPENED.
SIMPLEX P WITH TOBRAMYCIN 1 PACK | K014199 | BHA WAS PERFORMED FOR RIGHT FEMORAL NECK FRACTURE. TWO MINUTES AFTER FILLING THE INTRATHECAL CAVITY WITH CEMENT AND INSERTING THE STEM, THE OXYGEN SATURATION LEVEL BEGAN TO GRADUALLY DECREASE, FOLLOWED BY A DROP IN BLOOD PRESSURE AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS CONFIRMED DEAD ON THE SAME DAY. | CORRECTED DATA: COMPLAINT HISTORY REVIEW (H11). AN EVENT REGARDING BONE CEMENT IMPLANTATION SYNDROME (BCIS) INVOLVING SIMPLEX CEMENT MIX WAS REPORTED. THE EVENT WAS NOT CONFIRMED. METHOD & RESULTS: PRODUCT EVALUATION AND RESULTS: MATERIAL ANALYSIS, VISUAL, FUNCTIONAL AND DIMENSIONAL INSPECTIONS COULD NOT BE PERFORMED AS THE DEVICE REMAINS IMPLANTED. CLINICIAN REVIEW: NO MEDICAL RECORDS WERE RECEIVED FOR REVIEW WITH A CLINICAL CONSULTANT. PRODUCT HISTORY REVIEW: REVIEW OF THE DEVICE HISTORY RECORDS INDICATE TEN-PACKS WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW: THERE HAVE BEEN 1 OTHER SIMILAR EVENTS FOR THE LOT REFERENCED, WHICH RELATES TO THE SAME PATIENT/DEVICE THEREFORE NO FURTHER TRENDING IS REQUIRED. CONCLUSIONS: IT WAS REPORTED THAT TWO MINUTES AFTER FILLING THE INTRATHECAL CAVITY WITH CEMENT AND INSERTING THE STEM, THE OXYGEN SATURATION LEVEL BEGAN TO GRADUALLY DECREASE, FOLLOWED BY A DROP IN BLOOD PRESSURE AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT WAS CONFIRMED DEAD ON THE SAME DAY. THE EXACT CAUSE OF THE EVENT COULD NOT BE DETERMINED BECAUSE INSUFFICIENT INFORMATION WAS PROVIDED. FURTHER INFORMATION SUCH AS THE OPERATIVE REPORT AS WELL AS PATIENT HISTORY AND FOLLOW-UP NOTES ARE NEEDED TO COMPLETE THE INVESTIGATION FOR DETERMINING ROOT CAUSE. NO FURTHER INVESTIGATION FOR THIS EVENT IS POSSIBLE AT THIS TIME. IF ADDITIONAL INFORMATION BECOMES AVAILABLE TO INDICATE FURTHER EVALUATION IS WARRANTED, THIS RECORD WILL BE REOPENED.
VISUALASE SYSTEM | K081656 | DABECCO, R., GIGLIOTTI, M.J., MAO, G., MYERS, D., XU, L., LEE, P., RANJAN, T., AZIZ, K., YU, A. LASER INTERSTITIAL THERMAL THERAPY (LITT) FOR INTRACRANIAL LESIONS: A SINGLE-INSTITUTIONAL SERIES, OUTCOMES, AND REVIEW OF THE LITERATURE. BRITISH JOURNAL OF NEUROSURGERY. 2024. VOL. 38. NO. 3, 632¿638. HTTPS://DOI.ORG/10.1080/02688697.2021.1947972 INTRODUCTION: LASER INTERSTITIAL THERMAL THERAPY (LITT) IS A MINIMALLY INVASIVE TREATMENT METHOD IN MANAGING PRIMARY BRAIN NEOPLASMS, BRAIN METASTASES, RADIATION NECROSIS, AND EPILEPTOGENIC LESIONS, MANY OF WHICH ARE LOCATED IN OPERATIVE CORRIDORS THAT WOULD BE DIFFICULT TO ADDRESS. ALTHOUGH THE USE OF LASERS IS NOT A NEW CONCEPT IN NEUROSURGERY, ADVANCES IN TECHNOLOGY HAVE ENABLED SURGEONS TO PERFORM LASER TREATMENT WITH THE AID OF REAL-TIME MRI THERMOGRAPHY AS A GUIDE. IN THIS REPORT, WE PRESENT OUR INSTITUTIONAL SERIES AND OUTCOMES OF PATIENTS TREATED WITH LITT. METHODS: WE RETROSPECTIVELY EVALUATED 19 PATIENTS (AGE RANGE, 28¿77 YEARS) WHO UNDERWENT LITT AT ONE OR MORE TARGETS FROM 2015 TO 2019. PRIMARY ENDPOINT OBSERVED WAS MEAN PROGRESSION FREE SURVIVAL (PFS) AND OVERALL SURVIVAL (OS). RESULTS: SEVEN PATIENTS WITH GLIAL NEOPLASMS AND 12 PATIENTS WITH METASTATIC DISEASE WERE REVIEWED. AVERAGE HOSPITALIZATION WAS 2.4 DAYS. MEDIAN PFS WAS 7 AND 4 MONTHS IN THE METASTATIC GROUP AND PRIMARY GLIAL NEOPLASM GROUP, RESPECTIVELY (P ¼ 0.01). MEDIAN OS FROM TIME OF DIAGNOSIS WAS 41 AND 32 MONTHS (P ¼ 0.02) AND MEDIAN OS AFTER LITT THERAPY WAS 25 AND 24 MONTHS (P ¼ 0.02) FOR THE METASTATIC AND PRIMARY GLIAL NEOPLASM GROUPS, RESPECTIVELY. ONE PATIENT EXPERIENCED IMMEDIATE POST-PROCEDURAL MORBIDITY SECONDARY TO INCREASED INTRACEREBRAL EDEMA PERI-LESIONALLY WHILE ONE PATIENT EXPERIENCED POSTOPERATIVE MORTALITY AND EXPIRED SECONDARY TO HEMORRHAGE 1-MONTH POST-PROCEDURE. MEDIAN FOLLOW-UP WAS 10 MONTHS. CONCLUSION: LASER INTERSTITIAL THERMAL THERAPY (LITT) IS A SAFE, MINIMALLY INVASIVE TREATMENT METHOD THAT PROVIDES SURGEONS WITH CYT OREDUCTIVE TECHNIQUES TO TREAT NEUROSURGICAL CONDITIONS. BOTH PFS AND OS APPEAR TO BE MORE FAVORABLE AFTER LITT IN PATIENTS WITH METASTATIC DISEASE. IN PROPERLY SELECTED PATIENTS, THIS MODALITY OFFERS IMPROVED SURVIVAL OUTCOMES IN CONJUNCTION WITH OTHER SALVAGE THERAPIES. REPORTED EVENTS: PATIENT 13 EXPIRED 3 MONTHS AFTER LITT DUE TO CARDIAC ARREST PATIENT 15 EXPIRED 7 MONTHS AFTER LITT DUE TO UROSEPSIS ONE PATIENT (1/19, 5.3%) EXPERIENCED IMMEDIATE POST-OPERATIVE MORBIDITY PATIENT 10 DEVELOPED INCREASED EDEMA PERILESIONALLY ON POST-PROCEDURE DAY 2 (FIGURE 4) AND ULTIMATELY DEVELOPED RIGHT-SIDED FOOT DROP AND SYMPTOMS OF CHRONIC CAUDA EQUINA DUE TO PROGRESSION OF DISEASE 2 YEARS POST-PROCEDURALLY AFTER METASTATIC SPREAD OF NON-SMALL CELL LUNG CANCER TO THE SPINAL CORD. THE PATIENT ULTIMATELY EXPIRED 25 MONTHS AFTER UNDERGOING LITT. ONLY A SINGLE PATIENT (1/19; 5.3%) EXPERIENCEDPOST-OPERATIVE MORTALITY WITHIN THE FIRST 30 DAYS OF LITT THERAPY DUE TO POSTPROCEDURAL HEMORRHAGE. | A2) PATIENT AGE IS THE MEAN VALUE OF PATIENTS IN THE STUDY. A3) PATIENT GENDER IS THE MAJORITY VALUE OF PATIENT IN THE STUDY. A4) PATIENT WEIGHT NOT AVAILABLE FROM THE SITE. B3) EVENT DATE IS THE ONLINE PUBLISHING DATE OF THE LITERATURE ARTICLE. D4) DEVICE LOT NUMBER, OR SERIAL NUMBER, UNAVAILABLE. G4) 510(K) IS DEPENDENT UPON THE DEVICE MODEL NUMBER AND THEREFORE, UNAVAILABLE. H3,H6) NO PARTS HAVE BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. H4) DEVICE MANUFACTURING DATE IS DEPENDENT ON LOT NUMBER/SERIAL NUMBER, THEREFORE, UNAVAILABLE. MEDICAL ASSESSMENT WAS PERFORMED AND THEY FOUND THAT DUE TO THE AGGRESSIVE NATURE AND PRIOR SURGERY HISTORY, PATIENTS OFTEN HAVE A S URVIVAL RATE OF 1 YEAR EVEN PRIOR TO RECEIVING LITT PROCEDURE. THE MAIN GOAL WAS TO PROVIDE AN ADDITIONAL TREATMENT OPTION WHEN REVISION SURGERY WAS NOT AVAILABLE. MEDICAL SAFETY ASSESSES REPORTED DEATH OF 10 OUT OF 19 STUDY PARTICIPANTS WITHIN 2 YEARS POST LITT P ROCEDURE AND THE REPORTED SEVERITY IN THESE CASES AS INHERENT DISEASE PROGRESSION AND MEDICAL COMPLICATIONS ASSOCIATED WITH AGGRESSIVE PREVIOUSLY TREATED PRIMARY AND METASTATIC BRAIN TUMORS. ONE PATIENT EXPERIENCED INCREASED EDEMA SURROUNDING THE LASER ABLATION ZONE 2 DAYS POST PROCEDURE, AND ONE PATIENT EXPERIENCED HEMORRHAGE AS EVIDENCE ON MRI WITH DEATH OCCURRING WITHIN 30 DAYS POST PROCEDURE. MEDICAL SAFETY ASSESSES THE MORBIDITY OUTCOMES AND THEIR REPORTED SEVERITIES IN THESE CASES AS LABELED COMPLICATIONS. PER VISUALASE¿ MRI-GUIDED LASER ABLATION SYSTEM, WHICH STATES ¿POTENTIAL ADVERSE EVENTS ASSOCIATED WITH DAMAGE TO CRITICAL STRUCTURES MAY INCLUDE: INCREASED INTRACEREBRAL EDEMA OR PRESSURE. INTRACRANIAL BLEEDING¿. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A RENTAL DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA, LUNG DISEASE, AND REDUCED CARDIOPULMONARY RESERVE. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIBERATOR | K800742 | AS REPORTED: THE LEVEL INDICATOR SHOWED THAT THE DEVICE WAS STILL HALF FULL. THE DEVICE WAS DISCOVERED AROUND 11:00 P.M PUT INTO OPERATION (12 L/MIN). DURING AN INSPECTION (AROUND 1:30 P.M.) THE DEVICE STILL HAS OXYGEN DELIVERED AND THE RESIDENT LIVED. WHEN THE RELATIVES ENTERED THE ROOM ABOUT AN HOUR LATER, THERE WAS NO OXYGEN DELIVERY AND THE FILL LEVEL INDICATOR WAS RED. AT THIS POINT THE RESIDENT HAD ALREADY DIED. NOTE: CAIRE LIBERATOR LIQUID OXYGEN SYSTEMS PROVIDE SUPPLEMENTAL OXYGEN AND ARE NOT INTENDED OR MARKETED FOR LIFE SUPPORT FUNCTIONS PER THE DEVICE'S TECHNICAL DOCUMENTATION. A SECONDARY SOURCE OF SUPPLEMENTAL OXYGEN MUST BE ON HAND PER THE DEVICE IFU. | PURSUANT TO TITLE 21 - FOOD AND DRUGS, CHAPTER I - FOOD AND DRUG ADMINISTRATION DEPARTMENT OF HEALTH AND HUMAN SERVICES, SUBCHAPTER H -0 MEDICAL DEVICE, PART 803 - MEDICAL DEVICE REPORTING, SUBPART A - GENERAL PROVISIONS, SECTION 803.16, NEITHER THIS REPORT NOR ANY INFORMATION SUBMITTED HEREIN CONSTITUTES AN ADMISSION BY CAIRE INC. THAT THE DEVICE STATED IN THIS REPORT, CAIRE INC., OR CAIRE INC.'S EMPLOYEES, CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT STATED HEREIN. CAIRE CUSTOMER SERVICE HAS REQUESTED THE UNIT BE RETURNED TO (B)(6) FACILITY FOR EVALUATION. A FINAL REPORT WILL BE SUBMITTED WITH THE RESULTS OF THE INVESTIGATION IF THE UNIT BECOMES AVAILABLE FOR EVALUATION. SECTION D4 PRIMARY UNIQUE DEVICE IDENTIFIER (UDIS) # IS INTENTIONALLY BLANK PER RULE 21 CFR 801.30(1). THE DEVICE WAS MANUFACTURED AND LABELED IN 2012.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), CANCER, DEATH NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED TO FRESENIUS CUSTOMER SERVICE THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS FOUND UNRESPONSIVE BY A RELATIVE WHEN THEY WENT TO DISCONNECT THE PATIENT FROM THE LIBERTY SELECT CYCLER. THE PATIENT HAD REPORTEDLY PASSED AWAY. ADDITIONAL INFORMATION WAS PROVIDED UPON FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). IT WAS CONFIRMED THE PATIENT WAS CONNECTED TO THE CYCLER WHEN THEY WERE FOUND UNRESPONSIVE BY THE RELATIVE. THE PATIENT¿S TREATMENT HAD ALREADY BEEN COMPLETED. 911 WAS CALLED AND PARAMEDICS RESPONDED TO THE PATIENT¿S HOME. THE PATIENT WAS PRONOUNCED DECEASED UPON ARRIVAL OF EMERGENCY MEDICAL SERVICES (EMS). THERE WERE NO KNOWN ISSUES WITH THE LIBERTY SELECT CYCLER PRIOR TO THE TREATMENT. THE CAUSE OF DEATH IS UNKNOWN; HOWEVER, THE PATIENT HAD SIGNIFICANT CARDIAC HISTORY INCLUDING ATRIAL FIBRILLATION (A FIB), ISCHEMIC CARDIOMYOPATHY, CORONARY ARTERY DISEASE WITH STENT, AND HEART FAILURE WITH REDUCED LEFT VENTRICULAR EJECTION FRACTION (LVEF) = 30% WITH A HISTORY OF PLEURAL FLUID. THE PATIENT RECENTLY HAD A THORACENTESIS PROCEDURE DUE TO THIS REASON. THE PATIENT HAD BEEN INSTRUCTED TO USE 2.5% DIALYSATE SOLUTION FOR THREE DAYS AFTER BEING SEEN BY THEIR PRIMARY CARE PHYSICIAN (PCP). IT WAS THOUGHT THERE WAS EXTRA FLUID BUILDING UP IN THE LUNGS AGAIN DUE TO THE SAME ISSUE. IT WAS REPORTED THAT THE PATIENT¿S DEATH WAS NOT RELATED TO PD THERAPY OR THE USE OF ANY FRESENIUS PRODUCTS OR DEVICES. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT ON THE PUMP, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED A CASSETTE MEMBRANE. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST, A VALVE ACTUATION TEST, AND A TEACH PUMPS TEST. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE CYCLER¿S MUSHROOM HEADS. AN INTERNAL VISUAL INSPECTION IDENTIFIED EVIDENCE OF DRIED FLUID UNDER THE PUMP ASSEMBLY AND THE BOTTOM COVER. THE CAUSE OF THE DRIED FLUID COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
RENTAL-DREAMSTATION AUTO CPAP | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL-DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE; CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
POWERPORT ISP M.R.I. IMPLANTABLE PORT, SILICONE SINGLE-LUMEN, 9.6F | K073423 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT NINE MONTHS AND TWENTY EIGHT DAYS POST A PORT PLACEMENT FOR INTRAVENOUS ACCESS, THE PATIENT ALLEGEDLY DIAGNOSED WITH BACTEREMIA. IT WAS FURTHER REPORTED THAT PATIENT DEVELOPED SEPSIS AND SEPTIC SHOCK AND WAS TREATED WITH ANTIBIOTICS. HOWEVER, THE PATIENT WAS EXPIRED AND THE CAUSE OF DEATH WAS SEPTIC SHOCK AND METHICILLIN RESISTANT STAPHYLOCOCCUS AUREUS BACTEREMIA. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. THERE WAS NOTHING FOUND TO INDICATE THERE WAS A MANUFACTURING RELATED CAUSE FOR THIS EVENT. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. THE MEDICAL RECORDS REPORT THAT OVER 9 MONTHS AFTER PORT PLACEMENT, THE PATIENT REPORTED LEFT SIDE BACK PAIN THAT RADIATED DOWN THE LEFT LEG. PATIENT WAS FOUND TO HAVE FEVER OF UNKNOWN ORIGIN AND WAS DIAGNOSED WITH SEVERE SEPSIS SECONDARY TO MRSA BACTEREMIA. HE WAS STARTED ON INTRAVENOUS VANCOMYCIN WITH LOADING DOSE. CHEST X-RAY WAS PERFORMED FOR HISTORY OF SEPSIS. THE STUDY SHOWED LEFT SIDED PORT-A-CATH REMAINS. PATIENT EXPERIENCED BRIEF SVT AND HAD MODERATE PERICARDIAL EFFUSION. CODE BLUE CALLED FOR PEA NOTED ON CARDIAC MONITOR. PATIENT WAS BLUE IN COLOR, APNEIC, CHEST COMPRESSIONS WERE INITIATED, AND ACLS WAS STARTED. HEART RATE RETURNED TO LOW 100 AND FAMILY WAS NOTIFIED OF CODE EVENT AND SUCCESSFUL RESUSCITATION. ON (B)(6) 2024, THE INTENSIVIST MET WITH FAMILY FOR MEETING AND THE FAMILY DECIDED UPON WITHDRAWAL OF CARE. PATIENT WAS EXPIRED DUE TO SEPTIC SHOCK AND MRSA BACTEREMIA. THEREFORE, IT CAN BE CONFIRMED THAT THE PATIENT EXPERIENCED BACTEREMIA, SEPSIS, SEPTIC SHOCK, AND EXPIRED. HOWEVER, THE RELATIONSHIP TO THE PORT IS UNKNOWN. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE IS UNKNOWN. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: G3, H6 (METHOD). H11: D4 (UNIQUE IDENTIFIER (UDI) #), H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
MO.MA ULTRA | K092177 | LITERATURE WAS REVIEWED REGARDING TREATMENT RESULTS OF CAROTID ARTERY STENTING WITH AN OPEN-CELL STENT: ANALYSIS OF 734 CONSECUTIVE CASES AT A SINGLE CENTER. CAS PROCEDURES WERE PERFORMED FROM APRIL 2002 TO DECEMBER 2019. OF THESE 948 CAS PROCEDURES, 811 WERE PERFORMED WITH OPEN-CELL STENTS FROM 'MULTIPLE MANUFACTURER¿S, INCLUDING MEDTRONIC'S PROTÉGÉ RX. AN EMBOLIC PROTECTION SYSTEM WAS SELECTED FOR EVERY PATIENT. AN EMBOLIC PROTECTION SYSTEM WAS SELECTED FOR EVERY PATIENT. FROM 2010 ONWARD, THE MAIN METHOD WAS TRIPLE-BALLOON PROTECTION WITH THE PERCUSURGE GUARDWIRE AS THE DISTAL PROTECTION DEVICE AFTER LESION CROSSING AND THE MO.MA ULTRA FOR PROXIMAL BALLOON PROTECTION DURING LESION CROSSING. PERIPROCEDURAL COMPLICATIONS INCLUDED STOKE, IN-STENT PLAQUE PROTRUSION, HYPERFUSION SYNDROME, AND MORTALITY RATE OF 3 PATIENTS. THE CAUSES OF DEATH WERE CEREBRAL HEMORRHAGE DUE TO HYPERPERFUSION SYNDROME, ACUTE OCCLUSION AFTER CAS, AND HEMORRHAGIC SHOCK DUE TO INTRAPERITONEAL BLEEDING RELATED PUNCTURE. THE MEDIAN FOLLOW-UP PERIOD WAS 50 MONTHS AND OUTCOMES INCLUDING COMPLICATIONS BEYOND 30 DAYS AFTER TREATMENT WERE CEREBRAL INFARCTION, IPSILATERAL CEREBRAL INFARCTION, RESTENOSIS OR OCCLUSION, AND REINTERVENTION. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | LITERATURE REFERENCE: DOI.ORG/10.1016/J.WNEU.2024.04.108 A2: AVERAGE AGE A3A: MAJORITY SEX MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER SL 1 CURVE, 63 CM LENGTH, 8F | K052644 | THIS REPORT INCLUDES AN UPDATED DEVICE PROBLEM CODE. | CORRECTED DATA: B5, H2, H6. ADDITIONAL INFORMATION: D9, G3, H2, H3, H6. AN EVENT OF CARDIAC TAMPONADE AND DEATH WAS REPORTED. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED INCIDENT COULD NOT BE CONCLUSIVELY DETERMINED.
DREAMSTATION GO AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGED RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, CANCER AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED BY THE PATIENT. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
EMBOSHIELD NAV6 | K191173 | THIS IS FILED TO REPORT PATIENT DEATH. IT WAS REPORTED THROUGH THE PMCF (POST-MARKET CLINICAL FOLLOW-UP) REPORT, THAT ABBOTT OBTAINED DATA FROM THE VASCULAR QUALITY INITIATIVE (VQI) REGISTRY ACROSS UNITED STATES (US). OVERALL, THE DATA PRESENTED THROUGH THIS METHODOLOGY CONFIRMS THE SAFETY AND PERFORMANCE OF RX ACCULINK CAROTID STENT SYSTEM (CSS), X.ACT CSS, AND EMBOSHIELD NAV6 EMBOLIC PROTECTION SYSTEM (EPS). ADVERSE PATIENT EFFECTS FOR EMBOSHIELD NAV6 INCLUDE TRANSIENT ISCHEMIC ATTACK, STROKE, MYOCARDIAL INFARCTION, DEATH, PULMONARY COMPLICATIONS, RENAL COMPLICATIONS, ACCESS SITE COMPLICATIONS, THROMBOSIS, EMBOLIZATION, DISSECTION, PERFORATION, MEDICATION ADMINISTRATION, INTERVENTION, AMPUTATION / SURGERY, OCCLUSION, RE-HOSPITALIZATION AND CAROTID ARTERY SPASM. DEVICE ISSUES CAPTURED WERE UNABLE TO INSERT AND DIFFICULTY REMOVING THE FILTER. | THE DEVICE WAS NOT RETURNED FOR EVALUATION. CORRECTIVE AND PREVENTIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. THE LOT HISTORY RECORD (LHR) FOR THIS PRODUCT COULD NOT BE REVIEWED AND A SIMILAR COMPLAINT QUERY COULD NOT BE PERFORMED BECAUSE THE PRODUCT WAS NOT RETURNED FOR EVALUATION AND THE PART AND LOT NUMBERS WERE NOT REPORTED. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE EMBOSHIELD NAV6 INSTRUCTIONS FOR USE, AS A KNOWN ADVERSE EVENT POTENTIALLY ASSOCIATED WITH CAROTID STENTS AND EMBOLIC PROTECTION SYSTEMS. BASED ON THE REPORTED INFORMATION, THERE IS NO INDICATION THAT THE SUBSEQUENT PATIENT EFFECTS ARE RELATED TO A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. A CONCLUSIVE CAUSE FOR THE REPORTED PATIENT EFFECT AND THE RELATIONSHIP TO THE PRODUCT, IF ANY, CANNOT BE DETERMINED. B2, B3: ESTIMATED EVENT DATE IS 03/01/2024. D4- THE UDI IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. THE ADDITIONAL PATIENT EFFECTS AND MALFUNCTIONS LISTED IN THE REPORT ARE CAPTURED UNDER SEPARATE MEDWATCH REPORTS. ATTACHMENT ARTICLE TITLED: POST MARKET CLINICAL FOLLOW-UP EVALUATION REPORT CAROTID STENT AND EMBOLIC PROTECTION SYSTEMS.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS THIRTY-TWO ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT CORRECTED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA, KIDNEY DISEASE /TOXICITY, LUNG DISEASE, INFLAMMATORY RESPONSE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED AND BOX G: REPORT SOURCE - OTHER & 510K (RFB),BOX E : REPORTING INSTITUTION NAME UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING) LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DS2ADV AUTO CPAP DEVICE. THERE WAS AN ALLEGATION OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PLUS | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF NOSE IRRITATION, SKIN IRRITATION, DIZZINESS AND/OR HEADACHE, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
STARTER GUIDEWIRE | K935170 | EVENT DESCRIPTION: ECN EVENT DESCRIPTION: PHYSICIAN CALLED ME. SAID HE IS INVOLVED IN A MALPRACTICE SUIT. SAID HE USED A 'STARTER' WIRE DURING A HEART CATH. SAID HE WAS LOOKING FOR PRODUCT INFORMATION AND TO SEE IF THE WIRE COULD MAKE A PARTICULAR TURN. I ASKED IF HE COULD PROVIDE MORE DETAIL ON THE WIRE. HE SAID THE TIP WAS A 3MM SEMI-CIRCLE WIRE. I ADVISED HIM ON THE TIP SHAPES THAT WE SELL FOR OUR PERIPHERAL STARTER FAMILY OF WIRES: STRAIGHT TIP, J TIP, (B)(6). I ALSO SAID THE CLOSEST WIRE WE HAVE TO WHAT HE IS DESCRIBING IS A J WIRE. HE CLAIMS THAT IS NOT THE WIRE HE USED. I LISTED THE FIXED CORE J WIRE ON THIS REPORT, AS ITS THE WIRE THAT IS ORDERED MOST FREQUENTLY AT THE ACCOUNT. HE IS INTERESTED IN THE DEVICE SPECIFICATIONS OF THE STARTER WIRE, AND TORQUE RESPONSIVENESS AND GENERAL WIRE CAPABILITIES, AS HE BELIEVES THAT THE WIRE HE USED COULD NOT HAVE MADE THE TURN THAT WOULD HAVE CAUSED THE DEATH. PATIENT PRESENT AT TIME OF EVENT? YES, PATIENT COMPLICATIONS: DEATH IN THE PHYSICIAN'S OPINION, DID THE DEVICE OR PROCEDURE CAUSE OR CONTRIBUTE TO THE PATIENT COMPLICATION? NO MEDICAL OR SURGICAL INTERVENTIONS: NA PATIENT ADMITTED TO HOSPITAL BEYOND THE STANDARD OF CARE: UNKNOWN PATIENT OUTCOME: UNKNOWN PROCEDURE DATE-UNKNOWN TIME OF EVENT - OTHER: NOT. KNOWN TIME OF THE EVENT: OTHER. EVENT DATE: NO VALUE FOUND. AS REPORTED DEVICE CODES: 1163: DAMAGED/DEFECTIVE. AS REPORTED PATIENT CODES: 9398: NO CLINICAL SIGNS, SYMPTOMS OR CONDITIONS. DEVICE/PROCEDURE OUTCOME: UNABLE TO USE DEVICE - ATTEMPTED, PROCEDURE - NO INFORMATION AVAILABLE. PATIENT OUTCOME: F02: DEATH. 1. WAS THERE FAILURE MODE ATTRIBUTED TO THE DEVICE? A. PLEASE EXPLAIN FURTHER WHAT DOES "THE WIRE HE USED COULD NOT HAVE MADE THE TURN THAT WOULD HAVE CAUSED THE DEATH" MEANS. A: IT IS MY UNDERSTANDING OF THE PHYSICIAN'S STATEMENT - THAT HE CLAIMS THAT THE WIRE IN QUESTION COULD NOT HAVE MADE THE TURNS. NEEDED TO CANNULATE TO THE AREA THAT CAUSED THE PATIENT'S DEATH. 2. WAS THERE ANY DAMAGE NOTED TO THE PRODUCT PACKAGING UPON INSPECTION PRIOR TO USE? A: INFORMATION IS NOT KNOWN - AND NOTHING WAS STATED. 3. WAS THERE ANY REPORTED DIFFICULTY REMOVING THE PRODUCT FROM THE PACKAGING? A: INFORMATION IS NOT KNOWN - AND NOTHING WAS STATED. 4. WAS THE PRODUCT INSPECTED PRIOR TO USE AND APPEAR TO BE NORMAL? A: INFORMATION IS NOT KNOWN - AND NOTHING WAS STATED. 5. WILL THE DEVICE BE RETURNED FOR ANALYSIS? A: NO. 6. WHAT IS THE LOT# OR CATALOGUE# FOR THE DEVICE? A: LOT INFORMATION IS NOT KNOWN - UPN # IS: M001491211. 7. WHAT ACTIONS WERE PERFORMED ON THE GUIDEWIRE TO PREPARE THE GUIDEWIRE FOR USE? A: INFORMATION IS NOT KNOWN - AND NOTHING WAS STATED. 8. WERE ANY PROBLEMS NOTED DURING PREPARATION? A: INFORMATION IS NOT KNOWN - AND NOTHING WAS STATED. 9. WAS THE GUIDEWIRE PREPPED PROPERLY ACCORDING TO THE IFU? A: INFORMATION IS NOT KNOWN. 10. LIST/NAME ANY DEVICE THAT WAS USED WITH THE GUIDEWIRE. A: A 6F TERUMO GLIDE SHEATH WAS USED IN ACCORDANCE TO THE IMAGE ABOVE REGARDING THE PROCEDURE NOTES. OTHER DEVICES WERE USED IN THE PROCEDURE; HOWEVER, I DO NOT KNOW IF THIS DEVICE WAS USED ALONE OR IN CONJUNCTION WITH OTHER CATHETERS ETC. 11. WAS THERE DIFFICULTY ADVANCING THE DEVICES OVER THE GUIDE WIRE? A: INFORMATION IS NOT KNOWN - AND NOTHING WAS STATED. 12. WAS THERE ANY DIFFICULTY ADVANCING THE GUIDEWIRE TOWARDS THE TARGET LESION? A: INFORMATION IS NOT KNOWN. IT IS MY UNDERSTANDING AS THE PHYSICIAN DESCRIBES IT, THAT THERE IS NO WAY THAT THIS WIRE COULD TAKE THE TURN THAT WOULD BE REQUIRED TO LEAD THE WIRE TO THE VASCULAR BED THAT CAUSED THE PATIENT'S DEATH. 13. IS THERE ANY TARGET LESION/TARGET LESION CHARACTERISTIC INFORMATION AVAILABLE? A: INFORMATION IS NOT KNOWN ABOUT THE INTENDED TARGET LESION. THE PHYSICIAN DID DESCRIBE "A SPONTANEOUS POSTERIOR CEREBRAL ARTERY HEMORRHAGE". 14. WAS UNUSUAL FORCE EVER REQUIRED WHEN USING THE DEVICE? A: INFORMATION IS NOT KNOWN. NOTHING WAS STATED. 15. WAS THE DEVICE REMOVED EASILY AND INTACT FROM THE PATIENT? A: INFORMATION IS NOT KNOWN, NOTHING WAS STATED 16. WHAT IS THE DATE OF OCCURRENCE OF THE EVENT? A: INFORMATION WAS NOT DISCLOSED 17. WHAT IS THE PHYSICIAN'S OPINION OF THE RELATIONSHIP BETWEEN THE DEVICE AND PATIENT DEATH? A: THE PHYSICIAN BELIEVES THAT THIS WIRE COULD NOT HAVE CAUSED THE PERFORATION THAT CAUSED THE DEATH. | THERE WAS NOT ENOUGH INFORMATION AVAILABLE TO PROVIDE AN ACCURATE ACCOUNT ON WHAT MOVEMENTS WERE POSSIBLE THROUGHOUT THE SURGERY. FROM WHAT INFORMATION WAS GIVEN, IT CAN BE CONFIRMED THAT THE GUIDEWIRE WAS NOT DEFECTIVE AND WAS NOT THE CAUSE OF THE SURGERY'S OUTCOME. THE END USER DID NOT PROVIDE LOT TRACEABILITY; THEREFORE, LAKE REGION MEDICAL WAS UNABLE TO REVIEW THE DEVICE HISTORY RECORDS RELATIVE TO THE MANUFACTURING, INSPECTING, AND PACKAGING OF THIS PRODUCT.
BENTSON TYPE PTFE COATED GUIDEWIRE | K935170 | RELATED PRODUCT FAMILY: STANDARD TEFLON MATERIAL NUMBER: M0066201110 STERILE LOT NUMBER: NO VALUE FOUND SOLD TO ACCOUNT NUMBER: NO VALUE FOUND RETURNED PRODUCT EXPECTED: NO DEVICE/PROCEDURE OUTCOME: DEVICE - NO INFORMATION AVAILABLE, PROCEDURE - NO INFORMATION AVAILABLE PATIENT OUTCOME: F02: DEATH. EVENT DESCRIPTION: ECN EVENT DESCRIPTION: PHYSICIAN CALLED ME. SAID HE IS INVOLVED IN A MALPRACTICE SUIT. SAID HE USED A 'STARTER' WIRE DURING A HEART CATH. SAID HE WAS LOOKING FOR PRODUCT INFORMATION AND TO SEE IF THE WIRE COULD MAKE A PARTICULAR TURN. I ASKED IF HE COULD PROVIDE MORE DETAIL ON THE WIRE. HE SAID THE TIP WAS A 3MM SEMI-CIRCLE WIRE. I ADVISED HIM ON THE TIP SHAPES THAT WE SELL FOR OUR PERIPHERAL STARTER FAMILY OF WIRES: STRAIGHT TIP, J TIP, ROSEN, BENTSON AND NEWTON. I ALSO SAID THE CLOSEST WIRE WE HAVE TO WHAT HE IS DESCRIBING IS A J WIRE. HE CLAIMS THAT IS NOT THE WIRE HE USED. I LISTED THE FIXED CORE J WIRE ON THIS REPORT, AS ITS THE WIRE THAT IS ORDERED MOST FREQUENTLY AT THE ACCOUNT. HE IS INTERESTED IN THE DEVICE SPECIFICATIONS OF THE STARTER WIRE, AND TORQUE RESPONSIVENESS AND GENERAL WIRE CAPABILITIES, AS HE BELIEVES THAT THE WIRE HE USED COULD NOT HAVE MADE THE TURN THAT WOULD HAVE CAUSED THE DEATH. PATIENT PRESENT AT TIME OF EVENT? YES, PATIENT COMPLICATIONS: DEATH IN THE PHYSICIAN'S OPINION, DID THE DEVICE OR PROCEDURE CAUSE OR CONTRIBUTE TO THE PATIENT COMPLICATION: NO MEDICAL OR SURGICAL INTERVENTIONS: NA PATIENT ADMITTED TO HOSPITAL BEYOND THE STANDARD OF CARE: UNKNOWN PATIENT OUTCOME: UNKNOWN PROCEDURE DATE-UNKNOWN TIME OF EVENT - OTHER: NOT KNOWN TIME OF THE EVENT: OTHER EVENT DATE: NO VALUE FOUND AS REPORTED DEVICE CODES: 2456: USE OF DEVICE ISSUE - MISUSE AS REPORTED PATIENT CODES: NO VALUE FOUND. | THE END USER DID NOT PROVIDE LOT TRACEABILITY; THEREFORE, LAKE REGION MEDICAL WAS UNABLE TO REVIEW THE DEVICE HISTORY RECORDS RELATIVE TO THE MANUFACTURING, INSPECTING, AND PACKAGING OF THIS PRODUCT.
BIPAP AUTO BIFLEX, | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, INFLAMMATORY RESPONSE, CANCER AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, INFLAMMATORY RESPONSE, CANCER AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH DUE TO RESPIRATORY TRACT IRRITATION, LUNG DISEASE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K163584 | IT WAS REPORTED THE ALARMS WERE SILENCED AT THE CENTRAL STATION WITHOUT CLINICIAN INPUT. NURSES INDICATED NO ONE AT THE CENTRAL STATION SILENCED THE ALARMS OR TOUCHED THE MOUSE, YET THE ALARMS AUTO-SILENCED, PER THE CUSTOMER. THE CUSTOMER'S ALARM CONFIGURATION WAS UNKNOWN AT THE TIME OF THIS ASSESSMENT. GOOD FAITH EFFORTS ARE ONGOING. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. PATIENT CALLED REGARDING THE PATIENT LISTED AS CATHERINE. WANTS ALL COMMUNICATION STOPPED BECAUSE THE PATIENT PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA/VOMITING. THERE WAS A REPORT OF PATIENT DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA, VOMITING, ASTHMA, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA, LUNG DISEASE, DEATH AND CANCER. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, ASTHMA AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION REGARDING LUNG DISEASE, CANCER AND "DEATH". IN ADDITION, THE PATIENT ALSO ALLEGED RESPIRATORY TRACT IRRITATION, DIZZINESS AND HEADACHE. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | THE CUSTOMER REPORTED THAT THE DEVICE WILL BE RETURNED DUE TO THE USER PASSING AWAY CURRENTLY THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. MORE INFORMATION REGARDING THE EVENT IS BEING REQUESTED. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | BASED ON THE AVAILABLE INFORMATION, THIS COMPLAINT IS A DUPLICATE OF MFR #2518422-2024-100070. REFER TO MFR2518422-2024-100070 FOR ADDITIONAL INFORMATION.
CERELINK ICP PROBE 1L M BOLT | K173192 | N/A. | THE CERELINK SENSOR WAS NOT RETURNED FOR EVALUATION; THEREFORE, AN EVALUATION (AS PER CUSTOMER, PRODUCT NOT AVAILABLE) OF THE DEVICE COULD NOT BE PERFORMED. LOT NUMBER INFORMATION HAS BEEN PROVIDED; THEREFORE, DEVICE HISTORY RECORD (DHR) WAS REVIEWED, AND NO ANOMALIES WERE FOUND. THE ROOT CAUSE(S) OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE IN THE FUTURE, THIS COMPLAINT WILL BE REOPENED, AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR ISSUES. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED RESPIRATORY TRACT IRRITATION, REDUCED CARDIOPULMONARY. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. | THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTRE FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO PARTICLES IN AIR PATH. THERE WAS NO ERROR CODE FOUND, UNIT WAS CORRECTED. EVALUATION METHOD CODE, RESULT CODE AND CONCLUSION CODE HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. BOX H: PATIENT OUTCOME CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED. IN PREVIOUS REPORT REDUCED CARDIOPULMONARY RESERVE DIDN'T CAPTURED IN PATIENT OUTCOME CODE GRID. IN THIS HAS BEEN UPDATED.
FISHER & PAYKEL HEALTHCARE | K220703 | A HEALTHCARE FACILITY IN POLAND REPORTED, VIA A FISHER & PAYKEL HEALTHCARE (F&P) FIELD REPRESENTATIVE, AN EVENT THAT OCCURRED WHILE ADMINISTERING HIGH FREQUENCY OSCILLATORY VENTILATION (HFOV) TO A PREMATURE NEONATE OF 26 WEEKS OF GESTATIONAL AGE. THE PATIENT PASSED AWAY ON THE SEVENTH DAY OF VENTILATION. THE HEALTHCARE FACILITY REPORTED THAT THE FOLLOWING EQUIPMENT WAS USED TO ADMINISTER HFOV TO THE PATIENT: A NON-F&P VENTILATOR; A NON-F&P HME FILTER (PLACED AT THE EXPIRATORY PORT OF THE VENTILATOR); AN F&P 950AEU RESPIRATORY HUMIDIFIER (F&P 950AEU HUMIDIFIER), AN F&P 950N81 NEONATAL VENTILATOR DUAL HEATED CIRCUIT KIT (F&P 950N81 CIRCUIT). THE HEALTHCARE FACILITY REPORTED THE FOLLOWING: THROUGHOUT THE PERIOD OF TREATMENT, THIS PATIENT RECEIVED FIO2 OF 100%. CONDENSATION WAS OBSERVED IN THE EXPIRATORY LIMB OF THE F&P 950N81 CIRCUIT AND THE HME FILTER USED ON THE VENTILATOR'S EXHALATION VALVE REQUIRED FREQUENT REPLACEMENT. THE F&P 950N81 CIRCUIT NEXT TO THE PATIENT REMAINED DRY. VENTILATION PARAMETERS FOR THE PATIENT WERE INCREASED OVER SEVERAL DAYS. ON THE FIFTH DAY OF VENTILATION, THE SETTINGS ON THE VENTILATOR WERE CLOSE TO MAXIMUM AND THE PATIENT WAS REINTUBATED WITH A LARGER ENDOTRACHEAL TUBE, AND NITRIC OXIDE THERAPY WAS STARTED. ON THAT SAME DAY, THE F&P 950AEU HUMIDIFIER, F&P 950N81 BREATHING CIRCUIT, NON-F&P VENTILATOR EXPIRATORY VALVE, FLOW SENSORS AND HME FILTER WERE ALL REPLACED AS "WATER RESISTANCE IN THE EXHALATION VALVE DISTURBED THE PROPER OPERATION OF THE VENTILATOR". FOLLOWING THESE REPLACEMENTS, THE VENTILATOR PARAMETERS STABILIZED, AND THE HME FILTER REMAINED DRY FOR SIX HOURS. THE PATIENT WAS SUBSEQUENTLY TRANSITIONED TO CONVENTIONAL INVASIVE MECHANICAL VENTILATION DUE TO "EXTREME CIRCULATORY FAILURE". CONDENSATION WAS OBSERVED IN THE INSPIRATORY LIMB OF THE F&P 950N81 CIRCUIT. THE HEALTHCARE FACILITY CONFIRMED THAT THE PATIENT ULTIMATELY PASSED AWAY ON THE SEVENTH DAY OF VENTILATION. THE CAUSE OF DEATH IS REPORTED TO BE CARDIOVASCULAR RESPIRATORY FAILURE. FISHER & PAYKEL HEALTHCARE (F&P) HAS REQUESTED FURTHER INFORMATION REGARDING THE EVENT AND THE RETURN OF THE SUBJECT DEVICES FOR EVALUATION. | (B)(4). F&P DEVICES INVOLVED IN THE EVENT: F&P 950AEU RESPIRATORY HUMIDIFIER: SECTION D4: MODEL NUMBER - 950AEU, SECTION D4: CATALOG NUMBER - 950AEU, SECTION D4: SERIAL NUMBER - (B)(6). SECTION D4: LOT NUMBER - 2102042038. SECTION D4: UDI - (B)(4). SECTION G4: 510(K) NUMBER - K220703, THE 950AEU RESPIRATORY HUMIDIFIER IS NOT SOLD IN THE UNITED STATES OF AMERICA (USA) BUT INSTEAD A SIMILAR PRODUCT, 950JUS RESPIRATORY HUMIDIFIER IS SOLD IN THE USA. THE 510(K) FOR 950JUS IS K220703. SECTION H4: MANUFACTURE DATE - 18-FEB-2022. F&P 950SO2 SENSOR CARTRIDGE: SECTION D4: MODEL NUMBER - 950SO2, SECTION D4: CATALOG NUMBER - 950SO2, SECTION D4: SERIAL NUMBER - (B)(6), SECTION D4: LOT NUMBER - 2102260204, SECTION D4: UDI - (B)(4), SECTION G4: 510(K) NUMBER - K220703, SECTION H4: MANUFACTURE DATE - 13-JUL-2022. F&P 950N81 NEONATAL VENTILATOR DUAL HEATED CIRCUIT KIT: SECTION D4: MODEL NUMBER - 950N81, SECTION D4: CATALOG NUMBER - 950N81, SECTION D4: SERIAL NUMBER - NOT APPLICABLE, SECTION D4: LOT NUMBER - 2102790133, SECTION D4: UDI - (B)(4), SECTION G4: 510(K) NUMBER - K220703, THE 950N81 NEONATAL VENTILATOR DUAL HEATED CIRCUIT KIT IS NOT SOLD IN THE UNITED STATES OF AMERICA (USA) BUT INSTEAD A SIMILAR PRODUCT, 950N81J NEONATAL VENTILATOR DUAL HEATED CIRCUIT KIT IS SOLD IN THE USA. THE 510(K) FOR 950N81J IS K220703. SECTION H4: MANUFACTURE DATE - 18-SEP-2023. F&P 950AEU RESPIRATORY HUMIDIFIER, SECTION D4: MODEL NUMBER - 950AEU, SECTION D4: CATALOG NUMBER - 950AEU, SECTION D4: SERIAL NUMBER - (B)(6), SECTION D4: LOT NUMBER - 2102042038, SECTION D4: UDI - (B)(4), SECTION G4: 510(K) NUMBER - K220703, THE 950AEU RESPIRATORY HUMIDIFIER IS NOT SOLD IN THE UNITED STATES OF AMERICA (USA) BUT INSTEAD A SIMILAR PRODUCT, 950JUS RESPIRATORY HUMIDIFIER IS SOLD IN THE USA. THE 510(K) FOR 950JUS IS K220703. SECTION H4: MANUFACTURE DATE - 18-FEB-2022. F&P 950N81 NEONATAL VENTILATOR DUAL HEATED CIRCUIT KIT WITH UNKNOWN LOT NUMBER DETAILS. FISHER & PAYKEL HEALTHCARE (F&P) HAS REQUESTED FURTHER INFORMATION REGARDING THE EVENT AND THE RETURN OF ALL F&P DEVICES. WE ARE IN THE PROCESS OF COMPLETING OUR INVESTIGATION. WE ARE CURRENTLY AWAITING THE RETURN OF F&P 950AEU HUMIDIFIER (BATCH/ SERIAL (B)(6)) AND F&P 950N81 CIRCUITS FOR EVALUATION. WE WILL PROVIDE A FOLLOW-UP REPORT UPON COMPLETION OF OUR INVESTIGATION.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, LUNG DISEASE, AND BRAIN TUMOR. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | ON PREVIOUSLY SUBMITTED REPORT DESCRIBE EVENT OR PROBLEM HAS BEEN SUBMITTED WRONGLY. IN THIS REPORT UPDATED AS- THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER. THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
TEMPUS LS | K200849 | IT HAS BEEN REPORTED TO PHILIPS THE DEVICE FAILURE TO PACE WHILE IN USE. | UPDATE TO THE COMPONENT CODE GRID, EVALUATION RESULTS GRID AND THE CONCLUSION CODE GRID.
CERELINK ICP MONITOR | K183406 | 2 OF 2 REPORTS. OTHER MFG REPORT NUMBER: 3013886523-2024-00272. A FACILITY REPORTED: "SENSOR IMPLANTED ON (B)(6) DURING THE NIGHT 18-19 AUGUST, INTRA-VENTRICULAR PRESSURE OF PATIENT INCREASE SUDDENLY BUT THE VALUE OF ICP READ ON CERELINK MONITOR AND ON PATIENT MONITOR DID NOT MOVE, IT WAS STABLE. SENSOR EXPLANTED ON 19 AUGUST.". - CLINICAL CONSEQUENCE: "BRAIN INVOLVEMENT". - "THEY MONITOR ALTERNATIVELY THE INTRA-VENTRICULAR PRESSURE". - HOW DID YOU DETECT THAT INTRA-VENTRICULAR PRESSURE OF PATIENT INCREASE? "BILATERAL MYDRIASIS". - DID THE INCIDENT INCREASE THE PATIENT'S CARE TIME ? "THE INCIDENT IS RESPONSIBLE FOR THE PATIENT'S DEATH". ADDITIONAL INFORMATION HAS BEEN REQUESTED. A SUPPLEMENTAL REPORT WILL BE PROVIDED IF ADDITIONAL INFORMATION BECOMES AVAILABLE. | UPDATED FIELDS: D4, D9, G3, G6, H2, H3, H6, H11. THE CERELINK MONITOR WAS NOT RETURNED FOR EVALUATION AND LOT NUMBER INFORMATION HAS NOT BEEN PROVIDED; THEREFORE, AN EVALUATION OF THE DEVICE COULD NOT BE PERFORMED, AND DEVICE HISTORY RECORD (DHR) COULD NOT BE REVIEWED. THE ROOT CAUSE(S) OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE IN THE FUTURE, THIS COMPLAINT WILL BE REOPENED, AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR ISSUES. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED.
SOLITAIRE FR4 | K183022 | MEDTRONIC RECEIVED A REPORT THAT THE PATIENT UNDERWENT MECHANICAL THROMBECTOMY OF THE RIGHT MIDDLE CEREBRAL ARTERY, SEGMENT M1/M2. THE STENT-RETRIEVER WAS SELECTED ACCORDING TO THE MANUFACTURER'S RECOMMENDATIONS (3X40). USING THE STANDARD TECHNIQUE, IN ACCORDANCE WITH THE INSTRUCTIONS PROVIDED (IFU), THE DEVICE WAS PREPARED, DELIVERED AT THE LEVEL OF THE THROMBUS, AND THEN OPENED IN THE THROMBUS. SMOOTH COURSE, WITHOUT EXCESSIVE RESISTANCE. IN THE NEXT STEPS, THE DEVICE WAS REMOVED USING THE STANDARD TECHNIQUE USING SIMULTANEOUS ASPIRATION FROM THE DISTAL ACCESS CATHETER AND THE GUIDING CATHETER. THIS PART OF THE PROCEDURE ALSO PROCEEDED WITHOUT UNUSUAL RESISTANCE. A FRAGMENT OF THROMBUS WAS FOUND ON THE TOOL. THE ANGIOGRAPHY PERFORMED DID NOT REVEAL RECANALIZATION (I.E. ONLY A FRAGMENT OF THE THROMBUS WAS REMOVED). IN ACCORDANCE WITH THE MANUFACTURER'S INSTRUCTIONS, THE STENT-RETRIEVER WAS PREPARED FOR THE NEXT USE, DURING WHICH IT WAS FOUND THAT ONE OF THE TWO LEGS ATTACHING THE STENT-RETRIEVER TO ITS GUIDE (THE PUSHING WIRE IN THE MANUFACTURER'S NOMENCLATURE) WAS DETACHED AND BENT. IN THE NEXT STEPS, AFTER REPLACING THE STENT-RETRIEVER, THE THROMBUS WAS REMOVED, ACHIEVING RECANALIZATION OF THE VESSEL. A CONTROL CT SCAN OF THE HEAD REVEALED SIGNIFICANT SUBARACHNOID BLEEDING. THE PATIENT PASSED AWAY. ANCILLARY DEVICES INCLUDE A PHENOM 21/160 CATHETER, SYNCHRO2 SOFT 0.014IN X 300CM MICROWIRE AND SOFIA LUS 6F/125CM CATHETER. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CATHERA | K151638 | MEDTRONIC RECEIVED A REPORT THAT THE PATIENT UNDERWENT MECHANICAL THROMBECTOMY OF THE RIGHT MIDDLE CEREBRAL ARTERY, SEGMENT M1/M2. THE STENT-RETRIEVER WAS SELECTED ACCORDING TO THE MANUFACTURER'S RECOMMENDATIONS (3X40). USING THE STANDARD TECHNIQUE, IN ACCORDANCE WITH THE INSTRUCTIONS PROVIDED (IFU), THE DEVICE WAS PREPARED, DELIVERED AT THE LEVEL OF THE THROMBUS, AND THEN OPENED IN THE THROMBUS. SMOOTH COURSE, WITHOUT EXCESSIVE RESISTANCE. IN THE NEXT STEPS, THE DEVICE WAS REMOVED USING THE STANDARD TECHNIQUE USING SIMULTANEOUS ASPIRATION FROM THE DISTAL ACCESS CATHETER AND THE GUIDING CATHETER. THIS PART OF THE PROCEDURE ALSO PROCEEDED WITHOUT UNUSUAL RESISTANCE. A FRAGMENT OF THROMBUS WAS FOUND ON THE TOOL. THE ANGIOGRAPHY PERFORMED DID NOT REVEAL RECANALIZATION (I.E. ONLY A FRAGMENT OF THE THROMBUS WAS REMOVED). IN ACCORDANCE WITH THE MANUFACTURER'S INSTRUCTIONS, THE STENT-RETRIEVER WAS PREPARED FOR THE NEXT USE, DURING WHICH IT WAS FOUND THAT ONE OF THE TWO LEGS ATTACHING THE STENT-RETRIEVER TO ITS GUIDE (THE PUSHING WIRE IN THE MANUFACTURER'S NOMENCLATURE) WAS DETACHED AND BENT. IN THE NEXT STEPS, AFTER REPLACING THE STENT-RETRIEVER, THE THROMBUS WAS REMOVED, ACHIEVING RECANALIZATION OF THE VESSEL. A CONTROL CT SCAN OF THE HEAD REVEALED SIGNIFICANT SUBARACHNOID BLEEDING. THE PATIENT PASSED AWAY. ANCILLARY DEVICES INCLUDE A PHENOM 21/160 CATHETER, SYNCHRO2 SOFT 0.014IN X 300CM MICROWIRE AND SOFIA LUS 6F/125CM CATHETER. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER,DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT TRRITATION AND COPD. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF CANCER, NON- HODGKIN'S LYMPHOMA, PROSTATE CANCER AND PATIENT PASSING AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
CAPTIVATOR EMR | K140726 | IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A CAPTIVATOR ENDOSCOPIC MUCOSAL RESECTION DEVICE WAS USED IN THE ESOPHAGUS DURING A PROCEDURE PERFORMED ON AN UNKNOWN DATE. THE PROCEDURE WAS COMPLETED SUCCESSFULLY; HOWEVER, WHEN THE PATIENT WENT HOME, THE PATIENT STARTED COMPLAINING OF PAIN. THE PATIENT ENDED UP GOING BACK TO THE EMERGENCY ROOM, WHERE A PERFORATION WAS FOUND IN THE ESOPHAGUS. THE PATIENT ENDED UP PASSING AWAY, BUT THE EXACT CAUSE OF DEATH WAS NOT REPORTED. THE PHYSICIAN STATED THAT THE PLACEMENT OF THE BAND MAY HAVE BEEN INTO THE MUSCLE LAYER. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD FAITH EFFORTS. THE PHYSICIAN REPORTED THAT THE INCIDENT OCCURRED FEW YEARS AGO AND NOTED THAT INFORMATION REGARDING THE EVENT WAS LIMITED. | BLOCK B3: DATE OF EVENT: DATE OF EVENT WAS APPROXIMATED TO 08/01/2024 BASED ON THE DATE THE MANUFACTURER BECAME AWARE OF THE EVENT. BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO REPORT THE SUSPECTED DEVICE LOT NUMBER; THEREFORE, THE MANUFACTURE DATE AND EXPIRATION DATE ARE UNKNOWN. BLOCK H6: IMDRF PATIENT CODES E2330 AND E2114 CAPTURE THE REPORTABLE EVENTS OF PATIENT PAIN AND PERFORATION. IMDRF IMPACT CODES F08 AND F02 ARE BEING USED TO CAPTURE THE PROLONGED HOSPITALIZATION AND THE PATIENT'S DEATH.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, HEART ATTACK. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LIVER DISEASE/TOXICITY, CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, LUNG DISEASE, KIDNEY DISEASE, TOXICITY, COUGH, SINUS INFECTIONS, PNEUMONIA, ACUTE RESPIRATORY FAILURE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY FAILURE. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, THROAT, LUNG DISEASE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, ASTHMA, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND THE SECTION H11 SHOULD BE UPDATED AS, THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, ASTHMA, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION H6 HAS BEEN UPDATED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA /VOMITING, INFLAMMATORY RESPONSE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | THE VIDEO RECEIVED FOR THIS EVENT WAS REVIEWED BY A CLINICAL DEVELOPMENT ENGINEER (CDE). THE SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT FIRES A WHITE RELOAD ON A VESSEL, THEN EXPERIENCES A PAUSE FOR COMPRESSION WHEN THE STAPLER REACHES 20MM. A MESSAGE INDICATES THE TARGET TISSUE IS TOO THICK. WHEN THE STAPLER INSTRUMENT UNCLAMPS FROM THE VESSEL, MALFORMED STAPLES AND BLEEDING FROM THE VESSEL IS OBSERVED. THE SURGEON PLACES GAUZE OVER THE BLEED. UPON SLOWLY RELEASING THE GAUZE, THE BLEEDING CONTINUES. LATER IN THE VIDEO, THE STAPLER INSTRUMENT IS REINSTALLED WITH A WHITE RELOAD. THE STAPLER INSTRUMENT SUCCESSFULLY COMPLETES THE FIRE ON THE VESSEL. LATER IN THE VIDEO, THE STAPLER INSTRUMENT IS INSTALLED AGAIN WITH A GREEN RELOAD AND COMPLETES THE RELOAD FIRE ON ANOTHER LOCATION OF THE LUNG. PROFUSE BLEEDING IS OBSERVED IN THE BACKGROUND. THE ORIGIN OF THE BLEEDING IS NOT SHOWN IN THE VIDEO. SUCTIONING OF THE BLOOD WITH A LAPAROSCOPIC SUCTION IRRIGATOR IS PERFORMED. THE VIDEO ENDS. BASED ON THIS VIDEO REVIEW, THE TISSUE WAS TOO THICK TO BE TRANSECTED WITH THE WHITE RELOAD. THE DA VINCI SUREFORM USER MANUAL WARNS AGAINST IMPROPER SELECTION OF RELOAD COLOR AND INCLUDES A TABLE DETAILING THE APPROPRIATE TISSUE THICKNESS RANGE FOR EACH RELOAD COLOR.
SUREFORM | K173721 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED GASTRECTOMY WITH GASTROJEJUNOSTOMY PROCEDURE, THE PATIENT EXPERIENCED A DUODENAL LEAK REQUIRING THREE ADDITIONAL PROCEDURES; THE PATIENT ULTIMATELY EXPIRED ON POST-OPERATIVE DAY (POD) 23. THE INTUITIVE CLINICAL SALES REPRESENTATIVE (CSR) RECEIVED INFORMATION FROM THE SITE ROBOTICS COORDINATOR REPORTING THAT THE PATIENT RECEIVED UNSPECIFIED ADDITIONAL PROCEDURES ON POD 1, 7, AND 22 DUE TO A DUODENAL STUMP LEAK AND PLACEMENT OF AN ABDOMINAL NEGATIVE PRESSURE WOUND VAC DRESSING. THE SURGEON LATER REPORTED TO THE CSR THAT DURING THE DA VINCI-ASSISTED PROCEDURE, A SUREFORM 60 WHITE RELOAD WAS FIRED THAT RECEIVED PAUSES FOR COMPRESSION. THE SURGEON WAITED FOR THE PAUSES, AND THE RELOAD FIRE WAS THEN COMPLETED WITHOUT COMPLICATIONS. THE SURGEON REPORTED THAT EVERYTHING SEEMED FINE WITH THE RELOAD FIRE AND THE PROCEDURE WAS CONTINUED. THE CAUSE OF THE DUODENAL STUMP LEAK WAS STATED TO BE UNKNOWN. THERE WAS NO ALLEGATION OF ANY ISSUES AGAINST ANY OF THE DA VINCI PRODUCTS USED DURING THE PROCEDURE. ADDITIONAL ATTEMPTS TO CONTACT THE SURGEON HAVE NOT BEEN SUCCESSFUL. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS WERE OBSERVED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, VESSEL SEALER EXTEND, FENESTRATED BIPOLAR FORCEPS, TIP-UP FENESTRATED GRASPER, LARGE CLIP APPLIER, SUREFORM 60 STAPLER, MEGA SUTURECUT NEEDLE DRIVER, MONOPOLAR CURVED SCISSORS, PROGRASP FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THE PROCEDURE. REVIEW OF THE STAPLER LOGS FOUND THAT THAT THE SUREFORM 60 STAPLER (PART NUMBER 480460-09, LOT NUMBER L82240124-0064), WAS INSTALLED ON THE SYSTEM 7 TIMES AND FIRED 7 RELOADS (4 BLUE, 2 WHITE, 1 BLUE, IN THAT ORDER). ON 6 OF THE INSTALLS THE STAPLER FIRINGS COMPLETED WITH NO PAUSES FOR COMPRESSION. ON INSTALL 5, THE FIRST CLAMP WAS INCOMPLETE, STOPPING AT ABOUT 86% COMPLETION. THE NEXT CLAMP WAS SUCCESSFUL, AND THE FIRING COMPLETED WITH 9 PAUSES FOR COMPRESSION. AFTER THE 7TH FIRE, THE STAPLER WAS REMOVED AND NOT USED AGAIN. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DEVELOPED A LEAK FROM THE DUODENUM STAPLE LINE FOLLOWING A GASTRECTOMY WITH GASTROJEJUNOSTOMY RECONSTRUCTION. ALTHOUGH THE PROCEDURE SEEMED TO GO WELL WITH NO ISSUES NOTED DURING THE PROCEDURE, THE LEAK DEVELOPED IN THE FIRST DAY POSTOPERATIVELY. HOW THE LEAK MAY HAVE OCCURRED IS NOT PROVIDED. THE PATIENT UNDERWENT MULTIPLE ADDITIONAL OPERATIONS AND EVENTUALLY DIED SEVERAL WEEKS AFTER THE INITIAL PROCEDURE. THE CAUSE OF DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
ION | K182188 | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE THE PATIENT BEGAN BLEEDING, WENT INTO CARDIAC ARREST AND EXPIRED. THE PATIENT HAD A HISTORY OF CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE WAS INFORMED THAT AT THE TIME OF THE BIOPSY PROCEDURE, EXCESSIVE BLEEDING ORIGINATED FROM THE RIGHT MIDDLE LOBE AND THE PROCEDURE WAS ABORTED. IT WAS REPORTED THAT EPINEPHRINE MAY HAVE BEEN INSTILLED TO HELP SLOW DOWN THE BLEEDING. THE PATIENT SUBSEQUENTLY WENT INTO CARDIAC ARREST AND CPR WAS PERFORMED WITHOUT SUCCESS. THERE WERE NO REPORTED ERRORS OR MALFUNCTIONS OF THE SYSTEM OR ANY INSTRUMENTS IN USE DURING THE PROCEDURE. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN HAVE BEEN UNSUCCESSFUL. | A REVIEW OF THE ION SYSTEM LOGS FOR THE EVENT FOUND THERE WERE NO OBSERVED EVENTS THAT WOULD SUGGEST A PRODUCT ISSUE, AND LOGGED EVENTS WERE IN LINE WITH NORMAL SYSTEM FUNCTIONALITY. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL BIOPSY, DURING THE PROCEDURE BLEEDING WAS NOTED FROM THE RIGHT MIDDLE LOBE AND THE PROCEDURE WAS ABORTED. IT IS UNCLEAR WHAT INTERVENTIONS WERE TAKEN FOR THE BLEEDING. THE PATIENT SUFFERED A CARDIAC ARREST. RESUSCITATIVE EFFORTS WERE UNSUCCESSFUL AND THE PATIENT DIED. THERE WERE NO REPORTED MALFUNCTIONS OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. GIVEN THE AVAILABLE DATA THE REPORTED EVENT WAS PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. ---FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. ---FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. ---KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. ---BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. ---BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
REMSTAR PRO C-FLEX+ W/HUMID, SD CARD | K091319 | THE MANUFACTURER WAS NOTIFIED ABOUT A VOLUNTARY FIELD SAFETY NOTICE/RECALL RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED REPORTS ALLEGING THAT THE PATIENT EXPERIENCED THE FOLLOWING: "KIDNEY DISEASE/TOXICITY AND LUNG DISEASE. THE MANUFACTURER ALSO RECEIVED NOTIFICATION THAT THE PATIENT PASSED AWAY. NO SPECIFIC MEDICAL INTERVENTION WAS MENTIONED. THE MANUFACTURER WAS MADE AWARE OF THIS INFORMATION THROUGH THE CUSTOMER¿S LEGAL REPRESENTATIVE. DUE TO POTENTIAL LEGAL ISSUES SURROUNDING THIS CASE, NO FURTHER INVESTIGATION OR FOLLOW-UP CAN BE CONDUCTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE SUBMITTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING COPD, CAD, SUDDEN CARDIAC ARREST, DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE; CANCER; DEATH, LARGE-CELL LYMPH ARM PRIMARY VITREORETINAL LYMPH ARM PLEURAL FLUID. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPOR TED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL, DREAMSTATION ST30, GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING IRRITATION (SKIN), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, ASTHMA (NEW OR WORSENING), LUNG DISEASE (UNSPECIFIED). THERE WAS REPORT OF PATIENT DEATH. . DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
COCHLEAR OSIA OSI200 IMPLANT | K191921 | PER THE CLINIC, THE PATIENT UNDERWENT AN INITIAL IMPLANTATION OF TRANSCUTANEOUS OSIA IMPLANT SYSTEM AND A SUBSEQUENT MASTOIDECTOMY ON (B)(6) 2024, WITHIN THE SAME SURGERY. THERE WAS NO SURGICAL COMPLICATIONS AFTER THE INITIAL IMPLANTATION OF OSIA IMPLANT SYSTEM. HOWEVER, THE PATIENT HAD COMPLICATIONS IMMEDIATELY AFTER BOTH THE SURGERY WAS COMPLETED. SUBSEQUENTLY, THE PATIENT EXPIRED ON (B)(6) 2024, DUE TO BRAIN DEATH AFTER DAYS OF HOSPITALIZATION IN THE INTENSIVE CARE UNIT.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DIED DUE TO CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING LUNG CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE OBTAINING CLINICAL AUDIT LOGS AFTER A PATIENT DEATH. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. THE NURSE MANAGER REQUESTED VERIFICATION IF LEADS OFF ALARMS OCCURRED FOR THIS PATIENT BEING MONITORED ON TELEMETRY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | A PHILPS REMOTE SERVICE ENGINEER (RSE) AND A PHILIPS REMOTE CLINICAL SUPPORT (RCS) REPRESENTATIVE SPOKE WITH THE CUSTOMER. THE RSE RETRIEVED THE CLINICAL AUDIT LOGS AND PROVIDED THESE TO THE CUSTOMER. THE RSE AND RCS ASSISTED THE CUSTOMER IN REVIEWING LOG INFORMATION AND IDENTIFIED YELLOW LEVEL ECG ALARMS AND SPO2 ALARMS. THE RCS PROVIDED INFORMATION TO THE CUSTOMER ON HOW TO SEARCH THE PATIENT DATA BY THE PATIENT MEDICAL RECORD NUMBER. THE CUSTOMER INDICATED THERE WAS NO PRODUCT MALFUNCTION AND DECLINED ONSITE SERVICE. THE LOGS WERE PROVIDED TO A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) FOR TECHNICAL INVESTIGATION. THE PSE VERIFIED THAT ON (B)(6) 2024, ¿ECG LEADS OFF¿ (CYAN INOP) STARTED AT 05:06:32 AND THEN IT ESCALATED TO ¿!!ECG LEADS OFF¿ (YELLOW INOP) AT 05:06:36. THIS ESCALATED ¿!!ECG LEADS OFF¿ (YELLOW INOP) ENDED AT 06:16:50. THE REPORTED PROBLEM WAS NOT CONFIRMED. INFORMATION WAS PROVIDED TO THE CUSTOMER TO RESOLVE THE ISSUE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION CANCER, DEATH, LUNG CANCER. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
OBSIDIO | K213385 | IT WAS REPORTED THAT NON-TARGET EMBOLIZATION OCCURRED RESULTING IN PATIENT DEATH. THE PATIENT WAS ADMITTED TO THE HOSPITAL WITH VARICEAL BLEEDING. DUE TO ANATOMY AND THE COMPLEXITY OF THE PROCEDURE, THE DECISION WAS MADE TO PERFORM A TRANS JUGULAR INTRAHEPATIC PORTOSYSTEMIC SHUNT (TIPS) PROCEDURE. THE TIPS PROCEDURE WAS COMPLEX AND WAS ULTIMATELY COMPLETED. THE VARICEAL EMBOLIZATION COULD NOT BE PERFORMED AT THE SAME SETTING DUE TO INABILITY TO CATHETERIZE THE SHUNT AND DUE TO THE TIPS COMPLEXITY. THE PATIENT BECAME VERY ENCEPHALOPATHIC POST PROCEDURE. SUBSEQUENTLY, THE PATIENT WAS BROUGHT BACK TO IR FOR A SUCCESSFUL SHUNT AND VARICEAL EMBOLIZATION. ENCEPHALOPATHY IMPROVED. A FEW DAYS LATER, THE PATIENT STARTED TO EXPERIENCE HEMATEMESIS. THE PATIENT UNDERWENT ENDOSCOPY, WHICH DEMONSTRATED MULTIPLE STRESS ULCERS ON THE POSTERIOR WALL OF THE DUODENUM. ATTEMPTED ENDOSCOPIC TREATMENT WAS UNSUCCESSFUL. THE PATIENT THEN WAS BROUGHT TO IR FOR PLANNED ANGIOGRAPHY AND EMBOLIZATION WITH OBSIDIO IN THE GASTRODUODENAL ARTERY (GDA). THE GDA MEASURED APPROXIMATELY 2-3 MM IN DIAMETER. OBSIDIO WAS INJECTED VIA A MICROCATHETER USING THE STANDARD TECHNIQUE. IMMEDIATELY, THE OBSIDIO WAS NOTED TO FLOW THROUGH THE GDA OUT INTO THE RIGHT GASTROEPIPLOIC ARTERY MID TO DISTAL PORTION OF THE ARTERY. A SMALL CAST OF OBSIDIO WAS PRESENT IN THE DISTAL GDA. THE OBSIDIO SYRINGE WAS EMPTIED AND THE REMAINDER OF THE OBSIDIO IN THE MICROCATHETER WAS CLEARED WITH SALINE. THE FORCE OF THE SALINE INJECTION WAS NOT SLOW. SUBSEQUENTLY, ADDITIONAL GDA EMBOLIZATION WAS PERFORMED WITH COILS AND GELFOAM INCLUDING EMBOLIZATION OF THE SUPERIOR PANCREATICODUODENAL ARCADE. THE PATIENT WAS COMPLAINING OF ABDOMINAL PAIN FOLLOWING OBSIDIO EMBOLIZATION. THE PATIENT STOPPED BLEEDING POST PROCEDURE AND STABILIZED. DUE TO PERSISTENT PAIN, THE PATIENT UNDERWENT A COMPUTED TOMOGRAPHY (CT) SCAN DEMONSTRATING DEVASCULARIZATION OF THE PANCREATIC HEAD AND OBSIDIO IN SMALL PANCREATIC ARTERIES. THE PATIENT DEVELOPED NECROSIS OF THE PANCREATIC HEAD, SEVERE PANCREATITIS, ACUTE KIDNEY INJURY AND REQUIRED INTUBATION. THE PATIENT SUBSEQUENTLY DEVELOPED MULTIPLE PANCREATIC PSEUDOCYSTS, SUPPURATIVE PANCREATITIS WORSENED AND MULTIORGAN FAILURE. THE PATIENT EXPIRED SEVERAL WEEKS AFTER THE EMBOLIZATION PROCEDURE. | B3: DATE OF EVENT: NO DATE PROVIDED, USED THE FIRST DATE IN THE MONTH OF THE AWARE DATE.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REP DREAMSTATION AUTO CPAP THAT THE REPORTER ALLEGES THAT THE PATIENT HAS PASSED AWAY AND NO LONGER NEEDS THE DEVICE. DEVICE WAS NOT RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
UNKNOWN RENASYS TOUCH DEVICE | K153209 | IT WAS REPORTED THAT, DURING NPWT, THE PATIENT PASSED AWAY OVERNIGHT DUE TO CATASTROPHIC BLEED TO LEFT GROIN. THE PATIENT WAS USING A RENASYS TOUCH DEVICE IN SITU FOR MANAGEMENT OF EXUDATE. AN AMBULANCE ATTENDED THE INCIDENT. NO FURTHER INFORMATION IS AVAILABLE. | UPDATED RESULTS OF INVESTIGATION: THE DEVICE WAS NOT RETURNED FOR EVALUATION; THEREFORE, A DEVICE ANALYSIS COULD NOT BE PERFORMED. NO SUPPORTING DOCUMENTATION HAS BEEN RECEIVED AS OF THE DATE OF THIS INVESTIGATION THAT SUGGESTS THE USE OF THE NEGATIVE PRESSURE WOUND THERAPY (NPWT) DEVICE WAS IN ANY WAY INVOLVED IN THE PATIENT OUTCOME OF THE CATASTROPHIC BLEED TO LEFT GROIN. AT THIS TIME, WE HAVE NO REASON TO SUSPECT THAT THE PRODUCT FAILED TO MEET ANY PRODUCT SPECIFICATIONS AT THE TIME OF MANUFACTURE. DEVICE SPECIFIC IDENTIFIERS WERE NOT PROVIDED, THEREFORE, AN EVALUATION OF THE MANUFACTURING RECORDS AND PRIOR ACTIONS REVIEW COULD NOT BE PERFORMED. A REVIEW OF COMPLAINT HISTORY FOR THE PREVIOUS 14 MONTHS DID NOT REVEAL SIMILAR EVENTS FOR RENASYS DEVICES. THE PRODUCT¿S INSTRUCTIONS FOR USE (IFU) INDICATE THAT THE PATIENT, DEVICE, AND DRESSING SHOULD BE MONITORED FREQUENTLY TO DETERMINE IF THERE ARE ANY SIGNS OF BLEEDING. THE FREQUENCY SHOULD BE DETERMINED BY THE CLINICIAN BASED ON INDIVIDUAL CHARACTERISTICS OF THE PATIENT AND WOUND. THE ROOT CAUSE OF THE EVENT COULD NOT BE DETERMINED BASED ON THE LIMITED INFORMATION PROVIDED. HOWEVER, THERE IS NO INFORMATION TO SUGGEST THAT THE RENASYS DEVICE WAS A CAUSE OR CONTRIBUTOR TO THE REPORTED EVENT.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF LIVER DISEASE/TOXICITY, CANCER AND DEATH "SQUAMOUS ESOPHAGEAL CANCER, STAGE 4B". THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION REGARDING THYROID CANCER AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING SKIN IRRITATION, NAUSEA / VOMITING, CANCER, LIVER & KIDNEY RELATED ISSUE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO CANCER, HEADACHE, DIFFICULTY BREATHING. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE REPORTER ALLEGED RESPIRATORY TRACT IRRITATION, LUNG DISEASE, LUNG INFECTION, EXACERBATION LEADING TO DEATH. THE PATIENT WAS PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED A CALL FROM A SPOUSE ASKING THAT COMMUNICATIONS FROM PHILIPS BE STOPPED AS THE PATIENT HAS PASSED AWAY. NO ADDITIONAL INFORMATION WAS PROVIDED BY THE CALLER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SHORTNESS OF BREATH, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ILET BIONIC PANCREAS | K231485 | ON (B)(6) 2024 A MEDICAL DEVICE DISTRIBUTOR NOTIFIED BETA BIONICS THAT AN ILET USER HAD PASSED AWAY. THE DATE OF DEATH WAS NOT REPORTED. AT THE TIME OF THIS REPORT, ATTEMPTS BY BETA BIONICS TO DETERMINE IF THE USER WAS ON THE ILET AT THE TIME OF THE EVENT AS WELL AS CAUSE OF DEATH HAVE BEEN UNSUCCESSFUL. | NO PRODUCT WAS RETURNED FOR EVALUATION. ATTEMPTS TO DOWNLOAD THE DEVICE ENGINEERING LOGS WERE UNSUCCESSFUL, INDICATING THAT THE LOGS HAVE NOT BEEN SYNCED SINCE DEVICE WAS SHIPPED FROM MANUFACTURING. AS SUCH, A LOG REVIEW WAS UNABLE TO BE CONDUCTED. NO PRODUCT PERFORMANCE ISSUES CAN BE IDENTIFIED. IF THE PRODUCT IS RECEIVED AT A LATER DATE, OR THE ILET DATA IS UPLOADED, THE COMPLAINT WILL BE REOPENED AND INVESTIGATED ACCORDINGLY. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. WITHOUT DEVICE DATA OR A DATE OF DEATH, IT CANNOT BE DETERMINED IF THE USER WAS WEARING THE ILET AT THEIR TIME OF DEATH OR IF THE ILET WAS RELATED TO THEIR PASSING. HOWEVER, CASE NOTES INDICATE THE USER WAS 75 YEARS OLD, HAD A HISTORY OF CORONARY ARTERY DISEASE AND END-STAGE RENAL DISEASE AND WAS ON HEMODIALYSIS THREE TIMES PER WEEK. IT IS LIKELY THAT THESE COMORBID CONDITIONS CONTRIBUTED TO THEIR PASSING.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO RESPIRATORY FAILURE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF CANCER, DEATH AND ESOPHAGEAL CANCER RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS AND/OR HEADACHE; CANCER; DEATH, BRAIN CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION BIPAP AVAPS30 | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING MULTIPLE MYELOMA, CANCER, DEATH, HYPER- OR HYPOTHYROID DISEASE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
TEMPUS LS-MANUAL | K200849 | IT WAS REPORTED TO PHILIPS THAT THE CREW WAS ATTEMPTING TO PACE A PATIENT AND KEPT GETTING A "PACER ERROR" WHEN TRYING TO PACE THE PATIENT.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, METASTATIC BREAST CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION REGARDING RESPIRATORY FAILURE AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BREATHING ISSUES, BLOOD CLOTS, AND STROKE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
EVERFLO DOM | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A EVERFLO DOM NON- OPI DEVICE. THE PATIENT HAS ALLEGEDLY PASSED AWAY. NO MEDICAL INTERVENTION WAS REPORTED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IN PREVIOUS REPORT, BOX B: ADVERSE EVENT OR PRODUCT PROBLEM: DESCRIBE EVENT OR PROBLEM AND BOX H - (DEVICE) PROBLEM CODE GRID (1) WAS MENTIONED INCORRECTLY. THE DEVICE IS NOT PART OF RECALL. IN THIS REPORT, BOX B AND H HAS BEEN UPDATED OR CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. PATIENT CALLED STATING THAT HE WAS CALLING ABOUT THE PATIENT NAMED CATHERINE. WANTS ALL COMMUNICATION STOPPED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. PER THE REPORTER, ON (B)(6) 2024, THE CUSTOMER EXPERIENCED SEVERE HYPOGLYCEMIA AND WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU). A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. RELATED MEDSUN REPORT NUMBER #(B)(4). H3 OTHER TEXT: DEVICE NOT RETURNED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS AN ALLEGATION OF KIDNEY CANCER AND PATIENT PASSING AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING MULTIPLE MYELOMA. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION CPAP PRO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING BLADDER CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS COMPLETED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, CANCER, LIVER CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
AERO® TRACHEOBRONCHIAL STENT | K082284 | THE ACCOUNT ALLEGES THAT DURING A TRACHEOBRONCHIAL STENT DEPLOYMENT PROCEDURE, THE CLINICIAN WAS UNABLE TO DEPLOY A STENT BY SQUEEZING THE DEPLOYMENT TRIGGER. THEY SWITCHED TO A NEW STENT AND HAD A SIMILAR ISSUE BUT WERE SUCCESSFUL IN DEPLOYING THE STENT, BUT THE STENT HAD ENFOLDED [INVAGINATED]. THE 12X40 STENT WAS PLACED VIA DIRECT VISION. NO DAMAGE WAS NOTED TO THE STENT OR THE DEVICE BEFORE OR DURING THE DEPLOYMENT ATTEMPT. THE CLINICIAN TRIED TO ADJUST THE STENT USING A BALLOON INFLATION TECHNIQUE AND ENDED UP PERFORATING THE TRACHEOBRONCHIAL WALL. THE PHYSICIAN STATES THE PATIENT WAS NOT IN GOOD CONDITION BEFORE THE PROCEDURE. THE 58-YEAR-OLD FEMALE PATIENT WHO SUFFERED FROM A LARGE CANCEROUS ENDOBRONCHIAL TUMOR LIKELY CONTRIBUTED TO THE PATIENT'S TRACHEOBRONCHIAL PERFORATION BY WEAKENING THE TRACHEAL WALL DURING THE BALLOON EXPANSION TECHNIQUE RESULTED IN PERFORATION. A CODE WAS INITIATED PER HOSPITAL PROTOCOLS BUT WAS UNSUCCESSFUL. THE PATIENT EXPIRED AT 4:15 PM THE SAME DAY. COD MASSIVE HEMORRHAGE. THE PATIENT ALSO EXPERIENCED HEMOPTYSIS 1 WEEK BEFORE THE ACTUAL STENT PLACEMENT AND UNDERWENT PHOTODYNAMIC THERAPY (PDT) TREATMENT 4 WEEKS BEFORE THAT ACTUAL STENT PLACEMENT. AN AUTOPSY IS NOT LIKELY DUE TO THE PATIENT'S DEATH BEING ATTENDED. THE DEVICE IS RETURNING FOR ENGINEERING EVALUATION. IN-STENT BALLOONING IS ACCEPTABLE PER IFU. | THE SUSPECT DEVICE IS EXPECTED TO RETURN FOR EVALUATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE EVALUATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED METASTATIC BREATHING ISSUES. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALLEGED OF HAVING DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, THROAT AND BLADDER CANCER. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
CARDIOSAVE HYBRID, TYPE B PLUG | K112372 | IT WAS REPORTED THAT WHILE PUTTING A VERY UNSTABLE PATIENT ON THE CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) UNIT, THE PATIENT HAD VERY LOW PRESSURES 60/40 AND LESS. THE BALLOON PUMP WASN'T ABLE TO HELP THE PATIENT WITH THE PRESSURES WHERE THEY WERE. UNIT DID NOT DISPLAY ANY ALARM EXCEPT "BELOW AUGMENTATION LIMIT" DURING CODE. DISCONTINUED TREATMENT BECAUSE THE PATIENT PASSED AWAY. | UPDATED FIELDS; B4, D9, G3, G6, H2, H3, H4, H6, H11. THE CUSTOMER REQUESTED INSPECTION/PREVENTIVE MAINTENANCE (PM) ON THE UNIT TO ENSURE PROPER OPERATION. THE FSE STATED THAT THE UNIT DID NOT DISPLAY ANY ALARM EXCEPT "BELOW AUGMENTATION LIMIT". THE FSE INSPECTED THE LOGS AND FOUND NO ABNORMAL ALARMS INDICATING UNIT MALFUNCTION. THE FSE STATED THAT THE FAULT LOG DID NOT SHOW LOGS BACK FOR 8/22/24 SINCE UNIT WAS USED ON 9/6/24, AND LOGS WERE FULL OF ENTRIES FROM THAT DATE. THERE WAS NO INDICATION OF UNIT MALFUNCTION FOUND IN THE LOGS OR DURING TESTING/INSPECTION. THE FSE PERFORMED A COMPLETE PM WITH FULL CALIBRATION, FUNCTIONAL TESTING AND SAFETY CHECKS. THE UNIT PASSED ALL FUNCTIONAL AND SAFETY TESTS PER FACTORY SPECIFICATIONS. THE UNIT WAS RETURNED TO THE CUSTOMER.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER SERVICE CENTER FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED BY THE SERVICE TECHNICIAN. EVALUATION RESULT STATED THAT THE COMPLAINT WAS NOT CONFIRMED AND FOUND NO EVIDENCE OF VISIBLE FOAM DEGRADATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER SERVICE CENTER FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED BY THE SERVICE TECHNICIAN. EVALUATION RESULT STATED THAT THE COMPLAINT WAS NOT CONFIRMED AND FOUND NO EVIDENCE OF VISIBLE FOAM DEGRADATION. THE CORRECT EVENT DESCRIPTION SHOULD BE - THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR EVALUATION. THE DEVICE WAS VISUALLY INSPECTED BY THE SERVICE TECHNICIAN. EVALUATION RESULT STATED THAT THE DEVICE PASSED ALL TESTS. THE COMPLAINT WAS ORIGINALLY REPORTED UNDER THE RECALL, BUT AFTER FURTHER REVIEW IT WAS DETERMINED THIS COMPLAINT SHOULD NOT HAVE BEEN FILED UNDER THE RECALL. THEREFORE, BOX D, G AND H HAS BEEN UPDATED OR CORRECTED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER AND PROSTATE CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, KIDNEY DISEASE/TOXICITY AND REDUCED CARDIOPULMONARY RESERVE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT ALLEGEDLY PASSED AWAY. THE REPORTER ALSO ALLEGED OF HAVING CANCER, LUNG ADENOCARCINOMA. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND METASTATIC SQUAMOUS CELL CARCINOMA. THE PATIENT HAS BEEN PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO SKIN IRRITATION AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, BRAIN CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF LUNG DISEASE, CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
FISHER & PAYKEL HEALTHCARE | K131895 | A HEALTHCARE FACILITY REPORTED VIA A FISHER AND PAYKEL HEALTHCARE (F&P) FIELD REPRESENTATIVE THAT SEVERAL YEARS AGO, DURING THE COVID-19 PANDEMIC (NO DATE WAS SPECIFIED), TWO PATIENTS WERE RECEIVING THERAPY VIA F&P PT101 AIRVO 2 HUMIDIFIER AIRVO 2 (AIRVO 2) DEVICE(S) AND SUBSEQUENTLY PASSED AWAY. NO FURTHER INFORMATION REGARDING THESE REPORTED EVENTS WAS PROVIDED BY THE HEALTHCARE FACILITY. THE STAFF MEMBER FROM THE HEALTHCARE FACILITY WHO REPORTED THESE EVENTS STATED THAT THEY DID NOT SEE THESE EVENTS, WAS NOT ON SHIFT WHEN THESE EVENTS OCCURRED, AND DID NOT HAVE FIRSTHAND KNOWLEDGE OF THESE EVENTS. THIS REPORT IS RELATED TO MFR REPORT #: 9611451-2024-00656. | (B)(4). FISHER & PAYKEL HEALTHCARE (F&P) HAS MADE MULTIPLE REQUESTS FOR INFORMATION REGARDING THE DEVICE AND FOR THE DEVICE(S) TO BE RETURNED. NO RESPONSE HAS BEEN PROVIDED, HOWEVER, AND THE HEALTHCARE FACILITY HAS NOT RETURNED THE DEVICE(S). PRODUCT BACKGROUND: THE PT101 AIRVO 2 HUMIDIFIER (AIRVO 2) DEVICE IS AN ACTIVE HUMIDIFIER WITH INTEGRATED FLOW GENERATOR THAT DELIVERS HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES TO SPONTANEOUSLY BREATHING PATIENTS THROUGH A VARIETY OF PATIENT INTERFACES. IT'S INTENDED USE IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES. THE AIRVO 2 DEVICE SHOULD NOT BE USED FOR LIFE SUPPORT PURPOSES, AND APPROPRIATE PATIENT MONITORING MUST BE USED AT ALL TIMES. METHOD: THE SUBJECT AIRVO 2 DEVICE(S) WERE NOT RETURNED TO F&P FOR EVALUATION. OUR INVESTIGATION HAS NECESSARILY BEEN LIMITED THEREFORE TO THE INFORMATION REPORTED BY THE HEALTHCARE FACILITY AND OUR KNOWLEDGE OF THE PRODUCT. RESULTS: THE HEALTHCARE FACILITY HAS STATED THAT SEVERAL YEARS AGO, DURING THE COVID-19 PANDEMIC, TWO PATIENTS WERE RECEIVING THERAPY VIA AIRVO 2 DEVICE/S AND SUBSEQUENTLY DECEASED. NO FURTHER INFORMATION REGARDING THESE REPORTED EVENTS WAS PROVIDED BY THE HEALTHCARE FACILITY. THE STAFF MEMBER FROM THE HEALTHCARE FACILITY WHO REPORTED THESE EVENTS STATED THAT THEY DID NOT SEE THESE EVENTS, WAS NOT ON SHIFT WHEN THESE EVENTS OCCURRED, AND DID NOT HAVE FIRSTHAND KNOWLEDGE OF THESE EVENTS. CONCLUSION: OUR INVESTIGATION WAS UNABLE TO CONFIRM ANY DETAILS REGARDING THE REPORTED EVENT. NO FURTHER INFORMATION REGARDING THESE REPORTED EVENTS WAS PROVIDED BY THE HEALTHCARE FACILITY. THE STAFF MEMBER FROM THE HEALTHCARE FACILITY WHO REPORTED THESE EVENTS STATED THAT THEY DID NOT SEE THESE EVENTS, WAS NOT ON SHIFT WHEN THESE EVENTS OCCURRED, AND DID NOT HAVE FIRSTHAND KNOWLEDGE OF THESE EVENTS.
FISHER & PAYKEL HEALTHCARE | K131895 | A HEALTHCARE FACILITY REPORTED VIA A FISHER AND PAYKEL HEALTHCARE (F&P) FIELD REPRESENTATIVE THAT SEVERAL YEARS AGO, DURING THE COVID-19 PANDEMIC (NO DATE WAS SPECIFIED), TWO PATIENTS WERE RECEIVING THERAPY VIA F&P PT101 AIRVO 2 HUMIDIFIER AIRVO 2 (AIRVO 2) DEVICE(S) AND SUBSEQUENTLY PASSED AWAY. NO FURTHER INFORMATION REGARDING THESE REPORTED EVENTS WAS PROVIDED BY THE HEALTHCARE FACILITY. THE STAFF MEMBER FROM THE HEALTHCARE FACILITY WHO REPORTED THESE EVENTS STATED THAT THEY DID NOT SEE THESE EVENTS, WAS NOT ON SHIFT WHEN THESE EVENTS OCCURRED, AND DID NOT HAVE FIRSTHAND KNOWLEDGE OF THESE EVENTS. THIS REPORT IS RELATED TO MFR REPORT #: 9611451-2024-00655. | (B)(4). FISHER & PAYKEL HEALTHCARE (F&P) HAS MADE MULTIPLE REQUESTS FOR INFORMATION REGARDING THE DEVICE AND FOR THE DEVICE(S) TO BE RETURNED. NO RESPONSE HAS BEEN PROVIDED, HOWEVER, AND THE HEALTHCARE FACILITY HAS NOT RETURNED THE DEVICE(S). PRODUCT BACKGROUND: THE PT101 AIRVO 2 HUMIDIFIER (AIRVO 2) DEVICE IS AN ACTIVE HUMIDIFIER WITH INTEGRATED FLOW GENERATOR THAT DELIVERS HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES TO SPONTANEOUSLY BREATHING PATIENTS THROUGH A VARIETY OF PATIENT INTERFACES. IT'S INTENDED USE IS FOR THE TREATMENT OF SPONTANEOUSLY BREATHING PATIENTS WHO WOULD BENEFIT FROM RECEIVING HIGH FLOW WARMED AND HUMIDIFIED RESPIRATORY GASES. THE AIRVO 2 DEVICE SHOULD NOT BE USED FOR LIFE SUPPORT PURPOSES, AND APPROPRIATE PATIENT MONITORING MUST BE USED AT ALL TIMES. METHOD: THE SUBJECT AIRVO 2 DEVICE(S) WERE NOT RETURNED TO F&P FOR EVALUATION. OUR INVESTIGATION HAS NECESSARILY BEEN LIMITED THEREFORE TO THE INFORMATION REPORTED BY THE HEALTHCARE FACILITY AND OUR KNOWLEDGE OF THE PRODUCT. RESULTS: THE HEALTHCARE FACILITY HAS STATED THAT SEVERAL YEARS AGO, DURING THE COVID-19 PANDEMIC, TWO PATIENTS WERE RECEIVING THERAPY VIA AIRVO 2 DEVICE/S AND SUBSEQUENTLY DECEASED. NO FURTHER INFORMATION REGARDING THESE REPORTED EVENTS WAS PROVIDED BY THE HEALTHCARE FACILITY. THE STAFF MEMBER FROM THE HEALTHCARE FACILITY WHO REPORTED THESE EVENTS STATED THAT THEY DID NOT SEE THESE EVENTS, WAS NOT ON SHIFT WHEN THESE EVENTS OCCURRED, AND DID NOT HAVE FIRSTHAND KNOWLEDGE OF THESE EVENTS. CONCLUSION: OUR INVESTIGATION WAS UNABLE TO CONFIRM ANY DETAILS REGARDING THE REPORTED EVENT. NO FURTHER INFORMATION REGARDING THESE REPORTED EVENTS WAS PROVIDED BY THE HEALTHCARE FACILITY. THE STAFF MEMBER FROM THE HEALTHCARE FACILITY WHO REPORTED THESE EVENTS STATED THAT THEY DID NOT SEE THESE EVENTS, WAS NOT ON SHIFT WHEN THESE EVENTS OCCURRED, AND DID NOT HAVE FIRSTHAND KNOWLEDGE OF THESE EVENTS.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO CANCER STAGE IV LUNG DISEASE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS ,HEADACHE,HYPERSENSITIVITY,KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER DEATH, INJURIES INCLUDE LUNG DAMAGE, COPD, HEART DISEASE, RESPIRATORY FAILURE, AND CONGESTIVE HEART FAILURE. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND CANCER SARCOMA (SOFT TISSUE CANCER), SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING AND INFLAMMATORY RESPONSE. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PACIFIC PLUS | K123358 | A PACIFIC PLUS WAS BEING USED DURING A PERIPHERAL ANGIOPLASTY. THE DEVICE WAS PREPPED AS PER THE IFU WITH NO ISSUES IDENTIFIED. IT WAS REPORTED DURING BALLOON INFLATION THE PACIFIC PLUS RUPTURED AT AN UNKNOWN PRESSURE IN THE PATIENTS ARTERY AND WAS IMMEDIATELY REMOVED. AN ADMIRAL XTREME WAS ALSO USED, THE DEVICE WAS PREPPED AS PER THE IFU WITH NO ISSUES IDENTIFIED. IT WAS REPORTED THAT RESISTANCE WAS FELT DURING ADVANCEMENT AND IT WAS BLCOKING THE ARTERY. THE DEVICE WAS REMOVED. NO PATIENT INJURY WAS REPORTED. | CONTINUATION OF D10: PRODUCT ID SBI030120130 (LOT: 228536954); PRODUCT TYPE: 0674-PERIPHERAL PTA BALLOONS; IMPLANT DATE N/A; EXPLANT DATE N/A, MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
HEART LUNG MACHINE | K943803 | COMPLAINT ID: (B)(4). | NEW / ADDITIONAL INFORMATION WAS PROVIDED BY THE SSU (SALES AND SERVICE UNIT) DATED 2024-09-10 THAT WAS MISTAKENLY NOT ENTERED ON THE INITIAL IMPORTER EMDR (B)(4) THAT WAS SUBMITTED 2024-09-12.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND RESPIRATORY FAILURE. THE PATIENT HAS BEEN PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND UNSPECIFIED LUNG ISSUES. THE PATIENT HAS BEEN PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, SMALL CELL LUNG CANCER, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
HAMILTON-T1 | K181216 | THE FOLLOWING WAS REPORTED TO HAMILTON MEDICAL AG: A. SUMMATION OF THE EVENT: AN ANESTHESIA PROVIDER IN A DEPLOYED SETTING WAS USING THE HAMILTON-T1 VENTILATOR WITH THE UPAC VAPORIZER IN A PULL-THROUGH CONFIGURATION. AT CASE COMPLETION, THE PROVIDER INCREASED THE DELIVERED OXYGEN (HIGH PRESSURE PORT ATTACHED TO A REGULATOR-LIMITED H-CYLINDER) TO 100%. IN QUICK SUCCESSION, THE VENTILATOR INCREASED AIR FLOW TO >100 LPM, AND DELIVERED UNCONTROLLED HIGH PRESSURE TO THE PATIENT VIA THE ENDOTRACHEAL TUBE IN SITU. THIS WAS FOLLOWED BY A COMPLETE FAILURE OF THE HAMILTON-T1 VENTILATOR AND UNREGULATED VENTILATOR SHUT-DOWN. DURING TEST LUNG EXPERIMENTS PEAK AIRWAY PRESSURES AS HIGH AS 80MMHG WERE NOTED IMMEDIATELY PRECEDING THE T1 FAILURE. B. CAUSATIVE ELEMENTS: THE HAMILTON-T1 USED WAS EQUIPPED WITH THE NBC-COMPATIBLE HEPA AIR FILTER (HAMILTON PART NUMBER 161705), INTENDED TO BE USED IN CONTAMINATED ENVIRONMENTS IN CONCERT WITH A STANDARD NATO NBC CARTRIDGE FILTER ATTACHED TO THE AIR INLET ON THE BACK OF THE HAMILTON-T1 (NBC FILTER ADAPTER INSTRUCTION FOR USE, REF 161747, HAMILTON MEDICAL). THIS HEPA FILTER IMPROVES REMOVAL OF NBC CONTAMINATES FROM THE TRANSPORT VENTILATOR'S DELIVERED BREATHS TO THE PATIENT. LIKELY, DUE TO THIS INCREASED FILTRATION (RESISTANCE TO AIR FLOW) WHEN COUPLED WITH THE INCREASED DEMAND OF THE VENTILATOR ATTEMPTING TO DELIVER THE SET 100% OXYGEN CONCENTRATION, AND ON SEPTEMBER 10TH, 2024, HAMILTON MEDICAL AG ADDITIONALLY RECEIVED THE INFORMATION THAT THE INCIDENT RESULTED IN PATIENT DEATH. | HAMILTON MEDICAL AG CASE NUMBER IS: (B)(6) INVESTIGATION IS ONGOING. DUE TO THE MISSING SERIAL NUMBER, UDI INFORMATION IN SECTION D4 COULD NOT BE COMPLETED. THE IMPORTER REPORT NUMBER OF OUR IMPORTER HAMILTON MEDICAL INC IS (B)(4).
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, METASTATIC LIVER. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA AND VOMITING, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, SKIN CANCER, DEATH, LUNG, INFLAMMATION, PRE-DIABETIC, PULMONARY FIBROSIS, PNEUMONIA, STROKE, HEART FAILURE, QUADRUPLE BYPASS, CHRONIC KIDNEY FAILURE, STAGE 4; ACUTE RENAL FAILURE NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATIENT HAS PASSED AWAY DUE TO KIDNEY DISEASE/TOXICITY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, LUNG CANCER, KIDNEY CANCER, COPD. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
AERO® TRACHEOBRONCHIAL STENT | K082284 | THE ACCOUNT ALLEGES THAT DURING A TRACHEOBRONCHIAL STENT DEPLOYMENT PROCEDURE, THE CLINICIAN WAS UNABLE TO DEPLOY A STENT BY SQUEEZING THE DEPLOYMENT TRIGGER. THEY SWITCHED TO A NEW STENT AND HAD A SIMILAR ISSUE BUT WERE SUCCESSFUL IN DEPLOYING THE STENT, BUT THE STENT HAD ENFOLDED [INVAGINATED]. THE 12X40 STENT WAS PLACED VIA DIRECT VISION. NO DAMAGE WAS NOTED TO THE STENT OR THE DEVICE BEFORE OR DURING THE DEPLOYMENT ATTEMPT. THE CLINICIAN TRIED TO ADJUST THE STENT USING A BALLOON INFLATION TECHNIQUE AND ENDED UP PERFORATING THE TRACHEOBRONCHIAL WALL. THE PHYSICIAN STATES THE PATIENT WAS NOT IN GOOD CONDITION BEFORE THE PROCEDURE. THE 58-YEAR-OLD FEMALE PATIENT WHO SUFFERED FROM A LARGE CANCEROUS ENDOBRONCHIAL TUMOR LIKELY CONTRIBUTED TO THE PATIENT'S TRACHEOBRONCHIAL PERFORATION BY WEAKENING THE TRACHEAL WALL DURING THE BALLOON EXPANSION TECHNIQUE RESULTED IN PERFORATION. A CODE WAS INITIATED PER HOSPITAL PROTOCOLS BUT WAS UNSUCCESSFUL. THE PATIENT EXPIRED AT 4:15 PM THE SAME DAY. COD MASSIVE HEMORRHAGE. THE PATIENT ALSO EXPERIENCED HEMOPTYSIS 1 WEEK BEFORE THE ACTUAL STENT PLACEMENT AND UNDERWENT PHOTODYNAMIC THERAPY (PDT) TREATMENT 4 WEEKS BEFORE THAT ACTUAL STENT PLACEMENT. AN AUTOPSY IS NOT LIKELY DUE TO THE PATIENT'S DEATH BEING ATTENDED. THE DEVICE IS RETURNING FOR ENGINEERING EVALUATION. IN-STENT BALLOONING IS ACCEPTABLE PER IFU. | THE SUSPECT DEVICE IS EXPECTED TO RETURN FOR EVALUATION. A SUPPLEMENTAL REPORT WILL BE SUBMITTED ONCE THE EVALUATION IS COMPLETE.
REMSTAR PRO C-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS AND/OR HEADACHE, LUNG DISEASE, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER,DEATH AND LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, CANCER, AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, CANCER, AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT GOT CORRECTED. SECTION H6 UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ANAPLASTIC THYROID CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE AND CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DAMAGE, HEART ATTACK, PLEURAL EFFUSION AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO PATIENT INFORMATION WAS PROVIDED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW- UP REPORT WILL BE FILED
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
FISHER & PAYKEL HEALTHCARE | K103767 | FISHER & PAYKEL HEALTHCARE (FPH) WAS NOTIFIED BY FDA FORM 3500A THAT A HEALTHCARE FACILITY IN CALIFORNIA REPORTED THAT IN (B)(6) 2024, THE INSPIRATORY AND EXPIRATORY LIMBS OF AN RT265 INFANT DUAL HEATED EVAQUA2 BREATHING CIRCUIT WERE SWITCHED AND INCORRECTLY CONNECTED TO THE VENTILATOR FOLLOWING A ROUTINE CIRCUIT CHANGE. DUE TO THE INCORRECT CONNECTION OF THE CIRCUIT LIMBS TO THE VENTILATOR, THE DELIVERED GAS WAS NOT HUMIDIFIED, HOWEVER VENTILATION AND OXYGENATION REMAINED INTACT. THE HEALTHCARE FACILITY REPORTED THAT THE PATIENT HAD A CODE EVENT AND PASSED THE FOLLOWING MORNING. THE HEALTHCARE FACILITY REPORTED THAT THE DEVICE WAS FUNCTIONING PROPERLY AND DID NOT DIRECTLY CAUSE PATIENT HARM. IT WAS REPORTED THAT PRIOR TO THE CODE EVENT, THE COMPLEX PATIENT WAS AGITATED, TACHYCARDIC, AND HAD A LOW EJECTION FRACTION. THE HEALTHCARE FACILITY REPORTED THAT THE SET UP OF THE CIRCUIT TO THE VENTILATOR WAS CHECKED FOLLOWING THE CIRCUIT CHANGE, HOWEVER THE ISSUE WITH THE INCORRECT CONNECTION OF THE CIRCUIT WAS NOT IDENTIFIED UNTIL AFTER THE CODE EVENT. | (B)(4). SECTION D4: COMPLETE DEVICE IDENTIFICATION INFORMATION HAS BEEN REQUESTED BUT NOT PROVIDED BY THE HEALTHCARE FACILITY. SECTION H11: METHOD: FISHER & PAYKEL (F&P) HEALTHCARE REQUESTED THE RETURN OF THE SUBJECT RT265 INFANT DUAL-HEATED EVAQUA2 BREATHING CIRCUIT, AS WELL AS ADDITIONAL INFORMATION AND PHOTOGRAPHS. HOWEVER, NO RESPONSE TO F&P HEALTHCARE'S REQUESTS WAS RECEIVED. THEREFORE, OUR INVESTIGATION IS BASED ON THE INFORMATION INITIALLY PROVIDED AND OUR KNOWLEDGE OF THE PRODUCT. RESULTS: NO FURTHER INFORMATION WAS PROVIDED BY THE HEALTHCARE FACILITY REGARDING THE REPORTED EVENT. FROM THE INFORMATION THAT WAS INITIALLY PROVIDED, THERE WAS NO INDICATION OF A FAULT OR MALFUNCTION WITH THE SUBJECT DEVICE. THE HEALTHCARE FACILITY REPORTED THAT THE CIRCUIT WAS NOT CORRECTLY SET UP BY THE USER, AND STATED THAT THE INCORRECT SET UP 'DID NOT DIRECTLY CAUSE PATIENT HARM.' CONCLUSION: BASED ON THE INFORMATION PROVIDED BY THE HEALTHCARE FACILITY, THE REPORTED MISCONNECTION WAS DUE TO USER ERROR. WITH THE REPORTED MISCONNECTION, THE PATIENT WOULD CONTINUE TO RECEIVE GAS FLOW, HOWEVER NO HUMIDITY WOULD BE PROVIDED. THE USER INSTRUCTIONS THAT ACCOMPANY THE RT265 INFANT DUAL-HEATED EVAQUA2 BREATHING CIRCUIT SHOW IN PICTORIAL FORMAT THE CORRECT SET-UP OF THE CIRCUIT, AND ALSO STATE THE FOLLOWING: ENSURE APPROPRIATE VENTILATOR OR FLOW SOURCE ALARMS ARE SET BEFORE CONNECTING BREATHING SET TO PATIENT. APPROPRIATE PATIENT MONITORING (E.G. OXYGEN SATURATION) MUST BE USED AT ALL TIMES. FAILURE TO MONITOR THE PATIENT (E.G. IN THE EVENT OF AN INTERRUPTION TO GAS FLOW) MAY RESULT IN SERIOUS HARM, OR DEATH. REFER TO THE MR850 RESPIRATORY HUMIDIFIER INSTRUCTIONS FOR USE FOR ADDITIONAL SAFETY INFORMATION, CLINICAL CLAIMS, AND INSTRUCTIONS FOR SET UP AND USE.
VIZIGO | K170997 | IT WAS REPORTED THAT A PATIENT UNDERWENT A CARDIAC ABLATION WITH A VIZIGO AND THE PATIENT EXPERIENCED GROIN HEMATOMA / HEMORRHAGIC SHOCK / DISSEMINATED INTRAVASCULAR COAGULATION / ANEMIA / DYSPNEA REQUIRED COMPRESSION HEMOSTASIS, NORADRENALINE, VENTILATION, AND ECMO. THE PATIENT DIED. WIRE HAD DIFFICULTY PUNCTURING RIGHT FEMORAL VEIN AND MIGHT HAVE GRAZED THE ARTERY, SO IT WAS REMOVED FROM THE PATIENT¿S BODY. AS HEMATOMA HAD DEVELOPED, COMPRESSION HEMOSTASIS WAS PERFORMED. THE ABLATION ITSELF WAS PERFORMED BY REINSERTING THE SHEATH INTO THE LEFT FEMORAL VEIN AND PULMONARY VEIN ISOLATION WAS PERFORMED AS SCHEDULED. DURING THE PROCEDURE, THERE WAS NO INCREASE IN HEMATOMA IN THE RIGHT FEMORAL, BLOOD PRESSURE, ETC., WERE MAINTAINED, AND THE PATIENT WAS SEDATED BUT REMAINED CONSCIOUS. SINCE A DECREASE IN BLOOD PRESSURE WAS OBSERVED AFTER ADMINISTRATION OF PROTAMINE, NORADRENALINE WAS STARTED TO BE ADMINISTERED. SINCE THE PATIENT HAD ALREADY TAKEN STEROIDS BEFOREHAND, NO INTRAVENOUS ADMINISTRATION WAS PERFORMED. IN THE MORNING AFTER THE DAY OF THE PROCEDURE, THE PATIENT¿S CONDITION SUDDENLY CHANGED AND DIED. PATIENT HAD DYSPNEA AND COLD SWEATS AT NIGHT, SO BACK VALVE MASK WAS USED FOR ASSISTED VENTILATION. PATIENT'S PULSE WAS GRADUALLY INCREASING, BUT HEART RATE (HR) WAS IN THE 70'S WHEN THE IN-ROOM MONITOR WAS APPLIED. SPONTANEOUS RESPIRATION DISAPPEARED. IT WAS CHECKED WHETHER PERICARDIAL EFFUSION WAS PRESENT BY TRANSTHORACIC ECHOCARDIOGRAPHY, AND THERE WAS NO PERICARDIAL EFFUSION. ECMO WAS USED, AND CONTRAST-ENHANCED CT SCAN WAS PERFORMED TO SCRUTINIZE THE SOURCE OF BLEEDING. EXTRAVASATION WAS CONFIRMED AT THE RIGHT FEMORAL ARTERY PUNCTURE SITE BUT WHEN THE PATIENT ENTERED THE INTENSIVE CARE UNIT (ICU), HB REMAINED AT 10 ON BLOOD GAS ANALYSIS AND IT WAS NOT STRONGLY IMPLICATED IN THIS "CPA", BUT THERE WAS NO CONFIRMATION. HEMATOMA IN THE RIGHT FEMORAL AND RIGHT CERVICAL REGION WAS INCREASED AFTER CONTRAST-ENHANCED CT SCAN, BLOOD TESTS SHOWED FINDINGS OF DIC (DISSEMINATED INTRAVASCULAR COAGULATION). POSSIBLE CAUSES OF DYSPNEA INCLUDED EXACERBATION OF MYASTHENIA GRAVIS, SEVERE ALLERGIC REACTION TO DRUGS SUCH AS PROTAMINE, AND CIRCULATORY FAILURE DUE TO BLEEDING AT THE PUNCTURE SITE. PREOPERATIVELY, PROHIBITED OR CONCOMITANT MEDICATIONS WERE NOT USED FOR MYASTHENIA GRAVIS. THERE WERE NO FINDINGS TO SUGGEST ANAPHYLAXIS, THE DEGREE OF ANEMIA WAS MILD FOR CIRCULATORY FAILURE DUE TO HEMORRHAGE, AND NO OBVIOUS CAUSE COULD BE IDENTIFIED. PATHOLOGICAL AUTOPSY REVEALED THAT THE DIRECT CAUSE OF DEATH WAS FINDINGS OF HEMORRHAGIC SHOCK DUE TO DISSEMINATED INTRAVASCULAR COAGULATION SYNDROME (DIC). THE BRAIN, HEART, LUNGS, LIVER, AND GASTROINTESTINAL TRACT WERE ALSO EXAMINED, BUT THERE WERE NO OBVIOUS FINDINGS, AS THE CAUSE OF DIC. IN ADDITION, THE PATIENT HAD BEEN TAKING STEROIDS FOR A DECADE AND SHOWED ATROPHY OF THE ADRENAL GLANDS, WHICH SECRETE VARIOUS HORMONES, BUT IT WAS NOT CLEAR WHETHER THIS COULD BE THE DIRECT CAUSE OF THE SHOCK. PER PHYSICIAN'S OPINION, THIS ISSUE MIGHT HAVE BEEN CAUSED BY A PUNCTURE SITE IN THE FEMORAL VEIN, BUT SINCE NO PERICARDIAL EFFUSION WAS OBSERVED IN THE TTE, THE CAUSE WAS NOT CONSIDERED TO BE INTRACARDIAC. ADDITIONAL INFORMATION WAS RECEIVED ON 26-AUG-2024. THE PHYSICIAN'S OPINION ON THE CAUSE OF THE ADVERSE EVENT WAS PROCEDURE AND PATIENT CONDITION. THE POSSIBLE CAUSE OF THE HEMATOMA WAS THAT THE WIRE MIGHT HAVE INJURED THE ARTERY DUE TO MANIPULATING DIFFICULTY DURING GROIN PUNCTURE. POSSIBLE CAUSES OF DYSPNEA INCLUDED EXACERBATION OF MYASTHENIA GRAVIS, SEVERE ALLERGIC REACTION TO DRUGS SUCH AS PROTAMINE, AND CIRCULATORY FAILURE DUE TO BLEEDING AT THE PUNCTURE SITE, BUT NO OBVIOUS CAUSE COULD BE IDENTIFIED. | AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE NUMBER LOT 60000379 AND NO INTERNAL ACTION RELATED TO THE COMPLAINT WAS FOUND DURING THE REVIEW. IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, BIOSENSE WEBSTER, INC., OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. MANUFACTURER'S REFERENCE NUMBER: (B)(4).
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER, MULTIPLE MYELOMA AND PLASMA CELL LEUKEMIA. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO ,CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION/VISION TROUBLE/DRY EYE, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA, INFLAMMATORY RESPONSE, LUNG/LYMPH NODE/BRAIN CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF KIDNEY DISEASE/TOXICITY AND COPD. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PULMONARY FIBROSIS, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE, LIVER DISEASE, TOXICITY, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, AND MOUTH/LYMPH NODE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ROTAREX | K211738 | IT WAS REPORTED THROUGH A POST-MARKET CLINICAL FOLLOW-UP SURVEY ON ROTATIONAL CATHETER, THAT APPROXIMATELY TWO DAYS POST A THROMBECTOMY AND ATHERECTOMY FOR THE ACUTE OCCLUSION, THE PATIENT ALLEGEDLY DEVELOPED WITH REVASCULARIZATION SYNDROME DUE TO LONG VISCERAL AND LIMB ISCHEMIA. REPORTEDLY, THE PATIENT WAS EXPIRED AND THE RELATIONSHIP OF THE EVENT WAS MENTIONED AS NOT RELATED TO DEVICE, WAS NOT RELATED TO PROCEDURE, AND THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
ALARIS SYSTEM | K133532 | IT WAS INITIALLY REPORTED A REQUEST FOR LOG REVIEW ON THE PCA MODULE TO DETERMINE THE DOSES ADMINISTERED TO THE PATIENT FOR PAIN CONTROL. PER REPORT, THE PATIENT WAS IN POST-OP FOR KNEE SURGERY. THE CUSTOMER STATES THERE WAS NO DEVICE MALFUNCTION SUSPECTED. IT WAS REPORTED A "DEATH" OCCURRED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT THE PCA MODULE WAS USED FOR MORPHINE INFUSION. THE MEDICATION WAS CONTAINED IN TERUMO 60ML SYRINGE. THE INFUSION WAS PROGRAMMED AS PCA DOSE WITH CONTINUOUS INFUSION. PER FACILITY'S BIOMEDICAL ENGINEER, IT IS NOT BELIEVED THAT THE DEVICE CAUSED OR CONTRIBUTED TO PATIENT'S DEATH. PER REPORT, THE PRIMARY CAUSE OF DEATH WAS "CARDIAC ARREST POST KNEE SURGERY." THE PATIENT DIED ON (B)(6) 2024. ALTHOUGH ASKED, THE CUSTOMER DID NOT PROVIDE THE SPECIFIC REASON FOR WHY THE FACILITY WISHED TO REVIEW THE PCA DEVICE EVENT LOGS OTHER THAN THE CUSTOMER REQUESTED TO HAVE THE INFUSION EVENT DETAILS TO REVIEW INTERNALLY. THE CUSTOMER INDICATED THAT THEY ARE REQUIRED (BY THEIR FACILITY) TO CONDUCT EVENT LOG REVIEW IMMEDIATELY "DUE TO THE PATIENT BEING HIGH PROFILE." THE EVENT OCCURRED IN UNITED ARAB EMIRATES. MULTIPLE COMMUNICATIONS HAVE BEEN MADE BY BD REQUESTING FOR ADDITIONAL INFORMATION. THE CUSTOMER REFUSED TO PROVIDE FURTHER INFORMATION CITING CONFIDENTIALITY. | OMIT : C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. CORRECTION: MEDICAL DEVICE TYPE, MEDICAL DEVICE CATALOG #, UNIQUE IDENTIFIER (UDI) #, MEDICAL DEVICE SERIAL #, MEDICAL DEVICE MODEL #, CONCOMITANT MED PROD DATA, PMA / 510(K)#, DEVICE MANUFACTURE DATE. ADDITIONAL INFORMATION: IMDRF ANNEX B, C, D CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: DURING REVIEW OF THE DEVICE LOG FOR THIS INVESTIGATION NO ERRORS OR MALFUNCTIONS WERE IDENTIFIED, AND NO ISSUES OR ANOMALIES WERE OBSERVED DURING THE KEY PRESS REPLICATION PERFORMED ON LABORATORY EQUIPMENT. ¿ NO DEVICES WERE RETURNED FOR LABORATORY TESTING; THEREFORE, LOG REVIEW WAS PERFORMED USING THE PCA AND PCU LOG EVENTS PROVIDED BY THE CUSTOMER. ¿ A LOG REVIEW WAS PERFORMED FOR THE REQUESTED TIME FRAME FROM OF (B)(6) 2024 TO (B)(6) 2024. NO ERRORS OR MALFUNCTIONS WERE IDENTIFIED, AND NO ISSUES OR ANOMALIES WERE OBSERVED DURING THE KEY PRESS REPLICATION PERFORMED ON LABORATORY EQUIPMENT ¿ THERE WERE TWO CONFIRMED MORPHINE PROGRAMMING ACTIVITIES, 1ST ON (B)(6) 2024 AT 4:34:49 PM AND 2ND (B)(6) 2024 AT 2:45:26 PM, BOTH PROGRAMS HAD PARAMETERS SYRINGE SELECTION: TERUMO 50; INFUSION MODE PCA AND CONTINUOUS; CONTINUOUS RATE:1ML/H; DOSE:2MG/H, DOSE CONCENTRATION 2MG PROGRAMMED; LOCKOUT INTERVAL: 10 MINUTES. ¿ FROM THE EVENT LOG, THE INFUSION PROGRAMMED WAS MORPHINE WITH A CONCENTRATION OF 100MG/50ML AND A PCA DOSE OF 2 MG WITH A LOCKOUT PERIOD OF 10 MINUTES WAS CONFIRMED, NO MAX LIMIT WAS PROGRAMMED. ¿ THE EVENT LOG SHOWS A TOTAL ESTIMATED INFUSED VOLUME OF 96.64 ML FROM (B)(6) 2024 TO (B)(6) 2024, ¿ THERE WERE 293 UNSUCCESSFUL ATTEMPTS TO REQUEST A DOSE DURING LOCKOUT PERIOD WITHIN THE TIME FRAME REVIEWED. ROOT CAUSE: NO MALFUNCTION WAS ALLEGED ON THE DEVICE INVOLVED IN THE COMPLAINT; THE CUSTOMER REQUESTED AN EVENT LOG REVIEW.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NAUSEA, VOMITING, CANCER, ESOPHAGEAL CANCER, CHEST PRESSURE, AND RESPIRATORY TRACT IRRITATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF EYE IRRITATION, SKIN IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING METASTATIC PANCREATIC CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THE PATIENT ON THE MEDICAL TELEMETRY UNIT WAS FOUND UNRESPONSIVE WHILE THE TELEMETRY DEVICE WAS IN STANDBY MODE. THE PATIENT HAD BEEN OFF THE MONITOR WITH MX40 IN STANDBY MODE, WHICH WAS INITIATED BY THE TELEMETRY TECHNICIAN, FOR SEVERAL HOURS. THE SECOND TIME THE MX40 WAS PLACED IN STANDBY MODE, IT REMAINS UNKNOWN WHAT USER INITIATED THIS AS THE TECHNICIAN DID NOT REMEMBER PLACING THE MX40 IN STANDBY MODE THE SECOND TIME. THE PATIENT WAS FOUND UNRESPONSIVE, TRANSFERRED TO ICU, CODED, AND LATER EXPIRED. THERE WAS NO DEVICE MALFUNCTION FOUND UPON CUSTOMER INSPECTION. BASED ON THIS INFORMATION, THE LACK OF MONITORING DUE TO THE DEVICE BEING PLACED IN STANDBY MODE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT; HOWEVER, IT CANNOT BE DETERMINED HOW THE MX40 WAS PLACED IN STANDBY AND IT REMAINS UNKNOWN WHETHER THERE WAS A DEVICE MALFUNCTION, CLINICAL WORKFLOW OR USE ERROR THAT CONTRIBUTED. IF ADDITIONAL INFORMATION IS OBTAINED, PLEASE REQUEST RE-ASSESSMENT OF THE RECORD BY THE PMS CLINICAL EXPERT. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SMALL LUNG CANCER, HEART ATTACKS. THE PATIENT PASSED AWAY. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, HEART FAILURE, PLEURAL EFFUSION, AND RESPIRATORY FAILURE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, HEART FAILURE, PLEURAL EFFUSION, AND RESPIRATORY FAILURE, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION/BREAKDOWN WAS OBSERVED IN THIS DEVICE. PIL INITIATED THERAPY WITH THE DEVICE AND VERIFIED AIRFLOW, USING A KNOWN GOOD PIL SUPPLIED POWER SUPPLY AND AC POWER CORD. PIL ALSO OBSERVED CUSTOMERS HUMIDIFIER HEATER PLATE DID HEAT. PIL OBSERVED THE FOLLOWING SOUND ABATEMENT DEGRADED FOAM INDICATIONS: BLACK SUBSTANCE, CONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION, WAS OBSERVED ON THE OUTSIDE BOTTOM OF THE BLOWER BOX AND WAS STUCK TO IT. TINY BLACK PIECES OF A BLACK SUBSTANCE, CONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION, WERE FOUND ON THE BOTTOM OF THE ENCLOSURE. TINY BLACK PIECES OF A BLACK SUBSTANCE, CONSISTENT WITH SOUND ABATEMENT FOAM DEGRADATION, WERE FOUND IN THE HUMIDIFIER WATER TANK. DURING THE EVALUATION SECONDARY FINDINGS (32 ERRORS WERE LOGGED. DUST-LIKE CONTAMINATION AT THE AIR INLET, ON THE PCA, IN THE BLOWER BOX AND THROUGHOUT THE INSIDE ENCLOSURE. DUST-LIKE CONTAMINATION INSIDE HUMIDIFIER ENCLOSURE, ON AND UNDER THE HEATER PLATE, AND ON THE DRY BOX. POTENTIAL FLUID SPOTS IN BLOWER BOX AND BLOWER MOTOR INDICATING POSSIBLE WATER INGRESS. PIL CAN CONFIRM THE PRESENCE OF MULTIPLE CONTAMINANTS THAT ARE NOT CONSISTENT WITH DEGRADED SOUND ABATEMENT FOAM FROM THE SECONDARY FINDINGS. PIL WAS UNABLE TO ADDRESS THE SYMPTOMS SPECIFIED) WERE OBSERVED. THE ISSUE OF DEGRADED SOUND ABATEMENT FOAM IS ADDRESSED BY CAPA 7211. IN BOX D: DEVICE AVAILABLE FOR EVAL? AND DATE RETURNED TO MFG HAS BEEN UPDATED. IN BOX H: DEVICE EVAL. BY MFG? (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI ASSISTED APICO-DORSAL LEFT LUNG SEGMENTECTOMY, A BRANCH OF THE PULMONARY ARTERY (PA) WAS INJURED; MASSIVE BLEEDING OCCURRED RESULTING IN AN EMERGENT CONVERSION TO AN OPEN PROCEDURE AND THE PATIENT ULTIMATELY EXPIRED. THE INTUITIVE CLINICAL TERRITORY ASSOCIATE WAS PRESENT AT THE PROCEDURE AND REPORTED THAT DURING DISSECTION AROUND THE PA THE VESSEL RUPTURED AND BLOOD IMMEDIATELY WAS PROJECTED ONTO THE ENDOSCOPE, OBSTRUCTING VISION. THE ENDOSCOPE WAS REMOVED FOR CLEANING OF THE TIP TO RESTORE VISION AND ATTEMPTS TO REINSTALL THE ENDOSCOPE WERE UNSUCCESSFUL, LEADING TO AN EMERGENT CONVERSION TO OPEN THORACOTOMY. THE ESTIMATED BLOOD LOSS WAS MORE THAN 3L AND THE PATIENT WAS IN HEMORRHAGIC SHOCK. THE PATIENT WENT INTO CARDIAC ARREST, INTERNAL CARDIAC MASSAGE WAS PERFORMED, AND THE PA WAS CLAMPED. THE PATIENT WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) AND THE SEGMENTECTOMY PROCEDURE WAS COMPLETED OPEN. THE PATIENT EXPIRED ON POSTOPERATIVE DAY 10. | A VIDEO OF THE EVENT WAS PROVIDED BY THE CUSTOMER AND REVIEW SHOWED THE SURGEON WAS DISSECTING AND TRYING TO GO AROUND A BRANCH OF THE PA WHEN THE INJURED AREA STARTED TO OOZE. THE SURGICAL INSTRUMENT WAS MOVED AWAY FROM THE PA AND LATER REPOSITIONED IN SUCH A WAY THAT IT APPEARS TO PUSH DOWN ON THE MAIN BRANCH OF THE ARTERY, CAUSING THE MAJOR BLEEDING EVENT. THE ENDOSCOPE WAS REMOVED, CLEANED, AND REINSTALLED. THE ROBOT UNIVERSAL SURGICAL MANIPULATOR (USM 2), WHICH WAS HOUSING THE ENDOSCOPE, DISPLAYED A USER FACING MESSAGE SAYING, ¿ADVANCE ENDOSCOPE TO EARLIER POSITION¿. HOWEVER, THE BLEEDING FROM THE PA WAS OBSCURING THE VIEW OF THE SURGICAL WORKSPACE. THE ENDOSCOPE WAS REMOVED AGAIN AND REINSTALLED BACK ONTO USM 2 WHEN ¿ADVANCE ENDOSCOPE TO EARLIER POSITION¿ AND ¿DISENGAGEMENT OF THE ARM AT THE PSM¿ MESSAGING APPEARS. A SUCTION INSTRUMENT IS SEEN TRYING TO SUCTION THE BLOOD FROM THE SITE, FOLLOWED BY REMOVAL OF THE OTHER INSTRUMENTS AND ACCESSORIES, THEN REMOVAL OF THE SUCTION AND THE ENDOSCOPE. THE ACTIONS NOTED ABOVE OCCURRED OVER THE COURSE OF APPROXIMATELY TWO MINUTES. REVIEW OF THE DVSTREAM LOG AND KINEMATIC DATA SHOWED ENDOSCOPE REMOVAL, FOLLOWED BY RE-INSTALLATION. APPROXIMATELY 8 SECONDS LATER THE ENDOSCOPE WAS REMOVED AND REINSTALLED AGAIN, FOLLOWED BY A SHORT PERIOD OF ENDOSCOPE CONTROL ACTIVATION, THEN ENDOSCOPE REMOVAL WITHIN APPROXIMATELY 10 SECONDS. REVIEW OF THE SYSTEM LOGS FOUND THERE WERE NO ERRORS RELATED TO THE REPORTED PULMONARY ARTERY INJURY EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED FOUND THERE WERE NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THIS EVENT. VIDEO REVIEW COINCIDES WITH AN ENDOSCOPE THAT LOST ITS CONNECTION TO THE STERILE ADAPTER, PROMPTING THE USER WITH THE ¿ARM NOT READY. REMOVE THE ENDOSCOPE TO CONTINUE¿ MESSAGE. THIS PROMPT WOULD NOT COINCIDE WITH AN ERROR. THE FOLLOWING CONCOMITANT DEVICES WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, TIP-UP FENESTRATED GRASPER, MARYLAND BIPOLAR FORCEPS, FENESTRATED BIPOLAR FORCEPS, SUREFORM 45 CURVED TIP STAPLER. A SITE HISTORY REVIEW SHOWED THAT WITH THE EXCEPTION OF THE SINGLE-USE STAPLER, THE OTHER INSTRUMENTS AND THE ENDOSCOPE WERE ALL USED IN SUBSEQUENT PROCEDURES WITH NO REPORTED COMPLAINTS. REVIEW OF THE EVENT BY AN INTUITIVE MEDICAL SAFETY OFFICER CONCLUDED THAT THE PATIENT WAS UNDERGOING A PULMONARY SEGMENTECTOMY WHEN THERE WAS AN INJURY TO THE PULMONARY ARTERY. THE MOMENT THE PULMONARY ARTERY RUPTURED WAS CAPTURED ON VIDEO WHICH OCCURRED DURING DISSECTION OF THIS VESSEL AND THE SEGMENTAL BRANCHES. THIS RESULTED IN SPLASHING THE CAMERA WHICH REQUIRED IMMEDIATE CLEANING OF THE LENS. THE SURGICAL TEAM SUBSEQUENTLY HAD DIFFICULTY RE-INSTALLING THE CAMERA AFTER CLEANING THE LENS AS EVIDENCED BY THE LOG INFORMATION. THE PATIENT WAS CONVERTED TO AN OPEN THORACOTOMY AND LATER REQUIRED ECMO. THE PATIENT EVENTUALLY DIED 10 DAYS LATER. THE ROOT CAUSE OF THE DEATH IS RELATED TO THE BLEEDING FROM THE PULMONARY ARTERY. ALTHOUGH THIS OCCURRED DURING DISSECTION OF THE VESSEL, HOW THE VESSEL WAS INITIALLY WEAKENED AND SUBSEQUENTLY RUPTURED IS UNKNOWN. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION A2 AGE: THE PATIENT AGE WAS REPORTED AS "SIXTIES".
DAVINCI XI | K131861 | A PATIENT ENROLLED IN A POST-MARKET STUDY UNDERWENT A DA VINCI ASSISTED PULMONARY SEGMENTECTOMY FOR LEFT LOWER LOBE LUNG CANCER. THE MEDICAL RECORD DOCUMENTED THAT ADHESIONS WERE ENCOUNTERED DURING SURGERY. A CHEST TUBE DRAIN WAS PLACED POST-SURGERY AS STANDARD PROCEDURE AND WAS REMOVED ON POST-OPERATIVE DAY (POD) 4. THE PATIENT WAS DISCHARGED HOME THE SAME DAY WITHOUT COMPLICATIONS. THERE WAS NO REPORT OF ANY DA VINCI DEVICE MALFUNCTION DURING THE PROCEDURE. DURING THE 30-DAY POST-OPERATIVE STUDY VISIT, IT WAS FOUND THAT THE PATIENT HAD PASSED AWAY ON POD 5, THE DAY AFTER DISCHARGE FROM THE HOSPITAL. THE DETAILS OF THE PATIENT¿S COURSE AFTER DISCHARGE AND THE CAUSE OF DEATH WERE NOT PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED BUT HAS NOT BEEN MADE AVAILABLE. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD LIKELY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS RELATED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, FENESTRATED BIPOLAR FORCEPS, PERMANENT CAUTERY HOOK, TWO CADIERE FORCEPS, SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT. ALL MULTI-USE INSTRUMENTS HAVE BEEN USED IN SUBSEQUENT PROCEDURES. THE SUREFORM 45 CURVED-TIP STAPLER LOGS SHOWED THE INSTRUMENT FIRED 7 RELOADS. ALL FIRINGS WERE COMPLETED SUCCESSFULLY WITHOUT ISSUES. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. THE PATIENT IN THIS REPORT HAD MULTIPLE COMORBIDITIES AND UNDERWENT A PULMONARY SEGMENTECTOMY. THE PATIENT DIED 5 DAYS AFTER SURGERY, 1 DAY AFTER BEING DISCHARGED FROM THE HOSPITAL. NO COMPLICATIONS NOR ISSUES WITH ANY INTUITIVE SURGICAL PRODUCTS WERE REPORTED AND NO CAUSE OF DEATH IS PROVIDED.
TURNPIKE SPIRAL 135CM | K151981 | IT WAS REPORTED BY THE DIRECTOR AT (B)(6) HOSPITAL, THAT DURING A CTO LAD LESION SURGERY: "LAD WAS VERY TIGHT. COULD NOT PASS 1.5MM BALLOON SO TRIED TO ADVANCE A TURNPIKE THROUGH THE LESION. IT'S BELIEVED THAT THE TURNPIKE TIP MOVEMENT WAS CONSTRICTED, AND THE TURNPIKE CONTINUED TO BE ROTATED MORE THAN 2 FULL 360 DEGREE ROTATIONS. APPROX. 2-3 MM OF THE TURNPIKE TIP BROKE OFF IN THE PATIENT. AFTER THE TIP BROKE OFF, I WAS TOLD A STINGRAY WAS USED TO GO SUBINTIMAL, AND AT SOME POINT, AFTER THAT THE PATIENT STARTED TO DECLINE. THE SEPARATED TIP OCCLUDED FLOW DOWN THE LAD LEADING TO ST ELEVATIONS ON TELEMETRY AND SIGNIFICANT PATIENT PAIN. THE PATIENT WENT INTO CARDIOGENIC SHOCK AND VENTRICULAR FIBRILLATION ABOUT 1-2 HOURS LATER. OFFERED SURGERY, BUT PATIENT'S FAMILY WAS ADAMANT THAT SHE DID NOT WANT SURGERY. WERE UNABLE TO REMOVE THE DEVICE AS IT WAS LODGED IN THE LAD LESION. WERE NOT ABLE TO WIRE AROUND THE DEVICE. TRIED DISSECTION AND RE-ENTRY AROUND THE LESION BUT WAS UNSUCCESSFUL. CAUSE OF DEATH: VENTRICULAR FIBRILLATION, MYOCARDIAL INFARCTION." ADDITIONAL INFORMATION QUOTED: CARDIOVASCULAR LAB NOTES SUPPLIED BY (B)(6) HOSPITAL TO (B)(6): "THE PATIENT WAS BROUGHT TO THE LAB FOR A KNOWN HIGH-RISK PCI INVOLVING THE LEFT MAIN CORONARY ARTERY. DIAGNOSTICS TAKEN AT THE BEGINNING OF THE CASE BROUGHT TO LIGHT A TIGHT LESION IN THE MID LEFT ANTERIOR DESCENDING ARTERY. ATTEMPTS TO USE ANGIOPLASTY BALLOONING FAILED TO OPEN THE LESION ENOUGH TO PASS A STENT. A TURNPIKE MICROCATHETER WAS SELECTED TO HELP ADVANCE AND OPEN THE LESION. WHILE USING THE MICROCATHETER, A 1 CM PORTION OF THE CATHETER TIP EMBOLIZED AND WAS LODGED IN THE LAD RESULTING IN NO BLOOD FLOW DISTAL. THE DOCTOR EMERGENTLY CONSULTED HIS COLLEAGUE TO EVALUATE POTENTIAL FOR BYPASS OPTIONS. THE DOCTOR ALSO CONSULTED WITH FAMILY AND COMMUNICATED THE LIKELIHOOD OF A POOR OUTCOME WITHOUT SURGERY. THE FAMILY STATED IT WAS THE PATIENT'S WISH TO NOT HAVE CABG PERFORMED. ATTEMPTS AT REWIRING AND DISLODGING EVENTUALLY LED TO THE USE OF A MICROCATHETER TO FORM A NEW PASSAGEWAY FOR BLOOD FLOW. AN UNINTENDED DISSECTION PLANE RAPIDLY FORMED PROXIMALLY UNTIL IT WAS PRESUMED MOVING INTO THE ASCENDING AORTA. MULTIPLE VASOPRESSORS AND DEFIBRILLATIONS WERE STARTED AND DELIVERED TO TRY AND STABILIZE THE PATIENT, BUT THE PATIENT RAPIDLY DECLINED AND WAS UNABLE TO BE RESUSCITATED." | (B)(4)
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING BRAIN CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO BIPAP, DOM-RECRT | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
HAMILTON-C1 | K181216 | THE FOLLOWING WAS REPORTED TO HAMILTON MEDICAL AG: DIFFICULT PATIENT WAS VENTILATED WITH MASK NIV AT THE BEGINNING. AFTER INTUBATION, THE PATIENT WAS SWITCHED TO NIV-ST, AFTER WHICH THE PATIENT WAS DESATURATED (TURNED BLUE) AND WHEN 100% O2 WAS SET ON THE DEVICE, IT INDICATED "NO OXYGEN SUPPLY". THE PATIENT WAS THEN FURTHER VENTILATED WITH A RESUSCITATOR AND AN ALTERNATIVE OXYGEN SOURCE. THE DEVICE WAS BROUGHT INTO MEDICAL TECHNOLOGY. THE PATIENT DIED OF SEPSIS TWO DAYS LATER. THERE IS NO RELATION BETWEEN THE DESATURATION EVENT AS MENTIONED ABOVE AND THE REPORTED SEPSIS. THEREFORE, THE CODE F02 WAS NOT SELECTED BELOW. | HAMILTON MEDICAL AG CASE NUMBER IS: (B)(4). FOLLOW-UP NR. 1 INFORMATION: THE FOLLOWING SECTIONS WERE CHANGED, ADDED IN THIS FOLLOW-UP MDR: B4, G3, G6, H2, H6, H11 INVESTIGATION: THE INDICATED DEATH OF THE PATIENT WAS DUE TO SEPSIS. THERE IS NO RELATION BETWEEN THE DESATURATION EVENT AS MENTIONED ABOVE AND THE REPORTED SEPSIS. THE MOST PROBABLE ROOT CAUSE IS A USE ERROR. ACCORDING TO THE LOGFILE ANALYSIS IT SEEMS THAT THE USER DID NOT ADEQUATELY REACT TO THE ALARM BY EITHER CHECKING THE OXYGEN SUPPLY OR PROVIDING AN ALTERNATIVE SOURCE OF OXYGEN. WITH THIS INVESTIGATION IT HAS BEEN CONFIRMED THAT THE DEVICE DID NOT FAIL TO MEET ITS SPECIFICATIONS AT THE TIME OF THE EVENT WHILE THE VENTILATOR WAS IN USE.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024 DUE TO UNKNOWN CAUSES. REPORTEDLY, THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. H3 OTHER TEXT : DEVICE NOT RETURNED.
2008T HEMODIALYSIS SYSTEM W/BIBAG | K150708 | A USER FACILITY BIOMEDICAL TECHNICIAN REPORTED TO FRESENIUS THAT A HEMODIALYSIS (HD) PATIENT PASSED AWAY DURING THEIR TREATMENT ON A 2008T HD MACHINE. IT WAS REPORTED THAT THE PATIENT BLED OUT, WHICH WAS THE CAUSE OF THE INCIDENT. THE PATIENT SUBSEQUENTLY LOST CONSCIOUSNESS AND THE FACILITY ATTEMPTED TO REVIVE THE PATIENT BY GIVING THEM SALINE. THERE WERE NO REPORTED MACHINE ALARMS THAT OCCURRED. FOLLOWING THE EVENT, WHILE TESTING THE 2008T MACHINE, THE SYSTEM WAS DISPLAYING A ¿WD: FAIL LONG PULSE¿. ADDITIONAL INFORMATION WAS PROVIDED UPON FOLLOW-UP WITH THE FACILITY'S NURSE MANAGER (NM). THE NM CLARIFIED THAT THIS PATIENT DID NOT BLEED OUT AND REPORTED THAT THIS INFORMATION WAS ERRONEOUS. THE NM CONFIRMED THAT THE PATIENT DID LOSE CONSCIOUSNESS (BECAME UNRESPONSIVE) AND RECEIVED SALINE (DETAILS UNKNOWN). FURTHERMORE, THE NM STATES THE NURSING STAFF CALLED A CODE AND INITIATED A CALL TO 911. IT WAS REPORTED THAT THE PATIENT WAS TAKEN TO THE HOSPITAL WITH A PULSE. HOWEVER, THE PATIENT SUBSEQUENTLY EXPIRED IN THE HOSPITAL DUE TO CARDIAC ARREST (CLARIFYING THE PATIENT DID NOT EXPIRE DURING TREATMENT). THE NM STATED THE PATIENT HAD MULTIPLE COMORBID CONDITIONS WHICH WERE LIKELY ASSOCIATED WITH THE DEATH. PER THE NM, THE 2008T MACHINE WAS PUT THROUGH TESTING (AFTER THE EVENT) AND IT WAS CONFIRMED THAT THE MACHINE HAD NO ISSUES AND DID NOT CAUSE THE PATIENT EVENT. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HEMODIALYSIS (HD) WITH THE 2008T HD MACHINE AND THE PATIENT EVENT OF LOSS OF CONSCIOUSNESS REQUIRING 911 TRANSPORT TO THE HOSPITAL. THE PATIENT LATER EXPIRED AT THE HOSPITAL FROM CARDIAC ARREST. MANUFACTURER PRODUCT ANALYSIS WAS PENDING AT THE TIME THIS CLINICAL INVESTIGATION WAS PERFORMED. HOWEVER, THERE IS NO OBJECTIVE EVIDENCE THAT THE 2008T HD MACHINE HAD A MALFUNCTION THAT CAUSED THE EVENT. THE NURSE MANAGER FOR THE ASSOCIATED CLINIC STATED THE MACHINE WAS TESTED AND IT WAS CONFIRMED THAT THE MACHINE HAD NO ISSUES THAT CAUSED THE EVENT. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) HAVE INCREASED MORTALITY RISK FROM CARDIOVASCULAR EVENTS THAN THE GENERAL POPULATION. FURTHERMORE, THE NURSE MANAGER INDICATED THE PATIENT HAD MULTIPLE COMORBID CONDITIONS WHICH ARE LIKELY ASSOCIATED WITH THE PATIENT¿S DEATH.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS ZERO ERROR CODE FOUND AND DURING THE EVALUATION NO SECONDARY FINDINGS WAS OBSERVED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED. BOX G: MANUFACTURE DATE (RFB) UPDATED. BOX E WAS UPDATED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | ON 09/SEP/0224 IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6) 2024. IT IS UNKNOWN IF THE PATIENT HAD COMPLETED TREATMENT. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE UNSUCCESSFUL. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN USE OF THE LIBERTY SELECT CYCLER AND THE PATIENT DEATH. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THE EVENT. IT IS UNKNOWN IF THE PATIENT WAS FOUND DECEASED OR UNRESPONSIVE OR IF THE TREATMENT WAS COMPLETED AT THE TIME OF DEATH. THE MORTALITY RATES FOR PATIENTS WITH END-STAGE RENAL DISEASE ARE SIGNIFICANTLY HIGHER THAN THOSE WITHOUT THE DISEASE. CARDIOVASCULAR DISEASE ACCOUNTS FOR APPROXIMATELY 50% OF THOSE DEATHS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S), PLEASE RE-SUBMIT FOR A CLINICAL REVIEW AND THE NEED FOR A CLINICAL INVESTIGATION WILL BE REASSESSED ACCORDINGLY.
AVEIR¿ INTRODUCER KIT ¿ 50 CM | K211989 | IT WAS REPORTED THAT DURING AN IMPLANT PROCEDURE, THE PATIENT EXPERIENCED LOW BLOOD PRESSURE AND ARRHYTHMIA WHEN THE INTRODUCER WAS PLACED IN THE PATIENT. THE PATIENT STABILIZED INDEPENDENTLY FOLLOWING AN ECHOCARDIOGRAM AND CHEST X-RAY. AS THE DEVICE MAPPING PORTION OF THE PROCEDURE WAS PERFORMED, THE PATIENT'S BLOOD PRESSURE DROPPED AGAIN. AN ECHOCARDIOGRAM WAS PERFORMED AND THE LEFT VENTRICLE WAS FAILING. CARDIOPULMONARY RESUSCITATION WAS PERFORMED AND MEDICINE WAS ADMINISTERED IN ATTEMPT TO RESOLVE THE CARDIAC ARREST. THE PATIENT WAS THEN PUT ON BYPASS, BUT ULTIMATELY PASSED AWAY. THE PHYSICIAN SUSPECTED A PULMONARY HEMORRHAGE TO BE THE CAUSE OF THE PATIENT'S PASSING. THE PHYSICIAN DOES NOT ALLEGE A MALFUNCTION AGAINST THE DEVICE AND DOES NOT ATTRIBUTE THE DEATH TO THE DEVICE OR THE IMPLANT TOOLS. ADDITIONAL INFORMATION WAS REQUESTED, BUT NOT YET RECEIVED.
NXSTAGE SYSTEM ONE | K050525 | A REPORT WAS RECEIVED ON 23 AUG 2024 FROM THE HOME THERAPY NURSE (HTN) OF PATIENT ML, A 70 YEAR OLD FEMALE WITH A MEDICAL HISTORY INCLUDING MULTIPLE COMORBIDITIES AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON AN (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 26 AUG 2024 FROM THE HTN WHO STATED THE PATIENT WAS TREATING ALONE AT THE TIME OF THE EVENT. EMERGENCY MEDICAL SERVICES (EMS) WAS CALLED AND THE PATIENT WAS PRONOUNCED AT THE SCENE. THE CAUSE OF DEATH WAS NOT PROVIDED. | THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | K203365 | A SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED TO TREAT A LESION IN THE ILIAC ARTERIES PRIOR TO PLACING A 26 AORTIC VALVE. THE PHYSICIAN CHOSE THE M5+ 7.0 MM IVL CATHETER BASED ON DISEASE PROCESS, VESSEL SIZE, AND DIAMETER NEEDED TO ADVANCE THE AORTIC VALVE. THE LENGTH OF THE COMMON ILIAC AND EXTERNAL ILIAC ARTERIES WERE TREATED WITH ALL 300 PULSES OF THE IVL DEVICE. THE PHYSICIAN ADVANCED THE 16F DILATOR WITHOUT RESISTANCE, BUT MET RESISTANCE WHEN ATTEMPTING TO ADVANCE THE 18F DELIVERY SHEATH. THE PHYSICIAN ASKED FOR AN L6 8.0 MM IVL CATHETER; HOWEVER, WHILE THE L6 DEVICE WAS BEING OPENED AND PREPPED, THE PHYSICIAN MADE A SECOND ATTEMPT TO CROSS THE LESION WITH THE 18F SHEATH, AND THE SHEATH ADVANCED. AT THAT POINT, THE PHYSICIAN MADE THE DECISION NOT TO TREAT THE VESSEL ANY FURTHER AND TO PROCEED WITH THE INSERTION AND PLACEMENT OF THE VALVE. AS THE PHYSICIAN WAS ADVANCING THE VALVE THROUGH THE SHEATH, THE VALVE GOT STUCK NEAR THE AREA OF STENOSIS WITHIN THE SHEATH. THERE WAS DIFFICULTY ADVANCING OR REMOVING THE VALVE AS THE SHEATH HAD EXPANDED AND BECAME STUCK IN THE ARTERY. THE PHYSICIAN ATTEMPTED TO REMOVE THE SHEATH, AND THE PATIENT EXPERIENCED A DRASTIC DROP IN BLOOD PRESSURE. AN AORTIC BALLOON WAS EVENTUALLY PLACED TO CONTROL SUSPECTED BLEEDING. VASCULAR SURGERY WAS CALLED, AND THE SHEATH WAS REMOVED, WITH WHAT APPEARED TO BE INTIMAL TISSUE ON THE SHEATH. THE PATIENT CODED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. THE VASCULAR SURGEON PLACED COVERED STENTS THROUGHOUT THE COMMON AND EXTERNAL ILIAC ARTERIES, BUT THEY WERE UNABLE TO CONTROL THE PATIENT'S BLEEDING. COMPRESSIONS AND DEFIBRILLATION CONTINUED FOR OVER TWO HOURS. AN IMPELLA WAS EVENTUALLY PLACED, AND A LUCAS MECHANICAL COMPRESSION DEVICE WAS UTILIZED TO REMOVE THE PATIENT FROM THE LAB AND TRANSFER HIM TO THE INTENSIVE CARE UNIT (ICU). THE PLAN WAS TO DISCONTINUE COMPRESSIONS, WITH THE UNDERSTANDING THAT THE PATIENT WOULD EXPIRE. THE PATIENT EXPIRED THE SAME DAY AS THE PROCEDURE. | THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. BASED ON THE REPORTED INFORMATION, THERE WAS DIFFICULTY ADVANCING OR REMOVING THE VALVE, AND THE 18F SHEATH BECAME STUCK IN THE ARTERY WHICH MAY HAVE CONTRIBUTED TO THE DROP IN BLOOD PRESSURE AND SUBSEQUENT BLEEDING AND DEATH; HOWEVER, THIS COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THERE WAS NO IVL MALFUNCTION REPORTED FOR THE M5+ IVL CATHETER, AND THE L6 IVL CATHETER WAS NOT USED IN THE PATIENT. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI- CHOLECYSTECTOMY, THE PATIENT EXPERIENCED CARDIAC ARREST AND EXPIRED. THE INTUITIVE CLINICAL SALES REPRESENTATIVE (CSR) WAS INFORMED OF THE PATIENT EVENT AND TOLD THAT THE DETAILS ARE CONFIDENTIAL; THEREFORE, WAS NOT PROVIDED ANY SPECIFIC INFORMATION. THE CSR SPOKE WITH THE HEAD OF THE DEPARTMENT AND THE MEDICAL DIRECTOR OF THE HOSPITAL WHO STATED THAT THE PATIENT PROBABLY SUFFERED A HEART ATTACK AND DIED AFTER ATTEMPTED RESUSCITATION. THEY CONDUCTED THEIR OWN INTERNAL INVESTIGATION, AN AUTOPSY WAS CONDUCTED, AND STATED THAT NO FURTHER DETAILS WOULD BE DISCLOSED. THEY REPORTED THAT THE DA VINCI ROBOT HAD NO INFLUENCE ON THE INCIDENT, THERE IS NO CONNECTION BETWEEN DA VINCI AND THE INCIDENT, AND IT IS NOT INTUITIVE¿S CONCERN. | NO SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE WAS PROVIDED; THEREFORE, NO DAVINCI SYSTEM OR INSTRUMENT LOG FILES ARE AVAILABLE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DA VINCI SYSTEM USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED DURING A CHOLECYSTECTOMY REPORTEDLY FROM A MYOCARDIAL INFARCTION. HOW THAT DETERMINATION WAS MADE IS NOT PROVIDED. NO DEVICE RELATED ISSUES DURING THE CASE WERE REPORTED. THERE IS NO ALLEGATION OF ANY ISSUE WITH ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B3 DATE OF EVENT: EVENT DATE IS ESTIMATED AS THE INFORMATION PROVIDED ON 20-AUG-2024 WAS THAT THE EVENT OCCURRED "ABOUT 2 MONTHS AGO IN JUNE".
DREAMSTATION AUTO CPAP | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE TYPE OF REPORTED COMPLAINT WAS PREVIOUSLY MARKED INCORRECT AS SERIOUS INJURY. IN THIS REPORT IT HAS BEEN UPDATED AND CORRECTED AS DEATH.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP, JP | K131982 | THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARMS REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED HARM DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. IF ADDITIONAL INFORMATION IS OBTAINED ABOUT THIS EVENT.
REBAR | K093750 | ADDITIONAL INFORMATION RECEIVED REPORTED THE BASELINE LESION CHARACTERISTICS WAS ACOM ANEURYSM, 5.69MM X 5.12MM, THE ANEURYSM WAS FEEDING BY THE LEFT A1. THE DEATH WAS THER TO BE RELATED TO THE THERAPY. THERE WAS NO CLEAR CAUSE OF THE SUBARACHNOID HEMORRHAGE DETERMINED. THE PATIENT WAS ON ANTICOAGULANTS. ASS 100 AND CLOPIDOGREL 75MG. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A VENTILATOR'S SOUND ABATEMENT FOAM. THE REPORTER ALLEGED OF HAVING CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
NXSTAGE SYSTEM ONE | K050525 | A REPORT WAS RECEIVED ON 23 AUG 2024 FROM THE HOME THERAPY NURSE (HTN) OF PATIENT (B)(6), A 49 YEAR OLD FEMALE WITH A MEDICAL HISTORY INCLUDING MULTIPLE COMORBIDITIES AND END STAGE RENAL DISEASE APPROVED FOR PERFORMING SOLO HOME HEMODIALYSIS THERAPY, WHO STATED THE PATIENT WAS FOUND WITH THE ARTERIAL LINE DISCONNECTED AND AN UNSPECIFIED AMOUNT OF BLOOD LOST DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 27 AUG 2024 FROM THE HTN STATING THE PATIENT WAS FOUND UNRESPONSIVE WITH THE CENTRAL VENOUS CATHETER LEAKING. UPON EMERGENCY MEDICAL SERVICES (EMS) ARRIVAL THE PATIENT WAS BRADYCARDIC (NOS) WITH LOW BLOOD PRESSURE (NOS), FLUIDS (NOS) WERE ADMINISTERED. THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM (ER) AND CODED APPROXIMATELY 15 MINUTES UPON ARRIVAL. THE PATIENT WAS REVIVED BUT EXPIRED THE FOLLOWING MORNING, CAUSE OF DEATH HEART ATTACK/BLOOD LOSS. | THE CYCLER WAS NOT RECEIVED FOR EVALUATION AND NO DEFECT WAS FOUND DURING EVALUATION OF THE INVOLVED CARTRIDGE. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT PASSED AWAY AND THE DEVICE WAS GOING TO BE RETURNED AS IT WAS NO LONGER IN USE BY THE CUSTOMER. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT OUTCOME, HOWEVER LIMITED INFORMATION HAS BEEN PROVIDED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED THAT A PATIENT PASSED AWAY AND THE DEVICE WAS GOING TO BE RETURNED AS IT WAS NO LONGER IN USE BY THE CUSTOMER. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT OUTCOME, HOWEVER LIMITED INFORMATION HAS BEEN PROVIDED AT THIS TIME. THE DEVICE WAS RETURNED AND EVALUATED BY A THIRD-PARTY SERVICE CENTER. THE DEVICE WAS FOUND TO PASS ALL DEVICE TESTING. NO ABNORMALITIES WERE FOUND. THE DEVICE PERFORMED AS DESIGNED AND AS INTENDED. THE FOLLOWING CODES HAVE BEEN UPDATED AS FOLLOWS: EVALUATION METHOD CODE - TESTING OF ACTUAL/SUSPECTED DEVICE. EVALUATION RESULTS CODE - ANALYSIS OF INFORMATION PROVIDED BY USER/THIRD PARTY. COMPONENT GRID - NO INFORMATION. CONCLUSION CODE - NO PROBLEM DETECTED.
VIVO 45 LS | K193586 | ON AUGUST 26, 2024, BREAS MEDICAL WAS CONTACTED BY UNIFIED CARE SERVICES TO INSPECT A VIVO 45LS DEVICE, SN (B)(6), WHICH WAS BEING USED BY A PATIENT WHO DIED. THE DEATH OCCURRED ON (B)(6) 2024. NO FURTHER DETAILS WERE PROVIDED ABOUT THE EVEN AFTER MULTIPLE ATTEMPTS TO CONTACT THE INITIAL REPORTER. THERE IS NO ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. | BREAS MEDICAL RECEIVED THE DEVICE FOR INVESTIGATION ON SEPTEMBER 5, 2024. INVESTIGATION INCLUDED AN ANALYSIS OF THE LOG FILES, A VISUAL INSPECTION, AND FUNCTIONAL TESTING. ANALYSIS OF LOGS INDICATE THAT THE VIVO 45 LS ACCURATELY DELIVERED TREATMENT BASED ON THE MODES THAT WERE SET. THE DEVICE ALSO ALARMED ACCORDING TO SPECIFICATIONS WHEN MEASURED VALUES FELL OUT OF ALARM SETTING RANGES. THERE, BREAS MEDICAL HAS DETERMINED THAT THE DEVICE FUNCTIONED AS PROGRAMMED AND ALARMED AS DESIGNED. IT IS THEREFORE REASONABLE TO CONCLUDE THAT THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT DEATH.
FREEDOM PERIPHERAL NERVE STIMULATOR | K171366 | ON (B)(6) 2024, THE COMMERCIAL TEAM MEMBER ATTEMPTED TO FOLLOW-UP WITH THE PATIENT REGARDING THEIR TRIAL AND WAS INFORMED THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2024. THE COMMERCIAL TEAM MEMBER LAST SPOKE WITH THE PATIENT ON (B)(6) 2024 AND THEY WERE DOING WELL WITH THE TRIAL AND GETTING GREAT RELIEF IN THEIR BACK. UPON NOTIFYING THE PHYSICIANS OFFICE OF THE PATIENT'S DEATH, THE PHYSICIAN REPORTED THE PATIENT HAD A FUNGAL INFECTION THAT HAD CAUSED THE PATIENT TO LOSE A LOT OF WEIGHT. THE PATIENT WANTED TO PURSUE THE PNS TRIAL AS A MEANS OF TREATING THEIR BACK PAIN. | THE DEVICES ASSOCIATED WITH THIS COMPLAINT WERE NOT RETURNED AND COULD NOT BE ANALYZED BY ENGINEERING. ATTEMPTS TO RECOVER THE DEVICE WERE UNSUCCESSFUL, PREVENTING ENGINEERING FROM PERFORMING INVESTIGATION. THE OTHER ADVERSE EVENTS ISSUES QUESTIONNAIRE WAS REVIEWED FOR POTENTIAL CAUSES OF THE REPORTED ISSUE. BASED ON THIS REVIEW, IMPLANTING THE DEVICE IN AN OFF-LABEL LOCATION, IMPLANTING THE DEVICE TOO CLOSE TO THE TARGETED NERVE, AND INADVERTENTLY PUNCTURING THE BONE OR TISSUE DURING THE PROCEDURE HAVE BEEN RULED OUT AS POTENTIAL CAUSES. THE PATIENT HAD A FUNGAL INFECTION THAT HAD CAUSED THE PATIENT TO LOSE A LOT OF WEIGHT. IT IS NOT BELIEVED THAT THE DEVICE HAD ANYTHING TO DO WITH THE DEATH OF THE PATIENT. THE STIMULATOR IS USED TO TREAT PAIN. THE CAUSE OF THE REPORTED UNRELATED TO THE CURONIX DEVICE AS THE PATIENT HAD A PRE-EXISTING FUNGAL INFECTION (NO FAULT FOUND). RATES REVIEWED AT THE MOST RECENT COMPLAINT TRENDING MEETING DO NOT INDICATE A SIGNIFICANT INCREASE IN OTHER ADVERSE EVENTS ISSUES. OTHER ADVERSE EVENTS ISSUES RATES REMAIN ACCEPTABLY LOW; THUS, A CAPA IS NOT REQUIRED. OTHER ADVERSE EVENTS ISSUES RATES WILL CONTINUE TO BE TRACKED AND TRENDED.
POWERPORT ISP MRI 8CF INT. W/SP, ATT, SL | K063377 | IT WAS REPORTED THAT POST PORT PLACEMENT IN THE RIGHT SUBCLAVIAN VEIN WITH POSITION IN THE ATRIO-CAVAL JUNCTION, THE CATHETER ALLEGEDLY DISCONNECTED AND MIGRATED TO THE RIGHT CARDIAC CAVITIES WITH LARGE LEFT PLEURAL EFFUSION. IT WAS FURTHER REPORTED THAT THE PATIENT DEVELOPED ASPIRATIVE PNEUMONIA IN THE LEFT LUNG ACCOMPANIED BY PLEURAL EFFUSION OF THE ENTIRE HEMITHORAX. REPORTEDLY, REMOVAL OF THE PORT AND FOREIGN BODY LOCATED IN THE RIGHT ATRIUM WAS PERFORMED WITHOUT COMPLICATIONS. PATIENT DYSPNEA DID NOT IMPROVE. THE PATIENT REPORTEDLY EXPIRED. | H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE POWERPORT IMPLANTABLE PORT PRODUCTS THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE POWERPORT IMPLANTABLE PORT PRODUCTS ARE IDENTIFIED IN D2 AND G4. MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. HOWEVER, 5 MEDICAL IMAGES WERE REVIEWED. IMAGE REVIEW SHOWED THE ENTIRE CATHETER WITHIN THE RIGHT ATRIUM, IN ONE CASE, AND A SEGMENT OF THE CATHETER IN THE INFERIOR VENA CAVA IN ANOTHER. INFILTRATE DEVELOPING IN THE LEFT LUNG AND BILATERAL PNEUMOTHORACES WERE ALSO OBSERVED. BASED ON THE IMAGE REVIEW, THE REPORTED MIGRATION CAN BE CONFIRMED, AS WELL AS IDENTIFIED DISCONNECTION. PNEUMOTHORAX IS ALSO CONFIRMED. A DEFINITIVE ROOT CAUSE FOR THE DISCONNECTION AND MIGRATION COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. IT IS UNKNOWN WHETHER PATIENT AND/OR PROCEDURAL ISSUES CONTRIBUTED TO THE REPORTED EVENT. PATIENT DEATH AND PNEUMOTHORAX/RELATED EVENTS CANNOT BE VERIFIED TO BE RELATED TO THE PORT DEVICE. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE, A LABELING REVIEW IS NOT REQUIRED. H10: B5, G3, H6 (DEVICE, METHOD). H11: H6 (RESULT, CONCLUSION). H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
TRI-STAPLE 2.0 | K160176 | ACCORDING TO THE REPORTER, DURING A PROCEDURE, AFTER FIRING THE DEVICE AT THE TIME OF CREATING A PHARYNGEAL POUCH, IT WAS OBSERVED THAT THE STAPLES FIRED. HOWEVER, THE BLADE DID NOT CUT THROUGH THE TISSUE. THEREFORE, NONE OF THE POUCH WAS OPENED. TO RESOLVE THE ISSUE, THE SURGEON USED MICRO SCISSORS TO OPEN BETWEEN THE STAPLE LINE. FOUR TO FIVE DAYS FOLLOWING THE PROCEDURE THE PATIENT EXPIRED DUE TO CARDIAC ARREST AND WAS NOT DEVICE RELATED. | ADDITIONAL INFORMATION: B2 (DATE OF DEATH), B5, G3. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND PLEURAL EFFUSION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF CANCER, DEATH PROBLEMS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF CANCER, DEATH PROBLEMS RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. MEDICAL INTERVENTION WAS NOT SPECIFIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX G: DATE RECEIVED BY MFG (RFB) HAS BEEN UPDATED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION RESPIRATORY TRACT IRRITATION DIZZINESS AND/OR HEADACHE KIDNEY DISEASE/TOXICITY LUNG DISEASE DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LTV 1150 VENTILATOR | K051767 | IT WAS REPORTED TO VYAIRE MEDICAL THAT A PEDIATRIC PATIENT WAS ON THE LTV 1150 DEVICE VIA A TRACHEOSTOMY TUBE. ON THE MORNING OF (B)(6), THE DAY-SHIFT NURSE ARRIVED TO FIND THE TRACHEOTOMY HAD BECOME DISLODGED AT SOME POINT DURING THE EVENING AND WAS UNABLE TO GET A PULSE READING. | VYAIRE MEDICAL FILE IDENTIFICATION: (B)(4). H3: ROOT CAUSE FINDINGS: THE DEVICE EVENT LOG WAS REVIEWED AND THE INCIDENT WAS CONFIRMED DURING THE REVIEW. ON THE DATE OF THE INCIDENT IT WAS CONFIRMED THAT THE DEVICE WAS ALARMING AUDIBLY AND VISUALLY FOR MULTIPLE ALARMS FOR A TOTAL OF 18 MINUTES UNTIL THE DEVICE WAS MANUALLY POWERED OFF. THE DEVICE WAS TESTED AND SET TO VENTILATE WITH THE SAME PATIENT SETTINGS FOR A TOTAL OF 71 HOURS AND CONFIRMED TO BE FUNCTIONING AS INTENDED WITHOUT ANY FAULTS.
SENSATION 7FR. 40CC IAB | K063525 | IT WAS REPORTED THAT DURING INTRA-AORTIC BALLOON (IAB) THERAPY, LARGE AMOUNTS OF BLOOD WERE SEEN IN THE HELIUM TUBING AND ALL THE WAY BACKED UP TO THE CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP). THE IABP ALARMED WITH ¿GAS LOSS¿ AT THE TIME OF THE EVENT. IT WAS NOTED THAT THE PATIENT WAS INTUBATED AND SEDATED AND NOT MOVING AT THE TIME OF THE ALARM AND BACKFLOW OF BLOOD. THE PATIENT WAS STABLE AT THIS TIME. THEY WERE PREPARING THE PATIENT FOR IAB REMOVAL AND THE PHYSICIAN WAS AT THE BEDSIDE. THE PATIENT WAS TAKEN BACK DOWN TO THE CATH LAB FOR INSERTION OF A NEW IAB. IT WAS LATER REPORTED THAT THE PATIENT HAD EXPIRED. | UPDATED FIELD(S): B1, B2, B5, H1, H6. THE PRODUCT HAS BEEN RETURNED TO THE MANUFACTURER, BUT IS PENDING INVESTIGATION. ONCE THE INVESTIGATION IS COMPLETED A SUPPLEMENTAL REPORT WITH OUR FINDINGS WILL BE SUBMITTED. COMPLAINT RECORD ID # (B)(4).
HEMOCHRON ACTIVATED CLOTTING TIME PLUS (ACT+) TEST | K941007 | HEALTHCARE PROFESSIONAL REPORTED THAT A HEMOCHRON SIGNATURE ELITE ANALYZER AND ACT+ TEST DEVICE WAS USED TO MONITOR HEPARIN ANTICOAGULATION IN A PATIENT THAT WAS UNDERGOING A CORONARY ANGIOPLASTY PROCEDURE. ACT TEST RESULT OBTAINED AT BASELINE WAS NOT SPECIFIED. THE PATIENT RECEIVED CLOPIDOGREL BY MOUTH AND WAS GIVEN A 5000U INTRAVENOUS DOSE OF HEPARIN. THE RESULT OF AN ACT TEST OBTAINED 1 HOUR AFTER HEPARIN ADMINISTRATION WAS OUT-OF-RANGE HIGH. THE PATIENT EXPIRED DUE TO A SLOW HEART RATE AT AN UNSPECIFIED TIME. NO OTHER MEDICAL OR LABORATORY INFORMATION WAS SPECIFIED. ELECTRONIC QUALITY CONTROL (EQC) TESTING PERFORMED PRIOR TO THE PROCEDURE PASSED. | THIS MDR REFERENCES ACCRIVA DIAGNOSTICS' COMPLAINT NUMBER (B)(4) FOR HEMOCHRON JACT+ CUVETTE LOT D3JAC095 WITH A LABELED EXPIRATION DATE OF 07/31/2024. THE CUVETTE WAS USED WITH A HEMOCHRON SIGNATURE ELITE INSTRUMENT DURING THE PROCEDURE. THE SERIAL NUMBER OF THE SIGNATURE ELITE INSTRUMENT WAS NOT SPECIFIED. THE END USER SUSPECTS THE ISSUE MAY BE RELATED TO IMPROPER SAMPLE HANDLING OR COLLECTION PROCEDURES. DHR REVIEW OF JACT+ LOT D3JAC095 PRODUCTION RECORDS DID NOT REVEAL ANY NONCONFORMITIES OR ANOMALIES THAT COULD HAVE BEEN A FACTOR FOR OR A POTENTIAL ROOT CAUSE OF THE ALLEGATION. RELEASE TESTING OF THE CUVETTE LOT MET ALL SPECIFICATIONS. RETAINS OF JACT+ CUVETTE LOT D3JAC095 WERE TESTED UNDER LSR 2024-CSS-110 AND THE DEVICE PERFORMED WITHIN EXPECTATIONS. THE HEMOCHRON SIGNATURE ELITE INSTRUMENT WAS NOT RETURNED TO ACCRIVA DIAGNOSTICS, INC. FOR EVALUATION SO NO FURTHER INFORMATION IS AVAILABLE. HEALTH EFFECT CLINICAL CODE: 2493, ISCHEMIC HEART DISEASE. HEALTH EFFECT IMPACT CODE: 1802, DEATH. DEVICE PROBLEM CODE: 2993, ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM. COMPONENT CODE: 4765, CUVETTE. TYPE OF INVESTIGATION: 3331, ANALYSIS OF PRODUCTION RECORDS; 4114: DEVICE NOT RETURNED. INVESTIGATION FINDINGS: 213, NO DEVICE PROBLEM FOUND. INVESTIGATION CONCLUSIONS: 4310, CAUSE TRACED TO NON-DEVICE RELATED FACTORS. ACCRIVA DIAGNOSTICS, INC. HAS REQUESTED ALL DATA REQUIRED FOR FORM 3500A.
RESPIRONICS | K102985 | IT WAS REPORTED THAT A PATIENT HAD EXPIRED DURING CLINICAL AND THERAPEUTIC USE OF A V60 VENTILATOR. FURTHER INFORMATION STATES THAT THE PATIENT HAD ALLEGED DIFFICULTY BREATHING AND THE DEVICE, MASK, AND BREATHING CIRCUIT WERE SUBSEQUENTLY REPLACED. IT WAS FURTHER NOTED THAT THE DEVICE PROVIDED A HIGH RATE AND HIGH TIDAL VOLUME ALARM IN CONJUNCTION WITH A LEAK OF 10 L/MIN SOME TIME PRIOR TO THE PATIENT EXPIRING. NO OTHER INFORMATION WAS DISCLOSED. THE DEVICE AND ACCESSORIES HAVE BEEN DETERMINED TO BE IRRELEVANT TO THE PATIENT DEATH WITH NO CAUSE AND/OR CONTRIBUTION OF THE DEVICE TO THE PATIENT EXPIRATION NOTED OR ALLEGED. THE CUSTOMER COULD NOT PROVIDE THE DIAGNOSTIC REPORT OR THE EVENT LOG OF THE DEVICE. INFORMATION WAS PROVIDED THAT THE DEVICE WILL NOT BE REPAIRED AND THERE IS NO PLAN OF ACTION FOR REPAIR PER THE CUSTOMER. NO FURTHER SERVICE WAS PROVIDED BY PHILIPS. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED FOR FURTHER INVESTIGATION AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED AS APPLICABLE. | E1: REPORTER PHONE NUMBER - (B)(6).
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | ON 18/SEP/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2024. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2024. THE PATIENT¿S DAUGHTER DISCOVERED THE PATIENT HAD EXPIRED SEVERAL HOURS PRIOR (BASED ON RIGAMORTUS) WHEN SHE ARRIVED TO DISCONNECT THE PATIENT FROM THE LIBERTY SELECT CYCLER (AS SHE DID EVERY MORNING). THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES, AND THE PATIENT WAS TAKEN DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS ¿NATURAL CAUSES¿ BUT WOULD NOT ELABORATE ANY FURTHER. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE EVENTS OCCURRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS DUE TO NATURAL CAUSES AND WAS UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
EDWARDS TRANSFEMORAL BALLOON CATHETER | K140241 | AS REPORTED, IT WAS A CASE OF AN IMPLANT OF A 26MM SAPIEN 3 VALVE, IN AORTIC POSITION. DURING PROCEDURE, AFTER PERFORMING PRE-DILATION OF THE NATIVE AORTIC VALVE WITH AN 23MM EDWARDS BAV BALLOON, THE PATIENT PRESENTED HEMODYNAMIC INSTABILITY, FOLLOWED BY CARDIORESPIRATORY ARREST. CARDIOPULMONARY RESUSCITATION AND OROTRACHEAL INTUBATION WERE INITIATED. DESPITE THIS EVENT, IT WAS DECIDED TO PERFORM THE VALVE IMPLANT HOWEVER, IT WAS WITHOUT SUCCESS AFTER FORTY SEVEN (47) MINUTES IN CARDIORESPIRATORY ARREST. THE PERCEIVED ROOT CAUSE WAS THAT THE PATIENT WAS HEMODYNAMICALLY FRAGILE. | PER THE INSTRUCTIONS FOR USE (IFU), CARDIAC ARREST IS A KNOWN POTENTIAL ADVERSE EVENT ASSOCIATED WITH BALLOON VALVULOPLASTY, THE USE OF LOCAL AND/OR GENERAL ANESTHESIA, BIOPROSTHETIC HEART VALVES, AND THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE. CARDIAC ARREST OCCURS WHEN A MAJOR CARDIOVASCULAR, RESPIRATORY, OR METABOLIC EVENT RESULTS IN THE INABILITY OF CARDIAC MUSCLE TO GENERATE SUFFICIENT FORCE IN RESPONSE TO ELECTRICAL DEPOLARIZATION. IT MAY BE CAUSED BY A PROFOUND CARDIOVASCULAR INSULT. EXAMPLES INCLUDE SEVERE PROLONGED HYPOXIA OR ACIDOSIS, EXTREME HYPOVOLEMIA, FLOW-RESTRICTING PULMONARY EMBOLUS, CARDIAC TAMPONADE, THROMBOSIS (CORONARY OR PULMONARY). THESE EVENTS MAY BE RELATED TO THE EDWARDS DEVICE IF IT IS ASSOCIATED WITH CARDIAC TAMPONADE, ANNULAR RUPTURE, OR OTHER EDWARDS'S DEVICE-RELATED BLEED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. INVESTIGATION RESULTS SUGGEST/INDICATE AFTER PERFORMING PRE-DILATION OF THE NATIVE AORTIC VALVE WITH AN 23 MM EDWARDS BAV BALLOON, THE PATIENT PRESENTED HEMODYNAMIC INSTABILITY, FOLLOWED BY CARDIORESPIRATORY ARREST. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
EVIS EXERA III DUODENOVIDEOSCOPE | K220587 | IT WAS REPORTED IN A SOCIAL MEDIA POST THAT A PATIENT UNDERWENT AN ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP), PROCEDURE WITH AN OLYMPUS SCOPE BELIEVED TO BE JF AT A US FACILITY PASSED AWAY. THE POST STATED AN UNSTERILIZED OLYMPUS SCOPE WAS USED THAT WAS RECALLED. FURTHER FOLLOW-UP HAS BEEN CONDUCTED; HOWEVER, NO INFORMATION IS AVAILABLE REGARDING PATIENT DEATH AT THIS TIME. | THIS REPORT IS BEING SUPPLEMENTED TO PROVIDE ADDITIONAL INFORMATION BASED ON THE LEGAL MANUFACTURER'S (LM) FURTHER INVESTIGATION. THE DEVICE HISTORY RECORD WAS UNABLE TO BE REVIEWED FOR THIS DEVICE SINCE THE SERIAL AND/OR LOT NUMBER WAS NOT PROVIDED. HOWEVER, OLYMPUS ONLY RELEASES PRODUCTS TO MARKET THAT MEET ALL MANUFACTURING SPECIFICATIONS AND FINAL PRODUCT RELEASE CRITERIA. BASED ON THE RESULTS OF THE INVESTIGATION, A RELATIONSHIP BETWEEN THE SUBJECT DEVICE AND THE REPORTED DEATH COULD NOT BE IDENTIFIED. THEREFORE, THE ROOT CAUSE COULD NOT BE DETERMINED. THE EVENT CAN BE DETECTED/PREVENTED BY FOLLOWING THE INSTRUCTIONS FOR USE (IFU) IN SECTION: ¿AN INSUFFICIENTLY REPROCESSED ENDOSCOPE AND/OR ACCESSORY MAY POSE AN INFECTION CONTROL RISK TO THE PATIENTS AND/OR OPERATORS WHO TOUCH THEM.¿ OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
INTELLAMAP ORION HIGH RESOLUTION MAPPING CATHETER | K143481 | IT WAS REPORTED THAT APPROXIMATELY TWO WEEKS FOLLOWING A PULSE FIELD ABLATION (PFA) PROCEDURE USING A FARAWAVE PFA CATHETER THE PATIENT PASSED AWAY. THE ABLATION PROCEDURE WAS SUCCESSFULLY COMPLETED ON (B)(6) 2024, AND NO PATIENT COMPLICATIONS WERE NOTED AT THAT TIME. BOSTON SCIENTIFIC WAS THEN INFORMED THAT THE PATIENT PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS NOT REPORTED, AND NO AUTOPSY HAS BEEN PERFORMED AS OF YET. THE CATHETER WAS DISPOSED BY THE FACILITY AND WILL NOT BE RETURNED AS THE DEVICE PERFORMED AS EXPECTED DURING THE PROCEDURE AND THERE WERE NO REPORTS OF ANY PERFORMANCE CONCERNS. IT WAS FURTHER REPORTED THAT THE FARAWAVE CATHETER WAS USED TO PROVIDE 56 PFA LESIONS BETWEEN THE PULMONARY VEINS AND THE POSTERIOR WALL (OFF LABEL AS THE FARAWAVE IS ONLY INDICATED FOR TREATMENT OF THE PULMONARY VEINS). ACCORDING TO THE ACCOUNT, THE MOST LIKELY CAUSE OF DEATH WAS COMPLICATIONS RELATED TO ANESTHESIA. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
DREAM STATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY FAILURE, HEART FAILURE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAM STATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, CANCER, DEATH, OTHER (UNSPECIFIED EVENT), LUNG (NON-SMALL CELL CARCINOMA). THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
VERSYS FEMORAL STEM CEMENTED REVISION/CALCAR 12/14 NECK TAPER SIZE 13 250MM | K210842 | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN INITIAL HIP PROCEDURE ON AN UNKNOWN DATE USING A COMPETITOR'S IMPLANTS. SUBSEQUENTLY, THE PATIENT WAS REVISED FOR A PERIPROSTHETIC FRACTURE. DURING THE REVISION HEMI ARTHROPLASTY SURGERY, THE PATIENT'S CONDITION DETERIORATED, AND THE PATIENT DIED OF CARDIOPULMONARY ARREST. IT IS NOT KNOWN WHETHER THE PATIENT HAD ANY UNDERLYING MEDICAL CONDITIONS THAT MAY HAVE BEEN RELATED TO THE EVENT. THE PATIENT'S CARDIOPULMONARY ARREST OCCURRED AFTER THE BONE CEMENT AND STEM WERE INSERTED INTO THE PATIENT. THE CEMENT IS A COMPETITOR PRODUCT. THE CAUSE OF THE CARDIOPULMONARY ARREST WAS NOT DETERMINED; HOWEVER, NO CAUSAL RELATIONSHIP BETWEEN THE DEATH AND THE PRODUCT HAS BEEN NOTED. THERE WAS NO PATIENT INFORMATION OR MEDICAL RECORDS AVAILABLE. THOUGH REQUESTED, NO ADDITIONAL INFORMATION WAS AVAILABLE. | (B)(4). D10: CAT# UNKNOWN LOT# UNKNOWN DEPUY BONE CEMENT G2: FOREIGN: COUNTRY: JAPAN NO PRODUCT WAS RETURNED FOR EVALUATION. IT WAS REPORTED THAT A COMPETITOR PRODUCT WAS BEING REVISED IN A HIP PROCEDURE. AFTER COMPETITOR CEMENTATION AND STEM INSERTION, THE PATIENT'S CARDIAC CONDITION DETERIORATED LEADING TO CARDIAC ARREST WITHOUT RESUSCITATION. THE PATIENT'S AGE AND COMORBID CONDITIONS REMAIN UNKNOWN. PROCEDURAL RELATED COMPLICATIONS ARE INFLUENCED BY THE TYPE OF SURGERY, PATIENTS PRE-EXISTING COMORBID STATE, AND PERIOPERATIVE MANAGEMENT. SIDE EFFECTS OF THE ADMINISTRATION OF ANESTHESIA ARE A KNOWN CAUSE OF CARDIAC STRESS DURING THE INTRAOPERATIVE AND POSTOPERATIVE RECOVERY. MOST PATIENTS WITH A CARDIAC HISTORY WILL HAVE A CARDIAC WORKUP PRIOR TO CLEAR THEM FOR SURGERY TO DEMONSTRATE THE PATIENT'S CARDIAC STATUS IS FIT TO WITHSTAND THE STRESS OF ANESTHESIA, POSITIONING, AND SURGERY. AS THE REPORTED COMPLAINT INDICATED, A CARDIAC COMPLICATION DEVELOPED INTRAOPERATIVELY, BUT ANY FURTHER INTERVENTION WAS INSUFFICIENT TO PREVENT THE TERMINAL OUTCOME. AS THE REPORTED EVENT WAS UNRELATED TO THE IMPLANTED STEM, A REVIEW OF THE DEVICE HISTORY RECORDS WAS NOT PERFORMED. IF ANY FURTHER INFORMATION IS FOUND WHICH WOULD CHANGE OR ALTER ANY CONCLUSIONS OR INFORMATION, A SUPPLEMENTAL WILL BE FILED ACCORDINGLY. ZIMMER BIOMET WILL CONTINUE TO MONITOR FOR TRENDS.
DREAM STATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER, DEATH, PANCREATIC, NASAL DRIPS AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PULMONARY FIBROSIS, CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
WATCHMAN TRUSTEER ACCESS SYSTEM | K240018 | IT WAS REPORTED THAT RETROPERITONEAL HEMORRHAGE AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) IMAGING WAS UTILIZED. VENOUS GROIN ACCESS WAS OBTAINED USING ULTRASOUND. THE PATIENT HAD A NOTED INFERIOR VENA CAVA (IVC) FILTER, SO AN INTRODUCER SHEATH WAS USED. TRANSSEPTAL ACCESS WAS OBTAINED, AND A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS POSITIONED AT THE LAA. A 24MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY SYSTEM (WDS) WERE ADVANCED, DEPLOYED AND IMPLANTED SUCCESSFULLY. THE PROCEDURE WAS CONCLUDED, AND THE PATIENT WAS STABLE. THE PATIENT DIED BEFORE BEING DISCHARGED FROM THE HOSPITAL. ACCORDING TO THE PHYSICIANS, THE LIKELY CAUSE OF DEATH WAS A RETROPERITONEAL BLEED CAUSED BY ACCESS COMPLICATIONS DURING THE INDEX PROCEDURE, BUT NOT ASSOCIATED WITH THE WATCHMAN DEVICES. IT WAS FURTHER REPORTED THAT THE DEATH OCCURRED THE SAME DATE OF THE INDEX PROCEDURE. THE WAS KINKED DURING THE PROCEDURE DUE TO ANATOMICAL CHALLENGES NEAR THE DISTAL END OF THE SHEATH, AND A SECOND WAS OPENED TO COMPLETE THE PROCEDURE. AT THE END OF THE PROCEDURE, A NON-BOSTON SCIENTIFIC VASCULAR CLOSURE DEVICE WAS USED TO CLOSE THE GROIN ACCESS AND THE PATIENT HAD MINOR BLEEDING ISSUES POST PROCEDURE. FULL DOSE OF HEPARIN WAS GIVEN AFTER GETTING VENOUS GROIN ACCESS, DOSE GIVEN PER PATIENT WEIGHT. THE IVC FILTER WAS EASILY WIRED AND PER THE PHYSICIAN IT WAS NOT DISTURBED DURING THE PROCEDURE. | B2: DATE OF DEATH CORRECTED. B3: EVENT DATE CORRECTED. H6: CODE UPDATED FROM ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM TO MATERIAL DEFORMATION.
ESSENCE RENTAL | K042655 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO AN ESSENCE RENTAL DEVICE. THE PATIENT ALLEGEDLY PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED ABOUT THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ABOUT AN ALLEGATION OF CANCER AND DEATH. THERE WAS NO REPORT OF SPECIFIC MEDICAL INTERVENTIONS. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE CARRIED OUT. IF ANY ADDITIONAL INFORMATION IS RECEIVED A FOLLOW-UP REPORT WILL BE FILED. | CORRECTION TO PREVIOUS INITIAL SUBMISSION DUE DATE: ACTUAL DUE DATE IS 11NOV2023
DREAMSTATION ST30 | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR DEVICE EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION AND NO OTHERS CONTAMINATION FOUND. 0 ERROR LOG. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ESSENCE RENTAL | K042655 | THE MANUFACTURER RECEIVED INFORMATION RELATED TO AN ESSENCE NEBULIZER DEVICE. THE USER'S SON CONTACTED THE MANUFACTURER STATING THAT THE USER IS DECEASE AND THAT HE DOES NOT HAVE THE NEBULIZER FOR RETURN. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ESSENCE RENTAL WITH SSD | K042655 | THE MANUFACTURER RECEIVED INFORMATION REGARDING AN ESSENCE RENTAL WITH SSD. THE PATIENT HAS ALLEGEDLY PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ESSENCE RENTAL | K042655 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE ESSENCE RENTAL HAD STATES PATIENT HAS PASSED AWAY AND HIS WIFE HAS LOST THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY BECAME AWARE THAT A USER OF THE ESSENCE RENTAL HAD STATES PATIENT HAS PASSED AWAY AND HIS WIFE HAS LOST THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. IN BOX H: (DEVICE) PROBLEM CODE GRID HAS BEEN UPDATED IN THIS FOLLOW-UP.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REP DREAMSTATION AUTO BIPAP. THE REPORTER ALLEGED THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
CENTRIMAG BLOOD PUMP | K020271 | THE PATIENT WAS BEING TRANSITIONED TO COMFORT CARE AND WITHDRAWAL OF CARE WAS INITIATED. THE PATIENT PASSED AWAY. | MANUFACTURER'S INVESTIGATION CONCLUSION: A DIRECT CORRELATION BETWEEN THE CENTRIMAG BLOOD PUMP AND THE REPORTED PATIENT OUTCOME COULD NOT BE CONCLUSIVELY ESTABLISHED THROUGH THIS EVALUATION. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED BY THE ACCOUNT. THE LOT NUMBER OR OTHER IDENTIFYING INFORMATION OF THE PRODUCT WAS NOT REPORTED AND WAS NOT ABLE TO BE DETERMINED DURING THE INVESTIGATION. THE CENTRIMAG BLOOD PUMP INSTRUCTIONS FOR USE (IFU), REV. C IS CURRENTLY AVAILABLE. THE CENTRIMAG BLOOD PUMP IFU LISTS POTENTIAL ADVERSE EVENTS, INCLUDING DEATH, THAT MAY BE ASSOCIATED WITH THE USE OF THE CENTRIMAG CIRCULATORY SUPPORT SYSTEM. IFU WARNING #10: FREQUENT PATIENT AND DEVICE MONITORING IS RECOMMENDED. IFU CAUTION #2: THIS DEVICE SHOULD ONLY BE USED BY PERSONS THOROUGHLY TRAINED IN EXTRACORPOREAL CIRCULATION PROCEDURES. IFU CAUTION #15: ALWAYS HAVE A BACKUP CENTRIMAG PUMP, CONSOLE, MOTOR, AND ACCESSORIES AVAILABLE FOR USE. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
EVERFLO DOM | K061261 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE DEVICE WAS INVOLVED IN A HOUSE FIRE. BC1 ARRIVED ON SCENE OF WORKING FIRE INVOLVING A SINGLE-STORY RESIDENTIAL STRUCTURE. IT WAS REPORTED THE PATIENT'S BODY WAS FOUND IN THE HOME AFTER THE FIRE WAS CONTROLLED. THE PATIENT IS REPORTED TO HAVE UNDIAGNOSED DEMENTIA. | IN PREVIOUS REPORT, "BOX H - (DEVICE) PROBLEM CODE GRID (1) " WAS MENTIONED INCORRECTLY. THE DEVICE IS NOT PART OF RECALL. IN THIS REPORT, BOX H HAS BEEN UPDATED OR CORRECTED.
ESSENCE WITH SSD | K042655 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO AN ESSENCE DEVICE. THERE WAS AN ALLEGATION OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LINEAR 7.5FR. 40CC IAB (CHINA) | K041281 | IT WAS REPORTED THAT THE PATIENT WAS ADMITTED WITH ACUTE MYOCARDIAL INFARCTION (AMI) AND AN INTRA-AORTIC BALLOON (IAB) WAS IMMEDIATELY INSERTED. WHEN CONNECTED TO THE PUMP, AN UNKNOWN ALARM WAS GENERATED. IT WAS NOTED THAT THE BLOOD PRESSURE (BP) DISPLAYED BY THE PUMP WAS VERY DIFFERENT FROM THAT OF THE BEDSIDE MONITOR. THE PHYSICIAN IMMEDIATELY INSERTED A NEW IAB, BUT THEY THOUGHT THE BEST TIME FOR TREATMENT WAS MISSED. THE PATIENT REMAINED IN THE ICU. THE PATIENT PASSED AWAY TWO DAYS LATER. | EVENT SITE TELEPHONE: (B)(6). THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER SO WE ARE UNABLE TO COMPLETE AN EVALUATION. IF PROVIDED WE WILL SEND A SUPPLEMENTAL REPORT WITH OUR ADDITIONAL FINDINGS. COMPLAINT RECORD ID (B)(4).
REMSTAR PLUS C-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. DUE DATE HAS BEEN UPDATED. IN BOX G: DATE RECEIVED BY MFG HAS BEEN UPDATED. IN BOX H6: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAVE BEEN UPDATED.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THE PATIENT WENT INTO CARDIAC ARREST AND PASSED AWAY; HOWEVER, THE DEVICE DID NOT GENERATE AN ALARM. THE PATIENT WAS BEING MONITORED WITH TELEMETRY MX40 WHEN THE PATIENT DEVELOPED ASYSTOLE, FOR WHICH NO ALARM WAS GENERATED. THE NURSING STAFF WAS ALERTED TO THE CHANGE IN CONDITION BY THE PATIENT'S ROOMMATE. RESUSCITATION WAS ATTEMPTED BUT WAS UNSUCCESSFUL AND THE PATIENT SUBSEQUENTLY PASSED AWAY. GOOD FAITH EFFORTS ARE ONGOING. | A REVIEW OF THE DEVICE LOGS REVEALED THE DEVICE WAS WORKING AS DESIGNED. THE MX40 WORKED CORRECTLY AND SENT ALL INFORMATION TO THE CENTRAL. YELLOW AND RED ALARMS WERE GENERATED AT THE CENTRAL AND AUDIBLE TONES WERE GENERATED AS WELL. IN ADDITION, IT WAS NOTED THE ALARMS WERE ALSO SILENCED AT THE CENTRAL STATION. IT IS REITERATED THE NURSING STAFF DID INDICATE ALARMS WERE COMING THROUGH TO THEIR PHONES BUT THEY COULD NOT CONFIRM WHICH ALARMS. BASED ON THIS INFORMATION, IT REMAINS UNKNOWN WHETHER DEVICE BEHAVIOR, USE ERROR, ALARM MANAGEMENT, OR CLINICAL WORKFLOW MAY HAVE BEEN FACTORS IN THE EVENT. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND CHRONIC MYELOID LEUKEMIA. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT EXPIRED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND CHRONIC MYELOID LEUKEMIA. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT EXPIRED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX D AND H HAS BEEN UPDATED.
DEVILBISS HEALTHCARE | K071397 | DEVILBISS HEALTHCARE WAS NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY A PROVIDER WHO STATED THAT THE DEVICE WAS ALARMING. THE DEVICE IS OVER TWO YEARS OLD AND WAS RETURNED TO THE PROVIDER WHO TESTED THE DEVICE AND FOUND THE DEVICE TO HAVE LOW OXYGEN CONCENTRATION. THE PROVIDER REPORTED THAT THE DEVICE WAS VISUALLY AND AUDIBLY ALARMING AS PER SPECIFICATION. ROUTINE SERVICE IS REQUIRED FOR THE PROPER FUNCTION OF THE DEVICE. THE DEVICE IS NOT A LIFE SUPPORTING OR LIFE SUSTAINING DEVICE. THE LAST END USER TO UTILIZE THE DEVICE WAS REPORTED TO HAVE PASSED AWAY. IT IS UNKNOWN IF THE DEVICE CAUSED OR CONTRIBUTED TO THE END USER'S DEATH. DEVILBISS HEALTHCARE WILL FILE AN UPDATE IF ADDITIONAL INFORMATION BECOMES AVAILABLE. | THE PREVIOUSLY SUBMITTED REPORT HAD LACKING OR INCORRECT INFORMATION IN SECTION D4 UDI. SECTION D4 HAS BEEN UPDATED WITH THE CORRECT INFORMATION.
2008T HEMODIALYSIS SYS., WITH CDX | K150708 | ON (B)(6) 2024, IT WAS REPORTED THAT A PATIENT EXPERIENCED A CARDIAC EVENT DURING HEMODIALYSIS (HD) TREATMENT UTILIZING THE FRESENIUS 2008T MACHINE. THE PATIENT WAS SUBSEQUENTLY TRANSPORTED TO THE HOSPITAL AND PASSED AWAY. THE MACHINE WAS EVALUATED BY THE BIOMEDICAL TECHNICIAN AND THE ULTRAFILTRATION (UF) PUMP CALIBRATION WAS NOT WITHIN SPECIFICATIONS. THE MEASUREMENT WAS READING 23.3ML (23.90-24.10ML IS ACCEPTABLE RANGE). ADDITIONAL INFORMATION WAS PROVIDED THROUGH FOLLOW-UP WITH THE CLINICAL MANAGER (CM). THE PATIENT ARRIVED FROM THE COUNTY JAIL ACCOMPANIED BY A GUARD FOR A REGULARLY SCHEDULED HD TREATMENT ON (B)(6) 2024. THE PATIENT¿S 4-HOUR TREATMENT WAS INITIATED AT 6:49AM UTILIZING THE FRESENIUS 2008T MACHINE. THE PATIENT¿S BLOOD PRESSURE (BP) READINGS WERE NORMAL THROUGHOUT THE TREATMENT. THE PATIENT WAS TALKING AND LAUGHING WITH THE GUARD DURING THE TREATMENT. APPROXIMATELY TWO HOURS AND TWENTY-FIVE MINUTES INTO THE TREATMENT, THE PATIENT STATED THEY WERE GOING TO SLEEP AND PULLED THE BLANKET UP OVER HIS HEAD. THE DIALYSIS TECHNICIAN WHO WAS TAKING BP READINGS ON OTHER PATIENTS SAW THIS AND CALLED OUT TO THE PATIENT TO PULL THE BLANKET DOWN. THE PATIENT DID NOT RESPOND. THE TECH REPORTED HEARING WHAT SOUNDED LIKE SNORING COMING FROM THE PATIENT AND PROCEEDED TO REMOVE THE BLANKET FROM THEIR HEAD. TWO HOURS AND THIRTY-SIX MINUTES INTO THE TREATMENT, THE PATIENT CODED. 911 WAS CALLED AND A FULL CODE WAS RUN. APPROXIMATELY 8-10 MINUTES OF CARDIOPULMONARY RESUSCITATION (CPR) AND BAG MASK VENTILATION WERE ADMINISTERED PRIOR TO EMERGENCY MEDICAL SERVICES (EMS) ARRIVING. THE CM CONTINUED TO ASSIST THE EMTS AT THE CLINIC AS THEY DETERMINED IT NECESSARY TO INTUBATE THE PATIENT ON SCENE DUE TO THE PATIENT¿S MUSTACHE MAKING IT DIFFICULT TO OBTAIN A GOOD SEAL WITH THE BAG MASK VENTILATION. THE PATIENT¿S HEART RHYTHM WAS VENTRICULAR TACHYCARDIA (V TACH) ON THE DEFIBRILLATOR MONITOR. THIS WAS LATER CONFIRMED BY THE PHYSICIAN. THE PATIENT WAS ALSO ADMINISTERED EPINEPHRINE (DOSE UNKNOWN) BY THE EMT. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL WHERE HE WAS PRONOUNCED DECEASED. THE CORONER¿S REPORT DOES NOT LIST A CAUSE OF DEATH AND ONLY STATED THE TIME OF DEATH. THE CLINIC HAS NOT RECEIVED ANY PAPERWORK FROM THE HOSPITAL. NO FURTHER INFORMATION WAS AVAILABLE PERTAINING TO THE EVENT. THE CM STATED THAT THERE WERE NO ISSUES WITH THE 2008T MACHINE OR OTHER FRESENIUS PRODUCTS PRIOR TO OR DURING THE PATIENT¿S HD TREATMENT. THE CM AND CHARGE NURSE BOTH INSPECTED THE MACHINE AFTER THE EVENT AND TOOK A DIALYSATE SAMPLE. THERE WERE NO ABNORMAL FINDINGS. THE BIOMED TECHNICIAN COMPLETED AN INSPECTION OF THE MACHINE AND FOUND THE ULTRAFILTRATION (UF) PUMP CALIBRATION OUT OF RANGE. THE CM STATED IT WAS A SMALL DISCREPANCY AND WAS PULLING TOO LITTLE FLUID FROM THE PATIENT AS A RESULT. THE MACHINE REMAINS OUT OF SERVICE PENDING EVALUATION BY A FRESENIUS FIELD SERVICE TECHNICIAN. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN HEMODIALYSIS UTILIZING THE FRESENIUS 2008T MACHINE AND THE PATIENT EVENT OF CARDIAC ARREST WITH SUBSEQUENT DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE 2008T MACHINE. ADDITIONALLY, THERE IS NO ALLEGATION OF A MACHINE MALFUNCTION OR DEFICIENCY CAUSING THE PATIENT¿S ADVERSE EVENT. THE PATIENT WENT INTO CARDIAC ARREST EXHIBITED BY VENTRICULAR TACHYCARDIA (V TACH). ACCORDING TO THE UNITED STATES RENAL DATA SYSTEM (USRDS), THE LEADING CAUSE OF DEATH AMONG CKD PATIENTS UNDERGOING DIALYSIS IS RELATED TO CARDIAC ARRHYTHMIAS. SUDDEN CARDIAC DEATHS CAUSED BY VENTRICULAR TACHYARRHYTHMIAS ACCOUNT FOR APPROXIMATELY HALF OF THE DEATHS RELATED TO CARDIAC CAUSES. IT IS UNKNOWN IF THE PATIENT HAS A HISTORY OF CARDIAC COMORBIDITIES WHICH CAUSED OR CONTRIBUTED TO THE ADVERSE EVENT. IT IS ALSO UNKNOWN IF THE DIALYSIS PROCEDURE ITSELF (SHIFT IN ELECTROLYTES OR FLUID) CONTRIBUTED TO THE EVENT. ALTHOUGH THERE WAS A SLIGHT VARIANCE IN THE UF PUMP CALIBRATION BEING OUT OF SPECIFICATION BY 0.60ML THERE IS NO EVIDENCE THAT THIS SMALL DISCREPANCY IN FLUID REMOVAL CONTRIBUTED TO THE PATIENT EVENT. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OF A MALFUNCTION OR DEFICIENCY CAUSING OR CONTRIBUTING TO THE CARDIAC ARREST AND DEATH, THE FRESENIUS 2008T MACHINE CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
BIPAP AUTO BIFLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
BIPAP AUTO SERIES ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS MULTIPLE ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. IN THIS REPORT, IN BOX H METHOD CODE, RESULTS CODE AND CONCLUSION CODE HAS BEEN UPDATED/CORRECTED.
NEOTRACT UROLIFT 2 SYSTEM | K201837 | ON AUGUST 29, 2024, A PHYSICIAN REPORTED TO NEOTRACT THE FOLLOWING: A PATIENT HAD GONE TO THE EMERGENCY ROOM WITH LOWER ABDOMINAL PAIN THE EVENING AFTER HIS UROLIFT PROCEDURE ON (B)(6) 2024. THE PATIENT WAS ADMITTED TO EMERGENCY DEPARTMENT AFTER A CT SCAN SHOWED A SMALL PELVIC HEMATOMA. AN INFERIOR VENA CAVA (IVC) FILTER WAS PLACED BY INTERVENTIONAL RADIOLOGY ON (B)(6) 2024. EMERGENCY DEPARTMENT DISCHARGED THE PATIENT ON (B)(6) 2024. ON (B)(6) 2024, THE REPORTING PHYSICIAN TOLD NEOTRACT THAT THE PATIENT HAD DIED DUE TO A PULMONARY EMBOLISM (PE). THE FOLLOWING ADDITIONAL INFORMATION WAS RECEIVED BY NEOTRACT FROM THE REPORTING PHYSICIAN ON (B)(6) 2024: THE PATIENT HAD SUCCESSFULLY COMPLETED AN UNCOMPLICATED UROLIFT PROCEDURE ON (B)(6) 2024, WITH NO APPARENT COMPLICATIONS AND MINIMAL BLEEDING OBSERVED. AFTER COMPLETING THE UROLIFT, THE PHYSICIAN ALSO REMOVED KIDNEY STONES, AND A CATHETER WAS PLACED. THE PATIENT WAS 68 YEARS OLD WITH A PROSTATE SIZE OF 41CC. PROSTATE ANATOMY INCLUDED A HIGH MEDIAN BAR AND BULKY, ASYMMETRICAL TISSUE/LOBES THAT INDICATED THE USE OF AN ADVANCED TECHNIQUE CATHETER (ATC). THE PATIENT HAD BEEN HOSPITALIZED FROM (B)(6) 2024, FOR A PE AND HAD STARTED ELIQUIS TREATMENT, WHICH WAS STOPPED IN (B)(6) 2024 PRIOR TO THE UROLIFT PROCEDURE. ON SEPTEMBER 4, 2024, NEOTRACT WAS TOLD BY THE REPORTING PHYSICIAN THAT HE WAS UNCERTAIN ABOUT ANY CONNECTION BETWEEN THE UROLIFT PROCEDURE AND THE PATIENT¿S DEATH AND HE HAD NO WAY TO DETERMINE IF THE UROLIFT PROCEDURE CONTRIBUTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER, DEATH, LARGE B-CELL LYMPHOMA WITH HIGH GRADE FEATURES INVOLVING THE STOMACH, PROSTATE CANCER, AND NASAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON THE CYCLER. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT (MALE) EXPIRED PEACEFULLY AT HOME ON (B)(6) 2024 WHILE SLEEPING. THE PATIENT¿S SPOUSE DISCOVERED HER HUSBAND HAD EXPIRED SOMETIME DURING THE NIGHT. THE PATIENT¿S SPOUSE DECLINED HAVING AN AUTOPSY PERFORMED, AND THE PATIENT WAS TAKEN DIRECTLY TO THE FUNERAL HOME. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 SHOWED THE TREATMENT WAS COMPLETED DESPITE THE PATIENT¿S PASSING, AND NO ALARMS WERE NOTED. THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS A CARDIAC ARREST, SECONDARY TO MULTIPLE CARDIAC COMORBID CONDITIONS. THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER HAS NOT BEEN SCHEDULED FOR PICK-UP. THE PDRN CONFIRMED THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN HE EXPIRED; HOWEVER, THE PDRN REPORTED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). THE PDRN WAS UNABLE TO PROVIDE ANY PATIENT DEMOGRAPHICS, MEDICAL HISTORY, OR CCPD ORDERS DUE TO THE PATIENT¿S ELECTRONIC MEDICAL RECORD BEING CLOSED. ADDITIONALLY, THE PATIENT TRANSFERRED FROM AN UNKNOWN OUTPATIENT HOME DIALYSIS TO CLINIC TO (B)(6) ONLY A FEW DAYS PRIOR TO HIS EXPIRATION. | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE PATIENT¿S SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. PER THE PATIENT¿S PDRN, THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT¿S MULTIPLE CARDIAC COMORBID CONDITIONS. THE PDRN CONFIRMED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
EVIS EXERA III DUODENOVIDEOSCOPE | K220587 | IT WAS REPORTED THAT THERE WERE FOUR (4) POTENTIAL PATIENTS DIRECTLY IMPACTED BY ONE SUSPECTED MEDICAL DEVICE. THE PATIENTS UNDERWENT PROCEDURES INVOLVING THE SAME INDIVIDUAL SUSPECTED MEDICAL DEVICE BETWEEN (B)(6) 2024. THE PATIENT REFERENCED IN THIS REPORT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2024 FOR AN ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) PROCEDURE WITH SUSPECTED MEDICAL DEVICE FOR REMOVAL OF A STENT PREVIOUSLY PLACED IN (B)(6) 2023. THIS PATIENT HAD A PAST MEDICAL HISTORY OF UNTREATED STAGE IV LUNG CANCER, PANCREATIC ADENOCARCINOMA, AND GERO AND WAS ON PALLIATIVE CARES AT AN ASSISTED LIVING FACILITY FOR MULTIPLE MALIGNANCIES. THE PROCEDURE WAS UNABLE TO REPLACE STENTS AFTER 2 HOURS OF ATTEMPTS TO ADVANCE THE DUODENOSCOPE. THE PATIENT WAS DISCHARGED THE SAME DAY. THE PATIENT HAD AN INFORMATIONAL VISIT WITH HOSPICE ON (B)(6) 2024 WITH PLANS TO TRANSITION TO HOME HOSPICE. THE PATIENT WAS THEN ADMITTED ON (B)(6) 2024 FOR WORSENING ABDOMINAL PAIN AND UNDERWENT LNTERVENTIONAL RADIOLOGY (IR) CHOLANGIOGRAPHY AND PERMANENT STENTING OF A MALIGNANT STRICTURE OF THE COMMON BILE DUCT ON (B)(6) 2024. THE PATIENT HAD IMPROVING ABDOMINAL PAIN AND WAS DISCHARGED ON (B)(6) 2024. ON (B)(6) 2024, THE PATIENT WAS ADMITTED FOR ACUTE METABOLIC ENCEPHALOPATHY AND SEPTIC SHOCK. THE PATIENT HAD BEEN SET TO ADMIT TO HOSPICE AT ASSISTED LIVING FACILITY ON (B)(6) 2024, BUT PLANS CHANGED DUE TO HER RETURNING TO THE HOSPITAL. BLOOD CULTURE ON ADMISSION WAS POSITIVE FOR CARBAPENAMASE-PRODUCING CARBAPENEM RESISTANT E. COLI (CP-CRE) AND KLEBSIELLA PNEUMONIAE. THE FAMILY APPROVED NO ESCALATION OF CARES BEYOND INTRAVENOUS ANTIBIOTICS AND FLUIDS. THE PATIENT PASSED AWAY ON (B)(6) 2024. ON (B)(6) 2024, ERCP PROCEDURE INVOLVING SUSPECTED MEDICAL DEVICE WAS IDENTIFIED AS POSSIBLE LINK BETWEEN MULTIPLE CP-CRE INFECTIONS. ON (B)(6) 2024, SUSPECTED MEDICAL DEVICE WAS EXAMINED BY BORESCOPE DURING FACILITY INVESTIGATION. BORESCOPE REVEALED A SMALL TEAR WITHIN THE INNER SHEATH OF THE DEVICE. IT IS UNKNOWN WHEN THE TEAR APPEARED ON THE SCOPE, WHAT CAUSED IT, AND WHETHER IT WAS PRESENT FOR ERCP PROCEDURES INVOLVING AFFECTED PATIENTS. THE CUSTOMER FURTHER REPORTED THAT THE DEATH WAS POTENTIALLY CAUSED BY THIS SUBJECT DEVICE; HOWEVER, THE CUSTOMER WAS NOT SURE IF THIS DEVICE WAS THE MAIN CAUSE OF THE DEATH. THE DEVICE HAD AN INNER SHEATH TEAR WHICH THE CUSTOMER WAS NOT AWARE OF, AND THEY USED THE SCOPE ON SEVERAL PATIENTS. THE CUSTOMER FURTHER STATED THAT THE DEVICE MIGHT HAVE CAUSED A CROSS CONTAMINATION BETWEEN PATIENTS. REPORTEDLY, THE DEVICE HAS BEEN IN QUARANTINE IN A SECURE LOCATION. THE CUSTOMER CONFIRMED THAT THE DEVICE PASSED THE LEAK TEST PRIOR TO PROCEDURES BEING DONE. | THIS REPORT IS BEING SUPPLEMENTED TO PROVIDE ADDITIONAL INFORMATION BASED ON THE RESULTS OF THE LEGAL MANUFACTURER'S FINAL INVESTIGATION. A REVIEW OF THE DEVICE HISTORY RECORD FOUND NO DEVIATIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED ISSUE. THE FACILITY DID NOT PERFORM CULTURE TESTING OF THE DEVICE. ALSO, THE DEVICE WAS NOT RETURNED TO OLYMPUS FOR EVALUATION. BASED ON THE RESULTS OF THE INVESTIGATION, OLYMPUS CONFIRMED THAT THE FACILITY DID NOT DEVIATE FROM THE REPROCESSING STEPS OUTLINED IN THE INSTRUCTIONS FOR USE (IFU) MANUAL. THE RELATIONSHIP BETWEEN THE SUBJECT DEVICE AND THE REPORTED PATIENT DEATH COULD NOT BE IDENTIFIED. THEREFORE, THE ROOT CAUSE COULD NOT BE DETERMINED. THE FOLLOWING IS INCLUDED IN THE DEVICE IFU: ¿AN INSUFFICIENTLY REPROCESSED ENDOSCOPE AND/OR ACCESSORY MAY POSE AN INFECTION CONTROL RISK TO THE PATIENTS AND/OR OPERATORS WHO TOUCH THEM.¿ THIS SUPPLEMENTAL REPORT INCLUDES INFORMATION ADDED TO D8. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | K203365 | A SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED IN A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) CASE. IT WAS REPORTED THAT ACCESS WAS DIFFICULT AND THE IVL CATHETER WAS USED TO FACILITATE ACCESS AND DELIVERY OF THE TAVI SHEATH. THE TAVI SHEATH WAS PUT IN PLACE AFTER THE USE OF IVL THERAPY. A MAJOR DISSECTION OF UNKNOWN GRADE AND RUPTURE OF THE ILIAC AND AORTA OCCURRED IN WHICH THE PATIENT DID NOT SURVIVE. IT IS UNKNOWN EXACTLY WHEN THE DISSECTION AND DEATH OCCURRED. THERE WAS NO IVL DEVICE MALFUNCTION REPORTED. THERE IS NO AVAILABLE INFORMATION ON THE PATIENT, EVENT DATE, OR THE DEVICE RELATEDNESS. ATTEMPTS HAVE BEEN MADE TO OBTAIN ADDITIONAL INFORMATION. | THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. THE CAUSE OF THE DISSECTION, PERFORATION, AND SUBSEQUENT DEATH COULD NOT BE DEFINITIVELY DETERMINED BASED ON THE INFORMATION AVAILABLE. THERE WAS NO IVL DEVICE MALFUNCTION REPORTED. THE LOT NUMBER OF THE DEVICE WAS NOT PROVIDED. THEREFORE, REVIEW OF THE DEVICE MANUFACTURING AND TEST DOCUMENTATION COULD NOT BE COMPLETED. HOWEVER, SHOCKWAVE MEDICAL, INC. HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO DISTRIBUTION.
ION | K182188 | A PATIENT ENROLLED IN A REGISTRY STUDY UNDERWENT AN UNEVENTFUL ION BIOPSY PROCEDURE. THE PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS OR DEVICE MALFUNCTIONS. APPROXIMATELY TWO WEEKS AFTER THE PROCEDURE, THE PATIENT EXPERIENCED RESPIRATORY FAILURE THAT REQUIRED HOSPITALIZATION IN THE INTENSIVE CARE UNIT, AND PASSED AWAY 5 DAYS LATER (21 DAYS POST-PROCEDURE). THE CAUSE OF DEATH WAS DETERMINED BY THE PRINCIPLE INVESTIGATOR TO BE RESPIRATORY FAILURE FROM ACUTE PULMONARY EMBOLISM RELATED TO THE COMORBID CONDITION OF NON-SMALL CELL LUNG CANCER. THE EVENT WAS ASSESSED AS DEFINITELY RELATED TO THE PATIENT'S PRE-EXISTING CONDITION, BUT NOT RELATED TO THE ION DEVICES OR THE PROCEDURE. | REVIEW OF THE ION SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE AVAILABLE DATA IT IS POSSIBLE THE PROCEDURE MAY HAVE BEEN A CONTRIBUTING FACTOR TO THE VENOUS THROMBOEMBOLISM (VTE) EVENT AND THEREFORE THE PULMONARY EMBOLISM 2 WEEKS POST BIOPSY MAY HAVE BEEN PROCEDURE RELATED. THERE IS NO COMPELLING EVIDENCE THE EVENT WAS DEVICE RELATED. THE CAUSE OF DEATH WAS REPORTED TO BE DUE TO ACUTE RESPIRATORY FAILURE DUE TO PULMONARY EMBOLISM IN THE SETTING OF LUNG CANCER. THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES ASSOCIATED WITH THE REPORTED COMPLICATION. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 CASES REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. THERE WERE NO REPORTED EVENTS OF VENOUS THROMBOEMBOLISM (VTE) INCLUDING DEEP VENOUS THROMBOSIS OR PULMONARY EMBOLISM. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS AND NO VTE. HOWEVER, IT IS POSSIBLE THAT IMMOBILITY DURING GENERAL ANESTHESIA AND THE SMALL AMOUNT OF BLEEDING AND SUBSEQUENT CLOTTING ASSOCIATED WITH LUNG BIOPSIES MAY CONFER SOME DEGREE OF INCREASED RISK OF VTE EVENTS. LIKELY OF MORE RELEVANCE, IT IS ESTIMATED THAT CANCER CONFERS A 4-7X HIGHER RISK FOR VTE. THE HIGHER RATE OF VTE HAS BEEN REPORTED TO RANGE BETWEEN 3-13.9% IN PATIENTS WITH LUNG CANCER. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLOM JW, DOGGEN CJ, OSANTO S, ROSENDAAL FR. MALIGNANCIES, PROTHROMBOTIC MUTATIONS, AND THE RISK OF VENOUS THROMBOSIS. JAMA. 2005. BAGCHI A, KHAN MS, SARASWAT A, ANSARI A, NAI Q, IYER V, HAMOUDA D, KHUDER S, VERGHESE C. INCREASED INCIDENCE OF THROMBOTIC COMPLICATIONS WITH NON-SMALL CELL LUNG CANCER NECESSITATES CONSIDERATION OF PROPHYLACTIC ANTICOAGULATION IN YOUNG INDIVIDUALS. CUREUS. 2021. CHEW HK, DAVIES AM, WUN T, HARVEY D, ZHOU H, WHITE RH. THE INCIDENCE OF VENOUS THROMBOEMBOLISM AMONG PATIENTS WITH PRIMARY LUNG CANCER. J THROMB HAEMOST. 2008. CONNOLLY GC, DALAL M, LIN J, KHORANA AA. INCIDENCE AND PREDICTORS OF VENOUS THROMBOEMBOLISM (VTE) AMONG AMBULATORY PATIENTS WITH LUNG CANCER. LUNG CANCER. 2012.
ALARIS SYSTEM | K133532 | IT WAS INITIALLY REPORTED A FREE FLOW EVENT. NO FURTHER INFORMATION WAS PROVIDED. RECEIVED A COPY OF THE CUSTOMER'S MAUDE DATABASE REPORT FROM FDA WHICH STATES: "THE PATIENT WAS ADMITTED ON (B)(6) 2024 WITH ACUTE HYPOXIC RESPIRATORY FAILURE SECONDARY TO A CONGESTIVE HEART FAILURE EXACERBATION, NEW ONSET ATRIAL FIBRILLATION, AND CHEST PAIN. THE PATIENT WAS STARTED ON A HEPARIN DRIP IN THE EMERGENCY ROOM AND THE PATIENT RECEIVED 3 BAGS OF HEPARIN OVER 12 HOURS INSTEAD OF 26 HOURS. THE INTENDED INFUSION RATE WAS 9.3ML/HOUR. THE TOTAL INFUSION WAS 3 BAGS OF 250ML FOR A TOTAL OF 750ML. THE CONCENTRATION FOR THE INFUSION WAS 25,000 UNITS IN 5% DEXTROSE PER 250ML BAG. NO OTHER INFUSIONS WERE RUNNING. THE PUMP WAS PROGRAMMED VIA BAR CODE SCANNING. ON THE MORNING OF (B)(6) 2024, THE PATIENT WAS FOUND TO BE DIAPHORETIC, AGITATED, VOMITING, COMPLAINING OF HEADACHE AND "NOT FEELING GOOD". BLOOD PRESSURE AND HEART RATE WERE ELEVATED. LABORATORY RESULTS REVEALED HIGH ANTIXA LEVELS. THE PATIENT PROCEEDED TO HAVE EPISODES OF HEMATEMESIS. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) AND A CT (COMPUTED TOMOGRAPHY) OF THE HEAD REVEALED SIGNIFICANT BLEEDING IN MULTIPLE LOCATIONS WITH ACTIVE BLEEDING. NEUROSURGERY WAS CONSULTED AND THERE WAS NO INDICATION FOR SURGERY OR DECOMPRESSION, AND NO TREATMENT COULD BE OFFERED. THE PATIENT EXPERIENCED DECLINING MENTAL STATUS WITH AIRWAY COMPROMISE AND WAS INTUBATED. PHYSICIANS SPOKE WITH FAMILY AND UPDATED ON THE PATIENT'S DECLINE. A REPEAT HEAD CT WAS OBTAINED THAT REVEALED WORSENING BLEEDING IN THE BRAIN. PALLIATIVE CARE WAS CONSULTED, AND THE DECISION WAS MADE BY FAMILY TO WITHDRAW CARE. THE PATIENT WAS PRONOUNCED DEAD ON (B)(6) 2024 1725. POST-EVENT, A REVIEW OF THE PUMP DATA BY THE (B)(6) SYSTEM DIRECTOR DRUG POLICY AND MEDICATION SAFETY, FOUND THE CORRECT ADULT PROFILE WAS UTILIZED, THE CORRECT DRUG LIBRARY LINE ITEM WAS UTILIZED, THE CORRECT THERAPY NAME WAS USED, CORRECT UNITS WERE USED TO PROGRAM THE PUMP AND DOSES WERE PROGRAMMED WITHIN THE GUARDRAILS. IT WAS FURTHER NOTED THAT THE PUMP WAS PROGRAMMED WITHIN NORMAL RANGES AND IT WAS NOTED THAT THE RAPID ADMINISTRATION DOES NOT APPEAR TO BE THE RESULT OF A PUMP PROGRAMMING ERROR. NURSING STAFF REPORT NO PUMP ALARMS SOUNDED DURING INFUSIONS TO INDICATE PUMP WAS MALFUNCTIONING. ANNUAL PREVENTATIVE MAINTENANCE ASSESSMENT OF THE INVOLVED INFUSION PUMP MODULE WAS PERFORMED BY A GE BIOMED TECHNICIAN AT (B)(6) ON (B)(6) 2024. NO MAINTENANCE ISSUES WERE IDENTIFIED. ANNUAL PREVENTATIVE MAINTENANCE ASSESSMENT OF THE INVOLVED INFUSION PUMP BRAIN WAS PERFORMED BY A GE BIOMED TECHNICIAN AT (B)(6) ON (B)(6) 2024. NO MAINTENANCE ISSUES WERE IDENTIFIED. THE INFUSION PUMP TUBING (BD/ALARIS, LOT# 24055720) AND THREE EMPTY BAGS OF HEPARIN WERE NOT SEQUESTERED AND WERE DISCARDED. THE TEST WAS PERFORMED AT 1546 ON (B)(6) 2024 2243 ON (B)(6) 2024 0731 ON (B)(6) 2024 0952 ON (B)(6) 2024." MULTIPLE ATTEMPTS HAVE BEEN MADE REQUESTING THE CUSTOMER TO PROVIDE ADDITIONAL INFORMATION AND DEVICE RETURN FOR INVESTIGATION. | OMIT : MEDICAL DEVICE OTHER OPERATOR. ADDITIONAL INFORMATION : MANUFACTURER NARRATIVE. ROOT CAUSE : THE ROOT CAUSE FOR THE REPORTED ISSUE OF AN DEATH - FREE FLOW EVENT - HEPARIN WAS NOT DETERMINED BECAUSE NO PRODUCTS OR DEVICE LOGS WERE RETURNED.
SMART PORT CT | K101017 | AN END USER REPORTED AN ISSUE WITH A CT LOW PROFILE TITANIUM 6.6FR W/POLY CATH. DURING DEVICE REMOVAL, THE PORT WAS FOUND TO BE LEAKING CHEMOTHERAPY. IT WAS REPORTED THAT THE PATIENT LATER 'PASSED AWAY.' THERE WAS NO DIRECT ALLEGATION THAT THE PATIENT'S EXPIRATION WAS RELATED, IN ANY WAY, TO THE PORT MALFUNCTION. DESPITE NUMEROUS GOOD FAITH EFFORTS, ADDITIONAL INFORMATION HAS NOT BEEN PROVIDED. | THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PORT/CATHETER TUBING LEAKED IS CONFIRMED. INVESTIGATION OF THE RETURNED COMPLAINT SAMPLE SHOWED A SLICE/HOLE IN THE CATHETER TUBING APPROXIMATELY 1.5CM FROM THE PORT HOUSING. A DEVICE HISTORY RECORD REVIEW OF THE PACKAGING/ASSEMBLY/COMPONENT LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. ENGINEER EVALUATION: VISUAL INSPECTION OF SAMPLE: THE PORT WAS RECEIVED WITH THE LOCK AND APPROXIMATELY 15 CM OF CATHETER ATTACHED. THE SEPTUM HAD EVIDENCE OF MULTIPLE NEEDLE STICKS. NO VISUAL DEFECTS WERE NOTED TO THE SAMPLE. 40 PSI OF AIR WAS APPLIED TO THE PORT TO CHECK FOR LEAKS AND A SMALL SLICE IN THE CATHETER WAS FOUND BETWEEN THE 13 CM AND 14 CM MARK, APPROXIMATELY 1.5 CM FROM THE PORT BODY. THE DEFECT WAS VERY SMALL AND APPEARED TO BE A SLICE FROM A SHARP OBJECT. THE CATHETER WAS SECTIONED DISTAL TO THE SLICE AT APPROXIMATELY THE 12 CM MARK TO BE MEASURED. THE FOLLOWING CATHETER DIMENSIONS WERE VERIFIED PER 107482 REV. D: DIMENSION A (OD) - SAMPLE MET SPECIFICATION OF 0.085 ± 0.003." DIMENSION B (ID) - SAMPLE MET SPECIFICATION OF 0.056 ± 0.003." ROOT CAUSE/POTENTIAL ROOT CAUSE: THE CUSTOMER'S COMPLAINT DESCRIPTION OF A LEAK IS CONFIRMED TO BE CAUSED BY A SMALL SLICE IN THE CATHETER. NO MANUFACTURING DEFECTS WERE OBSERVED DURING SAMPLE REVIEW. THE MOST LIKELY ROOT CAUSE IS THAT THE SLICE OCCURRED DURING THE PORT REMOVAL PROCESS. THE APPEARANCE OF THE SLICE IS CONSISTENT WITH BEING CAUSED BY A SHARP OBJECT, SUCH AS A SCALPEL. LABELING REVIEW: THE INSTRUCTIONS FOR USE ITEM NUMBER: {16608102-01}, WHICH IS SUPPLIED TO THE USER WITH THIS ITEM NUMBER, CONTAINS THE FOLLOWING STATEMENTS: ABSENCE OF A BLOOD RETURN OR A POOR BLOOD RETURN CAN BE A SIGN OF A POTENTIAL COMPLICATION SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, FIBRIN FORMATION, THROMBOSIS OR MALPOSITION. THIS SHOULD BE EVALUATED PRIOR TO DEVICE USAGE. A BLOOD RETURN SHOULD BE PRESENT PRIOR TO USAGE OF DEVICE FOR ANY THERAPY OR TESTING. IF THE PATIENT COMPLAINS OF PAIN, OR IF THERE IS SWELLING WHEN THE DEVICE IS FLUSHED OR WHEN MEDICATION OR CONTRAST MEDIA IS ADMINISTERED, EVALUATE THE DEVICE FOR INFILTRATION, PROPER NEEDLE PLACEMENT, AND POTENTIAL COMPLICATIONS SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, THROMBOSIS OR MALPOSITION. FAILURE TO ASSESS THESE COMPLAINTS OR OBSERVATIONS CAN LEAD TO DEVICE FAILURE. CAUTION: AVOID PIERCING CATHETER WITH SUTURE NEEDLE. POTENTIAL COMPLICATIONS: CATHETER FRAGMENTATION AND CATHETER PINCH-OFF. CATHETER PLACEMENT CONSIDERATIONS: WARNING: AVOID MEDIAL CATHETER PLACEMENT INTO SUBCLAVIAN VEIN THROUGH PERCUTANEOUS TECHNIQUE. THIS PLACEMENT COULD LEAD TO CATHETER OCCLUSION, DAMAGE, RUPTURE, SHEARING, OR FRAGMENTATION DUE TO COMPRESSION OF THE CATHETER BETWEEN THE FIRST RIB AND CLAVICLE. CATHETER SHEARING HAS BEEN REPORTED WHEN THE CATHETER IS INSERTED VIA A MORE MEDIAL ROUTE IN THE SUBCLAVIAN VEIN. PINCH-OFF SYNDROME: PINCH-OFF SYNDROME SIGNS MAY INCLUDE DIFFICULTY IN ASPIRATING BLOOD, RESISTANCE TO FLUSHING OR INFUSION OF MEDICATIONS OR FLUIDS THAT IMPROVES WITH POSITION CHANGES, INFRACLAVICULAR PAIN AND/OR SWELLING WITH CATHETER FLUSHING OR INFUSION PALPITATIONS, SUDDEN ONSET CHEST PAIN, CARDIAC ARRHYTHMIAS, EXTRA HEART SOUND, CHEST WALL SWELLING AT THE PORT POCKET, VEIN INSERTION SITE, PAIN IN SHOULDER OR PORT AREA NOT ASSOCIATED WITH SWELLING, COUGH, PARESTHESIA OF ARM ON SIDE OF CATHETER WITHDRAWAL OCCLUSION OR SWISHING SOUND WITH CATHETER FLUSHING. WARNING: AVOID MEDIAL CATHETER PLACEMENT INTO SUBCLAVIAN VEIN THROUGH PERCUTANEOUS TECHNIQUE. THIS PLACEMENT COULD LEAD TO CATHETER OCCLUSION, DAMAGE, RUPTURE, SHEARING, OR FRAGMENTATION DUE TO COMPRESSION OF THE CATHETER BETWEEN THE FIRST RIB AND CLAVICLE. CATHETER SHEARING HAS BEEN REPORTED WHEN THE CATHETER IS INSERTED VIA A MORE MEDIAL ROUTE IN THE SUBCLAVIAN VEIN. NOTE: IF INFUSION OR ASPIRATION IS SUCCESSFUL UPON LIFTING ARM ABOVE THE HEAD AND TURNING THE HEAD, CONSIDER PINCH-OFF SYNDROME AS A POSSIBLE CAUSE. THE LINE SHOULD BE RADIOLOGICALLY EVALUATED IF PINCH-OFF SYNDROME IS SUSPECTED. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED FOR TRENDS. REFERENCE: (B)(4).
SOLITAIRE FR4 | K183022 | ADDITIONAL INFORMATION RECEIVED REPORTED THAT AFTER THE THROMBECTOMY, THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL ON (B)(6) 2024. THE PATIENT HAD INHALATION PNEUMOPATHIA THAT IMPROVED UNDER DUAL ANTIBIOTHERAPY, BUT THE PATIENT'S RESPIRATORY CONDITION WORSENED. AFTER 24 HOURS OF NON-INVASIVE VENTILATION, IT WAS NOT POSSIBLE TO OBTAIN HEMATOSIS. THE PATIENT CONDITION CONTINUED TO WORSEN. DEEP SEDATION WAS INITIATED, AND THE PATIENT WAS DECEASED ON (B)(6) 2024. THE SITE'S CAUSALITY ASSESSMENT OF RESPIRATORY INFECTION AND HEMORRHAGIC TRANSFORMATION STROKE PH2 WAS NOT RELATED TO PROCEDURE, DEVICES AND UNDERLYING CONDITION OR DISEASE, CAUSAL RELATIONSHIP TO DISEASE UNDER STUDY; THE RATIONALE FOR THE RELATIONSHIP TO SYSTEM OR PROCEDURE: NO RELATIONSHIP. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REACT CATHETER | K182097 | ADDITIONAL INFORMATION RECEIVED REPORTED THAT AFTER THE THROMBECTOMY, THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL ON (B)(6) 2024. THE PATIENT HAD INHALATION PNEUMOPATHIA THAT IMPROVED UNDER DUAL ANTIBIOTHERAPY, BUT THE PATIENT'S RESPIRATORY CONDITION WORSENED. AFTER 24 HOURS OF NON-INVASIVE VENTILATION, IT WAS NOT POSSIBLE TO OBTAIN HEMATOSIS. THE PATIENT CONDITION CONTINUED TO WORSEN. DEEP SEDATION WAS INITIATED, AND THE PATIENT WAS DECEASED ON (B)(6) 2024. THE SITE'S CAUSALITY ASSESSMENT OF RESPIRATORY INFECTION AND HEMORRHAGIC TRANSFORMATION STROKE PH2 WAS NOT RELATED TO PROCEDURE, DEVICES AND UNDERLYING CONDITION OR DISEASE, CAUSAL RELATIONSHIP TO DISEASE UNDER STUDY; THE RATIONALE FOR THE RELATIONSHIP TO SYSTEM OR PROCEDURE: NO RELATIONSHIP. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K122899 | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME POST A PORT PLACEMENT, THE PATIENT ALLEGEDLY DEVELOPED AN INFECTION. REPORTEDLY, THE PATIENT WAS EXPIRED AND THE PRIMARY CAUSE OF DEATH WAS NOT PROVIDED, AND THE RELATIONSHIP OF DEATH WITH THE DEVICE WAS UNKNOWN. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
RESUSCITAIRE RADIANT WARMER | K120642 | THE CUSTOMER REPORTED THAT A PREMATURE INFANT REQUIRED IMMEDIATE MEDICAL INTERVENTION USING A RESUSCITAIRE. REPORTEDLY, BOTH THE MATERNITY AND NEONATAL TEAMS HAD CHECKED THE EQUIPMENT BEFORE USE, CONFIRMING THAT IT WAS FUNCTIONING PROPERLY. DURING THE PROCESS OF GIVING BREATHS AND SUCTION TO THE NEWBORN, THE RESUSITAIRE SUDDENLY STOPPED WORKING. THE MATERNITY TEAM REPORTED THAT EITHER THE OXYGEN OR AIR CYLINDER HAD NOT BEEN TURNED ON. MEANWHILE, AND THE NEONATAL TEAM REPORTED THAT ONE OF THE CYLINDERS WAS EMPTY. THE INFANT WAS STABILIZED AND BROUGHT TO THE NEONATAL INTENSIVE CARE UNIT. THE PATIENT'S CONDITION DETERIORATED OVERNIGHT, AND THE PATIENT PASSED AWAY. | A COMPLAINT WAS RECEIVED REGARDING THE RESUSCITAIRE DURING USE OF SUCTION TO THE NEWBORN, SUDDENLY STOPPED WORKING. THE SITE INVESTIGATION IDENTIFIED THE CAUSE AS EITHER THE OXYGEN OR AIR CYLINDER WAS NOT TURNED ON, OR ONE OF THE CYLINDERS (OXYGEN OR AIR) WAS EMPTY. FURTHER INFORMATION WAS REQUESTED BUT NOT PROVIDED. BASED ON THE INFORMATION PROVIDED, THERE WAS NO DEVICE MALFUNCTION. THE RESUSCITAIRE INFORMATION FOR USE HAS A WARNING: IF EITHER THE AIR OR OXYGEN GAS SOURCE PRESSURE IS REDUCED, CREATING A PRESSURE DIFFERENTIAL OF 207 KPA (30 PSI), THE MICROBLENDER ALARM SOUNDS. THIS CONDITION CAN ALTER THE FIO2 AND FLOW OUTPUT FROM THE MICROBLENDER. PATIENT INJURY COULD OCCUR.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. THE PATIENT CONTACT REQUESTED THAT THE SUPPLIES BE PICKED UP. IN ADDITIONAL FOLLOW-UP, IT WAS CONFIRMED WITH THE PATIENT¿S PD NURSE THAT ON (B)(6) 2024 THE PATIENT EXPIRED WHILE CONNECTED TO THE FRESENIUS CYCLER DURING AN UNKNOWN PHASE OF PD TREATMENT. THE NURSE REPORTED THAT THERE WERE NO ISSUES WITH THE FRESENIUS CYCLER IN RELATION TO THE DEATH. THE NURSE PROVIDED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS THAT INCLUDED END STAGE RENAL DISEASE (ESRD), CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), OBESITY, AND HYPERTENSION (HTN). THE CAUSE OF THE DEATH WAS REPORTED AS UNKNOWN AND THE ESRD DEATH CERTIFICATE WAS NOT AVAILABLE. THE PATIENT HAS BEEN DISCHARGED FROM THE CLINIC SYSTEM AND NO OTHER INFORMATION WAS AVAILABLE. THE FRESENIUS CYCLER HAS BEEN RETURNED TO MANUFACTURER (VIA ASSOCIATED CLINIC) AT THE TIME FOLLOW-UP WAS COMPLETED. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A (AS RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED ON THE CYCLER AND PASSED. THE CYCLER UNDERWENT AND PASSED A TOUCHSCREEN TEST, VOLTAGE VERIFICATION CHECK, TEACH PUMPS, SYSTEM AIR LEAK TEST, VALVE ACTUATION TEST, AND PATIENT SENSOR CHECK. AN INVESTIGATION OF THE CYCLER MUSHROOM HEADS VERIFIED THAT THE SURFACE CONDITIONS AND ALIGNMENTS WERE WITHIN SPECIFICATION. AN INTERNAL VISUAL INSPECTION IDENTIFIED EVIDENCE OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE DRIED FLUID COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DREAMSTATION HUMIDIFIER, GB-RENT | K131982 | THE CUSTOMER REPORTED THE PATIENT PASSED AWAY AND THE DREAMSTATION HUMIDIFIER DEVICE WILL BE RETURNED. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
INOGEN ONE G5 PORTABLE OXYGEN CONCENTRATOR | K032818 | THE PATIENT'S WIFE CALLED TO REPORT ISSUE WITH UNIT REGARDING IT NOT BEING ABLE TO HOLD A CHARGE FOR LONG AND IS PUTTING OUT LOW AMOUNT OF O2, WHICH RESULTS IN PT'S O2 LEVELS DROPPING. ACCORDING TO THE PATIENT'S WIFE, THEY WENT TO THE PATIENTS FOLLOW-UP APPOINTMENT WITH THE SURGEON AFTER HAVING SOME TOES AMPUTATED DUE TO DIABETES WHEN THE PATIENT BECAME SHORT OF BREATH. THE DOCTOR SAID THE POC WAS NOT WORKING, THE PATIENTS O2 LEVEL WAS 59% AND TO TAKE HIM TO THE HOSPITAL ACROSS THE STREET. THEY WENT TO ER, CALLED INOGEN, AND HE WAS ADMITTED DUE TO LOW OXYGEN. HE WAS ADMITTED FROM THE ER TO (B)(6) HOSPITAL ON (B)(6) 2024 WHERE THE PATIENT LATER DIED ON (B)(6) 2024. | CONTAMINATED COLUMN AND CRACKED SENSOR BLOCK CANNULA RECEPTACLE WERE THE CAUSE OF LOW EXTERNAL O2 22.8%. THE DIAGNOSTIC SCREEN SHOWED HIGH PRESSURE AND LOW SIEVE LIFE DUE TO CONTAMINATED COLUMN.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS CONFIRMED THE PATIENT HAD DESATURATED AND THE NURSES WERE NOT AWARE OF THE CHANGE IN CONDITION. THE PATIENT WAS EMERGENTLY INTUBATED AND TRANSFERRED TO CVICU WHERE SHE LATER PASSED AWAY. IT WAS INDICATED THE ALARM WAS NOT AUDIBLE AT THE BEDSIDE OR CENTRAL STATION. THE CUSTOMER WAS ALSO NOT SURE WHETHER THE ALARMS WERE SET AND FAILED OR NOT PROPERLY SET UP. WHEN STAFF REVIEWED THE WAVEFORMS, IT WAS NOTED THE SPO2 HAD BEEN TRENDING DOWNWARD FOR APPROXIMATELY 10 MINUTES PRIOR TO THE BRADYCARDIA ALARM. A PRELIMINARY REVIEW OF THE CLINICAL AUDIT LOGS CONFIRMED THE LOW HEART RATE AND BRADYCARDIA ALARM WAS GENERATED AT THE TIME OF THE EVENT BUT NO SPO2 ALARMS WERE NOTED AT THAT TIME. | INITIAL REVIEW OF THE LOGS DETERMINED THAT THERE WERE TECHNICAL INOPS FROM 11:05 UNTIL 11:55 BUT THE LOGS ARE STILL UNDER FURTHER REVIEW. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, DEATH, OTHER, SWELLING IN ABDOMEN, INSOMNIA AND SEIZURES. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. WE ARE SENDING A FINAL REPORT SINCE NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME.
ENDO GIA | K111825 | ACCORDING TO THE REPORTER, POST-OPERATIVE OF A DUODENAL RESECTION, THE STAPLE LINE LEAKED. THE PATIENT HAD FISTULA AND ABDOMINAL INFECTION. THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU), EXTENDING THE PATIENT'S HOSPITALIZATION. THREE DAYS FOLLOWING THE INITIAL PROCEDURE, THE PATIENT UNDERWENT A SECOND PROCEDURE TO RESECT AND RE-STAPLE THE TISSUE. THE PATIENT DIED ON THE SAME DAY DUE TO DUODENAL PERFORATION. | ADDITIONAL INFORMATION: B5, D9, G3, H3, H6 H3 EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. THE DEVICE WAS AVAILABLE FOR EVALUATION. THE EVALUATION FOUND NO POTENTIALLY CONTRIBUTING FACTORS, AND THE SAMPLE MET ALL RELATED SPECIFICATIONS. IT WAS REPORTED THAT THE STAPLE LINE CONTENT HAS LEAK. THE REPORTED ISSUE COULD NOT BE CONFIRMED. THE MOST LIKELY CAUSE COULD NOT BE IDENTIFIED BECAUSE NO RELATED PROBLEM WAS DETECTED WITH THE DEVICE. THE MANUFACTURING RECORDS FOR EACH DEVICE ARE THOROUGHLY REVIEWED PRIOR TO RELEASE TO ENSURE THAT IT MEETS ALL MEDTRONIC QUALITY SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, KIDNEY DISEASE/TOXICITY, CANCER, LUNG AND BREAST CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, DEATH, OTHER ACUTE PULMONARY FIBROSIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN SECTION H- TYPE OF REPORTED COMPLAINT AS ONLY ADVERSE EVENT. AFTER REVIEW, IT WAS DETERMINED THAT SECTION H- TYPE OF REPORTED COMPLAINT SHOULD BE FILED AS DEATH. SECTION H- TYPE OF REPORTED COMPLAINT CORRECTED IN THIS REPORT.
BIPAP AUTOSV ADV | K090248 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
2008T MACHINE | K150708 | IT WAS REPORTED TO FRESENIUS THAT THIS UNKNOWN PATIENT WITH RENAL FAILURE (RF) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING HD THERAPY (DATE NOT PROVIDED). NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., DISCHARGE SUMMARY, DEATH CERTIFICATE, TREATMENT RECORD, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | PLANT INVESTIGATION: THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION AND THE SERIAL NUMBER COULD NOT BE OBTAINED. AS A SERIAL NUMBER COULD NOT BE DETERMINED, DEVICE HISTORY AND MANUFACTURING RECORDS COULD NOT BE REVIEWED. AS A PHYSICAL EVALUATION COULD NOT BE PERFORMED, A DEFINITIVE CONCLUSION REGARDING THE REPORTED INCIDENT COULD NOT BE REACHED AND A CAUSE COULD NOT BE CONFIRMED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A REGISTERED NURSE (RN) REPORTED TO FRESENIUS THAT THIS UNKNOWN PERITONEAL DIALYSIS (PD) PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY EXPIRED ON THE PRIOR NIGHT WHILE CONNECTED TO A HOSPITAL-PROVIDED LIBERTY SELECT CYCLER. IT WAS STATED THE PATIENT¿S DEATH WAS NOT CYCLER-RELATED AND THE CALL TO FRESENIUS WAS INTENDED FOR CYCLER PROCESSING FOLLOWING THE PATIENT DEATH. NO FURTHER INFORMATION WAS PROVIDED. THERE WAS NO SPECIFIC ALLEGATION THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. MULTIPLE ATTEMPTS WERE MADE TO ACQUIRE THE PATIENT¿S IDENTITY AND DETAILS CONCERNING THIS REPORTED EVENT; HOWEVER, NO ADDITIONAL INFORMATION WAS OBTAINED. AS A RESULT, PATIENT DEMOGRAPHICS AND CAUSE OF DEATH REMAIN UNKNOWN. | PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | UDI NUMBER IS UPDATED IN THIS REPORT.
BIPAP AUTOSV ADVANCED SYSTEM ONE | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
BIPAP PRO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THERE WAS REPORT OF PATIENT DEATH. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR | K130077 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE, TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX + | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX + | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, SKIN IRRITATION, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX D: COMMON DEVICE NAME WAS CORRECTED.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
ACHIEVE MAPPING CATHETER - 20 MM | K102588 | IT WAS REPORTED THAT DURING A CRYO ABLATION PROCEDURE WHEN THE FIRST ABLATION WAS PERFORMED, THE PATIENT'S BLOOD PRESSURE DECREASED. THE PROCEDURE WAS DISCONTINUED. A SYMPATHOMIMETIC VASOCONSTRICTOR MEDICATION WAS ADMINISTERED BUT THE BLOOD PRESSURE DID NOT RECOVER. THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT) AND DIRECT CURRENT WAS PERFORMED. THE VENTRICULAR TACHYCARDIA (VT) DID NOT RESOLVE, AND THE PATIENT WAS STARTED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT DID NOT RECOVER FROM CARDIAC ARREST AND ECMO WAS LATER STOPPED. THE PATIENT DIED. NO PERICARDIAL EFFUSION WAS OBSERVED DURING THE PROCEDURE AND A POST-MORTEM COMPUTERIZED TOMOGRAPHY (CT) SCAN REVEALED THE PATIENT HAD AN AORTIC DISSECTION AND CALCIFIED LESIONS IN THE AORTA. THE DISSECTION WAS NOTED TO BE AWAY FROM THE TRANSSEPTAL AND ABLATION SITES. NO AUTOPSY WAS PERFORMED. | CONTINUATION OF D10: PRODUCT ID: 4FC12 (0012228207); PRODUCT TYPE: 0629-FLEXCATH SHEATH; PRODUCT ID: AFAPRO28 (19288); PRODUCT TYPE: 0624-ARCTIC FRONT CATHETER; PRODUCT ID: EP003994S (UNKNOWN LOT); PRODUCT TYPE: TRANSSEPTAL NEEDLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
FLEXCATH ADVANCE STEERABLE SHEATH | K123591 | IT WAS REPORTED THAT DURING A CRYOABLATION PROCEDURE WHEN THE FIRST ABLATION WAS PERFORMED, THE PATIENT'S BLOOD PRESSURE DECREASED. THE PROCEDURE WAS DISCONTINUED. A SYMPATHOMIMETIC VASOCONSTRICTOR MEDICATION WAS ADMINISTERED BUT THE BLOOD PRESSURE DID NOT RECOVER. THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT) AND DIRECT CURRENT WAS PERFORMED. THE VENTRICULAR TACHYCARDIA (VT) DID NOT RESOLVE, AND THE PATIENT WAS STARTED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO). THE PATIENT DID NOT RECOVER FROM CARDIAC ARREST AND ECMO WAS LATER STOPPED. THE PATIENT DIED. NO PERICARDIAL EFFUSION WAS OBSERVED DURING THE PROCEDURE AND A POST-MORTEM COMPUTERIZED TOMOGRAPHY (CT) SCAN REVEALED THE PATIENT HAD AN AORTIC DISSECTION AND CALCIFIED LESIONS IN THE AORTA. THE DISSECTION WAS NOTED TO BE AWAY FROM THE TRANSSEPTAL AND ABLATION SITES. NO AUTOPSY WAS PERFORMED. | CONTINUATION OF D10: PRODUCT ID 990063-020 (228826709); PRODUCT TYPE: 0623-ACHIEVE CATHETER; PRODUCT ID AFAPRO28 (19288); PRODUCT TYPE: 0624-ARCTIC FRONT CATHETER; PRODUCT ID EP003994S (UNKNOWN LOT); PRODUCT TYPE: TRANSSEPTAL NEEDLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI XI | K131861 | A REVIEW OF A LITERATURE ARTICLE DESCRIBING A RETROSPECTIVE ANALYSIS OF ALL PATIENTS UNDERGOING DA VINCI ASSISTED ROBOTIC DISTAL PANCREATECTOMY (DP) AND PANCREATICODUODENECTOMY (PD) PROCEDURES PERFORMED BY A SINGLE SURGEON, WAS COMPLETED. THE STUDY EVALUATED THE SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREATECTOMY PROGRAM. THE STUDY INCLUDED SURGERIES FROM MAY 2014 TO DECEMBER 2020, DURING WHICH 62 PATIENTS UNDERWENT ROBOTIC PANCREATECTOMY, 34 PATIENTS WERE IN THE PD GROUP AND 28 PATIENTS WERE IN DP GROUP. ONE PATIENT WHO UNDERWENT PD EXPERIENCED A CARDIAC ARREST FOLLOWING A PULMONARY EMBOLISM ON POST-OPERATIVE DAY 5 AND COULD NOT BE REVIVED. THERE WERE NO DA VINCI DEVICE ISSUES REPORTED IN THE ARTICLE. ADDITIONAL INFORMATION WAS REQUESTED FROM THE DESIGNATED AUTHOR; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | THIS MEDWATCH REPORT REFLECTS ONE OF TWO DEATHS THAT OCCURRED, BOTH IN THE PD GROUP. SEE RELATED REPORT NUMBER 1 FOR THE OTHER MORTALITY IDENTIFIED IN THE ARTICLE. NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR THE DA VINCI SYSTEM IDENTIFIERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. A REVIEW OF THE ARTICLE PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF A PULMONARY EMBOLUS. NO ALLEGATION OF ANY ISSUE WITH ANY INTUITIVE SURGICAL PRODUCTS WAS MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. CITATION: MCKAY B, BROUGH D, KILBURN D, CAVALLUCCI D. SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREAS PROGRAM IN THE AUSTRALIAN SETTING: A CASE SERIES AND NARRATIVE REVIEW. ANZ J SURG. 2024 JUL-AUG;94(7-8):1247-1253. DOI: 10.1111/ANS.18998. EPUB 2024 MAR 26. PMID: 38529778. SECTION A - PATIENT INFORMATION - NO SPECIFIC PATIENT DEMOGRAPHICS WERE PROVIDED FOR THIS PATIENT. STUDY DEMOGRAPHICS IN THE PD GROUP INCLUDED 34 PATIENTS WITH A MEDIAN AGE OF 68 YEARS (RANGE 42-84); THE GENDER IS CITED AS 13 FEMALE, PRESUME 21 MALE. SECTION D - SUSPECT MEDICAL DEVICE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE DA VINCI SYSTEM ASSOCIATED WITH THE ADVERSE EVENT, A GENERIC SYSTEM MATERIAL NUMBER WAS USED. BLANK MDR FIELDS: THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELDS G5 AND G7 ARE NOT APPLICABLE.
DAVINCI XI | K131861 | A STUDY TO EVALUATE THE SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREATECTOMY PROGRAM, WAS SUMMARIZED IN A LITERATURE ARTICLE DESCRIBING A RETROSPECTIVE ANALYSIS OF ALL PATIENTS UNDERGOING DA VINCI ASSISTED ROBOTIC DISTAL PANCREATECTOMY (DP) AND PANCREATICODUODENECTOMY (PD) PROCEDURES PERFORMED BY A SINGLE SURGEON. THE STUDY INCLUDED SURGERIES FROM MAY 2014 TO DECEMBER 2020, DURING WHICH 62 PATIENTS UNDERWENT ROBOTIC PANCREATECTOMY, 34 PATIENTS WERE IN THE PD GROUP AND 28 PATIENTS WERE IN DP GROUP. ONE PATIENT WHO UNDERWENT PD EXPERIENCED GRADE C POST-OPERATIVE PANCREATIC FISTULA AND GASTRODUODENAL ARTERY HEMORRHAGE. THE PATIENT DIED FROM MULTIORGAN FAILURE ON POST-OPERATIVE DAY 18. THERE WERE NO DA VINCI DEVICE ISSUES REPORTED IN THE ARTICLE. ADDITIONAL INFORMATION WAS REQUESTED FROM THE DESIGNATED AUTHOR; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | THIS MEDWATCH REPORT REFLECTS ONE OF TWO DEATHS THAT OCCURRED, BOTH IN THE PD GROUP. SEE RELATED REPORT NUMBER 1 FOR THE OTHER MORTALITY IDENTIFIED IN THE ARTICLE. NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR THE DA VINCI SYSTEM IDENTIFIERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. A REVIEW OF THE ARTICLE PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD DIED AS A RESULT OF PANCREATIC FISTULA. THE DETAILS OF HOW THIS MAY HAVE OCCURRED ARE NOT PROVIDED. NO ALLEGATION OF ANY ISSUE WITH ANY INTUITIVE SURGICAL PRODUCTS WAS MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. CITATION: MCKAY B, BROUGH D, KILBURN D, CAVALLUCCI D. SAFETY AND FEASIBILITY OF INSTITUTING A ROBOTIC PANCREAS PROGRAM IN THE AUSTRALIAN SETTING: A CASE SERIES AND NARRATIVE REVIEW. ANZ J SURG. 2024 JUL-AUG;94(7-8):1247-1253. DOI: 10.1111/ANS.18998. EPUB 2024 MAR 26. PMID: 38529778. SECTION A - PATIENT INFORMATION - NO SPECIFIC PATIENT DEMOGRAPHICS WERE PROVIDED FOR THIS PATIENT. STUDY DEMOGRAPHICS IN THE PD GROUP INCLUDED 34 PATIENTS WITH A MEDIAN AGE OF 68 YEARS (RANGE 42-84); THE GENDER IS CITED AS 13 FEMALE, PRESUME 21 MALE. SECTION D - SUSPECT MEDICAL DEVICE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE DA VINCI SYSTEM ASSOCIATED WITH THE ADVERSE EVENT, A GENERIC SYSTEM MATERIAL NUMBER WAS USED.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY AND THE EQUIPMENT WAS BEING RETURNED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO THE SERVICE CENTER FOR INVESTIGATION. THE DEVICE WAS VISUALLY INSPECTED. NO FAULTS WERE REPORTED. THERE WERE 632 BLOWER HOURS WITH NO LOGGED ERRORS. THE USER INTERFACE PANEL MAIN AND ALARM RIBBON CABLE HAD DAMAGE. THE USER INTERFACE PANEL WAS REPLACED. THE FINE AND POLLEN FILTERS WERE REPLACED. THE DEVICE WAS CLEANED AND PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT A FEW DAYS AFTER A DA VINCI-ASSISTED SIGMOID COLECTOMY PROCEDURE, THE PATIENT EXPIRED. THE DISTRIBUTOR CLINICAL SUPPORT SPECIALIST WAS INFORMED THAT THE PATIENT HAD AN UNSPECIFIED SERIOUS COMPLICATION AFTER THE PROCEDURE AND WAS TOLD THAT THE SURGEON DOES NOT THINK THAT A DA VINCI SYSTEM, INSTRUMENT, AND/OR ACCESSORY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. NO SPECIFIC INFORMATION WAS AVAILABLE. ADDITIONAL INFORMATION WAS REQUESTED BUT HAS NOT BEEN PROVIDED. | NO INTUITIVE PRODUCT WILL BE RETURNED FOR INVESTIGATION, AND NO PHOTOS OR VIDEO OF THE PROCEDURE ARE AVAILABLE. THE DA VINCI SYSTEM SERIAL NUMBER WAS PROVIDED AND REVIEW OF THE DEVICE HISTORY RECORD FOUND NO NON-CONFORMANCES WERE IDENTIFIED THAT WOULD BE RELATED TO AN ADVERSE EVENT. THE PROCEDURE DATE WAS REQUESTED BUT WAS NOT PROVIDED. REVIEW OF THE AVAILABLE SURGEON¿S PROCEDURE HISTORY FOUND THAT MULTIPLE SIGMOID COLECTOMY PROCEDURES HAVE BEEN PERFORMED BY THIS SURGEON. THUS, THE SPECIFIC PROCEDURE FOR THE REPORTED EVENT COULD NOT BE DETERMINED, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED A FEW DAYS AFTER A SIGMOID COLECTOMY. NO INFORMATION IS PROVIDED REGARDING THE DETAILS OF THE PROCEDURE, POST OPERATIVE COURSE, OR THE CIRCUMSTANCES LEADING TO THE DEATH. NO CAUSE OF DEATH IS PROVIDED. BASED ON THE LIMITED DETAILS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING. ADDITIONALLY, PATIENT HAS PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
INTELLIVUE MX800 PATIENT MONITOR | K150310 | IT WAS REPORTED THE PATIENT PASSED AWAY AND THERE WERE GAPS IN THE DATA NOTED UPON RETROSPECTIVE REVIEW. THE CUSTOMER REQUESTED ASSISTANCE AFTER THE PATIENT'S DEATH IN DETERMINING THE REASON FOR MISSING DATA AND WAVEFORMS. IT WAS UNKNOWN WHETHER THIS ISSUE WAS RELATED TO THE PATIENT DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) EVALUATED THE DEVICE LOGS AND CONFIRMED THAT THE DEVICE WAS WORKING AS DESIGNED. THE MONITORING SYSTEM AND THE CENTRAL STATION CORRECTLY ALARMED WITH ¿ECG LEADS OFF¿ INOP (¿EKG ELEKTRDN AB¿) ON (B)(6) 2024 FOR BED. DUE TO THE INOP CONDITION, THE MONITORING SYSTEM WAS NOT ABLE TO MEASURE AN ECG SIGNAL OR PROVIDE A PHYSIOLOGICAL ALARM, THUS THE GAPS IN THE ECG SIGNAL. INOPS ARE TECHNICAL ALARMS, THEY INDICATE THAT THE MONITOR CANNOT MEASURE OR DETECT ALARM CONDITIONS RELIABLY. IF AN INOP INTERRUPTS MONITORING AND ALARM DETECTION (FOR EXAMPLE, LEADS OFF), THE MONITOR PLACES A QUESTION MARK IN PLACE OF THE MEASUREMENT NUMERIC AND AN AUDIBLE INDICATOR TONE WILL BE SOUNDED. OTHER INOPS THAT MIGHT HAVE BEEN PRESENT AT THE SAME TIME AT THE PATIENT MONITOR AND THE CENTRAL STATION ARE NOT LOGGED BY THE CENTRAL STATION VERSION USED AT THIS CUSTOMER SITE. THE ABOVE INFORMATION SUPPORTS THAT THE MONITORING SYSTEM AND THE CENTRAL STATION WORKED ACCORDING TO SPECIFICATION. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM COULD NOT BE CONFIRMED, THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. E1; REPORTER INSTITUTION PHONE NUMBER (B)(6). E1: REPORTER PHONE NUMBER (B)(6).
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ABOUT AN ALLEGATION OF A PATIENT DEATH. IN ADDITION, THERE WAS ALSO ALLEGATIONS OF LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
BIVONA TRACHEOSTOMY TUBE, NEONATAL/PEDIATRIC | K914088 | ICU MEDICAL RECEIVED A SUMMONS ON 10 SEPTEMBER 2024 FROM THE (B)(6) FIRM REGARDING THE DEATH OF A 12 MONTH OLD CHILD WHO WAS DEPENDENT ON SUPPLEMENTAL OXYGEN. THE COMPLAINT IN THE LAWSUIT ALLEGES THE FOLLOWING. THE PATIENT HAD A TRACHEAL STOMA IN PLACE AND REQUIRED HOME HEALTH SERVICES TO MONITOR THEIR RESPIRATORY CARE. THE ORDER WAS FOR A ¿3.5 BIVONA FLEXI 53MM WITH 3ML NS INFLATES CUFF Q.¿ AT 1PM ON (B)(6) 2023, THE PATIENT BEGAN TO COUGH AND THEN VOMIT. THE NURSE ATTEMPTED TO CHANGE THE TRACH TUBE AND INSERT WHAT IS DESCRIBED AS A 55 MM AND A SMALLER 45 MM TRACH WITH NO SUCCESS. THE PATIENT CRIED AND SHOWED SIGNS OF RESPIRATORY DISTRESS. THE MOTHER ATTEMPTED TO INSERT THE TRACH WITHOUT SUCCESS AND ATTEMPTED TO BAG THE PATIENT TO PROVIDE EMERGENCY AIRWAY SUPPORT. EMERGENCY MEDICAL SERVICES ARRIVED AT APPROXIMATELY 1:40 PM AND PROVIDED CARDIOPULMONARY RESUSCITATION. THE PATIENT WAS TRANSPORTED TO (B)(6) HOSPITAL AND WAS ADMITTED TO THE PEDIATRIC INTENSIVE CARE UNIT. THE REPORTER STATES THAT THE RESPIRATORY DISTRESS LED TO BRAIN DAMAGE AND THE PATIENT WAS THEN DECLARED BRAIN DEAD ON (B)(6) 2023 AT 12:35 PM. THE ICU MEDICAL AFFAIRS TEAM PROVIDED THE FOLLOWING MEDICAL ASSESSMENT: ICU MEDICAL WAS INFORMED ON (B)(6) 2024 OF A COMPLAINT THAT REPORTED THE DEATH OF A 12-MONTH-OLD MALE PATIENT. THE REPORT DESCRIBES A PATIENT THAT HAD A HISTORY OF REQUIRING SUPPLEMENTAL OXYGEN VIA A TRACHEOSTOMY TUBE, THE CHILD HAD A TRACHEAL STOMA AND REQUIRED HOME HEALTH CARE. IT IS REPORTED THAT ON (B)(6) 2023 THE PATIENT HAD A COUGH FOLLOWED WITH EMESIS, IT IS NOT DESCRIBED WHETHER A TRACHEOSTOMY TUBE WAS IN PLACE OR IF THE BALLOON WAS INFLATED, THEREBY PROVIDING A PROTECTED AIRWAY. IT IS REPORTED THAT THE NURSE ATTEMPTED TO CHANGE THE 3.5 BIVONA FLEXI TRACHEOSTOMY TUBE BUT WAS NOT SUCCESSFUL. THE 3.5 TRACHEOSTOMY TUBE WAS REMOVED BUT THE REPLACEMENT TRACHEOSTOMY TUBE COULD NOT BE REINSERTED INTO THE TRACHEA. IT IS REPORTED THAT THE PATIENT DEMONSTRATED RESPIRATORY DISTRESS AND EMS ARRIVED 40 MINUTES AFTER THE INITIATION OF COUGH AND EMESIS. THE PATIENT WAS TRANSFERRED TO (B)(6) HOSPITAL AND THE PATIENT WAS DECLARED DECEASED ON (B)(6) 2023. LACK OF DETAILS REGARDING THIS CLINICAL EVENT PRECLUDES A COMPREHENSIVE MEDICAL ASSESSMENT. BASED ON A REVIEW OF THE EVENTS IN THIS CASE, THE CUMULATIVE EVIDENCE STATES THAT THE PATIENT HAD A PROTECTED AIRWAY WITH A CUFFED TRACHEOSTOMY TUBE IN PLACE AND FOLLOWING AN EPISODE OF COUGHING AND EMESIS A CLINICAL DECISION, NOT DOCUMENTED IN THE COMPLAINT, WAS MADE TO CHANGE THE TRACHEOSTOMY TUBE. COUGH AND EMESIS ARE NOT UNUSUAL EVENTS IN PATIENTS WITH A TRACHEOSTOMY TUBE, AND A TRACHEOSTOMY TUBE CAN BE USED TO PROTECT THE AIRWAY. THERE IS NO DESCRIPTION THAT REPORTS WHY THE SECONDARY TUBE COULD NOT BE PLACED. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THIS MEDICAL ASSESSMENT WILL BE REVISED AS NEEDED. | INVESTIGATION SUMMARY: INVESTIGATION IS UNDERWAY AND ONCE THE INVESTIGATION HAS BEEN COMPLETED THE RESULTS WILL BE SENT IN A FOLLOW UP REPORT.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | "THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED."
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE MANUFACTURER RECEIVED INFORMATION REGARDING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. ADDITIONALLY, FOAM PARTICLES WERE NOT OBSERVED DURING THE EVALUATION OF THE DEVICE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, INFLAMMATORY RESPONSE AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THERE WAS 1 ERROR CODE FOUND. THE SECONDARY FINDINGS WERE OBSERVED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. IN BOX G: DATE RECEIVED BY MFG (RFB) HAS BEEN UPDATED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION REGARDING CANCER AND "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
BIPAP AUTO | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
STRATTICE LAPAROSCOPIC 10CM X 16CM | K070560 | THIS IS FOLLOW UP#1 TO REPORT ON 08/JAN/2025, PMQA RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT UNDERWENT A ¿VENTRAL¿ HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT ¿SP200065-011¿ ON (B)(6) 2019. THE PATIENT UNDERWENT AN ¿EXAM UNDER SEDATION, PAIN CONTROL WITH US GUIDED INJECTION OF MARCAINE¿ ON (B)(6) 2019. THE PATIENT ALSO UNDERWENT AN ¿ABDOMINAL WALL INCISION & DEBRIDEMENT, ENTEROCUTANEOUS FISTULA, PE¿ ON (B)(6) 2019. THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE ¿OPEN VENTRAL HERNIA REPAIR WITH MESH (STRATTICE), SMALL BOWEL RESECTION¿ BETWEEN (B)(6) 2019. THE PATIENT ALSO UNDERWENT AN ¿EXAM UNDER SEDATION, PAIN CONTROL WITH US GUIDED INJECTION OF MARCAINE¿ ON OR ABOUT (B)(6) 2019. THE PATIENT WAS TREATED FOR ¿ABDOMINAL WALL ABSCESS, PE, ENTEROCUTANEOUS FISTULA - DRAINED ABDOMINAL ABSCESS; ABDOMINAL WALL INCISION & DEBRIDEMENT¿ BETWEEN (B)(6) 2019. THE PATIENT WAS SEEN FOR ¿ABDOMINAL PAIN¿ BETWEEN 2019 AND 2023. THE PPF FORM ALSO INDICATES THAT THE PATIENT WAS IMPLANTED WITH A NON-ABBVIE DEVICE, ¿BARD VENTRALX ST, LOT #HUYC1262,¿ ON 22/OCT/2014. THE FORM INDICATES THAT THE DEVICE WAS REMOVED OR REVISED ON (B)(6) 2017 DUE TO ¿HERNIA REPAIR.¿ THE PATIENT WAS IMPLANTED WITH AN UNKNOWN NON-ABBVIE DEVICE ON THE SAME DATE. THE FORM INDICATES THAT THE DEVICE WAS REVISED OR REMOVED ON (B)(6) 2017 DUE TO ¿HERNIA REPAIR. THE PATIENT WAS IMPLANTED WITH A SECOND UNKNOWN NON-ABBVIE DEVICE ON THE SAME DATE. THE FORM INDICATES THAT THE DEVICE WAS REVISED OR REMOVED ON THE FOLLOWING DATES: (B)(6) 2019 DUE TO ¿EXAM UNDER SEDATION, US, PAIN, US GUIDED INJECTION WITH MARCAINE;¿ (B)(6) 2019 DUE TO ¿ABDOMINAL WALL INCISION & DEBRIDEMENT, ENTEROCUTANEOUS FISTULA, PE.¿ AS REPORTED IN THE INITIAL: LIMITED INFORMATION WAS REPORTED THROUGH A LEGAL EVENT THAT A PATIENT WAS IMPLANTED WITH STRATTICE ON (B)(6) 2019. THE FORM ALSO INDICATES THAT ON AN UNSPECIFIED DATE, THE STRATTICE IMPLANT HAD BEEN ¿REMOVED OR REVISED¿. NO INFORMATION WAS PROVIDED ON THE INDICATION FOR THE INITIAL SURGERY OR THE REASON FOR THE REMOVAL OR REVISION. PATIENT DIED ON (B)(6) 2023. THE CAUSE OF DEATH IS UNKNOWN. | A REVIEW OF THE DEVICE HISTORY RECORD FOR STRATTICE LOT SP200065 HAS BEEN INITIATED. IF ANY NEW, CHANGED OR CORRECTED INFORMATION IS NOTED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. THIS IS A KNOWN POTENTIAL ADVERSE EVENT ADDRESSED IN THE PRODUCT LABELING. WITHOUT RELEVANT PATIENT FACTORS, A RELATIONSHIP BETWEEN THE STRATTICE AND THIS EVENT COULD NOT BE DETERMINED. DUE TO THE LEGAL PROCESS, IF ADDITIONAL INFORMATION IS MADE AVAILABLE DURING LEGAL PROCEEDINGS, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. ADDITIONAL AND/OR CORRECTED DATA: A2, B3, B5, B7, D3, D4, H4, H6.
REMSTAR AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS ALLEGED EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER. THE PATIENT WAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. DUE TO POTENTIAL LITIGATION SURROUNDING THIS CASE, NO FOLLOW UP CAN BE PERFORMED AT THIS TIME. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
POWERPORT VUE M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K063377 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT APPROXIMATELY ONE YEAR TWO MONTHS AND SEVEN DAYS A POST PORT PLACEMENT THROUGH RIGHT SUBCLAVIAN VEIN FOR ADMINISTRATION OF CHEMOTHERAPY TO TREAT ADENOCARCINOMA, THE PATIENT ALLEGEDLY DEVELOPED REDNESS AROUND THE PORT SITE. IT WAS FURTHER REPORTED THAT THE PORT CATHETER ALLEGEDLY FRACTURED AND NEEDED TO BE REMOVED. REPORTEDLY, THE PORT WAS REMOVED BUT PATIENT CONDITION WAS WORSENED AND EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: D4 (EXPIRATION DATE: 05/2021). H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
REMSTAR | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN SECTION H1 AS SERIOUS INJURY. AFTER REVIEW, IT WAS DETERMINED THAT SECTION H1 SHOULD BE FILED AS DEATH. SECTION H1 CORRECTED IN THIS REPORT.
REMSTAR PRO C-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PREVIOUSLY ALLEGED CANCER AND DEATH. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS, AND DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS THREE ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTRE CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS CORRECTED. IN BOX H: (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH AND SHORTNESS OF BREATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO PARTICLES IN AIR PATH. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS ZERO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTRE CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT CORRECTED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, KIDNEY, BLADDER AND PROSTATE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTO BIFLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RECURRING PNEUMONIA, BRONCHITIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, NAUSEA / VOMITING, LIVER DISEASE/TOXICITY, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BRILLIANCE | K033326 | ON (B)(6) 2024, WITH THE INTENT TO REPLACE AN INDWELLING CHEST TUBE (DUE TO ONGOING EMPYEMA AND ITS FAILURE TO DRAIN) UNDER CT GUIDANCE, A FEMALE PATIENT WAS TRANSPORTED FROM THE INTENSIVE CARE UNIT (ICU) TO A CT ROOM ON HIGH FLOW OXYGEN. THE PATIENT'S INITIAL BLOOD PRESSURE AND HEART RATE WERE REPORTED AS WITHIN NORMAL LIMITS. THE PATIENT WAS RESPONSIVE; HOWEVER, SHE WAS CONFUSED AND WAS NOT ABLE TO CONSENT. DURING THE PROCEDURE, NO SEDATION WAS USED, ONLY FENTANYL FOR PAIN. THE PATIENT HAD THE INITIAL SURVIEW SCAN; HOWEVER, THE PROCEDURE COULD NOT START DUE TO AN IMAGING ISSUE . THE TEAM TRIED TO SEE IF THE IMAGING COULD WORK, BUT AFTER SEVERAL MINUTES, THE DECISION WAS MADE TO MOVE PATIENT TO ANOTHER SCANNER ACROSS THE HALL. THE PATIENT WAS THEN TRANSFERRED TO THE SCANNER ACROSS THE HALL (TAKING ONLY A COUPLE MINUTES) WHERE THE SURVIEW IMAGE WAS PERFORMED, AND THE CHEST TUBE PROCEDURE WAS STARTED. IT WAS REPORTED BY THE ATTENDING PHYSICIAN THAT PRIOR TO HER TRANSPORT TO THE SECOND SCANNER DURING TRANSPORT, AND UPON HER ARRIVAL TO THE SECONDARY SCANNER, VITAL SIGNS AND OXYGEN REQUIREMENTS WERE SIMILAR TO THOSE NOTED UPON HER ARRIVAL AT THE FIRST SCANNER. THE SECOND DEVICE WORKED AS INTENDED. EARLY DURING THE IMAGING PHASE OF THE SECOND PROCEDURE, THE PATIENT HAD SUDDEN DETERIORATION OF HER RESPIRATORY STATUS. PATIENT WAS ADMINISTERED NARCAN (TO REVERSE FENTANYL) BUT THERE WAS NO CHANGE IN HER STATUS. A RAPID RESPONSE WAS CALLED; HOWEVER, GIVEN HER DO NOT RESUSCITATE (DNR) STATUS, CODE/RESUSCITATION WAS NOT INITIATED. THE ATTENDING PHYSICIAN NOTED THAT, SINCE THE PATIENT HAD A PACEMAKER, SHE HAD ELECTRICAL ACTIVITY BUT NO PULSE (I.E., PEA) AND WAS PRONOUNCED DEAD. PER THE ATTENDING PHYSICIAN, THE MOST LIKELY CAUSE OF DEATH WAS RELATED TO THE PATIENT¿S UNDERLYING COMORBIDITIES AND HER POOR CARDIAC/RESPIRATORY STATUS. THE HOSPITAL PHYSICIAN POSITED THAT CAUSE OF DEATH WAS LIKELY EITHER A MYOCARDIAL INFARCTION (MI) OR A PULMONARY EMBOLISM (PE); HOWEVER, THAT ASSUMPTION CANNOT BE CONFIRMED, AS AN AUTOPSY WAS NOT CONDUCTED. FROM A MEDICAL PERSPECTIVE, IT IS UNLIKELY THAT THE DELAYED INSERTION OF THE CHEST TUBE, IS LINKED TO THE IMAGING FAILURE ON FIRST CT THAT REQUIRED THE PATIENT TO BE TRANSFERRED TO THE SECOND CT, CONTRIBUTED TO OR CAUSED THE DEATH OF THE PATIENT. THE PATIENT WAS CRITICALLY ILL, WITH SEVERAL CO-MORBIDITIES , PRIOR TO BEING TRANSFERRED TO FIRST CT AND THEN HAD SUDDEN RESPIRATORY DETERIORATION AFTER THE PROCEDURE STARTED IN THE SECOND SCANNER. A NEW CHEST TUBE WAS NOT PLACED; HOWEVER, DUE TO HER RAPID DECLINE AND THE PATIENT¿S CO-MORBIDITIES, IT IS UNLIKELY THAT THE OUTCOME WAS AFFECTED BY THE DELAY ATTRIBUTED TO THE FIRST DEVICE. FURTHERMORE, AS THE PATIENT WAS DNI/DNR DUE TO HER UNDERLYING POOR PROGNOSIS, A RAPID RESPONSE WAS CALLED; HOWEVER, THERE WAS NO ATTEMPT TO RESUSCITATE ONCE SHE DETERIORATED CLINICALLY. PHILIPS FIELD SERVICE ENGINEER (FSE) ARRIVED ON SITE ON SEPTEMBER 10, 2024, TO EVALUATE THE BRILLIANCE CT DEVICE AND DEVICE LOGS. BASED ON AN INITIAL INVESTIGATION THE COMMON IMAGE RECONSTRUCTION SERVER (CIRS) GRAPHICS PROCESSOR BOARD (GPU) WAS ORDERED TO ADDRESS THE CUSTOMER¿S COMPLAINT. ON SEPTEMBER 12, 2024, THE FSE REPLACED THE CIRS GPU. THE DEVICE WAS TESTED AND RETURNED TO USE. THE CIRS GPU WAS SENT TO PHILIPS ON SEPTEMBER 23, 2024, FOR ADDITIONAL INVESTIGATION. A PHYSICIAN-TO-PHYSICIAN PHONE CONVERSATION BETWEEN PHILIPS AND THE HOSPITAL ATTENDING PHYSICIAN WAS HELD ON (B)(6) 2024, TO OBTAIN THE AFOREMENTIONED INFORMATION RELATED TO PATIENT STATUS, PAST MEDICAL HISTORY, AND EVENTS LEADING UP TO THE PATIENT¿S DEATH. THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE COMPLETED UPON INVESTIGATION OF THE RETURNED COMPONENT. | FINAL REPORT: THE INVESTIGATION INTO THE FAILURE OF THE BRILLIANCE CT SYSTEM DETERMINED THAT THE ROOT CAUSE OF THE IMAGING ISSUE WAS A FAILURE IN THE GPU (GRAPHICS PROCESSING UNIT). THE AFFECTED GPU WAS REPLACED ON SEPTEMBER 12, 2024, AFTER A PHILIPS FIELD SERVICE ENGINEER (FSE) CONDUCTED AN ONSITE EVALUATION. SUBSEQUENT TESTING CONFIRMED THAT THE DEVICE WAS FUNCTIONING AS INTENDED POST-REPLACEMENT. THE FAULTY COMPONENT WAS SENT TO PHILIPS FOR FURTHER ANALYSIS ON SEPTEMBER 23, 2024, AND A PHYSICIAN-TO-PHYSICIAN CONVERSATION WAS CONDUCTED ON SEPTEMBER 26, 2024, TO CLARIFY PATIENT DETAILS AND EVENTS LEADING UP TO THE PATIENT¿S DEATH. THE FAILURE OF THE GPU COULD LEAD TO A LOSS OF FUNCTION DURING CLINICAL USE, RESULTING IN MISSING OR PARTIAL IMAGES DURING A PROCEDURE. THIS COULD NECESSITATE A RESCAN, POTENTIALLY EXPOSING THE PATIENT TO ADDITIONAL RADIATION. HOWEVER, THE SEVERITY OF HARM IS CONSIDERED NEGLIGIBLE TO MINOR BASED ON THE NATURE OF THE IMAGING ISSUE AND RADIATION EXPOSURE. THE ISSUE RELATED TO THE GPU FAILURE HAS BEEN INVESTIGATED, CORRECTED, AND EVALUATED FOR RISK. MITIGATION MEASURES AND RISK MANAGEMENT ANALYSES CONFIRM THAT THE IMPACT ON PATIENT SAFETY IS MINIMAL. THE SYSTEM HAS BEEN RESTORED TO FULL FUNCTIONALITY, AND NO FURTHER CORRECTIVE ACTIONS ARE REQUIRED. THE RISK ASSOCIATED WITH THIS INCIDENT REMAINS WITHIN ACCEPTABLE LEVELS, AND THE FAILURE IS UNLIKELY TO HAVE DIRECTLY CONTRIBUTED TO THE ADVERSE PATIENT OUTCOME. CORRECTIONS: THE REPORTING INSTITUTION ADDRESSES WAS UPDATED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING COLON CANCER, DEATH, OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTOSV ADV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
PACEL BIPOLAR PACING CATHETER | K152784 | AFTER USING THE TEMPORARY PACING CATHETER, THE PATIENT EXPIRED. THE CAUSE OF DEATH REMAINS UNCERTAIN. | ADDITIONAL INFORMATION: H6: ONE 6F PACEL BIPOLAR PACING CATHETER (TORQUE CONTROL) WAS RECEIVED FOR ANALYSIS. ELECTRICAL TESTING REVEALED BOTH THE TIP AND RING ELECTRODE CIRCUITS MET SPECIFICATIONS OF ACCEPTABLE RESISTANCE VALUES WITH NO OPEN OR SHORT CIRCUITS DETECTED. THE CATHETER SHAFT OUTSIDE DIAMETER MEASUREMENT WAS WITHIN SPECIFICATIONS. NO RESISTANCE AND NO ANOMALIES WERE NOTED WHEN THE RETURNED CATHETER WAS INSERTED THROUGH A 6F SHEATH FROM CURRENT INVENTORY. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. THE CAUSE OF THE REPORTED PATIENT DEATH REMAINS UNKNOWN.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED AND DUST LIKE CONTAMINATION WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS THIRTY-TWO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. DEVICE PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID WAS UPDATED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS AND DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS 32 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. BOX H: PROBLEM CODE GRID, EVALUATION METHOD CODE, EVALUATION RESULTS CODE, AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT IS ALLEGING CANCER. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 07/19/2024, AND SECTION H 11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT IS ALLEGING CANCER. IN ADDITION, THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
ION | K182188 | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNEVENTFUL ION ENDOLUMINAL PROCEDURE. AFTER THE ION PROCEDURE, ENDOBRONCHIAL ULTRASOUND (EBUS) WAS PERFORMED WITH A BRONCHOSCOPE. THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE WAS INFORMED THAT THE PATIENT CODED AND EXPIRED WHILE IN POST-PROCEDURE RECOVERY. THERE WERE NO REPORTED INTRAPROCEDURAL ION SYSTEM MALFUNCTIONS/ISSUES. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. | A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS PERFORMED FOR THIS PROCEDURE AND NO RELATED NON-CONFORMANCES WERE IDENTIFIED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE UNDERWENT AN ION ENDOLUMINAL BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. IN RECOVERY, THE PATIENT DIED AFTER SUFFERING A CODE. NO FURTHER INFORMATION IS AVAILABLE BUT IT IS LIKELY PROCEDURE RELATED GIVEN TEMPORAL ASSOCIATION. THERE IS INSUFFICIENT INFORMATION TO DETERMINE A RELATIONSHIP TO THE DEVICES, BUT THERE WAS NO REPORTED MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. ATTEMPTS AT OBTAINING FURTHER INFORMATION HAVE BEEN UNSUCCESSFUL. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). --OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019.
REMSTAR PLUS C-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR SE | K130077 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR SE | K130077 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNEVENTFUL DA VINCI-ASSISTED SACROCOLPOPEXY WITH HYSTERECTOMY AND SALPINGO-OOPHORECTOMY PROCEDURE. THE PROCEDURE WAS COMPLETED WITH A GYNECOLOGY SURGEON AND A SECONDARY GENERAL SURGEON WAS PRESENT TO PERFORM LYSIS OF ADHESIONS. THE PATIENT REQUIRED A SECOND PROCEDURE FOR OPEN SURGERY DUE TO AN UNSPECIFIED BOWEL INJURY. THE SECOND PROCEDURE WAS COMPLETED AND THE PATIENT EXPIRED ON POD 2. THERE WAS NO ALLEGATION OF A DA VINCI PRODUCT ISSUE RELATED TO THIS EVENT. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL TO DATE. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, CADIERE FORCEPS, MONOPOLAR CURVED SCISSORS, TIP-UP FENESTRATED GRASPER, MARYLAND BIPOLAR FORCEPS, LARGE NEEDLE DRIVER, LARGE SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A ROBOTIC SACROCOLPOPEXY AND WAS RE-EXPLORED ON POD 1 FOR A BOWEL INJURY. THE SECTION OF BOWEL AND HOW THE BOWEL INJURY MAY HAVE OCCURRED ARE NOT PROVIDED. THE INTRAOPERATIVE FINDINGS AT THE TIME OF THE RE-EXPLORATION WERE NOT PROVIDED. NEITHER THE DETAILS OF THE HOSPITAL COURSE AFTER THE INITIAL OPERATION NOR AFTER THE RE-EXPLORATION ARE PROVIDED. THE PATIENT DIED ON POD 2 AND NO DETAILS AS TO HOW THEY DIED ARE PROVIDED. BASED ON THE INFORMATION PROVIDED, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
INTELLIVUE MP30 | K122439 | THE CUSTOMER REPORTED INTERFERENCE ON THE ECG WHEN SPECIFIC LEADS ARE PLUGGED INTO THE MONITOR. THE DEVICE WAS IN USE ON PATIENT AT TIME OF EVENT, THERE WAS AN ADVERSE EVENT REPORTED. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
RING PROFILE 3D 30MM | K073324 | MEDTRONIC RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT ON THE SAME DAY AS THE IMPLANT DATE OF THIS 30MM ANNULOPLASTY RING, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE 30MM ANNULOPLASTY RING COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR PRO C-FLEX+ DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF KIDNEY DISEASE/TOXICITY. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON THE CYCLER. UPON FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN), IT WAS REPORTED THIS PATIENT EXPIRED AT HOME DUE TO AN UNREPORTED CAUSE ON (B)(6) 2024. IT WAS AFFIRMED THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF DEATH. THE PATIENT WAS FOUND PULSELESS AND UNRESPONSIVE ON THE MORNING OF (B)(6) 2024 WHILE CONNECTED TO HER CYCLER. EMERGENCY SERVICES WERE NOT ACTIVATED, AND THE PATIENT WAS PRONOUNCED DECEASED IN HER HOME. THOUGH THE EXACT CAUSE OF DEATH WAS NOT REPORTED, IT WAS CONFIRMED THE PATIENT¿S DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF NAY FRESENIUS PRODUCT(S) OR DEVICE(S). | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CANNOT BE DETERMINED; HOWEVER, THERE WAS NO INDICATION THIS PATIENT¿S DEATH WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ABOUT AN ALLEGATION OF A PATIENT DEATH. IN ADDITION, THERE WAS ALSO ALLEGATIONS OF LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURER INSPECTED EXTERNALLY AND INTERNALLY, OBSERVED INITIATED THERAPY TO THE DEVICE, THERE WAS A SLIGHT WHINY NOISE, POTENTIALLY DUE TO THE DUST LIKE CONTAMINATE ON THE INSIDE OF THE BLOWER. DUST LIKE CONTAMINATE ON THE PHILIPS WHITE ULTRA FINE FILTER, EXTERIOR AND INTERIOR OF THE TOP ENCLOSURE, PCA (PRINTED CIRCUIT ASSEMBLY), FLOW MANIFOLD, EXTERIOR AND INTERIOR OF THE BLOWER HOUSING, BLOWER GROMMET, BLOWER, RIGHT SIDE ASSEMBLY, HUMIDIFIER CABLE, AIR INLET SEAL, EXTERIOR AND INTERIOR OF THE BOTTOM ENCLOSURE, ISO PORT (INTERNATIONAL ORGANIZATION OF STANDARDIZATION), FLIP LID SEAL, DRY BOX, DRY BOX SEAL, DRY BOX INLET SEAL, TOP HOUSING, EXTERIOR AND INTERIOR OF THE HUMIDIFIER BOTTOM ENCLOSURE, HEATER PLATE, HEATER PLATE SPRING, EXTERIOR AND INTERIOR OF THE HUMIDIFIER LOWER BASE, AND THE SOUND ABATEMENT FOAM. MISSING THE UI KNOB (USER INTERFACE), SD CARD, AND THE SD CARD COVER. HAIRLIKE FIBERS ON THE FLIP LID SEAL AND THE HUMIDIFIER LOWER BASE. AN UNKNOWN CONTAMINATE ON THE EXTERIOR OF THE TOP ENCLOSURE WHERE THE UI KNOB WOULD SIT, EXTERIOR OF THE SIDE PANEL, AND THE EXTERIOR OF THE HUMIDIFIER SIDE PANEL. A YELLOWISH TINT ON THE FLOW MANIFOLD, BLOWER GROMMET, BLOWER BELLOW, AIR INLET SEAL, FLIP LID SEAL, DRY BOX SEAL, AND THE DRY BOX INLET SEAL. A PAPER-LIKE SUBSTANCE ON THE INTERIOR OF THE WATER TANK. POTENTIAL BIOCONTAMINATION ON THE BLOWER GROMMET, BLOWER BELLOW, AND THE AIR INLET SEAL. POTENTIAL LIQUID INGRESS WITH A WHITE POWDER CONSISTENT WITH MINERAL DEPOSITS ON THE EXTERIOR AND INTERIOR OF THE HUMIDIFIER FLIP LID, FLIP LID SEAL, WATER TANK, INTERIOR OF THE TOP ENCLOSURE, PCA, BLOWER HOUSING, BOTTOM ENCLOSURE, ISO PORT, DRY BOX SEAL, DRY BOX, AND DRY BOX INLET SEAL. THE SCREW BOSS WAS STRIPPED ON THE RIGHT-SIDE ASSEMBLY. THERMAL DAMAGE TO THE J9 AND ON THE HUMIDIFIER CABLE. THERMAL DAMAGE TO THE J9 HAS BEEN PREVIOUSLY INVESTIGATED AS DOCUMENTED. THE MANUFACTURER USED A FITT (FOAM INTEGRITY TEST TOOL) AND WAS NOT ABLE TO CONFIRM THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER CONCLUDES THERE WAS NO EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE AND THEY WERE UNABLE TO DIRECTLY ADDRESS THE SYMPTOMS.
BIPAP AUTOSV ADV SYS ONE | K090539 | "THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED."
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR PRO C-FLEX+ DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S ATTORNEY REPORTING ALLEGATIONS OF CANCER. THERE WAS A REPORT OF PATIENT DEATH HAVE OCCURRED. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PROXIMATE*75MM LINEAR CUTTER | K020779 | IT WAS REPORTED THAT POST-OP A PROCEDURE, THERE WAS AN ANASTOMOTIC LEAK. REOPERATION WAS REQUIRED. EVENT TIMELINE: 9/11 - CABG 9/15 - RIGHT, ASCENDING COLECTOMY DUE TO ISCHEMIC BOWEL (B)(6)- SMALL BOWEL RESECTION ¿10 DAYS LATER¿ (APPROX. 9/25)- CT SCAN, NO ¿OVER THE WEEKEND¿ (APPROX. 9/28 OR 9/29) - PATIENT BECAME SEPTIC (B)(6) - REOPERATION TO REPAIR STAPLE LINE LEAK. | (B)(4). DATE SENT: 10/9/2024 D4: BATCH # UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF THE POST-MARKET SURVEILLANCE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH DUE TO CANCER, REDUCED CARDIOPULMONARY RESERVE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE DEVICE RETURNED FOR EVALUATION AND THERE WAS NO EVIDENCE OF FOAM PARTICLE OBSERVED. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
TIGHTRAIL ROTATING DILATOR SHEATH | K150360 | A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT ATRIAL (RA) AND A RIGHT VENTRICULAR (RV) LEAD IMPLANTED IN 1990, AND AN RA AND RV LEAD, IMPLANTED IN 2019. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH, THE RA AND RV LEADS (IMPLANTED IN 2019) WERE EXTRACTED SUCCESSFULLY. NEXT, UPSIZING TO A 16F GLIDELIGHT AND THEN SWITCHING TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH, THE RV LEAD (IMPLANTED IN 1990) WAS REMOVED. THEN, USING THE 13F TIGHTRAIL ON THE REMAINING RA LEAD (IMPLANTED IN 1990), ADVANCEMENT WAS MADE THROUGH THE SUBCLAVIAN AND INNOMINATE REGIONS, AND PROGRESSED FARTHER THROUGH BINDING SITES IN THE SUPERIOR VENA CAVA (SVC). ONCE THROUGH THE BINDING, THE RA LEAD DISLODGED, AND THE PATIENT'S BLOOD PRESSURE DROPPED. RESCUE EFFORTS BEGAN, INCLUDING RESCUE BALLOON, MEDICATIONS, AND STERNOTOMY. A 1 CM PERFORATION TO THE RIGHT ATRIUM WAS DISCOVERED AND REPAIRED (MDR #3007284006-2024-00204). HOWEVER, WHILE CLOSING THE PATIENT'S CHEST, A HEMATOMA AND TWO SMALL PERFORATIONS IN THE SVC WERE DETECTED (MDR #3007284006-2024-00205). ATTEMPTS WERE MADE TO REPAIR THE TWO PERFORATIONS; HOWEVER, THE PATIENT'S TISSUES WERE FRIABLE, AND EXTENDED INTO ONE LARGE PERFORATION. THE PATIENT WENT INTO DISSEMINATED INTRAVASCULAR COAGULATION (DIC), AND FURTHER INTERVENTION COMMENCED, INCLUDING ADMINISTRATION OF BLOOD AND PLACING THE PATIENT ON PUMP. THE SVC PERFORATION KEPT EXTENDING WHILE REPAIR ATTEMPTS CONTINUED, AND DESPITE EXTENSIVE RESCUE EFFORTS, THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE 13F TIGHTRAIL IN USE WHEN THE SVC PERFORATIONS OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. | H3): THE TIGHTRAIL WAS NOT RETURNED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6): GREAT VESSEL PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH THE USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS DUST WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS 32 ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED. SECTION H6 WAS UPDATED AND REPAIR EVALUATION CAPTURED.
DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | K213919 | IT WAS REPORTED BY THE PATIENT'S DAUGHTER THAT THE PATIENT HAD PASSED ON AWAY (B)(6) 2024. THE REPORTER CONTACTED DEXCOM TO INQUIRE IF DEXCOM HAD MORE DETAILS ABOUT THE PATIENT'S CGM READINGS DURING A PERIOD OF SENSOR ERROR PRIOR TO THE PATIENT'S PASSING. THE REPORTER NOTED A BREAK IN THE CGM READINGS BETWEEN 4:00PM-5:00PM THE DAY OF THE PATIENT'S DEATH. THE REPORTER INQUIRED WHY THE DEXCOM WAS NOT READING FOR A CERTAIN AMOUNT OF TIME. SHE MENTIONED THAT SHE COULD SEE THE ALERTS AND IT NEVER ALERTED THE PATIENT THAT THE SENSOR WAS OFF/NOT READING AND DID NOT ALERT HER TO HYPOGLYCEMIA. WHEN CGM READINGS RETURNED, THE VALUE WAS REPORTEDLY "LOW", BUT SHE HAD REPORTEDLY ALREADY PASSED AWAY BY THAT POINT WHEN SHE RECEIVED THAT ALERT. THE REPORTER INQUIRED HOW LONG THE PATIENT HAD BEEN IN SENSOR ERROR PRIOR TO THE "LOW" CGM READING. THE PATIENT WAS NOTED TO HAVE ATTEMPTED TO TREAT HERSELF WITH PEANUT BUTTER. THE PATIENT'S CAUSE OF DEATH WAS BELIEVED TO BE HYPOGLYCEMIA, ALTHOUGH NO AUTOPSY WAS PERFORMED. THE REPORTER ALLUDED TO USING THE FOLLOW APP AND STATED THAT NO ONE WAS ALERTED AFTER THE CGM READING WAS 100 MG/DL . THE REPORTER NOTED THAT NO ONE WAS ALERTED TO THE SENSOR ERROR AND ALSO NO ONE WAS ALERTED TO THE HYPOGLYCEMIA. THE REPORTER NOTED THAT SHE WAS NOT ALERTED UNTIL THE CGM READING WAS 39 MG/DL AND THE PATIENT HAD ALREADY PASSED AWAY. THE REPORTER NOTED AN ABSENCE OF AN ALERT FOR 15 MINUTES WHEN REVIEWING THE PATIENT'S APP. THERE WAS REPORTEDLY NO FURTHER INFORMATION SINCE THE CALLER WANTED TO CHECK IF THE DEXCOM SYSTEM COULD HAVE RECORDED THE SENSOR READINGS DURING THE TIME FRAME. NO PRODUCT WAS PROVIDED FOR INVESTIGATION. DATA WAS RECEIVED BUT IS PENDING EVALUATION. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. | (B)(4). DIABETES MELLITUS IS A KNOWN CAUSE OF HYPOGLYCEMIA.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE SERVICE CENTER FOR EVALUATION. DURING SERVICING OF THE DEVICE, NO EVIDENCE OF FOAM PARTICLES WAS OBSERVED. THE DEVICE HAS BEEN REPAIRED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
EDWARDS ESHEATH+ INTRODUCER SET | K200258 | AS REPORTED FROM GERMANY, A 16FR ESHEATH LINER PUNCTURE AND SUBSEQUENT PATIENT DEATH OCCURRED DURING A TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. DURING THE PROCEDURE, THERE WAS RESISTANCE WHILE TRYING TO ADVANCE THE COMMANDER DELIVERY SYSTEM AND 29MM SAPIEN 3 VALVE THROUGH THE ESHEATH. THE PATIENT'S BLOOD PRESSURE WAS FALLING; CPR AND INTUBATION BEGAN. THE DELIVERY SYSTEM EXITED THROUGH THE SIDE OF THE ESHEATH. A PERFORATION OF THE ABDOMINAL AORTA WAS SEEN. IN ORDER TO BE ABLE TO INTRODUCE A CODA BALLOON TO BLOCK THE AORTA, AN ATTEMPT WAS MADE TO REMOVE THE VALVE CATHETER FROM THE SHEATH. THE VALVE WAS LOST WITHIN THE SHEATH. THE SHEATH AND DELIVERY SYSTEM WERE WITHDRAWN UNDER FLUOROSCOPY. HOWEVER, THE VALVE COULD NOT BE WITHDRAWN FROM THE RIGHT COMMON FEMORAL ARTERY. THE RIGHT COMMON FEMORAL ARTERY WAS OPENED USING VASCULAR SURGERY AND THE VALVE WAS RETRIEVED. UPON INSPECTION OF THE VALVE, IT WAS NOTED THAT ONE VALVE STRUT WAS BENT AND HAD EXITED THE SHEATH. DESPITE BLOCKING THE AORTA WITH A BALLOON, HEMODYNAMIC STABILIZATION WAS NOT ACHIEVED. IN THE CASE OF UNCONTROLLED BLEEDING AND CPR FOR APPROXIMATELY 45 MINUTES, AN INTERDISCIPLINARY DECISION WAS MADE AGAINST VA-ECMO IMPLANTATION WITH CENTRAL CANNULATION. THE PATIENT DIED. | INVESTIGATION IS ONGOING.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. REPORTEDLY, THE CAUSE OF DEATH WAS DETERMINED TO BE NATURAL CAUSES; HOWEVER, THE REPORTER ALSO ALLEGED THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE EVENT. NO ADDITIONAL INFORMATION WAS PROVIDED. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. H3 OTHER TEXT : DEVICE NOT RETURNED.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024 AFTER EXPERIENCING DIABETIC KETOACIDOSIS; HOWEVER, THIS COULD NOT BE CONFIRMED AT THE TIME OF THIS REPORT, AS PUMP DATA IS NOT AVAILABLE FOR REVIEW. REPORTEDLY, THE LAST INTERACTION FOUND IN THE PUMP HISTORY WAS ON (B)(6) 2024 WHEN THE CUSTOMER DELIVERED A MANUAL BOLUS. THE REPORT FROM THE SWEDISH MEDICAL PRODUCTS AGENCY INDICATES THAT THE DECEDENT DID NOT HAVE A WORKING SENSOR ON HIS CONTINUOUS GLUCOSE MONITOR (CGM) IN HIS LAST DAYS OF LIFE AND AS SUCH, BLOOD GLUCOSE LEVELS WERE NOT AVAILABLE. TANDEM CAN ONLY CONFIRM THAT THERE IS NO CGM DATA FOR THE DATES (B)(6) 2024. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. H3 OTHER TEXT: DEVICE NOT RETURNED.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATIONRESPIRATORY TRACT IRRITATIONDIZZINESS AND/OR HEADACHEINFLAMMATORY RESPONSEKIDNEY DISEASE/TOXICITYREDUCED CARDIOPULMONARY RESERVEDEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN THIS REPORT, IN BOX G REPORT SOURCE - OTHER HAS BEEN UPDATED/CORRECTED. IN BOX H METHOD CODE, RESULTS CODE AND CONCLUSION CODE HAS BEEN UPDATED/CORRECTED.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA (NEW OR WORSENING); CANCER; DEATH, LUNG CANCER. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
BIPAP PRO BI-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER,DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | BOX D IS UPDATED IN THIS REPORT.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED LUNG DISEASE, LIVER DISEASE, EYE IRRITATION, NOSE IRRITATION, CANCER. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED LUNG DISEASE, LIVER DISEASE/TOXICITY, EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING AND CANCER. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: PATIENT OUTCOME CODE GRID EVALUATION METHOD CODE GRID EVALUATION RESULTS CODE GRID CONCLUSION CODE GRID HAS BEEN UPDATED IN PREVIOUS REPORT PRODUCT BRAND NAME CAPTURED WRONGLY, IN THIS REPORT HAS BEEN UPDATED CORRECTLY.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
CADD-SOLIS VIP AMBULATORY INFUSION PUMP | K111275 | IT WAS REPORTED THAT A PATIENT DIED AT HOME IN THE PRESENCE OF AMBULANCE VITORIA. PER REPORTER THE DEATH OCCURRED WHILE USING THE PUMP FOR CHEMOTHERAPY WITH A HICKMAN CATHETER INTO THE HEART. HOWEVER, REPORTER STATED THAT THERE WAS NOT YET ANY EVIDENCE THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO THE DEATH. PER REPORTER THE PATIENT WAS SEEING A MEDICAL TEAM TWICE WEEKLY AND HAVING THE BAG CHANGED AND HAD RECENTLY BEEN VIEWED AT "AH WITH NO ISSUES." PER A MEDICAL REVIEW LACK OF DETAILS REGARDING THE EVENT PRECLUDED A COMPREHENSIVE MEDICAL ASSESSMENT. | THE PUMP WAS RETURNED FOR EVALUATION. VISUAL INSPECTION FOUND THE PUMP WAS IN GOOD CONDITION. REVIEW OF THE EVENT HISTORY LOG FOUND NO UNUSUAL ACTIVITY. THE DEVICE PASSED FUNCTIONAL TESTING AND NO PROBLEMS WERE FOUND. THE MANUFACTURING DEVICE HISTORY RECORD REVIEW AND THE SERVICE HISTORY REVIEW FOUND NO EVIDENCE OF AN ISSUE WITH THE DEVICE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER; DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. DUE DATE HAS BEEN UPDATED. IN BOX G: DATE RECEIVED BY MFG HAS BEEN UPDATED. IN BOX H6: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAVE BEEN UPDATED.
TABLO HEMODIALYSIS SYSTEM | K190793 | IT WAS REPORTED THAT THERE WAS A CODE BLUE, BLOOD WAS NOT RETURNED AND PATIENT EXPIRED DURING TREATMENT. THE TREATMENT STARTED AT 8:00PM AND PATIENT CODED AT 11:00PM, NO ADDITIONAL INFORMATION WAS PROVIDED. | H3/H6: FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION OR NONCONFORMANCE, THAT CONTRIBUTED TO THE REPORTED EVENT. POTENTIAL ADVERSE EVENTS IN THE INSTRUCTIONS FOR USE (IFU) WITH THE TABLO SYSTEM INCLUDES, BUT ARE NOT LIMITED TO, OTHER, MORE SERIOUS, COMPLICATIONS ARISING FROM DIALYSIS, SUCH AS HEMORRHAGE, AIR EMBOLISM, ACIDOSIS, ALKALOSIS OR HEMOLYSIS, CAN CAUSE SERIOUS PATIENT INJURY OR DEATH. FROM THE INFORMATION PROVIDED, EVENT CAUSALITY COULD NOT BE DETERMINED. MULTIPLE FOLLOW UP ATTEMPTS TO GATHER ADDITIONAL INFORMATION FROM THE INITIAL REPORTER WERE UNSUCCESSFUL; THEREFORE, ROOT CAUSE OF THE EVENT WAS NOT ABLE TO BE CONFIRMED. HOWEVER IS IT NOT SUSPECTED TO BE PRODUCT RELATED. OUTSET MEDICAL, INC. TECHNICAL SUPPORT ENGINEER (TSE) REVIEWED SITE SYSTEM LOG WITH A PROCEDURE DATE OF (B)(6) 2024 AND VERIFIED THAT THERE WAS NO ISSUE WITH THE SYSTEM WHICH CAUSED THE PATIENT EVENT. A REVIEW OF PRODUCTION RECORDS FOR THIS SERIAL NUMBER DID NOT NOTE ANY RELATED MANUFACTURING NONCONFORMANCES THAT WOULD CONTRIBUTE TO THIS EVENT.
REMSTAR LITE | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED RESPIRATORY FAILURE, REPARATORY TRACT IRRITATION, EYE IRRITATION, NOSE IRRITATION. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
TRILOGY | K083526 | THE MANUFACTURER WAS MADE AWARE THAT A PATIENT THAT USES A TRILOGY 100 VENTILATOR PASSED AWAY WHILE ON THE DEVICE. AN INVESTIGATION OF THIS COMPLAINT IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. AS PART OF THE COMPLAINT HANDLING PROCESS, THE MANUFACTURER IS MAKING ATTEMPTS TO RETRIEVE THE DEVICE FOR INVESTIGATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
MARK 5 NUVO LITE | K123738 | ON SEPTEMBER 19, 2024, NIDEK MEDICAL PRODUCTS WAS CONTACTED BY A DISTRIBUTOR IN GERMANY THAT A POTENTIAL INCIDENT INVOLVING ONE OF OUR CONCENTRATORS AND A PATIENT SMOKING OCCURRED, AND THAT INCIDENT RESULTED IN THE PATIENT'S DEATH. ACCORDING TO THE DISTRIBUTOR, THE PATIENT HAD BEEN ON THERAPY SINCE (B)(6) 2024 AND HAD BEEN ADVISED IN THE PRESENCE OF THE CAREGIVER TO NOT USE THE DEVICE WHILE SMOKING. THEY FURTHER INCLUDED THAT ADDDITIONALLY ON (B)(6) 2024 THE DEVICE HAD BEEN EQUIPPED WITH A FIRESAFE FIRE PROTECTION VALVE AND HAD PROVIDED THE IFU TO THE PATIENT AND CAREGIVER. | ON SEPTEMBER 19, 2024 NIDEK MEDICAL PRODUCTS WAS CONTACTED BY A GERMAN DISTRIBUTOR CONCERNING A FIRE INVOLVING A NIDEK MEDICAL NUVO LITE OXYGEN CONCENTRATOR, SERIAL NUMBER (B)(6). THE END USER OF THE DEVICE DIED IN THE FIRE. IN THEIR INITIAL REPORT, THE DISTRIBUTOR INDICATED THAT THE END USER WAS IN A FACILITY RECEIVING LONG-TERM CARE AND WAS CAUTIONED TO NOT SMOKE WHILE OPERATING THE DEVICE. ON SEPTEMBER 19, 2024 NIDEK MEDICAL PRODUCTS WAS CONTACTED BY A GERMAN DISTRIBUTOR CONCERNING A FIRE INVOLVING A NIDEK MEDICAL NUVO LITE OXYGEN CONCENTRATOR, SERIAL NUMBER (B)(6). THE END USER OF THE DEVICE DIED IN THE FIRE. IN THEIR INITIAL REPORT, THE DISTRIBUTOR INDICATED THAT THE END USER WAS IN A FACILITY RECEIVING LONG-TERM CARE AND WAS CAUTIONED TO NOT SMOKE WHILE OPERATING THE DEVICE. THE DISTRIBUTOR INCLUDED A STATEMENT FROM THE NURSE WHO WAS CARING FOR THE END USER. THE NURSE SAID, "THE PATIENT TOUCHED THE NASAL CANNULA WITH A LIT CIGARETTE AND THIS CAUSED THE DEVICE TO BURN BACK AND BE DAMAGED." PHOTOS PROVIDED BY THE DISTRIBUTOR SUPPORTED THE NURSE'S CLAIM. THE DEVICE ARRIVED AT THE MANUFACTURER'S FACILITY ON SEPTEMBER 30, 2024. THE FRONT HOUSING OF THE DEVICE IS THE ONLY NON-ACCESSORY COMPONENT OF THE DEVICE THAT SUSTAINED DAMAGE; ALL COMPONENTS ON THE INSIDE OF THE DEVICE, INCLUDING ELECTRICAL COMPONENTS, ARE IN ACCEPTABLE CONDITIONS AND FUNCTION CORRECTLY. BASED ON THE PHOTOS AND STATEMENTS PROVIDED, ALONG WITH THE MANUFACTURER'S INVESTIGATION, IT IS DETERMINED THAT THE FIRE ORIGINATED OUTSIDE THE DEVICE, CAUSED BY THE END USER OF THE DEVICE SMOKING WHILE USING THE DEVICE. THE NUVO LITE INSTRUCTIONS FOR USE (IFU) CLEARLY STATE THAT OXYGEN PROMOTES RAPID BURNING AND STRICTLY PROHIBIS SMOKING OR THE PRESENCE OF OPEN FLAMES NEAR THE DEVICE OR ITS ACCESSORIES. FURTHERMORE, THE IFU HIGHLIGHS THE RISKS ASSOCIATED WITH SMOKING WHILE USING THE CONCENTRATOR, CATEGORIZING IT AS A CONTRAINDICATION DUE TO THE INCREASED RISK OF FIRE. THE IFU ALSO DETAILS THE INCLUDSION OF SAFETY FEATURES SUCH AS A METAL FIRE BREAK AT THE OXYGEN OUTLET TO PREVENT FIRE FROM ENTERING THE DEVICE. NIDEK MEDICAL HAS DETERMINED THAT THE ROOT CAUSE OF THIS INCIDENT WAS USER NON-COMPLIANCE WITH THESE INSTRUCTIONS AND WARNINGS.
HEARTSTART XL+ DEFIBRILLATOR/MONITOR | K110825 | PHILIPS RECEIVED A COMPLAINT ON THE HEARTSTART XL+ DEFIBRILLATOR INDICATING NO RESPONSE TO DEFIBRILLATION FOR 30-MINUTE PAUSE. THE DEVICE WAS REPORTED TO BE IN USE ON A DEATH PATIENT. | PATIENT OUTCOME CODE HAS BEEN UPDATED. CORRECTION: THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS REMOTE SERVICE PERSONNEL AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. THE HEART HAD ARRESTED FOR 30 MINUTES BEFORE DEFIBRILLATION TREATMENT, THEREFORE IT IS NOT SUITABLE TO USE A DEFIBRILLATOR FOR RESUSCITATION TREATMENT. THE DEVICE WAS EVALUATED REMOTELY. THE DEFIBRILLATOR DID NOT HAVE ANY MALFUNCTION. THE DEVICE REMAINS AT THE CUSTOMER'S SITE AND NO FURTHER EVALUATION IS WARRANTED AT THIS TIME. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS USE ERROR. THE REPORTED PROBLEM WAS CONFIRMED. THE CUSTOMER WAS EXPLAINED THE PRODUCT FUNCTIONALITY. IT HAS BEEN CONCLUDED THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH MW5159309 WITH INFORMATION THAT A PATIENT PASSED AWAY WHILE ON A TRILOGY EVO VENTILATOR. A VENTILATOR DEPENDENT PATIENT WITH A TRACHEOSTOMY WAS IN STABLE CONDITION AND BEING SEEN BY A PULMONOLOGIST DURING A SOFTWARE UPDATE ON THE DEVICE. THE SETTINGS WERE CONFIRMED AND UNCHANGED. THE UPDATE WAS COMPLETE WITHOUT ANY CONCERN. THE PATIENT'S MOTHER CALLED THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH MW5159309 WITH INFORMATION THAT A PATIENT PASSED AWAY WHILE ON A TRILOGY EVO VENTILATOR. A VENTILATOR DEPENDENT PATIENT WITH A TRACHEOSTOMY WAS IN STABLE CONDITION AND BEING SEEN BY A PULMONOLOGIST DURING A SOFTWARE UPDATE ON THE DEVICE. THE SETTINGS WERE CONFIRMED AND UNCHANGED. THE UPDATE WAS COMPLETE WITHOUT ANY CONCERN. THE PATIENT'S MOTHER CALLED THE FOLLOWING MORNING TO REPORT THAT THE PATIENT PASSED AWAY UNEXPECTEDLY ON THE NIGHT OF THE SOFTWARE UPDATE. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED, ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED DEATH. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL REPORT WILL BE FILED. THE FOLLOWING MORNING TO REPORT THAT THE PATIENT PASSED AWAY UNEXPECTEDLY ON THE NIGHT OF THE SOFTWARE UPDATE. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED, ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED DEATH. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL REPORT WILL BE FILED.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED A VOLUNTARY MEDWATCH MW5159308 WITH INFORMATION A PATIENT PASSED AWAY WHILE ON A TRILOGY EVO DEVICE. ON (B)(6) 2024 A SOFTWARE UPDATE WAS PERFORMED ON THE DEVICE. THE SETTINGS OF SIMV-PC IPAP 13, EPAP 5, PS 13, RR 14, IT 1, RISE TIME 1 WERE CONFIRMED AND NOT CHANGED. THE UPDATE WAS COMPLETE WITHOUT ANY CONCERN AND THE PATIENT WAS IN STABLE CONDITION. A STAFF MEMBER WAS NOTIFIED THAT THE PATIENT PASSED AWAY UNEXPECTEDLY, MAY HAVE BEEN A HEART ATTACK, HEART STOPPED SUDDENLY. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED DEATH. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. UPON COMPLETION OF THE INVESTIGATION, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, DEATH AND OTHER. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE CUSTOMER (DISTRIBUTOR) DECLINED TO PROVIDE ADDITIONAL INFORMATION DURING A GOOD FAITH EFFORT ATTEMPT. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
FLAT CUT TALAR DOME SZ 2 INFINITY ADAPTIS | K191393 | AS REPORTED: "SUBJECT: (B)(6) INFINITY¿ WITH ADAPTIS¿ TOTAL ANKLE REPLACEMENT FOLLOW-UP (ITAR2) ON (B)(6) 2022: FLEXOR HALLUCIS LONGUS LACERATION. INTRA-OPERATIVE LACERATION. AFTER TIBIAL RESECTION COMPLETED AND TIBIAL RESECTION COMPLETED, FHL WAS NOTED TO HAVE SUSTAINED COMPLETE LACERATION AND WAS NOT REPAIRED. (B)(6) 2023: NO COMPLAINTS OR SYMPTOMS VOICED." | AFTER A COMPLETE REVIEW OF THE DATA DURING INVESTIGATION, THE DEVICE CATALOG # 33680032 HAS BEEN DEEMED CONCOMITANT. THEREFORE, MFR REPORT # 3010667733-2024-00719 IS NO LONGER APPLICABLE AND THE RECORD WILL CANCELED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K220414 | IT WAS REPORTED A PATIENT DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE INDICATED FOR A CASE OF ATRIAL FIBRILLATION (A FIB), A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR USE. THE PHYSICIAN SUCCESSFULLY COMPLETED THE TRANSSEPTAL PUNCTURE AND REMOVED THE VERSACROSS DILATOR FROM THE WATCHMAN SHEATH (ITS VALVE WAS OPENED). AT THIS TIME, THE PATIENT BEGAN EXPERIENCING AN ST SEGMENT ELEVATION AND BECAME HEMODYNAMICALLY UNSTABLE. THE PHYSICIAN GAVE THE PATIENT MEDICATION TO TREAT THE ST SEGMENT ELEVATION. THE PATIENT WENT INTO CARDIAC ARREST AND REQUIRED CARDIOPULMONARY RESUSCITATION (CPR). AN IC PHYSICIAN ARRIVED, AND FEMORAL ACCESS WAS OBTAINED AND A NON-BOSTON SCIENTIFIC 6FR DIAGNOSTIC CATHETER WAS INTRODUCED. AN ANGIOGRAM OF THE RIGHT CORONARY ARTERY SHOWED THAT THERE WAS AN AIR EMBOLISM PRESENT IN THE PATIENT. AIR WAS VISUALIZED DURING ASPIRATION OF THE SIDE PORT ON THE WATCHMAN SHEATH. THE PHYSICIAN USED A NON-BOSTON SCIENTIFIC CATHETER TO REMOVE THE AIR FROM THE PATIENT. AN INTRA-AORTIC BALLOON PUMP (IABP) WAS PLACED, AND THE PATIENT WAS SENT TO INTENSIVE CARE UNIT. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED THE NEXT DAY. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). NO AIR WAS VISUALIZED WAS WHEN REMOVING THE DILATOR. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, NOSE IRRITATION AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
OMNIPOD 5 POD | K203768 | IT HAS BEEN REPORTED BY THE PATIENT'S NEPHEW THAT THE PATIENT WAS TRANSPORTED TO HOSPITAL BY EMERGENCY MEDICAL SERVICES. THE PATIENT HAD HIGH BLOOD GLUCOSE (BG) LEVELS, ALTHOUGH NO SPECIFIC BG LEVELS WERE PROVIDED. THE PATIENT LATER PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH HAS NOT BEEN CONFIRMED AT THE TIME OF THIS REPORT. THE REPORTER STATES THE NURSE AT THE HOSPITAL TOLD HIM THE HIGH BG'S HAD LED TO OTHER COMPLICATIONS IN HER BODY. THIS IS REPORT 2 OF 2 CONCERNING PATIENT REFERENCE (B)(4). | THE PATIENT CONTACTED INSULET ON (B)(6) 2024 FOR SUPPORT (REF (B)(4)) ACTIVATING A POD AND REPORTED BG LEVELS OF 381 MG/DL. ON (B)(6) 2024, DURING AN INSULET INITIATED FOLLOW UP CALL, INSULET WAS MADE AWARE THAT THE PATIENT HAD SOUGHT MEDICAL TREATMENT FOR HYPERGLYCEMIA ON (B)(6) 2024 AND SUBSEQUENTLY EXPIRED IN THE HOSPITAL ON (B)(6) 2024. NO FURTHER INFORMATION IS AVAILABLE. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K150709 | IT WAS REPORTED A CARDIAC PERFORATION, PERICARDIAL EFFUSION AND DEATH OCCURRED. PROCEDURE WAS CANCELLED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT FOR LAAC KIT WAS SELECTED FOR USE. THE PHYSICIAN CROSSED THE SEPTUM MID ON THE BICAVAL VIEW AND MID-ANTERIOR ON THE SHORT ACCESS VIEW ON INTRACARDIAC ECHOCARDIOGRAPHY (ICE). WHEN PHYSICIAN BUZZED ACROSS, A SUCCESSFUL DEPLOYMENT OF THE VERSACROSS RF WIRE INTO THE LEFT ATRIUM (LA) OCCURRED. THE CATHETER COULD NOT BE VISUALIZED ON ICE HENCE FLUOROSCOPY IMAGING WAS CHECKED, AND THE WATCHMAN SHEATH AND VERSACROSS DILATOR WERE SEEN ALL THE WAY TO THE POSTERIOR SIDE ON THE BACK WALL OF THE LEFT ATRIUM AND THE VERSACROSS RF WIRE WAS HALFWAY INSIDE OF THE DILATOR TIP. AT THIS TIME, AN EFFUSION CHECK WAS PERFORMED USING TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). TEE IMAGING CONFIRMED THERE WAS A PERICARDIAL EFFUSION BELIEVED TO BE LOCATED AT THE POSTERIOR ASPECT OF THE LEFT ATRIUM. IT WAS SUSPECTED THAT THE VERSACROSS DILATOR WAS PUSHED THROUGH THE BACK WALL OF THE LEFT ATRIUM WHILE TRYING TO FLOSS ACROSS THE SEPTUM AS THE SEPTUM WAS SLIGHTLY ANEURISMAL. ONCE THE EFFUSION WAS CONFIRMED, THE PHYSICIAN TAPPED THE PERICARDIAL SPACE, REMOVED FLUIDS, AND GAVE BLOOD. SURGICAL BACKUP WAS CALLED AND APPROXIMATELY TEN TO TWENTY MINUTES LATER, SURGICAL BACKUP ARRIVED AND AT THIS TIME THE PATIENT HAD NO BLOOD CIRCULATING, AND NO CHEST COMPRESSIONS WERE DONE. THE PATIENT PASSED AWAY. IN THE PHYSICIAN'S OPINION, PERICARDIAL EFFUSION OCCURRED AFTER TRANSSEPTAL PUNCTURE WHICH CONTRIBUTE TO THE PATIENT COMPLICATIONS. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. NO PERICARDIAL EFFUSION WAS NOTED PRIOR TO THE PROCEDURE. IN THE PHYSICIAN'S OPINION, THE DEVICE HAS CONTRIBUTED TO THE EVENT, SINCE THE PERICARDIAL EFFUSION WAS AFTER TRANSSEPTAL PUNCTURE. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS DILATOR. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC YET. A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION IS IN PROGRESS, BUT IT WAS NOT AVAILABLE STILL.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K220414 | IT WAS REPORTED A CARDIAC PERFORATION, PERICARDIAL EFFUSION AND DEATH OCCURRED. PROCEDURE WAS CANCELLED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT FOR LAAC KIT WAS SELECTED FOR USE. THE PHYSICIAN CROSSED THE SEPTUM MID ON THE BICAVAL VIEW AND MID-ANTERIOR ON THE SHORT ACCESS VIEW ON INTRACARDIAC ECHOCARDIOGRAPHY (ICE). WHEN PHYSICIAN BUZZED ACROSS, A SUCCESSFUL DEPLOYMENT OF THE VERSACROSS RF WIRE INTO THE LEFT ATRIUM (LA) OCCURRED. THE CATHETER COULD NOT BE VISUALIZED ON ICE HENCE FLUOROSCOPY IMAGING WAS CHECKED, AND THE WATCHMAN SHEATH AND VERSACROSS DILATOR WERE SEEN ALL THE WAY TO THE POSTERIOR SIDE ON THE BACK WALL OF THE LEFT ATRIUM AND THE VERSACROSS RF WIRE WAS HALFWAY INSIDE OF THE DILATOR TIP. AT THIS TIME, AN EFFUSION CHECK WAS PERFORMED USING TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). TEE IMAGING CONFIRMED THERE WAS A PERICARDIAL EFFUSION BELIEVED TO BE LOCATED AT THE POSTERIOR ASPECT OF THE LEFT ATRIUM. IT WAS SUSPECTED THAT THE VERSACROSS DILATOR WAS PUSHED THROUGH THE BACK WALL OF THE LEFT ATRIUM WHILE TRYING TO FLOSS ACROSS THE SEPTUM AS THE SEPTUM WAS SLIGHTLY ANEURISMAL. ONCE THE EFFUSION WAS CONFIRMED, THE PHYSICIAN TAPPED THE PERICARDIAL SPACE, REMOVED FLUIDS, AND GAVE BLOOD. SURGICAL BACKUP WAS CALLED AND APPROXIMATELY TEN TO TWENTY MINUTES LATER, SURGICAL BACKUP ARRIVED AND AT THIS TIME THE PATIENT HAD NO BLOOD CIRCULATING, AND NO CHEST COMPRESSIONS WERE DONE. THE PATIENT PASSED AWAY. IN THE PHYSICIAN'S OPINION, PERICARDIAL EFFUSION OCCURRED AFTER TRANSSEPTAL PUNCTURE WHICH CONTRIBUTE TO THE PATIENT COMPLICATIONS. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. NO PERICARDIAL EFFUSION WAS NOTED PRIOR TO THE PROCEDURE. IN THE PHYSICIAN'S OPINION, THE DEVICE HAS CONTRIBUTED TO THE EVENT, SINCE THE PERICARDIAL EFFUSION WAS AFTER TRANSSEPTAL PUNCTURE. IT WAS FURTHER REPORTED THAT THE WATCHMAN TRUSTEER WAS USED WITH THE VERSACROSS RF WIRE. THE PHYSICIAN WAS ADVANCING THE DILATOR WITHOUT VERSACROSS RF WIRE FULLY DEPLOYED. | CORRECTION TO THE INITIAL MDR: THE FIELDS B5 (DESCRIBE EVENT OR PROBLEM), D3 (MANUFACTURER NAME), D4 (MODEL NUMBER, LOT NUMBER, CATALOG NUMBER, EXPIRATION DATE AND UNIQUE IDENTIFIER (UDI), AND G1 (MFR SITE FACILITY NAME, MFR SITE ADDRESS 1, MFR SITE CITY, MFR SITE ZIP/POST CODE, MFR SITE COUNTRY) WERE UPDATED. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WHO HAS THE MACHINE HAS PASSED AWAY. CALLER WANTED TO RETURN THE MACHINE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED TO FRESENIUS THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., DISCHARGE SUMMARY, DEATH CERTIFICATE, TREATMENT RECORDS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. ADDITIONALLY, NO ON-SITE EVALUATION WAS PERFORMED BY A FRESENIUS FIELD SERVICE TECHNICIAN (FST). A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER DUE TO HEART FAILURE. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PDRN STATED THAT THE PATIENT LIVED WITH THE GRANDDAUGHTER. THE GRANDDAUGHTER CONTACTED THE PDRN AND REPORTED THAT ON THE MORNING OF (B)(6) 2024 911 WAS CALLED. THE REASON FOR CALLING 911 WAS NOT PROVIDED TO THE PDRN. THE PATIENT WAS IN TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER AT THE TIME OF THE EVENT (UNKNOWN PHASE AND CYCLE). PARAMEDICS ARRIVED AND DISCONNECTED THE PATIENT FROM TREATMENT. THE PATIENT WAS TRANSPORTED TO THE HOSPITAL AND WAS LATER PRONOUNCED DECEASED. THE GRANDDAUGHTER STATED THE PATIENT DID NOT HAVE A DO NOT RESUSCITATE (DNR) ORDER AT THE TIME OF THE EVENT; HOWEVER, THE PATIENT¿S CARDIOLOGIST ADVISED HER HEART WAS NOT WORKING ANYMORE AND IT WAS JUST A MATTER OF TIME BEFORE SHE PASSED. THE PATIENT¿S FAMILY SAID HEART FAILURE WAS THE CAUSE OF DEATH. THE PATIENT REPORTEDLY WAS NOT STRONG ENOUGH TO SURVIVE ANY SURGERY. THE PATIENT HAD TWO PRIOR HEART VALVES OPERATED ON WHEN SHE WAS YOUNGER. THERE WERE NO REPORTED ISSUES WITH THE LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS AVAILABLE. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER WITH NO PHYSICAL DAMAGE NOTED. THERE WAS NO EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT AND THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. A SIMULATED TREATMENT WITH REDUCED DWELL WAS PERFORMED ON THE CYCLER AND PASSED. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST AND VALVE ACTUATION TEST. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
SENDIT DELIVERY DEVICE | K211654 | THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE IN THE L M1 SEGMENT OF THE MIDDLE CEREBRAL ARTERY (MCA) USING A PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER (RED72), A SENDIT DELIVERY DEVICE (SENDIT), A BENCHMARK BMX96 ACCESS SYSTEM (BMX96), A PENUMBRA SYSTEM RED 68 REPERFUSION CATHETER (RED68), A PENUMBRA SYSTEM RED 43 REPERFUSION CATHETER (RED43), A DIAGNOSTIC CATHETER (5F), GUIDEWIRES, A CLOSURE DEVICE, AND A SHEATH (8FR) VIA THE RIGHT GROIN. DURING THE PROCEDURE, THE RED72 WAS POSITIONED IN THE PROXIMAL L M1 SEGMENT (PROXIMAL TO THE CLOT) OVER THE SENDIT AND GUIDEWIRE. THE SENDIT AND GUIDEWIRE WERE THEN WITHDRAWN. THE PHYSICIAN ATTEMPTED TO ADVANCE THE RED72 INTO THE CLOT; HOWEVER, SIGNIFICANT RESISTANCE WAS ENCOUNTERED, AND THE CLOT COULD NOT BE REACHED. NEXT, THE RED72 WAS WITHDRAWN INTO THE DISTAL L SUPRACLINOID INTERNAL CAROTID ARTERY (ICA). A ROADMAP ANGIOGRAPHY WAS PERFORMED WHICH DEMONSTRATED NEW OCCLUSION OF THE DISTAL L POSTERIOR COMMUNICATING ARTERY (PCA). THE RED72 WAS ADVANCED TO THE CLOT OVER A GUIDEWIRE INTO THE DISTAL L MCA. TWO PASSES WERE COMPLETED. A CONTROL RUN DEMONSTRATED NEAR-COMPLETE REVASCULARIZATION OF THE L M2 INFERIOR DIVISION WITH PERSISTENT OCCLUSION OF THE L M2 SUPERIOR DIVISION. UNDER ROADMAP ANGIOGRAPHY, A RED68 WAS POSITIONED IN THE L M2 SUPERIOR DIVISION OVER A RED43 AND GUIDEWIRE. THE RED43 AND GUIDEWIRE WERE REMOVED, AND THE RED68 WAS ADVANCED INTO THE CLOT. TWO PASSES WERE COMPLETED. CONTROL RUNS CONFIRMED COMPLETE REVASCULARIZATION WITH THROMBOLYSIS IN CEREBRAL INFARCTION (TICI 2B) WITH FILLING OF BOTH SUPERIOR AND INFERIOR DIVISIONS. IT WAS NOTED THAT A PERFUSION DEFECT IN THE REGION OF PREDICTED CORE AND A SMALL, SUSPECTED, NON-FLOW-LIMITING DISSECTION FLAP IN THE PROXIMAL L M1 SEGMENT (IN THE REGION OF PRIOR RESISTANCE OF THE RED72) AND NEW CLOT OR INJURY IN THE L PCA (PROXIMAL TO THE DISTAL OCCLUSION) WERE CONFIRMED WITH ANGIOGRAPHY. SUBSEQUENT RUNS DEMONSTRATED A WORSENING APPEARANCE OF THE L M1 WITH A SUSPECTED DISSECTION, RE-OCCLUSION OF THE L M2 (SUPERIOR DIVISION), AND COMPLETE OCCLUSION OF THE FETAL L PCA FROM ITS ICA ORIGIN. THE ORIGINAL OCCLUSION SPONTANEOUSLY RECANALIZED ON SUBSEQUENT RUNS. A VASCULAR NEUROLOGIST WAS CONSULTED, AND A DECISION WAS MADE TO PLACE A STENT IN THE L M1 SEGMENT TO THE L SUPRACLINOID ICA. ALTHOUGH THE STENT OPENED THE VESSEL, THE L M1 SEGMENT OCCLUDED COMPLETELY (TICI 0). THE USE OF ANTIPLATELET AGENTS WAS CONSIDERED AND DISCUSSED WITH THE VASCULAR NEUROLOGIST. DUE TO THE SIZE OF THE PATIENT¿S STROKE, THE PHYSICIANS DECIDED TO DEFER ANTIPLATELETS UNTIL REPEAT IMAGING THE NEXT DAY. A DYNACT WAS PERFORMED AND CONFIRMED NO EVIDENCE OF A SIGNIFICANT HEMORRHAGE. THE PROCEDURE ENDED AT THIS POINT. ALL DEVICES WERE REMOVED, AND THE ACCESS SITE WAS SEALED WITH A CLOSURE DEVICE. THERE WERE NO APPARENT PERI-PROCEDURAL COMPLICATIONS, AND THE PATIENT WAS TRANSFERRED TO THE NEURO ICU INTUBATED BUT IN STABLE CONDITION. IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2024. ON 27-SEP-2024, ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEFT M1 DISSECTION WAS REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE PENUMBRA SYSTEM AND A DEFINITE RELATIONSHIP TO THE INDEX PROCEDURE. ON 27-SEP-2024, ADDITIONAL INFORMATION WAS RECEIVED THAT THE NEW THROMBOEMBOLISM TO THE FETAL L PCA WAS REPORTED TO BE A NON-SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE PENUMBRA SYSTEM AND A PROBABLE RELATIONSHIP TO THE INDEX PROCEDURE. | THE PRODUCT WAS NOT RETURNED FOR EVALUATION. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS. THIS REPORT IS ASSOCIATED WITH MFR REPORT NUMBER: 1. 3005168196-2024-00338.
PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER | K211654 | THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE IN THE L M1 SEGMENT OF THE MIDDLE CEREBRAL ARTERY (MCA) USING A PENUMBRA SYSTEM RED 72 REPERFUSION CATHETER (RED72), A SENDIT DELIVERY DEVICE (SENDIT), A BENCHMARK BMX96 ACCESS SYSTEM (BMX96), A PENUMBRA SYSTEM RED 68 REPERFUSION CATHETER (RED68), A PENUMBRA SYSTEM RED 43 REPERFUSION CATHETER (RED43), A DIAGNOSTIC CATHETER (5F TEMPO), GUIDEWIRES, A CLOSURE DEVICE AND A SHEATH (8FR) VIA THE RIGHT GROIN. DURING THE PROCEDURE, THE RED72 WAS POSITIONED IN THE PROXIMAL L M1 SEGMENT (PROXIMAL TO THE CLOT) OVER THE SENDIT AND GUIDEWIRE. THE SENDIT AND GUIDEWIRE WERE THEN WITHDRAWN. THE PHYSICIAN ATTEMPTED TO ADVANCE THE RED72 INTO THE CLOT; HOWEVER, SIGNIFICANT RESISTANCE WAS ENCOUNTERED, AND THE CLOT COULD NOT BE REACHED. NEXT, THE RED72 WAS WITHDRAWN INTO THE DISTAL L SUPRACLINOID INTERNAL CAROTID ARTERY (ICA). A ROADMAP ANGIOGRAPHY WAS PERFORMED WHICH DEMONSTRATED NEW OCCLUSION OF THE DISTAL L POSTERIOR COMMUNICATING ARTERY (PCA). THE RED72 WAS ADVANCED TO THE CLOT OVER A GUIDEWIRE INTO THE DISTAL L MCA. TWO PASSES WERE COMPLETED. A CONTROL RUN DEMONSTRATED NEAR-COMPLETE REVASCULARIZATION OF THE L M2 INFERIOR DIVISION WITH PERSISTENT OCCLUSION OF THE L M2 SUPERIOR DIVISION. UNDER ROADMAP ANGIOGRAPHY, A RED68 WAS POSITIONED IN THE L M2 SUPERIOR DIVISION OVER A RED43 AND GUIDEWIRE. THE RED43 AND GUIDEWIRE WERE REMOVED, AND THE RED68 WAS ADVANCED INTO THE CLOT. TWO PASSES WERE COMPLETED. CONTROL RUNS CONFIRMED COMPLETE REVASCULARIZATION WITH THROMBOLYSIS IN CEREBRAL INFARCTION (TICI 2B) WITH FILLING OF BOTH SUPERIOR AND INFERIOR DIVISIONS. IT WAS NOTED THAT A PERFUSION DEFECT IN THE REGION OF PREDICTED CORE AND A SMALL, SUSPECTED, NON-FLOW-LIMITING DISSECTION FLAP IN THE PROXIMAL L M1 SEGMENT (IN THE REGION OF PRIOR RESISTANCE OF THE RED72) AND NEW CLOT OR INJURY IN THE L PCA (PROXIMAL TO THE DISTAL OCCLUSION) WERE CONFIRMED WITH ANGIOGRAPHY. SUBSEQUENT RUNS DEMONSTRATED A WORSENING APPEARANCE OF THE L M1 WITH A SUSPECTED DISSECTION, RE-OCCLUSION OF THE L M2 (SUPERIOR DIVISION), AND COMPLETE OCCLUSION OF THE FETAL L PCA FROM ITS ICA ORIGIN. THE ORIGINAL OCCLUSION SPONTANEOUSLY RECANALIZED ON SUBSEQUENT RUNS. A VASCULAR NEUROLOGIST WAS CONSULTED, AND A DECISION WAS MADE TO PLACE A STENT IN THE L M1 SEGMENT TO THE L SUPRACLINOID ICA. ALTHOUGH THE STENT OPENED THE VESSEL, THE L M1 SEGMENT OCCLUDED COMPLETELY (TICI 0). THE USE OF ANTIPLATELET AGENTS WAS CONSIDERED AND DISCUSSED WITH THE VASCULAR NEUROLOGIST. DUE TO THE SIZE OF THE PATIENT¿S STROKE, THE PHYSICIANS DECIDED TO DEFER ANTIPLATELETS UNTIL REPEAT IMAGING THE NEXT DAY. A DYNACT WAS PERFORMED AND CONFIRMED NO EVIDENCE OF A SIGNIFICANT HEMORRHAGE. THE PROCEDURE ENDED AT THIS POINT. ALL DEVICES WERE REMOVED, AND THE ACCESS SITE WAS SEALED WITH A CLOSURE DEVICE. THERE WERE NO APPARENT PERI-PROCEDURAL COMPLICATIONS, AND THE PATIENT WAS TRANSFERRED TO THE NEURO ICU INTUBATED BUT IN STABLE CONDITION. IT WAS REPORTED THAT THE PATIENT EXPIRED ON (B)(6) 2024. ON 27-SEP-2024, ADDITIONAL INFORMATION WAS RECEIVED THAT THE LEFT M1 DISSECTION WAS REPORTED TO BE A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE PENUMBRA SYSTEM AND A DEFINITE RELATIONSHIP TO THE INDEX PROCEDURE. ON 27-SEP-2024, ADDITIONAL INFORMATION WAS RECEIVED THAT THE NEW THROMBOEMBOLISM TO THE FETAL L PCA WAS REPORTED TO BE AN ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE PENUMBRA SYSTEM AND A PROBABLE RELATIONSHIP TO THE INDEX PROCEDURE. ON 12-OCT-2024, ADDITIONAL INFORMATION WAS RECEIVED THAT A 24-HOUR COMPUTED TOMOGRAPHY (CT) REVEALED A LARGE EVOLVING ACUTE LEFT MCA AND PCA TERRITORY INFARCTION WITH CYTOTOXIC EDEMA. THE CEC CHAIR ADJUDICATED EVOLVING ACUTE LEFT MCA AND PCA TERRITORY INFARCTION WITH CYTOTOXIC EDEMA AS A SERIOUS ADVERSE EVENT WITH A PROBABLE RELATIONSHIP TO THE INDEX STROKE AND A POSSIBLE RELATIONSHIP TO THE INDEX PROCEDURE AND A POSSIBLE RELATIONSHIP TO THE PENUMBRA SYSTEM. ON 18-OCT-2024, ADDITIONAL INFORMATION WAS RECEIVED THAT THE CEC CHAIR ADJUDICATED THE NEW THROMBOEMBOLISM TO THE FETAL L PCA AS A SERIOUS ADVERSE EVENT WITH A POSSIBLE RELATIONSHIP TO THE INDEX STROKE, A PROBABLE RELATIONSHIP TO THE INDEX PROCEDURE, AND A POSSIBLE RELATIONSHIP TO THE PENUMBRA SYSTEM. | PLEASE NOTE THAT THE FOLLOWING SECTIONS ARE BEING UPDATED BASED ON ADDITIONAL INFORMATION PROVIDED BY THE PENUMBRA CLINICAL TEAM ON 10/12/2024: 1. SECTION B. BOX 5. DESCRIBE EVENT OR PROBLEM. 2. SECTION H. BOX 6. PATIENT CODE 1, 2, AND 3 H3 OTHER TEXT : PLACEHOLDER.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED MEDIASTINAL MASS RESECTION (THYMECTOMY), ABOUT 2 HOURS INTO THE PROCEDURE, WHILE THE SURGEON WAS DISSECTING TISSUE A MAJOR VESSEL WAS INJURED AND BLEEDING OCCURRED; THE DIRECT CAUSE OF THE INJURY IS UNKNOWN. THE INTUITIVE CLINICAL TERRITORY ASSOCIATE WAS INFORMED BY THE ROBOTICS COORDINATOR THAT BLOOD SQUIRTED AND COVERED THE ENDOSCOPE LENS, OBSTRUCTING VISION. THE PROCEDURE WAS CONVERTED TO AN OPEN PROCEDURE TO ADDRESS THE BLEEDING. AT SOME POINT IN TIME DURING THE PROCEDURE THE PATIENT CODED. THE PATIENT EXPIRED ON THE OPERATING ROOM TABLE ABOUT ONE HOUR INTO THE UNSPECIFIED INTERVENTIONS. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS (RECOVERABLE OR NON-RECOVERABLE FAULTS) OCCURRED DURING THOSE PROCEDURES. A REVIEW OF THE INSTRUMENT LOG FOUND THAT THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, CADIERE FORCEPS, LONG BIPOLAR GRASPER, TIP-UP FENESTRATED GRASPER, SUREFORM STAPLER 60, SYNCHROSEAL, SUCTION IRRIGATOR. ALL MULTI-USE INSTRUMENTS USED IN THE CASE HAVE BEEN USED IN SUBSEQUENT PROCEDURES. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THESE INSTRUMENTS. REVIEW OF THE DEVICE HISTORY RECORDS (DHR) FOUND THERE WERE NO NON-CONFORMANCES IDENTIFIED TO BE RELATED TO THE DEVICES USED DURING THE EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT WAS UNDERGOING A ROBOTIC THYMECTOMY WHEN THEY GOT INTO BLEEDING. THERE ARE SEVERAL MAJOR VESSELS IN CLOSE PROXIMITY TO THE THYMUS. WHICH VESSEL WAS INVOLVED AND HOW THE SURGEON GOT INTO THE BLEEDING WAS NOT PROVIDED. THE CASE WAS CONVERTED TO OPEN BUT THE PATIENT DIED DURING THE PROCEDURE. NO OTHER DETAILS AS TO HOW THE INJURY OCCURRED NOR THE CAUSE OF DEATH WERE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE CUSTOMER (DISTRIBUTOR) DECLINED TO PROVIDE ADDITIONAL INFORMATION DURING A GOOD FAITH EFFORT ATTEMPT. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR W/SD CARD | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY MISSED INFORMATION IN SECTION E COUNTRY IS NOT ADDED. AFTER REVIEW, IT WAS DETERMINED THAT SECTION E SHOULD BE FILED WITH COUNTRY NAME. SECTION E IS CORRECTED IN THIS REPORT.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED
REMSTAR AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING CANCER AND DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING CANCER AND DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S QUALITY PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. THE MANUFACTURE INSPECTED EXTERNALLY OBSERVED THAT AN UNKNOWN DUST-LIKE CONTAMINANT INSIDE THE UI DISPLAY AND INSIDE THE AIR INLET. THE BOTTOM OF THE DEVICE WAS OBSERVED TO HAVE A SLICE ACROSS THE ENCLOSURE AND THE WARNING LABEL. A BLUISH-GRAY HARD UNKNOWN SUBSTANCE WAS OBSERVED ON THE EXTERIOR OF THE BOTTOM ENCLOSURE AND ALONG THE EDGES OF THE WARNING LABEL. AN UNKNOWN DUST-LIKE CONTAMINANT WAS OBSERVED ON THE INTERIOR OF THE TOP COVER, ALONG WITH WHAT APPEARED TO BE BURN MARKS. THE MANUFACTURER CONCLUDES THERE WAS SOUND ABATEMENT FOAM WAS OBSERVED TO HAVE DUST-LIKE CONTAMINANT ON IT, AND EVIDENCE OF DEGRADATION/BREAKDOWN WAS CONFIRMED WITH IT LEAVING SMALL PIECES ALONG THE EDGES OF THE BOTTOM OF THE BLOWER HOUSING BOTTOM AND WAS EASILY PULLING APART UPON PICKING IT UP OUT OF THE ENCLOSURE. IN THIS REPORT DEVICE EVALUATION HAS BEEN UPDATED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
BIPAP PRO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REP DREAMSTATION AUTO BIPAP. THE PATIENT PASSED AWAY 2 YEARS AGO. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED RECEIVING INFORMATION ALLEGING A REP DREAMSTATION AUTO BIPAP. THE PATIENT PASSED AWAY 2 YEARS AGO. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IN THE PREVIOUS REPORT, SECTION D: MODEL NUMBER, CATALOG ITEM IDENTIFIER, SERIAL NUMBER AND PRODUCT UDI WAS UPDATED/CORRECTED IN THIS REPORT.
SMART PORT CT | K062414 | ON OR ABOUT (B)(6) 2022, PLAINTIFF'S DECEDENT UNDERWENT PLACEMENT OF AN ANGIODYNAMICS SMARTPORT PRODUCT, REFERENCE NUMBER H787CT80STPD0, LOT NUMBER 5733844. THE DEVICE WAS IMPLANTED BY DR. (B)(6), M.D., (B)(6) HOSPITAL IN (B)(6), FOR ADMINISTRATION OF CHEMOTHERAPY. DECEDENT PRESENTED TO (B)(6) HOSPITAL WITH COMPLAINTS OF FEVER AND WEAKNESS. IT WAS DETERMINED BY DECEDENT'S MEDICAL TEAM THAT HE HAD AN INFECTION CAUSED BY THE SMARTPORT AND THAT THE DEFECTIVE DEVICE HAD TO BE REMOVED. DECEDENT'S DEFECTIVE SMARTPORT WAS REMOVED. ON OR ABOUT (B)(6), 2022, DECEDENT PASSED AWAY. THE DEATH CERTIFICATE LISTS SEPSIS AS THE CAUSE OF DEATH. | THE REPORTED DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4).
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REMSTAR PRO C-FLEX+ DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REMSTAR PRO C-FLEX+ DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED THAT THE DEVICE WILL BE AVAILABLE FOR EVALUATION, BUT IT HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. AFTER DEVICE RETURN AND COMPLETION OF EVALUATION, A FOLLOW-UP REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, DEATH, KIDNEY DAMAGE, LUNG DAMAGE AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
AVALON FM30 FETAL MONITOR | K140535 | IT WAS REPORTED THE MATERNAL HEART RATE WAS INTERPRETED AS THE FETAL HEART RATE. THE FETAL MONITOR WAS CONNECTED AS THE PATIENT WAS WAITING FOR DELIVERY; HOWEVER, THE MATERNAL HEART RATE AND BLOOD PRESSURE WAS NOT MEASURED USING THE SAME SYSTEM. THE MONITOR RECORDING PAPER (CTG) DISPLAYED 'CCV ? ON ALL MEASUREMENTS AND ALARM WAS GENERATED VIA CTG MEASUREMENTS. A CHECK BY THE PHYSICIAN IN THE MORNING INDICATED THE FETAL HEART RATE HAD DROPPED SO THE PATIENT WAS SCHEDULED FOR A C-SECTION AT 3PM. AN ULTRASOUND WAS PERFORMED AT 2:45 PM, REVEALING NO FETAL HEART RATE. THE BABY WAS DELIVERED VIA C-SECTION AND RESUSCITATION WAS PERFORMED FOR MORE THAN 20 MINUTES; HOWEVER, THE RESUSCITATION WAS UNSUCCESSFUL. IT WAS DETERMINED THE BABY HAD PASSED PRIOR TO DELIVERY. THE STAFF WAS AWARE OF THE CCV (COINCIDENCE) FUNCTIONALITY, BUT THE MOTHER WAS NOT MONITORED WITH THE FETAL MONITOR AND THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. | IT WAS STATED THAT THE MONITOR RECORDING PAPER (CTG) DISPLAYED 'CCV ? ON ALL MEASUREMENTS AND ALARM WAS GENERATED VIA CTG MEASUREMENTS. A CHECK BY THE PHYSICIAN IN THE MORNING INDICATED THE FETAL HEART RATE HAD DROPPED SO THE PATIENT WAS SCHEDULED FOR A C-SECTION AT 3PM. AN ULTRASOUND WAS PERFORMED AT 2:45 PM, REVEALING NO FETAL HEART RATE. THE BABY WAS DELIVERED VIA C-SECTION AND RESUSCITATION WAS PERFORMED FOR MORE THAN 20 MINUTES; HOWEVER, THE RESUSCITATION WAS UNSUCCESSFUL. IT WAS DETERMINED THE BABY HAD PASSED PRIOR TO DELIVERY. THE STAFF WAS AWARE OF THE CCV (COINCIDENCE) FUNCTIONALITY, BUT THE MOTHER WAS NOT MONITORED WITH THE FETAL MONITOR. THE DISTRIBUTOR (B)(4) INTERVIEWED THE CUSTOMER AND PROVIDE THE INFORMATION AND SPECIFICATION REGARDING THE FM20/30 MONITOR, THE CUSTOMER ACKNOWLEDGED THE INCIDENT WAS NOT RELATED TO AVALON FM MONITOR. THE CUSTOMER STATED THAT THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION. BASED ON THE INFORMATION IT HAS INDICATED THE PATIENT DEATH WAS NOT RELATED TO THE PHILIPS DEVICE, AS STATED BY THE HEALTHCARE SAFETY MANAGER. THE CUSTOMER WAS PROVIDED WITH INFORMATION REGARDING THE FM MONITOR AND SPECIFICATIONS. NO FURTHER INFORMATION WAS PROVIDED; THEREFORE, THE DEVICE OPERATION WAS NOT ABLE TO BE CONFIRMED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. DUE TO THE FACT THAT THIS DEVICE DIDN'T CONTRIBUTE TO THE DEATH, THIS WOULD NOT BE A REPORTABLE EVENT WITH PHILIPS BASED ON WHEN A PARAMETER FAILS TO MEASURE THE DEVICE IS DESIGNED TO GENERATE AND ANNUNCIATE AUDIBLE AND VISUAL PATIENT ALARMS TO ALERT THE STAFF TO A CHANGE IN THE PATIENT CONDITION. THIS WOULD BE IMMEDIATELY DETECTABLE TO USERS DURING CLOSE OBSERVATION AS DESCRIBED IN THE PRODUCT LABELING, SHIPPED WITH THE DEVICE, AND WOULD PROMPT THE CLINICAL STAFF TO ATTEND MORE CLOSELY TO THE PATIENT. THIS WOULD ALLOW THE USER TO IMPLEMENT ALTERNATIVE METHODS OF MONITORING PER HOSPITAL PROTOCOL, AND/OR TO TROUBLESHOOT THE MONITOR IF NECESSARY. NOT REPORTABLE.
RECERT, DREAMSTATION AUTO, FR | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THE PATIENT REPRESENTATIVE ALLEGES THAT THE DECEASED WAS NOT USING HIS CPAP. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTARAUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | ADDITIONAL INFORMATION WAS RECEIVED, AND SECTION H 11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WERE TWO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
BIPAP PRO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO KIDNEY DISEASE/TOXICITY AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
BIPAP AUTO BI FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH, OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP PRO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
BIPAP AUTO SV ADV SYSTEM ONE | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, SARCOMA CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, HYPERSENSITIVITY, ASTHMA (NEW OR WORSENING), CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE /TOXICITY, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP PRO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY; DEATH THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
MX40 PATIENT WEARABLE MONITOR | K113125 | IT WAS REPORTED THERE WAS NO TELEMETRY DATA FOR TWO HOURS FOR A POST-SURGICAL PACEMAKER IMPLANT TELEMETRY PATIENT WHO PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | ADDITIONAL INFORMATION WAS REVIEWED BY THE PRODUCT SUPPORT ENGINEER AND BASED ON THIS INFORMATION, THERE WAS NO DEVICE MALFUNCTION WHICH CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. IT REMAINS UNKNOWN WHAT FACTORS CAUSED OR CONTRIBUTED TO THE DEATH. FUNCTIONAL TESTING/SERVICE REPAIR/TECHNICAL INVESTIGATION: THE FOLLOWING FUNCTIONAL TESTS WERE PERFORMED: THE REMOTE SERVICE ENGINEER COLLECTED THE LOG FILES AND PERFORMED A FIRST ANALYSIS WITH THE CUSTOMER BIOMED. THE ANALYSIS DONE CONFIRMED THERE ARE SUCCESSIONS OF ALARMS "ECG FAULT CONTACT", WHICH IS WHY THE ECG COULD NOT BE RECORDED DURING THIS PERIOD. IN ADDITION, THERE IS ALSO A "TRANSMITTER DEACTIVATED" ALARM, WHICH IS THE CONSEQUENCE OF THE SUCCESSION OF ECG MESSAGES FOR 10MIN. IN SUMMARY, THE TELEMETRY DID TRANSMIT DATA TO THE CONTROL PANEL, BUT THE CONDITION (ECG FAULT CONTACT) MEANT THAT NO ECG TRACE COULD BE RECORDED. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION BY THE PRODUCT SUPPORT ENGINEER (PSE) AND THE RESULTS FROM LOGFILE REVIEW INDICATE THAT BETWEEN 9:10 AND 9:15 AS WELL AS BETWEEN 9:26 AND 10:09, THERE WAS NO CONNECTION OF THE MX40 TO THE CENTRAL STATION. THE PSE CONFIRMED THE MX40 WAS CONNECTED TO THE CENTRAL STATION STARTING (B)(6) 2024 23:41 TO (B)(6) 2024 9:10. THEN THE MX40 WAS PUT TO STANDBY AT 9:10 AND WAS TAKEN OUT OF STANDBY AT 9:15, SO THE TIME THE MX40 WAS NOT ONLINE WAS CAUSED BY A USER PLACING THE MX40 IN STANDBY. AT 9:26 THE MX40 WAS AGAIN PUT INTO STANDBY. AT 10:09 THE MX40 WAS TAKEN OUT OF STANDBY (BY REMOVING AND REINSERTING THE BATTERY). THIS IS DOCUMENTED BY A ¿REBOOT¿. A NORMAL ¿STANDBY LEFT¿ IS EITHER INITIATED FROM THE CENTRAL OR THE MX40 ITSELF AND IS DOCUMENTED WITH A ¿RESTART¿ AS SHOWN AT 9:15. SO ACCORDING TO THE LOGS, THE USER PUT THE MX40 IN STANDBY MODE BETWEEN 9:10 TO 9:15 AND BETWEEN 9:26 TO 10:09 ALREADY STARTING AT 8:56 THERE HAVE BEEN SEVERAL INOP¿S SHOWING THAT THE ECG LEADS ARE DETACHED. THIS WAS ANNOUNCED BY A TECHNICAL INOP AT THE CENTRAL AND THIS INOP IS ALSO CONFIGURED AS YELLOW ALARM (CAN BE IDENTIFIED BY THE TWO !! IN FRONT OF THE ALARM MESSAGE). IN ADDITION, THE ALARMS HAVE BEEN SILENCED SEVERAL TIMES FROM THE CENTRAL ¿ECG LEADS OFF ENDED¿, AND THEN THE SYSTEM GENERATED NEW ALARMS/INOP¿S. SO TECHNICAL AND YELLOW ALARMS HAVE BEEN ANNOUNCED FROM 8:56 UNTIL 9:10 (WHEN THE MX40 WAS PUT TO STANDBY). OVERALL THE MX40 WAS WORKING CORRECTLY AND ANNOUNCED THE ALARMS AND THE INOP¿S (ECG LEADS OFF). FINALLY, THE MX40 WAS SWITCHED TO STANDBY, WHERE NO MONITORING IS POSSIBLE. THE INVESTIGATION DETERMINED THAT THE DEVICE FUNCTIONED AS DESIGNED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS DUE TO USER KNOWLEDGE. THE REPORTED PROBLEM WAS CONFIRMED THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. IT REMAINS UNKNOWN WHAT FACTORS CAUSED OR CONTRIBUTED TO THE DEATH. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTOSV ADV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
WATCHMAN TRUSTEER ACCESS SYSTEM | K240018 | IT WAS REPORTED A PERICARDIAL EFFUSION AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A WATCHMAN LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. PHYSICIAN WAS USING INTRACARDIAC ECHO (ICE) TO PERFORM THE TRANSEPTAL PUNCTURE WITH A VERSACROSS CONNECT AND THE WATCHMAN TRUSTEER SHEATH. THE PHYSICIAN CROSSED THE SEPTUM MID ON THE BICAVAL VIEW AND MID-ANTERIOR ON THE SHORT ACCESS VIEW ON ICE. WHEN PHYSICIAN BUZZED ACROSS, A SUCCESSFUL DEPLOYMENT OF THE PIGTAIL VERSACROSS WIRE INTO THE LEFT ATRIUM (LA) OCCURRED. THE CATHETER COULD NOT BE VISUALIZED ON ICE, BUT THE WAS SHEATH WAS VIEWED. FLUOROSCOPY IMAGING WAS CHECKED, AND THE WAS SHEATH AND DILATOR WERE SEEN ALL THE WAY TO THE POSTERIOR SIDE ON THE BACK WALL OF THE LEFT ATRIUM AND THE WIRE WAS HALFWAY INSIDE OF THE DILATOR TIP. AT THIS TIME AN EFFUSION CHECK WAS PERFORMED USING TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). TEE IMAGING CONFIRMED THERE WAS A PERICARDIAL EFFUSION. IT WAS SUSPECTED THAT THE DILATOR WAS PUSHED THROUGH THE BACK WALL OF THE LEFT ATRIUM WHILE TRYING TO FLOSS ACROSS THE SEPTUM AS THE SEPTUM WAS SLIGHTLY ANEURISMAL. ONCE THE EFFUSION WAS CONFIRMED, THE PHYSICIAN TAPPED THE PERICARDIAL SPACE, REMOVED FLUIDS, AND GAVE BLOOD. SURGICAL BACKUP WAS CALLED. APPROXIMATELY 10-20 MINUTES LATER, SURGICAL BACKUP ARRIVED AND AT THIS TIME THE PATIENT HAD NO BLOOD CIRCULATING, AND NO CHEST COMPRESSIONS WERE DONE. THE PATIENT PASSED AWAY.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING) LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REM AUTO A-FLEX | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REM AUTO A-FLEX SYSONE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS ELDERLY PATIENT ON PERITONEAL DIALYSIS (PD) EXPIRED ON (B)(6) 2024 WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN FACT, IT WAS REPORTED THAT THE DEATH WAS NOT PD RELATED. IN ADDITIONAL FOLLOW-UP, IT WAS CONFIRMED WITH THE PATIENT¿S PD NURSE THAT ON (B)(6) 2024 THE PATIENT EXPIRED WHILE CONNECTED TO THE FRESENIUS CYCLER DURING AN UNKNOWN PHASE OF PD TREATMENT. THE NURSE REPORTED THAT THERE WERE NO MALFUNCTIONS OR ISSUES WITH THE PD TREATMENT IN RELATION TO THE DEATH. ADDITIONALLY, THE NURSE STATED THAT THE PATIENT WAS COMPLETING PD TREATMENTS WITHOUT ANY ADVERSE EFFECTS PRIOR TO DEATH. THE NURSE PROVIDED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS THAT INCLUDED END STAGE RENAL DISEASE (ESRD), CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), DIABETES MELLITUS, HEART FAILURE AND EMPHYSEMA. THE CAUSE OF THE DEATH WAS REPORTED AS UNKNOWN AND THE ESRD DEATH CERTIFICATE WAS NOT AVAILABLE. HOWEVER, THE PD NURSE INDICATED THE DEATH IS SUSPECTED TO BE ASSOCIATED WITH THE PATIENT¿S PRE-EXISTING MEDICAL CONDITIONS. THE FRESENIUS CYCLER WAS PENDING RETURN TO MANUFACTURER AT THE TIME FOLLOW-UP WAS COMPLETED. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP MAY EXIST BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. BASED ON THE REPORTED INFORMATION, THE PATIENT¿S CAUSE OF DEATH CANNOT BE DETERMINED AND THE ESRD DEATH CERTIFICATE WAS NOT AVAILABLE. CURRENTLY, THERE IS NO ALLEGATION THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED PRE-EXISTING COPD, DIABETES MELLITUS, HEART FAILURE AND EMPHYSEMA WHICH ALL INCREASE MORTALITY RISK. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
CORTRAK 2 ENTERAL ACCESS SYSTEM | K220588 | IT WAS REPORTED VIA FDA MEDWATCH MW5159602, ¿CORTRAK CAUSE PNEUMOTHORAX IN GENTLEMAN. CHEST TUBE.¿ | THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. A ROOT CAUSE COULD NOT BE DETERMINED. ALL INFORMATION REASONABLY KNOWN AS OF 14 OCT 2024 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL INC. AVANOS MEDICAL INC. HAS NO INDEPENDENT KNOWLEDGE OF THE REPORTED EVENT BUT IS RELAYING THE INFORMATION PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE IS CURRENTLY AT THE SERVICE CENTER AWAITING EVALUATION. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER HAS RECEIVED THE DEVICE EVALUATION FROM THE SERVICE CENTER.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS UNKNOWN. THE CONTACT REPORTED THAT THE CUSTOMER HAD AN INFUSION SET SITE INFECTION AND THE CUSTOMER WAS FIGHTING A COLD OR FLU PRIOR TO THE EVENT. WHILE IN MEDICAL CARE, THE CUSTOMER HAD A STROKE. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
FLEXCATH ADVANCE STEERABLE SHEATH | K123591 | IT WAS REPORTED THAT DURING A CRYO ABLATION PROCEDURE, AFTER THE FIRST AND ONLY DELIVERY ON THE LEFT SUPERIOR PULMONARY VEIN, THE PATIENT EXPERIENCED A FEELING OF NAUSEA AND CHEST PAIN. THE PHYSICIAN STOPPED THE CRYOABLATION DELIVERY. ST ELEVATION WAS OBSERVED, AND THEN THE PATIENT HAD MULTIPLE VENTRICULAR FIBRILLATION EPISODES WHICH WERE TREATED WITH SEVERAL ELECTRICAL SHOCK DEFIBRILLATIONS AND CARDIAC MASSAGE. THE PATIENT DIED AFTER CARDIOPULMONARY RESUSCITATION (CPR). IT WAS NOTED THAT THE PHYSICIAN STATED THAT THE DEATH WAS DUE TO A HEART ATTACK (MYOCARDIAL INFARCTION) AND THAT THE ADVERSE EVENT WAS NOT RELATED TO THE DEVICE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
EDWARDS ESHEATH INTRODUCER SET | K200258 | THROUGH REVIEW OF MEDICAL ARTICLE "COMPUTED TOMOGRAPHY DEFINED FEMORAL ARTERY PLAQUE COMPOSITION PREDICTS VASCULAR COMPLICATIONS DURING TRANSCATHETER AORTIC VALVE IMPLANTATION," THE FOLLOWING EVENT WAS IDENTIFIED AS PERTAINING TO AN EDWARDS DEVICE: ONE PATIENT WITH A 23MM SAPIEN 3 VALVE IMPLANTED IN AORTIC POSITION BY TRANSFEMORAL APPROACH HAD A MAJOR VASCULAR COMPLICATION DUE TO A RETROPERITONEAL HEMATOMA RESULTING IN VARC TYPE 4 BLEEDING. | ARTICLE REFERENCE: CARANDE, E. J., SALIM, T. S., CHASE, A., SEKAR, B., ALDALATI, O., HAILAN, A., KHURANA, A., SMITH, D., & OBAID, D. R. (2023). COMPUTED TOMOGRAPHY DEFINED FEMORAL ARTERY PLAQUE COMPOSITION PREDICTS VASCULAR COMPLICATIONS DURING TRANSCATHETER AORTIC VALVE IMPLANTATION. THE BRITISH JOURNAL OF RADIOLOGY, 96(1152), 20230296. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. A PROCEDURAL OR POST PROCEDURAL BLEED/ HEMORRHAGE WITHOUT AN OBVIOUS SOURCE OF BLEEDING IS A RARE BUT POTENTIALLY LIFE-THREATENING COMPLICATION OF CORONARY/CARDIAC INTERVENTIONAL PROCEDURES AND TENDS TO OCCUR MORE FREQUENTLY IN THE PRESENCE OF MORE AGGRESSIVE ANTICOAGULATION REGIMENS. POSSIBLE SOURCES INCLUDE PUNCTURE OF THE FEMORAL ARTERY ABOVE THE INGUINAL LIGAMENT AND ABOVE THE INFERIOR EPIGASTRIC ARTERY, ALLOWING THE RESULTANT BLEEDING TO EXTEND INTO THE RETROPERITONEAL SPACE, OR INADVERTENT TRAUMA OR PERFORATION OF THE ANNULAR STRUCTURE OR VENTRICLES DURING THE PROCEDURE. THIS TYPE OF BLEEDING MAY NOT BE RECOGNIZED UNTIL THE POST PROCEDURAL PERIOD. EDWARDS EXTENSIVELY TRAINS PHYSICIANS BEFORE THEY ARE QUALIFIED TO USE THE TRANSCATHETER HEART VALVE (THV). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE CAUSE OF THE EVENT REMAINS UNKNOWN BUT FACTORS SUCH AS INCREASED SIZE OF OUTER SHEATH DIAMETER AND GREATER NECROTIC CORE VOLUME IN THE COMMON FEMORAL ARTERY MAY HAVE CAUSED OR CONTRIBUTED TO THIS EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
SPIDER FX | K063204 | LITERATURE WAS REVIEWED REGARDING: THE CLINICAL OUTCOMES OF PERCUTANEOUS TRANSCATHETER RELEASE OF STUCK MECHANICAL MITRAL VALVE WITH PETROS PERCUTANEOUS TRANSCATHETER RELEASE OF STUCK MECHANICAL VALVE (PETROS) WITH CEREBRAL EMBOLIC PROTECTION (CEP). THE TIME FRAME OF THIS STUDY WAS: FROM 2020 TO 2023 MULTIPLE MANUFACTURER¿S DEVICES WERE USED IN THE STUDY POPULATION, WHICH INCLUDED MEDTRONIC SPIDERFX EMBOLIC PROTECTION DEVICE USED IN EACH ICA THROUGH THE RIGHT AND LEFT COMMON FEMORAL ARTERIES. THE CEP WAS ONLY PARTIAL AS THE VERTEBRAL ARTERIES WERE NOT PROTECTED. AN INTERVENTIONAL RADIOLOGIST WAS ON CALL TO INTERVENE IN CASE OF VERTEBRAL ARTERY EMBOLISM BUT WAS NEVER REQUIRED. THE FILTERS SHOWED DEBRIS IN THE RIGHT INTERNAL CAROTID ARTERY (2.50%; N=3), LEFT INTERNAL CAROTID ARTERY (4.20%; N=1), AND BOTH FILTERS (4.20%; N=1). NO DEBRIS WAS FOUND IN 79.20% (N=19) OF CASES. THREE DEATHS WERE REPORTED IN THE POPULATION. ONE PATIENT DIED DUE TO SEVERE CHRONIC OBSTRUCTIVE PULMONARY DISEASE WITH TYPE-II RESPIRATORY FAILURE, DESPITE A TECHNICAL SUCCESS, AND A 2ND PATIENT HAD A SUDDEN DEATH ON THE 64TH DAY OF POST-PETROS. FOR THE THIRD PATIENT, DURING THE ATTEMPT TO REMOVE THE GUIDE CATHETER, THE THROMBUS ACCIDENTALLY EMBOLIZED AND CLOGGED THE FILTER, AND THERE WAS NO FLOW BEYOND THE FILTER. WHILE REMOVING THE FILTER, THE THROMBUS SPILLED OUT, AND BOTH THE LEFT MIDDLE CEREBRAL ARTERY AND ANTERIOR CEREBRAL ARTERY WERE OCCLUDED. IMMEDIATE ASPIRATION WAS DONE BUT WAS ONLY PARTIALLY SUCCESSFUL WITH SLOW FILLING OF THE DISTAL VESSELS. INTRA-ARTERIAL RETEPLASE BOLUS OF 5 UNITS WAS GIVEN, AND THE PATIENT WAS SHIFTED TO CARDIAC INTENSIVE CARE AFTER ENDOTRACHEAL INTUBATION AND MECHANICAL VENTILATION. MITRAL PETROS WAS DEFERRED. THE PATIENT DEVELOPED DENSE HEMIPLEGIA WITH APHASIA. SERIAL CT BRAIN SHOWED MASSIVE LEFT MIDDLE CEREBRAL ARTERY AND ANTERIOR CEREBRAL ARTERY TERRITORY INFARCT WITH MIDLINE SHIFT. THE PATIENT DIED ON THE SEVENTH DAY AFTER THE PROCEDURE. | LITERATURE REFERENCE:DOI: 10.1161/CIRCINTERVENTIONS.124.01404. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REMSTAR AUTO A FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY DUE TO LUNG AND LIVER CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND LIVER DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER, DEATH THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A DS2ADV AUTO CPAP WITH AN ALLEGATION OF LUNG DISEASE; CANCER; DEATH THIS DEVICE IS NOT PART OF THE RECALL. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. THE CAUSE OF DEATH WAS NOT YET IDENTIFIED; HOWEVER, IT WAS REPORTED THAT THE CUSTOMER EXPERIENCED LOW BLOOD GLUCOSE LEVELS. AT THE TIME OF THIS REPORT, PUMP DATA IS NOT AVAILABLE FOR REVIEW. MULTIPLE ATTEMPTS WERE MADE BY TANDEM TECHNICAL SUPPORT TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO ADDITIONAL INFORMATION COULD BE OBTAINED. SHOULD TANDEM OBTAIN ADDITIONAL INFORMATION, AN UPDATE REPORT WILL BE MADE. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
XENOSURE BIOLOGIC PATCH | K190882 | IT WAS REPORTED THAT FOR TWO PATIENTS RECEIVING XENOSURE PATCH IMPLANTS, AFTER THE DOCTOR CLOSED THE ATRIAL SEPTAL DEFECT (ASD) PROCEDURE ON THE PATIENT, EVALUATION WAS PERFORMED WITH A TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) MACHINE FOUND BUBBLE FORMATION IN THE LEFT ATRIAL CHAMBER. THE BUBBLES FORMED WHILE THE TEE MACHINE WAS REMOVED OR WHILE LEFT ALONE. THEY TRIED TO DEAIRING ALL THE BUBBLES, THEN TRIED TO WEAN THE PATIENT FROM HEART/LUNG MACHINE. IT LOOKED OKAY IN THE BEGINNING, BUT THE PATIENT HAD VENTRICULAR FIBRILLATION. THEY DECIDED TO PUT THE PATIENT IN THE HEART/LUNG MACHINE AGAIN, AND EVALUATED WITH TEE MACHINE, AND THE BUBBLES FORMED AGAIN. ADDITIONALLY, IT WAS REPORTED THAT ONE OF THE PATIENT PASSED AWAY IN THE ICU. HOWEVER, IT IS UNKNOWN WHICH PATIENT/XENOSURE LOT WAS INVOLVED IN THE REPORTED DEATH. ADDITIONALLY, IT WAS STATED THE DEATH DOES NOT APPEAR RELATED TO THE XENOSURE PATCH. NOTE: THIS IS REPORT 1 OF 2 RELATED TO THE FIRST PRODUCT USED IN THE EVENT. REPORT 1220948-2024-00208 WAS SUBMITTED FOR THE SECOND DEVICE INVOLVED IN THIS EVENT. | THE PRODUCT WAS NOT AVAILABLE FOR RETURN. THEREFORE, THE EXACT CAUSE OF THE REPORTED ISSUE IS UNKNOWN. HOWEVER, THE XENOSURE PRODUCT CONSISTS OF ONE PIECE OF BOVINE PERICARDIAL TISSUE THAT HAS BEEN SELECTED FOR MINIMAL TISSUE BLEMISHES. THERE ARE NO ADDITIONAL COMPONENTS OF THE PRODUCT THAT WOULD CREATE AIR BUBBLES OR CAUSE THE REPORTED ISSUE. IT IS LIKELY THAT THE SURGICAL TECHNIQUE AND PATIENT CONDITIONS CONTRIBUTED TO THE REPORTED ISSUE. WHILE THE EXACT CAUSE OF THE ONE INSTANCE OF PATIENT DEATH REMAINS UNKNOWN, IT WAS STATED THE XENOSURE PATCH IS NOT RELATED TO THE PATIENT DEATH. THE LOT HISTORY RECORD FOR THE DEVICE WAS REVIEWED, NO ISSUES WERE FOUND DURING MANUFACTURING OR PACKAGING THAT WOULD CAUSE OR CONTRIBUTE TO THE REPORTED EVENT. WE HAVE NOT RECEIVED ANY SIMILAR COMPLAINTS RELATED TO THIS LOT NUMBER. NOTE: THIS IS REPORT 1 OF 2 RELATED TO THE FIRST PRODUCT USED IN THE EVENT. REPORT 1220948-2024-00208 WAS SUBMITTED FOR THE SECOND DEVICE INVOLVED IN THIS EVENT.
TRI-STAPLE 2.0 | K160176 | ACCORDING TO THE REPORTER, DURING A VIDEO-ASSISTED THORACOSCOPIC SURGERY TO THE PATIENT THAT WAS CARDIOLOGICALLY BURDENED, AFTER FIRING TWO 45-MILLIMETER (MM) RELOADS, A REINFORCED 60-MM RELOAD WAS USED, BUT THE STAPLES DID NOT FORM A B-SHAPE. THE STAPLES WERE OPEN WITH STRAIGHT LEGS. TO RESOLVE THE ISSUE, A NEW 45MM RELOAD WAS USED WITH THE SAME HANDLE. IT WAS NOTED THAT AFTER DISSECTING THE SPECIMEN, STRONG BLEEDING WAS PRESENT. THE PROCEDURE WAS CONVERTED TO OPEN, AND A THREE-MM HOLE IN THE PULMONARY TRUNK WAS DISCOVERED AND SUPPORTED WITH VASCULAR SUTURE. THE PATIENT LOST A LOT OF BLOOD AND DIED DUE TO CARDIOLOGY STATUS. AFTER THE PROCEDURE, THE ANALYSIS TEAM CONCLUDED THAT THE ONLY WAY TO MAKE A HOLE WAS THE OPEN STAPLES FROM THE SPECIMEN, WHICH WERE MOVING AROUND THE OPERATION PLACE TO REACH GOOD VISIBILITY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ASTHMA AND LUNG DISEASE, DIZZINESS/HEADACHE AND EYE, SKIN, NOSE AND RESPIRATORY TRACT IRRITATION. THERE WAS REPORT OF PATIENT DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
EDWARDS ESHEATH+ INTRODUCER SET | K200258 | AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE WITH A 23 MM SAPIEN 3 ULTRA VALVE, THERE WAS DIFFICULTY ADVANCING THE 14FR ESHEATH+. A 16FR DILATOR WAS USED. THE RIGHT ILIAC WAS BALLOONED WITH A 7 MM BALLOON FOLLOWED BY AN 8 MM BALLOON. THE ESHEATH+ TIP WAS THEN ADVANCED TO THE ILIAC BIFURCATION AND JUST INTO THE AORTA. IN THE PROCESS OF ADVANCING THE DELIVERY SYSTEM WITH VALVE THROUGH THE 14FR ESHEATH+ UNDER GREAT RESISTANCE, THE ILIAC PERFORATED AND THE PATIENT PASSED AWAY. IT WAS NOTED PRIOR TO THE CASE THAT ACCESS HAD BEEN DISCUSSED. IT WAS INDICATED THAT THE RIGHT ILIAC DOES NOT APPEAR TO BE AN OPTION DUE TO THE HEAVILY CALCIFIED AREA IN THE COMMON ILIAC ARTERY. THE LEFT ILIAC WAS POSSIBLE, BUT IT WAS SUGGESTED THAT A CUTDOWN BE PERFORM FOR CONTROLLED ACCESS AND CLOSURE DUE TO THE CALCIFICATION AND POSSIBLE SHOCKWAVE OF THE EXTERNAL ILIAC. IT WAS ALSO INDICATED THAT THE LEFT CAROTID WAS GOOD FOR ALTERNATIVE ACCESS AND THE RIGHT SUBCLAVIAN MAY BE DIFFICULT BASED ON THE ACUTE BEND IN THE ARTERY. SUBSEQUENTLY, AN INTERVENTION WAS PERFORMED TO THE ILIAC ARTERIES PRIOR TO THE CASE. THE INTERVENTION PERFORMED WAS EXTENSIVE BALLOONING AND SHOCKWAVE TO THE RIGHT ILIAC ALONG WITH PLACEMENT OF A STENT IN THE LEFT ILIAC. ADDITIONAL INFORMATION WAS PROVIDED BY THE FCS. PER THE FCS, THE VALVE BECAME LODGED IN THE ESHEATH+ AND WAS NOT ABLE TO ADVANCE. AN ATTEMPT WAS MADE TO PULL THE VALVE OUT, BUT IT WAS PULLED TOO HARD THAT THE ESHEATH+ CAME OUT. THE VALVE WAS PARTIALLY DEPLOYED IN ORDER TO REMOVE THE DELIVERY SYSTEM NOSECONE. DURING THE ATTEMPT TO REMOVE THE DELIVERY SYSTEM, IT WAS NOTED THAT THE BALLOON HAD SEPARATED FROM THE SHAFT SLIGHTLY AFTER IT WAS PARTIALLY INFLATED TO GET THE NOSECONE THROUGH THE VALVE. RESISTANCE WAS FELT IN THE ESHEATH+ AT THE COMMON ILIAC ARTERY. THE ESHEATH+ APPEARED VERY KINKED PAST THE NON-EXPANDABLE PORTION. AFTER THE ESHEATH+ WAS REMOVED, THE VALVE WAS EVALUATED, AND IT APPEARED TO HAVE A BENT STRUT AT THE OUTFLOW OF THE VALVE FRAME. AN ATTEMPT WAS MADE TO REPAIR THE LACERATED ILIAC ARTERY BY PERFORMING AN OCCLUSION OF THE AORTA WITH A BALLOON AND PLACING COVERED STENTS, CROSSING OVER THE BIFURCATION OF THE RIGHT ILIAC BIFURCATION. HOWEVER, IT APPEARED THE DISSECTION HAD CARRIED INTO THE DESCENDING AORTA. IT WAS NOTED THAT THE CAUSE OF THE DISSECTION WAS NOT DUE TO THE VALVE BEING OUTSIDE OF THE ESHEATH+ BUT WAS DUE TO PUSHING TOO HARD THAT THE DELIVERY SYSTEM AND ESHEATH+ BENT AND APPEARED TO LACERATE THE ILIAC ARTERY. | THIS IS ONE OF THREE MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE. THE INVESTIGATION IS ONGOING.
DREAMSTATION AUTO BIPAP | K131982 | DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. COSMETIC DEFECT FOUND DEBRIS IN LCD SCREEN. THE DEVICE WAS SCRAPPED. THE MANUFACTURER RECEIVED INFORMATION CANCER; DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, LUNG DISEASE, CANCER, DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER A DA VINCI ASSISTED RADICAL PROSTATECTOMY (WITHOUT LYMPHADENECTOMY), THE PATIENT REQUIRED A POST-OPERATIVE OPEN PROCEDURE AT AN UNSPECIFIED TIME DUE TO BLEEDING AND LATER EXPIRED DUE TO A MYOCARDIAL INFARCTION. THE SURGEON REPORTED THAT WHILE THERE WAS SOME INTRA-OPERATIVE BLEEDING DURING THE DA VINCI-ASSISTED PROCEDURE, THE BLEEDING WAS CONTROLLED AND THE PROCEDURE WAS COMPLETED. THE SURGEON STATED THAT THERE WAS NO CONCERN FOR ANY DA VINCI ROBOTIC ISSUES. POST-OPERATIVELY, THE PATIENT'S HEMOGLOBIN WAS STABLE AND CHECKED TWICE AT 8.7 AND 8.8G/DL FOUR HOURS APART AND THE PATIENT HAD NORMAL VITAL SIGNS AND ABDOMINAL EXAM. HOWEVER, THE PATIENT EXPERIENCED A VENTRICULAR FIBRILLATION ARREST AND WAS CODED FOR 25 MINUTES AFTER WHICH THE PATIENT¿S ABDOMEN WAS NOTED TO BE DISTENDED. THE SURGEON STATED THAT THE DISTENTION WAS LIKELY DUE TO THE BAGGING EFFORTS DURING CPR. DUE TO CONCERN FOR POSSIBLE BLEEDING AN EMERGENT EXPLORATORY LAPAROSCOPY WAS PERFORMED WHERE AN UNSPECIFIED AMOUNT OF BLOOD WAS FOUND IN THE ABDOMEN AND EVACUATED; THE BLOOD WAS REPORTED TO LIKELY BE OLD. THE SURGEON STATED THAT THE PATIENT EXPIRED DUE TO A MYOCARDIAL INFARCTION. NO AUTOPSY WAS PERFORMED. | A REVIEW OF THE DA VINCI SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, 30 DEGREE ENDOSCOPE PLUS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, TWO MONOPOLAR CURVED SCISSORS, TWO LARGE NEEDLE DRIVERS. A SITE HISTORY REVIEW FOUND NO SUBSEQUENT COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT A ROBOTIC PROSTATECTOMY. BLEEDING WAS NOTED TO HAVE OCCURRED INTRAOPERATIVELY BUT THE SURGEON STATED THERE WAS NO CONCERN FOR A ROBOTIC ISSUE. THE DETAILS OF WHAT BLED OR HOW THE BLEEDING MAY HAVE OCCURRED WERE NOT PROVIDED. THE HEMOGLOBIN WAS STABLE ON TWO OCCASIONS IN THE POST OPERATIVE PERIOD BUT THE TIMELINE AS TO WHEN THOSE LEVELS WERE DRAWN WAS NOT PROVIDED. VITALS WERE ALSO REPORTEDLY NORMAL. HOWEVER, THE PATIENT DEVELOPED A DYSRHYTHMIA, CODED, AND WAS NOTED TO BE DISTENDED. THE PATIENT WAS EVENTUALLY RE-EXPLORED AT WHICH TIME BLOOD WAS FOUND. NEITHER THE VOLUME OF BLOOD DISCOVERED AT RE-EXPLORATION, THE SOURCE OF THE BLOOD LOSS, NOR ANY ADDITIONAL FINDINGS WERE PROVIDED. ALTHOUGH THERE WAS NO AUTOPSY, THE SURGEON REPORTED THE PATIENT DIED AS A RESULT OF A MYOCARDIAL INFARCTION. NO ADDITIONAL SUPPORTING EVIDENCE FOR THIS CONCLUSION WAS PROVIDED. BASED ON THE INFORMATION PROVIDED, INSUFFICIENT EVIDENCE IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION B2 - DATE OF DEATH IS ESTIMATED AS SAME DATE OF PROCEDURE AS COMPLICATION OCCURRED IN THE PACU POST-PROCEDURE. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
ION | K182188 | A PATIENT ENROLLED IN A REGISTRY STUDY UNDERWENT AN UNEVENTFUL ION BIOPSY PROCEDURE AND WAS DISCHARGED HOME 3.5 HOURS POST-PROCEDURE. NO ION SYSTEM MALFUNCTIONS OR OTHER ISSUES OCCURRED DURING THE BIOPSY PROCEDURE. IT WAS REPORTED THAT 8 DAYS POST-DISCHARGE THE PATIENT BEGAN TO EXPERIENCE RESPIRATORY DISTRESS IN THE EVENING AT HOME. THE PATIENT WENT INTO CARDIAC ARREST AND WAS UNABLE TO BE REVIVED. THE STUDY PRINCIPAL INVESTIGATOR CLASSIFIED THE EVENT AS NOT RELATED TO THE ION DEVICES, POSSIBLY RELATED TO THE ION PROCEDURE, AND PROBABLY RELATED TO THE PATIENT¿S PRE-EXISTING COMORBIDITIES. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. CONCOMITANT PRODUCTS: TWO LESIONS IN THE LEFT UPPER LOBE WERE BIOPSIED: ONE LOCATED 0 MM FROM THE PLEURA (18 X 24 X 20 MM) ¿ PRODUCTS USED WERE A 19G FLEXISION NEEDLE, 1.1 MM - CRYOPROBE, CYTOLOGY BRUSH; AND ONE LOCATED 13 MM FROM THE PLEURA (17 X 15 X 18 MM) ¿ 23G FLEXISION NEEDLE; 19G FLEXISION NEEDLE, 1.1 - MM CRYOPROBE, CYTOLOGY BRUSH. A REVIEW OF THE EVENT WAS CONDUCTED BY AN INTUITIVE SURGICAL MEDICAL OFFICER WHO CONCLUDED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL BIOPSY AS PART OF A REGISTRY CLINICAL STUDY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS DISCHARGED HOME A FEW HOURS AFTER THE PROCEDURE. THE PATIENT DEVELOPED RESPIRATORY DISTRESS AT HOME ABOUT 1 WEEK LATER PROMPTING EMERGENCY SERVICES TO BE CONTACTED. THE PATIENT WENT INTO CARDIAC ARREST. ATTEMPTS AT RESUSCITATION WERE UNSUCCESSFUL AND THE PATIENT DIED. THE PATIENT HAD SIGNIFICANT MEDICAL CO-MORBIDITIES INCLUDING SEVERE PULMONARY FIBROSIS/INTERSTITIAL PNEUMONITIS, CHRONIC BRONCHITIS, PULMONARY HYPERTENSION, RHEUMATOID ARTHRITIS, DIABETES, HYPERTENSION AND A 90 PACK-YEAR SMOKING HISTORY. BASED ON THE AVAILABLE DATA THE EVENT WAS POSSIBLY PROCEDURE RELATED IN THE SETTING OF MULTIPLE SIGNIFICANT CO-MORBIDITIES AND WAS NOT DEVICE RELATED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF (B)(4) BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE (B)(4) WITH (B)(4) TOTAL DEATHS (B)(4). A MULTICENTER STUDY REPORTED (B)(4) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH (B)(4) CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF (B)(4) REPORTED (B)(4). ANOTHER SURVEY-BASED STUDY OF (B)(4) BRONCHOSCOPIES DESCRIBED (B)(4) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN (B)(4) PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF (B)(4). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING (B)(4) CASES REPORTED NO DEATHS. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF (B)(4) BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). --OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY. THERE IS NOT ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE DREAMSTATION AUTO CPAP DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. DURING THE EVALUATION, TOBACCO CONTAMINATION WAS FOUND IN THE DEVICE, AND IT WAS RECOMMENDED THAT THE DEVICE BE SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
BIPAP AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, CANCER, DEATH, PANCREATIC AND METASTISIZED TO LIVER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE, CANCER AND IDIOPATHIC PULMONARY FIBROSIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
SERVO-I UNIT BASE | K123149 | MANUFACTURER'S REF.#: (B)(4). | A CORRECTION OF FIELD # B2 OUTCOMES ATTRIBUTED TO ADVERSE EVENT: PREVIOUS OUTCOMES ATTRIBUTED TO ADVERSE EVENT: LIFE THREATENING. CORRECTED OUTCOMES ATTRIBUTED TO ADVERSE EVENT: DEATH. A CORRECTION OF FIELD # H1 TYPE OF REPORTABLE EVENT: PREVIOUS TYPE OF REPORTABLE EVENT: SERIOUS INJURY. CORRECTED TYPE OF REPORTABLE EVENT: DEATH. A CORRECTION OF FIELD # H6 HEALTH EFFECT-IMPACT CODE: PREVIOUS HEALTH EFFECT-IMPACT CODE: 2199. CORRECTED HEALTH EFFECT-IMPACT CODE: 1802.
SERVO-I UNIT BASE | K123149 | MANUFACTURER'S REF.#: (B)(4). | A CORRECTION OF FIELD # B2 OUTCOMES ATTRIBUTED TO ADVERSE EVENT: PREVIOUS OUTCOMES ATTRIBUTED TO ADVERSE EVENT: LIFE THREATENING. CORRECTED OUTCOMES ATTRIBUTED TO ADVERSE EVENT: DEATH. A CORRECTION OF FIELD # H1 TYPE OF REPORTABLE EVENT: PREVIOUS TYPE OF REPORTABLE EVENT: SERIOUS INJURY. CORRECTED TYPE OF REPORTABLE EVENT: DEATH. A CORRECTION OF FIELD # H6 HEALTH EFFECT-IMPACT CODE: PREVIOUS HEALTH EFFECT-IMPACT CODE: 2199. CORRECTED HEALTH EFFECT-IMPACT CODE: 1802.
FLEXCATH ADVANCE STEERABLE SHEATH | K123591 | IT WAS REPORTED THAT DURING A CRYO ABLATION PROCEDURE WHEN THE SHEATH AND INTEGRATED DILATOR/NEEDLE WERE ADVANCED THROUGH THE GROIN, RESISTANCE WAS FELT. THE GUIDEWIRE WAS REMOVED AND IT WAS NOTED TO BE BENT. A NEW GUIDEWIRE FROM ANOTHER MANUFACTURER WAS INSERTED AND RESISTANCE WAS FELT AGAIN. THE GUIDEWIRE WAS REPLACED AGAIN AND THE SHEATH WAS ALSO REPLACED WITH ANOTHER MANUFACTURER'S DEVICE. IT WAS NOTED THAT THE PATIENT'S BLOOD PRESSURE DROPPED, AND THEN THE SEDATION WAS REDUCED. THE PATIENT'S CIRCULATION BECAME UNSTABLE AND A VASOPRESSOR DRUG WAS ADMINISTERED. THE GROIN WAS EXAMINED WITH AN ULTRASOUND, AND A GROIN BLEED WAS DETECTED. A VASCULAR SURGEON ATTEMPTED TO CLOSE THE PUNCTURE SITE WITH A BALLOON, BUT IT WAS UNSUCCESSFUL. THE CASE WAS ABORTED. THE PATIENT WAS NOT UNDER GENERAL ANESTHESIA. REANIMATION MEASURES WHICH INCLUDED CARDIOPULMONARY RESUSCITATION (CPR) WERE UNSUCCESSFUL AND THE PATIENT DIED IN THE CARDIAC CATHETERIZATION LAB. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
FLEXCATH ADVANCE STEERABLE SHEATH | K123591 | IT WAS LATER REPORTED THAT DURING MAPPING, TRANSEPTAL ACCESS WAS LOST WITH THE SHEATH; HOWEVER, THE SHEATH WAS ABLE TO BE READVANCED ACROSS THE SEPTUM. AFTER TWO ABLATIONS WERE PERFORMED, THE PATIENT EXPERIENCED BRADYCARDIA AND AN EFFUSION WAS NOTED FROM AN ULTRASOUND. THE PATIENT "CODED", A PERICARDIAL TAP WAS PERFORMED TO REMOVE THE EFFUSION, AND THEN TIME OF DEATH WAS CALLED. | CONTINUATION OF D10: PRODUCT ID: NON-MEDTRONIC; PRODUCT TYPE: MAPPING CATHETER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
ACHIEVE MAPPING CATHETER - 20 MM | K102588 | IT WAS LATER REPORTED THAT DURING MAPPING, TRANSEPTAL ACCESS WAS LOST WITH THE SHEATH; HOWEVER, THE SHEATH WAS ABLE TO BE READVANCED ACROSS THE SEPTUM. AFTER TWO ABLATIONS WERE PERFORMED, THE PATIENT EXPERIENCED BRADYCARDIA AND AN EFFUSION WAS NOTED FROM AN ULTRASOUND. THE PATIENT "CODED", A PERICARDIAL TAP WAS PERFORMED TO REMOVE THE EFFUSION, AND THEN TIME OF DEATH WAS CALLED. | CONTINUATION OF D10: PRODUCT ID: NON-MEDTRONIC; PRODUCT TYPE: MAPPING CATHETER MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
VUE PACS | K110919 | PHILIPS COMPANY RECEIVED A COMPLAINT ON 26-SEP-2024 REGARDING THE VUE PACS SYSTEM. THE REPORT STATED THAT A RADIOLOGIST EXPERIENCED A LOADING ERROR (BLACK SCREEN) WHILE ATTEMPTING TO ACCESS IMAGES FROM A CTA (CT ANGIOGRAPHY) OF THE ABDOMEN/PELVIS FOR A PATIENT SUSPECTED OF AN OCCULT GI BLEED. UNFORTUNATELY, THE PATIENT PASSED AWAY SHORTLY AFTER THE SCAN. INVESTIGATION SUMMARY: SYSTEM FUNCTIONALITY AND USE: INVESTIGATION AND REVIEW OF LOGS FROM THE INCIDENT CONCLUDED THAT: THE CTA (COMPUTED TOMOGRAPHY ANGIOGRAPHY) IMAGES ARRIVED IN THE VUE PACS AND APPEARED IN THE WORKLIST WITH THE CORRECT PRIORITY. THE ON-CALL INTERVENTIONAL RADIOLOGIST ACCESSED THE STUDY ON THE FIRST ATTEMPT APPROXIMATELY 15 MINUTES AFTER UPLOAD, REVIEWED THE IMAGES, AND PERFORMED LINE MEASUREMENTS AS PER SYSTEM LOGS. THIS WAS CONFIRMED BY THE CUSTOMER AND SHOWN IN THE LOG. THE CUSTOMER CONFIRMED THAT THE PATIENT¿S DETERIORATION WAS ALREADY KNOWN, AND PATIENT CARE WAS PROVIDED. A SECOND OFF-SITE RADIOLOGIST ENCOUNTERED A LOADING ERROR AND BLACK SCREEN WHEN ATTEMPTING TO VIEW THE SAME STUDY 40 MINUTES AFTER. EFFORTS TO OBTAIN ADDITIONAL INFORMATION: THE COMPANY MADE MULTIPLE GOOD-FAITH EFFORTS (GFES) TO REQUEST ADDITIONAL INFORMATION FROM THE CUSTOMER, INCLUDING CLINICAL QUESTIONS, WORKSTATION DETAILS, AND CONTEXT ON THE ACCESS ISSUE. THE CUSTOMER PROVIDED ONLY LIMITED INFORMATION AND DID NOT COMPLETELY RESPOND TO FOLLOW-UP REQUESTS, PREVENTING A FULL ANALYSIS. CONCLUSION: THE INVESTIGATION DETERMINED THAT: THE VUE PACS SYSTEM FUNCTIONED AS EXPECTED FOR THE ON-CALL RADIOLOGIST, WHO ACCESSED AND REVIEWED THE STUDY WITHOUT ISSUES. THE INABILITY OF THE SECOND RADIOLOGIST TO ACCESS THE IMAGES DID NOT RESULT IN A DELAY IN PATIENT CARE. DESPITE REPEATED GFES TO OBTAIN ADDITIONAL INFORMATION, THE INVESTIGATION REMAINS INCONCLUSIVE DUE TO THE LACK OF A DETAILED RESPONSE FROM THE CUSTOMER. IF ANY ADDITIONAL RELEVANT INFORMATION IS RECEIVED, THE CASE WILL BE REOPENED, AND THE CONCLUSIONS WILL BE RE-EVALUATED ACCORDINGLY. | PHILIPS INVESTIGATION IS ON-GOING. PHILIPS WILL PROVIDE ANOTHER REPORT IN 30 DAYS.
REMSTAR AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS ZERO ERROR CODE FOUND AND DURING THE EVALUATION NO SECONDARY FINDINGS WERE OBSERVED. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LIFEPAK® 15 DEFIBRILLATOR/MONITOR | K142430 | THE CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE WAS NOT PROVIDING AUDIO PROMPTS DURING USE ON A PATIENT. IN THIS STATE THE DEVICE MAY NOT BE ABLE TO DELIVER DEFIBRILLATION THERAPY IF NEEDED. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. | DEVICE EVALUATION: STRYKER EVALUATED THE DEVICE AND THE REPORTED ISSUE WAS UNABLE TO BE VERIFIED AND UNABLE TO BE DUPLICATED. NO FAULT WAS FOUND WITH THE DEVICE. THE DEVICE PASSED FUNCTIONAL AND PERFORMANCE TESTING AND WAS RETURNED TO THE CUSTOMER. THE CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. THE CUSTOMER REPORTED THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT OUTCOME. CORRECTIONS: SECTION B1 OF THE INITIAL MEDWATCH REPORT INDICATES: ADVERSE EVENT AND PRODUCT PROBLEM. SECTION B1 OF THE INITIAL MEDWATCH REPORT SHOULD INDICATE: PRODUCT PROBLEM. SECTION B2 OF THE INITIAL MEDWATCH REPORT INDICATES: DEATH. SECTION B2 OF THE INITIAL MEDWATCH REPORT SHOULD BE BLANK. SECTION H1 OF THE INITIAL MEDWATCH REPORT INDICATES: DEATH. SECTION H1 OF THE INITIAL MEDWATCH REPORT SHOULD INDICATE: MALFUNCTION. SECTION H6 HEALTH IMPACT CODE GRID OF THE INITIAL MEDWATCH REPORT INDICATES: DEATH. SECTION H6 HEALTH IMPACT CODE GRID OF THE INITIAL MEDWATCH REPORT SHOULD INDICATE: NO HEALTH CONSEQUENCES OR IMPACT.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CHRONIC OBSTRUCTIVE PULMONARY DISEASE AND ACUTE RESPIRATORY DISTRESS SYNDROME. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY FAILURE AND HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION HUMID | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | MDR 2518422-2024-41716 IS THE ORIGINAL REPORT ASSOCIATED WITH THIS DEVICE. A DUPLICATE REPORT HAS BEEN SUBMITTED FOR THIS DEVICE UNDER REPORT MDR 2518422-2024-66137. THIS FOLLOW UP REPORT IS BEING FILED FOR AWARENESS OF THIS DUPLICATE REPORT.
REMSTAR PRO C-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH AND INTESTINAL METAPLASIA THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER PREVIOUSLY MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. USING A KNOWN GOOD POWER SUPPLY AND POWER CORD,THE MANUFACTURER WAS ABLE TO CONFIRM THE DEVICE POWERS ON AND PROVIDED AIRFLOW. THE MANUFACTURER CONFIRMED THE OPERATION OF THE HEATER PLATE. DURING REVIEW OF THE ERROR LOGS, THE MANUFACTURER DETERMINED THAT THE DEVICE HAD 13923.6 MACHINE HOURS, 13923.6 BLOWER HOURS AND 13899.1 THERAPY HOURS. THE ERROR LOG SHOWED 1 ERROR LOGGED. 1 INSTANCE OF E-53 (ERR_COMP_LOG_SEM_TIMEOUT) A CONTINUE ERROR. DURING A CHANGE TO THE COMPLIANCE LOGGER, THE SEMAPHORE TIMER EXPIRES, AND THE SEMAPHORE IS RELEASED. THIS CAN OCCUR OF THE MODEM IS REMOVED WHILE THE COMPLIANCE LOGGER HAS CONTROL OF THE SEMAPHORE. CARE ORCHESTRATOR DATA SHOWS THE PATIENT HAD A COMPLIANCE RATE OF (B)(4) AND USED A SYSTEM ONE HUMIDIFICATION MODE, WITH A HUMIDIFIER SETTING OF 3. DURING THE INVESTIGATION THE MANUFACTURER OBSERVED DRIED LIQUID SPOTS ON THE TOP ENCLOSURE. AN UNKNOWN PIECE OF DEBRIS WAS OBSERVED ON THE BACKSIDE OF THE UI PANEL. POTENTIAL DRIED LIQUID SPOTS WERE OBSERVED ON THE BLOWER HOUSING, ON THE BOTTOM OF AND INSIDE THE BLOWER MOTOR. THE MANUFACTURER OBSERVED A WHITISH POWDERY DUST-LIKE CONTAMINATION IN THE AIR INLET, INSIDE THE BLOWER MOTOR, ON THE BOTTOM OF THE BLOWER HOUSING, ON THE SOUND ABATEMENT FOAM AND THROUGHOUT THE BOTTOM ENCLOSURE. DUST-LIKE CONTAMINATION WAS OBSERVED ON THE RIGHT-SIDE PANEL, IN THE ISO PORT, ON THE INTERIOR SIDE OF THE TOP ENCLOSURE ON THE PCA AND ON THE BLOWER CAP. THE MANUFACTURER OBSERVED A SMALL AMOUNT OF POTENTIAL DEGRADED SOUND ABATEMENT FOAM ON THE BOTTOM OF THE BLOWER HOUSING. THE MANUFACTURER CONFIRMED THE PRESENCE OF DEGRADED SOUND ABATEMENT FOAM. (DS6HFLG) THE MANUFACTURER OBSERVED POTENTIAL DRIED LIQUID SPOTS ON THE WATER TANK, INTERIOR SIDE OF THE FLIP LID ASSEMBLY AND IN THE LOWER BASE. DUST-LIKE CONTAMINATION WAS OBSERVED THROUGHOUT THE WATER TANK, ON THE INTERIOR SIDE OF THE FLIP LID SEAL, IN THE BOTTOM ENCLOSURE, ON THE DRY BOX, AND IN THE LOWER BASE. A BROWNISH CONTAMINATION WAS OBSERVED ON THE FLIP LID SEAL. THE MANUFACTURER CONCLUDES THERE WAS EVIDENCE OF SOUND ABATEMENT FOAM DEGRADATION IN THE DEVICE AND DUST CONTAMINATION WAS ABLE TO CONFIRM THE CUSTOMER'S COMPLAINT. IN BOX G: DATE RECEIVED BY MFG. HAS BEEN UPDATED. IN BOX H: (DEVICE) PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER AND LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING DEATH AND CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 02/01/2025, AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING DEATH AND CANCER. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING EVALUATION SECONDARY FINDINGS AND DUST CONTAMINATION WERE OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WERE 11 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD NOT CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. IN BOX H: DEVICE PROBLEM CODE, EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, DEATH, CANCER AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. | IN PREVIOUS REPORT, AE OUTCOME WAS CAPTURED INCORRECTLY, AND IT IS CAPTURED CORRECTLY IN THIS REPORT. DEATH.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO SERIES AFLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, PNEUMONIA AND HEART ATTACK. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, PANCREATIC. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ENDO ANCHOR SYSTEM - HELI-FX AAA | K140036 | IT WAS REPORTED THAT HELI-FX ENDOANCHORS WERE IMPLANTED IN THE PATIENT FOR THE PREVENTION OF NECK DILATATION, THREE ENDURANT II STENT GRAFTS WERE ALSO IMPLANTED. IT WAS REPORTED THAT APPROXIMATELY 2 WEEKS LATER THE PATIENT SUFFERED MULTI-SYSTEM ORGAN FAILURE RESULTING IN DEATH. SITE ASSESSED EVENT AS RELATED TO INDEX PROCEDURE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), HYPERSENSITIVITY, NAUSEA, VOMITING, KIDNEY DISEASE/TOXICITY, LUNG, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY CANCER, BILATERAL PULMONARY EMBOLISM AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THERE WAS A DEATH REPORTED OR ALLEGED TO HAVE OCCURRED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
SJM RIGID SADDLE RING | K042734 | THE ARTICLE, "CLINICAL AND ECHOCARDIOGRAPHIC PREDICTORS OF THE ANTERIOR MITRAL LEAFLET REPAIR FAILURE", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO ANALYZE PREOPERATIVE CLINICAL AND ECHOCARDIOGRAPHIC CHARACTERISTICS IN PATIENTS WITH ISOLATED ANTERIOR MITRAL LEAFLET PROLAPSE AND TO DEFINE POSSIBLE PREDICTORS THAT ARE AFFECTING REPAIR DURABILITY IN PATIENTS WITH AML PROLAPSE. DEVICES INCLUDED IN THE ARTICLE WERE SJM SADDLE RING, SJM SEGUIN RING, CE PHYSIO I RING, CE CLASSIC RING, MEDTRONIC DURAN ANCORE, AND SJM TAILOR RING. THE ARTICLE CONCLUDED THAT BETTER LONG-TERM REPAIR DURABILITY IN PATIENTS WITH ANTERIOR MITRAL LEAFLET PROLAPSE AND PRESERVED SINUS RHYTHM AND LEFT-VENTRICLE DIAMETERS JUSTIFIES EARLY RECONSTRUCTIVE APPROACH. [THE PRIMARY AND CORRESPONDING AUTHOR WAS BOGDAN OKILJEVIC, CLINIC FOR CARDIAC SURGERY, INSTITUTE OF CARDIOVASCULAR DISEASES DEDINJE, 11000 BELGRADE, SERBIA, WITH CORRESPONDING EMAIL: BOGDANOKILJEVIC@GMAIL.COM] THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2010 TO MARCH 2021. A TOTAL OF 129 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 90 (69.8%) RECEIVED AN ABBOTT DEVICE. IN THE AML GROUP, THE AVERAGE AGE WAS 54.78 YEARS AND THE MAJORITY GENDER WAS MALE. IN THE POSTERIOR MITRAL LEAFLET (PML) GROUP, THE AVERAGE AGE WAS 55.4 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DIABETES MELLITUS, HYPERLIPOPROTEINEMIA, HYPERTENSION, SMOKING HISTORY, HEART FAILURE. | B2 - DATE OF DEATH WAS ESTIMATED TO END OF STUDY RANGE. B3: THE EVENT DATE WAS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLED "CLINICAL AND ECHOCARDIOGRAPHIC PREDICTORS OF THE ANTERIOR MITRAL LEAFLET REPAIR FAILURE". INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
SJM SÃ©GUIN SEMI-RIGID ANNULOPLASTY RING | K092310 | THE ARTICLE, "CLINICAL AND ECHOCARDIOGRAPHIC PREDICTORS OF THE ANTERIOR MITRAL LEAFLET REPAIR FAILURE", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO ANALYZE PREOPERATIVE CLINICAL AND ECHOCARDIOGRAPHIC CHARACTERISTICS IN PATIENTS WITH ISOLATED ANTERIOR MITRAL LEAFLET PROLAPSE AND TO DEFINE POSSIBLE PREDICTORS THAT ARE AFFECTING REPAIR DURABILITY IN PATIENTS WITH AML PROLAPSE. DEVICES INCLUDED IN THE ARTICLE WERE SJM SADDLE RING, SJM SEGUIN RING, CE PHYSIO I RING, CE CLASSIC RING, MEDTRONIC DURAN ANCORE, AND SJM TAILOR RING. THE ARTICLE CONCLUDED THAT BETTER LONG-TERM REPAIR DURABILITY IN PATIENTS WITH ANTERIOR MITRAL LEAFLET PROLAPSE AND PRESERVED SINUS RHYTHM AND LEFT-VENTRICLE DIAMETERS JUSTIFIES EARLY RECONSTRUCTIVE APPROACH. [THE PRIMARY AND CORRESPONDING AUTHOR WAS BOGDAN OKILJEVIC, CLINIC FOR CARDIAC SURGERY, INSTITUTE OF CARDIOVASCULAR DISEASES DEDINJE, 11000 BELGRADE, SERBIA, WITH CORRESPONDING EMAIL: BOGDANOKILJEVIC@GMAIL.COM]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2010 TO MARCH 2021. A TOTAL OF 129 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 90 (69.8%) RECEIVED AN ABBOTT DEVICE. IN THE AML GROUP, THE AVERAGE AGE WAS 54.78 YEARS AND THE MAJORITY GENDER WAS MALE. IN THE POSTERIOR MITRAL LEAFLET (PML) GROUP, THE AVERAGE AGE WAS 55.4 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED ATRIAL FIBRILLATION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, DIABETES MELLITUS, HYPERLIPOPROTEINEMIA, HYPERTENSION, SMOKING HISTORY, HEART FAILURE. | B2 - DATE OF DEATH WAS ESTIMATED AS THE END OF STUDY RANGE. B3: DATE OF EVENT WAS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: ARTICLE TITLED "CLINICAL AND ECHOCARDIOGRAPHIC PREDICTORS OF THE ANTERIOR MITRAL LEAFLET REPAIR FAILURE". INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
EMBOTRAP | K193063 | THE HEALTHCARE PROFESSIONAL REPORTED THAT THE PATIENT WHO HAS BEEN HOSPITALIZED TWICE PREVIOUSLY DUE TO CEREBRAL HEMORRHAGE UNDERWENT A MECHANICAL THROMBECTOMY PROCEDURE THAT TARGETED AN OCCLUSION IN THE RIGHT INTERNAL CAROTID (IC) TOP TO THE RIGHT M1 SEGMENT OF THE MIDDLE CEREBRAL ARTERY (MCA). THE PROCEDURE WAS PERFORMED VIA THE COMBINED TECHNIQUE USING A 5MM X 37MM EMBOTRAP III REVASCULARIZATION DEVICE (ET309537 / 23L089AV) AND AN AXS VECTA 46 INTERMEDIATE CATHETER (STRYKER). THE THROMBUS OCCLUDED THE RIGHT IC TOP NEAR THE RIGHT M1 SEGMENT. A GUIDEWIRE AND MICROCATHETER (UNSPECIFIED BRAND) WERE DELIVERED AND CROSSED THE LESION AND CONTINUED TO ADVANCE UP TO THE RIGHT M2 SEGMENT. THE EMBOTRAP III DEVICE AND THE AXS VECTA 46 WERE DELIVERED AND THE EMBOTRAP III WAS DEPLOYED. THE MICROCATHETER WAS REMOVED FIRST AND THEN THE THROMBUS WAS WITHDRAWN WITH THE EMBOTRAP III AND THE AXS VECTA 46. ANGIOGRAPHY CONFIRMED THAT THE THROMBUS REMAINED AT THE IC TOP, PROMPTING A SECOND PASS USING THE SAME SYSTEM SET UP. THE GUIDEWIRE AND THE MICROCATHETER CROSSED THE LESION OF THE THROMBUS AND REACHED THE M1. THE EMBOTRAP III DEVICE AND THE AXS VECTA 46 WERE DELIVERED AND THEN THE EMBOTRAP III WAS DEPLOYED AND THE THROMBUS RETRIEVED. ANGIOGRAPHY CONFIRMED THAT AN APPROXIMATE MTICI SCORE OF 2B WAS ACHIEVED. THE THROMBECTOMY WAS COMPLETED AND THE WHOLE SYSTEM (ALL DEVICES) WERE REMOVED. TREATMENT WAS CONSIDERED COMPLETED. A HEMORRHAGE FROM THE LEFT INTERNAL CAROTID ARTERY (ICA) WAS OBSERVED, WHICH WAS ON THE OPPOSITE SIDE OF THE TREATED VESSEL (THE THROMBECTOMY WAS PERFORMED ON THE RIGHT ICA). THE PATIENT SUBSEQUENTLY EXPIRED. THE PHYSICIAN COMMENTED THAT THE PATIENT HAD BEEN TWICE HOSPITALIZED PREVIOUSLY DUE TO CEREBRAL HEMORRHAGES, AND THE BLOOD VESSELS WERE PRONE TO RUPTURE. ¿THE BLOOD PRESSURE IN THE ENTIRE CEREBRAL VESSEL MIGHT HAVE INCREASED DUE TO THE MANIPULATION OF THE DEVICES (MICROCATHETER, STENT, ASPIRATION CATHETER) DURING THE THROMBECTOMY, AND THIS COULD CAUSE THE HEMORRHAGE FROM THE VESSEL EVEN THOUGH AT THE OPPOSITE VESSEL WHICH DID NOT DIRECTLY CONTACT WITH THE DEVICES. ALSO, FURTHER HEMORRHAGE MIGHT HAVE OCCURRED DUE TO THE INTRAVENOUS HEPARIN. RATHER THAN THE EFFECT OF THE EMBOTRAP III ITSELF, THE THROMBECTOMY PROCEDURE PUT A CERTAIN LOAD TO THE CEREBRAL VESSELS, AND THE HEMORRHAGE MIGHT HAVE OCCURRED. THE THROMBECTOMY PROCEDURE ITSELF WAS COMPLETED, BUT THE OUTCOME WAS REGRETTABLE.¿ ON 16-OCT-2024, ADDITIONAL INFORMATION WAS RECEIVED. THE INFORMATION CONFIRMED TWO PASSES WERE MADE WITH THE EMBOTRAP III DEVICE. ON 21-OCT-2024, ADDITIONAL INFORMATION WAS RECEIVED. THE DATE OF DEATH WAS (B)(6) 2024. THE PATIENT¿S OFFICIAL CAUSE OF DEATH WAS DUE TO BRAIN HEMORRHAGE. THE INFORMATION INDICATED THAT DETAILS RELATED TO WHETHER THERE WERE ANY EXCESSIVE VESSEL TORTUOSITY OR ACUTE BENDS IN THE TARGET VESSEL COULD NOT BE OBTAINED, BUT ¿THE THROMBUS RETRIEVAL WAS PERFORMED SMOOTHLY.¿ THERE WAS NO DEVICE PERFORMANCE ISSUE RELATED TO THE EMBOTRAP III DEVICE SUCH AS WITHDRAWAL DIFFICULTY DURING THE PROCEDURE. NO FURTHER INFORMATION IS AVAILABLE. | MANUFACTURER¿S REF. NO: (B)(4). INFORMATION REGARDING PATIENT IDENTIFIER, DATE OF BIRTH, AGE, SEX, GENDER, WEIGHT, RACE, AND ETHNICITY WERE NOT PROVIDED. SECTION E.1: THE INITIAL REPORTER PHONE: (B)(6) . THE INITIAL REPORTER EMAIL ADDRESS WAS NOT AVAILABLE / REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE IS NOT AVAILABLE TO BE RETURNED FOR ANALYSIS. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT (23L089AV) TOP AND SUB ASSEMBLY PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESSES RELATED TO THE REPORTED COMPLAINT. THERE WERE NO NON-CONFORMANCES RELATED TO DEVICE MANUFACTURE OR INSPECTION. PRODUCT ANALYSIS CANNOT BE CONDUCTED AS THE PRODUCT WAS NOT RETURNED FOR ANALYSIS. NO DETERMINATION OF CAUSES AND POSSIBLE CONTRIBUTING FACTORS COULD BE MADE. AS SUCH, THE INVESTIGATION WILL BE CLOSED. ON (B)(6) 2024, THIS COMPLAINT WAS THOROUGHLY DISCUSSED WITH THE SENIOR MEDICAL SAFETY OFFICER (MSO). PER THE MSO, THERE IS A POSSIBILITY THAT THE STRESS OF THE PROCEDURE ON THE CEREBRAL CIRCULATION COULD AFFECT PRESSURE, AS THE PHYSICIAN STATED. HOWEVER, IT IS VERY UNLIKELY THAT THE PROCEDURE ITSELF WAS THE TRIGGER FOR A CEREBRAL HEMORRHAGE IN THE CONTRALATERAL SIDE OF THE TREATMENT. THEREFORE, THE DEVICE NOR THE PROCEDURE CAUSED OR CONTRIBUTED TO THE DEATH OF THE PATIENT. PER THE ASSESSMENT FROM THE SENIOR MEDICAL SAFETY OFFICER (MSO), RECEIVED ON 16-OCT-2024, ALTHOUGH THERE IS A POSSIBILITY THAT THE STRESS OF THE PROCEDURE ON THE CEREBRAL CIRCULATION COULD AFFECT PRESSURE, IT IS VERY UNLIKELY THAT THE PROCEDURE ITSELF WAS THE TRIGGER FOR A CEREBRAL HEMORRHAGE IN THE CONTRALATERAL SIDE OF THE TREATMENT. BASED ON THIS INFORMATION, THE DEVICE NOR PROCEDURE CAUSED OR CONTRIBUTED TO THE EVENT OF A HEMORRHAGE. FURTHER, THERE WERE MANY EXTERNAL FACTORS BEYOND THE USE OF THE DEVICE THAT WOULD LEAD TO THE BLEED, INCLUDING THE PATIENT¿S SUSCEPTIBILITY FOR VESSEL RUPTURE (AS THE PATIENT HAD A HISTORY OF TWO CEREBRAL HEMORRHAGES) AND THE ADMINISTERED INTRAVENOUS ANTICOAGULATION MEDICATION (HEPARIN). THIS COMPLAINT WAS THOROUGHLY DISCUSSED WITH CERENOVUS SENIOR MEDICAL AFFAIRS DIRECTOR AND NEURO-INTERVENTIONALIST, DR. (B)(6), ON (B)(6) 2024. PER THE SENIOR MEDICAL DIRECTOR, THE EVENT HAS NO CORRELATION WITH EMBOTRAP III DEVICE. IN ADDITION, INCREASED BLOOD PRESSURE FROM DEVICE MANIPULATION WITHIN THE VESSEL IS NOT SEEN IN USUAL NEUROINTERVENTIONAL PROCEDURES. BLOOD PRESSURE RISE IS, HOWEVER, A COMMON PHENOMENON SEEN IN STROKE PATIENTS, AS THIS IS A NORMAL COMPENSATORY MECHANISM TO OVERCOME ISCHEMIA RESULTING FROM THE INDEX STROKE EVENT. PER THE ADDITIONAL INFORMATION RECEIVED ON 21-OCT-2024 AND THE ASSESSMENT FROM THE CERENOVUS SENIOR MEDICAL DIRECTOR, RECEIVED ON 18-OCT-2024; IT WAS CONFIRMED THERE WAS NO DEVICE PERFORMANCE ISSUE RELATED TO THE EMBOTRAP III DEVICE DURING THE PROCEDURE, SUCH AS DEPLOYMENT OR WITHDRAWAL DIFFICULTY, AND THE THROMBUS RETRIEVAL WAS PERFORMED SMOOTHLY. AS EXPLAINED IN THE REFERENCED MEDICAL LITERATURE, THE EXERTION OF ABRASIVE FORCES BY A STENT-RETRIEVER ON THE ENDOTHELIUM IS A KNOWN OCCURRENCE DURING THROMBECTOMY PROCEDURES; HOWEVER, THESE FORCES ARE EXERTED IN A CIRCUMFERENTIAL AND LINEAR PATTERN. MECHANICAL DISRUPTION OF VESSEL WALL INTEGRITY BY THE USE OF THE EMBOTRAP III DEVICE WOULD FOLLOW THIS CHARACTERISTIC PATTERN, AND A HEMORRHAGE WOULD BE OBSERVED AT A CLOSER SITE TO WHERE THE DEVICES WAS USED, I.E., THE RIGHT INTERNAL CAROTID ARTERY OR RIGHT MIDDLE CEREBRAL ARTERY. THE NOTION THAT THE DEVICE CAUSED OR CONTRIBUTED TO A HEMORRHAGE ON THE CONTRALATERAL SIDE OF THE TREATED VESSEL IS NOT PLAUSIBLE. FURTHER, AN INCREASED BLOOD PRESSURE DURING/AFTER STROKE IS RELATED TO THE BODY¿S NORMAL DEFENSE MECHANISM TO OVERCOMPENSATE FOR LACK OF BLOOD FLOW AND IS NEEDED/SOMETIMES MEDICALLY FACILITATED (PERMISSIVE HYPERTENSION GUIDELINE: SYSTOLIC > 140 TO 180 MMHG), AS IT¿S ASSOCIATED WITH BETTER OUTCOMES FOR STROKE PATIENTS (SEE REFERENCED LITERATURE). BASED ON THE ASSESSMENT OF THE SENIOR MEDICAL DIRECTOR, THE REPORTED INCREASED BLOOD PRESSURE WITHIN THE CEREBRAL CIRCULATION WOULD BE ATTRIBUTED TO THE PATIENT¿S BASELINE INJURY/STROKE AND THERE IS NO CORRELATION BETWEEN THIS EVENT AND THE USED DEVICE NOR PROCEDURE AS A CONTRIBUTING FACTOR. REFERENCES: DE GEORGIA M., BOWEN T., DUNCAN KR., CHEBL AB. BLOOD PRESSURE MANAGEMENT IN ISCHEMIC STROKE PATIENTS UNDERGOING MECHANICAL THROMBECTOMY. NEUROL RES PRACT. 2023 MAR 30;5(1):12. DOI: 10.1186/S42466-023-00238-8. PMID: 36991520; PMCID: PMC10061853. KATZ JM, HAKOUN AM, DEHDASHTI AR, CHEBL AB, JANARDHAN V, JANARDHAN V. UNDERSTANDING THE RADIAL FORCE OF STROKE THROMBECTOMY DEVICES TO MINIMIZE VESSEL WALL INJURY: MECHANICAL BENCH TESTING OF THE RADIAL FORCE GENERATED BY A NOVEL BRAIDED THROMBECTOMY ASSIST DEVICE COMPARED TO LASER-CUT STENT RETRIEVERS IN SIMULATED MCA VESSEL DIAMETERS. INTERV NEUROL. 2020 JAN;8(2-6):206-214. DOI: 10.1159/000501080. EPUB 2019 AUG 5. PMID: 32508903; PMCID: PMC7253859. TENG D, PANNELL JS, RENNERT RC, LI J, LI YS, WONG VW, CHIEN S, KHALESSI AA. ENDOTHELIAL TRAUMA FROM MECHANICAL THROMBECTOMY IN ACUTE STROKE: IN VITRO LIVE-CELL PLATFORM WITH ANIMAL VALIDATION. STROKE. 2015 APR;46(4):1099-106. DOI: 10.1161/STROKEAHA.114.007494. EPUB 2015 FEB 24. PMID: 25712942. ON 29-OCT-2024, THIS FILE WAS REEXAMINED, AND THE DETERMINATION HAS BEEN REVISED. THE ASSESSMENTS OF THE MSO AND SENIOR MEDICAL DIRECTOR HAS DETERMINED THE EMBOTRAP III DEVICE NOR PROCEDURE CAUSED OR CONTRIBUTED TO THE EVENT/DEATH; HOWEVER, THE VOICE OF THE CUSTOMER TAKES PRECEDENCE. BASED ON THE TREATING PHYSICIAN¿S STATEMENT, (¿THE BLOOD PRESSURE IN THE ENTIRE CEREBRAL VESSEL MIGHT HAVE INCREASED DUE TO THE MANIPULATION OF THE DEVICES (MICROCATHETER, STENT, ASPIRATION CATHETER) DURING THE THROMBECTOMY, AND THIS COULD CAUSE THE HEMORRHAGE FROM THE VESSEL EVEN THOUGH AT THE OPPOSITE VESSEL WHICH DID NOT DIRECTLY CONTACT WITH THE DEVICES¿), THE CORRELATION BETWEEN THE EVENT TO THE USED EMBOTRAP III DEVICE AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT ENTIRELY. THEREFORE, THIS EVENT WILL BE CONSERVATIVELY REPORTED TO THE US FDA REPORTING UNDER CRITERIA 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED FOR STATISTICAL SIGNALS AND CORRECTIVE / PREVENTIVE ACTION MAY BE TRIGGERED AT A LATER TIME. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE WAS REQUESTING A RETURN LABEL TO RETURN THE DEVICE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
FREEDOM PERIPHERAL NERVE STIMULATOR | K171366 | A PATIENT WAS IMPLANTED WITH A CURONIX PERIPHERAL NERVE STIMULATOR FOR TRIGEMINAL CANCER PAIN ON (B)(6) 2024. THE PATIENT PASSED AWAY ON (B)(6) 2024 DUE TO CANCER. | THE OTHER ADVERSE EVENTS ISSUES QUESTIONNAIRE WAS REVIEWED. THE PATIENT DID NOT EXPERIENCE ANY ADVERSE EVENTS AND WAS RECEIVING GOOD BENEFIT FROM THEIR DEVICE. ALTHOUGH THE PATIENTS DEATH WAS UNRELATED TO THE CURONIX DEVICE, THE DEVICE WAS IMPLANTED IN AN OFF-LABEL LOCATION AS IT WAS IMPLANTED IN THE CRANIOFACIAL REGION TO TREAT FACIAL PAIN. THE TRANSMITTERS ASSOCIATED WITH THE COMPLAINT WERE NOT SENT TO CURONIX HEADQUARTERS FOR ENGINEERING INVESTIGATION. A CURONIX REPRESENTATIVE CONDUCTED A REVIEW OF STERILIZATION AND PACKAGING RECORDS FOR THE RESPECTIVE PRODUCT LOT; CURONIX HAS CONFIRMED THAT THE PRODUCT WAS DELIVERED STERILE, VALIDATED STERILIZATION PARAMETERS WERE USED, AND STERILE BARRIERS WERE VERIFIED TO BE INTACT FOLLOWING PACKAGING. THE STIMULATOR IS USED TO TREAT PAIN. THE CAUSE OF THE PATIENTS DEATH IS UNRELATED TO THE CURONIX DEVICE AS THE PATIENT PASSED AWAY DUE TO CANCER. THE PROVIDED INFORMATION DOES NOT EVIDENCE THAT THE CURONIX DEVICE CONTRIBUTED TO THE ISSUE (NO FAULT FOUND). RATES REVIEWED AT THE MOST RECENT COMPLAINT TRENDING MEETING DO NOT INDICATE A SIGNIFICANT INCREASE IN OTHER ADVERSE EVENTS ISSUES. OTHER ADVERSE EVENTS ISSUES RATES REMAIN ACCEPTABLY LOW; THUS, CAPA IS NOT REQUIRED. OTHER ADVERSE EVENTS ISSUES RATES WILL CONTINUE TO BE TRACKED AND TRENDED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED UNKNOWN INFORMATION. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
SIMPLYGO | K111885 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A SIMPLY GO DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
SIMPLYGO MINI | K111885 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A SIMPLY GO DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER AN ELECTIVE DA VINCI-ASSISTED INGUINAL HERNIA REPAIR, THE PATIENT ¿WOULD NOT WAKE UP FROM ANESTHESIA¿ AND PASSED AWAY. THE PATIENT HAD REPORTEDLY EXPERIENCED CHEST PAIN PRIOR TO THE PROCEDURE. THE CAUSE OF THE CHEST PAIN AND THE PATIENT¿S MEDICAL HISTORY WERE NOT PROVIDED. THE SITE ROBOTICS COORDINATOR INFORMED THE INTUITIVE CLINICAL SALES ASSOCIATE THAT THE PROCEDURE WAS COMPLETED WITHOUT COMPLICATIONS AND THE DA VINCI ROBOT WAS UNDOCKED. THE HOSPITAL OR STAFF ATTEMPTED TO WAKE THE PATIENT BUT WERE UNSUCCESSFUL. CARDIOCEREBRAL RESUSCITATION (CCR) WAS ADMINISTERED AND AN INTERVENTION TEAM WAS CALLED TO ASSIST; HOWEVER, THE PATIENT EXPIRED. THE ROBOTICS COORDINATOR REPORTED THAT THE EVENT HAD NOTHING TO DO WITH THE DA VINCI PRODUCTS. ADDITIONAL INFORMATION WAS REQUESTED BUT HAS NOT BEEN PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, FENESTRATED BIPOLAR FORCEPS, MONOPOLAR CURVED SCISSORS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED SHORTLY AFTER AN INGUINAL HERNIA REPAIR. THE PATIENT HAD CHEST PAINS PRIOR TO SURGERY BUT THE PROCEDURE WAS CARRIED OUT ANYWAY. NO ISSUES OR OTHER COMPLICATIONS OCCURRED DURING THE PROCEDURE. THE CAUSE OF DEATH IS NOT GIVEN BUT THE PREOPERATIVE CHEST PAIN IS CONCERNING FOR A CARDIAC EVENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
RENTAL-DREAMSTATION AUTO CPAP | K131982 | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A RENTAL-DREAMSTATION AUTO CPAP DEVICE STATING PATIENT PASSED AWAY AND NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THERE WAS NO SERIOUS PATIENT HARM OR INJURY. THERE WAS NO MEDICAL INTERVENTION REPORTED. THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THIS NOTIFICATION AND NO INCREASED RISK TO THE INTENDED USER. THIS DEVICE MEETS UL AND IEC REQUIREMENTS FOR FLAMMABILITY. BASED ON THE INFORMATION AVAILABLE, NO MDR WILL BE FILED. IF UPON EVALUATION OF THE DEVICE A REPORTABLE ISSUE/MALFUNCTION IS IDENTIFIED, THIS DECISION WILL BE RE-EVALUATED USING THE DATE OF THE REPAIR EVALUATION AS THE ""BECOME AWARE"" DATE IF A REPORTABLE ISSUE/MALFUNCTION IS IDENTIFIED.
RENTAL-DREAMSTATION AUTO CPAP | K131982 | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A RENTAL-DREAMSTATION AUTO CPAP DEVICE STATING PATIENT PASSED AWAY AND NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THERE WAS NO SERIOUS PATIENT HARM OR INJURY. THERE WAS NO MEDICAL INTERVENTION REPORTED. THE DEVICE WAS NOT RETURNED TO THE MANUFACTURER FOR EVALUATION. THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THIS NOTIFICATION AND NO INCREASED RISK TO THE INTENDED USER. THIS DEVICE MEETS UL AND IEC REQUIREMENTS FOR FLAMMABILITY. BASED ON THE INFORMATION AVAILABLE, NO MDR WILL BE FILED. IF UPON EVALUATION OF THE DEVICE A REPORTABLE ISSUE/MALFUNCTION IS IDENTIFIED, THIS DECISION WILL BE RE-EVALUATED USING THE DATE OF THE REPAIR EVALUATION AS THE ""BECOME AWARE"" DATE IF A REPORTABLE ISSUE/MALFUNCTION IS IDENTIFIED.
MC3 NAUTILUS ECMO OXYGENATOR | K191935 | MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF AN MC3 NAUTILUS EXTRA CORPOREAL MEMBRANE OXYGENATION (ECMO) OXYGENATOR, IT WAS REPORTED THAT THERE WAS A HIGH-PRESSURE EXCURSION EXPERIENCED ON THE MEMBRANE CAUSING LOW FLOW TO PATIENT WHICH RESULTED IN THE PATIENT'S DEATH. ON (B)(6), THE PATIENT WAS PLACED ON VENOARTERIAL (VA) ECMO FOR AN ANTERIOR ST-ELEVATION MYOCARDIAL INFARCTION (STEMI) AT RCH. THE PATIENT EXPERIENCED A PULSELESS ELECTRICAL ACTIVITY (PEA) ARREST. THE TOTAL DOWNTIME WAS 40 MINUTES. THE PATIENT HAD A LEFT ANTERIOR DESCENDING (LAD) BALLOON ANGIOPLASTY. ON (B)(6), THE PATIENT WAS TRANSFERRED TO SPH. ON (B)(6), THEY ATTEMPTED TO WEAN AND IT FAILED. ON (B)(6), THE PATIENT WAS SENT TO THE OPERATING ROOM (OR) FOR A HEARTMATE 3 (HM3) INSERT. THE PATIENT BECAME ACIDOTIC AND THIS LED TO RIGHT VENTRICULAR (RV) FAILURE. A RIGHT VENTRICULAR ASSIST DEVICE (RVAD) WAS INSERTED; HOWEVER, THE PATIENT THEN BECAME HYPOXIC. A RVAD AND OXYGENATOR WERE INSERTED. THE PATIENT WAS OFF HEPARIN SINCE (B)(6) AT 8.40A.M. THE RVAD WAS CHATTERING WITH A FLOW OF 2.0 LPM AND THE HM3 WAS FLOWING AT 2.9 LPM. AT 16.04P.M, THE FLOW WAS INCREASED TO 2.5 LPM. THE REVOLUTIONS PER MINUTE (RPM) WAS AT 2000. THE PRE-OXYGENATOR PRESSURE WAS 175 MMHG AND THE POST-OXYGENATOR PRESSURE WAS 145 MMHG. THE RVAD FRACTION OF INSPIRED OXYGEN (FIO2) INCREASED FROM 60% TO 70%. THE HM3 FLOW WAS AT 3.1 LPM. AT 16.05P.M, THE CHATTER AND RVAD FLOWS DROPPED TO 2.0-2.1 LPM. THEY REVIEWED THE CHART AND NOTICED THE INCREASE IN DELTA P. THEY RE-ZEROED THE PRE-MEMBRANE AND POST-MEMBRANE PRESSURE TRANSDUCERS WITH NO CHANGE IN PRESSURE. AT 16.16P.M, THE DELTA P'S WERE IN THE SIXTIES. THE PRE-OXYGENATOR PRESSURE WAS AT 178MMHG AND THE POST-OXYGENATOR PRESSURE WAS 118MMHG. AT 16.18P.M, THE DELTA P WAS 80 AND THE RVAD FLOW WAS 1.8 LPM. THE HM3 HAD A LOW FLOW ALARM. AT 16.20P.M, THEY CALLED THE CUSTOMER BACK FOR A CIRCUIT CHANGE. THE RPM INCREASED TO 2300 WITH NO CHATTER. AT 16.21P.M, THE PRE-OXYGENATOR PRESSURE WAS 280 MMHG AND THE POST-OXYGENATOR PRESSURE WAS 78MMHG. THE FLOW WAS 1.4 LPM WITH NO CHATTER. THE PRE-OXYGENATOR PRESSURE WAS IN THE LOW 300'S. THE RVAD FLOW WAS AT 0.8 LPM. THEY PREPPED THE LINES AND EQUIPMENT FOR A CIRCUIT CHANGE. THE RVAD FLOW WAS AT 0.6 LPM. THE PHYSICIAN (DR. (B)(6)) ENTERED AND INFORMED THE TEAM TO NOT TO CHANGE CIRCUIT AS PER THE OTHER PHYSICIAN (DR. (B)(6) DUE TO FUTILITY. AT 16.28P.M, THE RVAD AND ECMO WERE TURNED OFF AND THE PATIENT DIED. THERE WERE MULTIPLE TRANSFUSIONS ADMINISTRATED TO THE PATIENT FROM 15.25P.M TILL 1P.M. THE NEXT DAY. THE PATIENT WAS GIVEN 40ML OF PROTHROMBIN COMPLEX 6 TIMES, DIFFERENT AMOUNTS OF PLATELETS AT MULTIPLE TIMES, 50ML FIBRINOGEN TWICE, 180 ML PLATELETS, MULTIPLE MLS OF (RED BLOOD CELLS) RBC AND 500/250ML OF ALBUMIN 5% MULTIPLE TIMES. THE PATIENT'S INTERNATIONAL NORMALIZED RATIO (INR) INCREASED TO 3.3, THEIR PARTIAL THROMBOPLASTIN TIME (PTT) WAS GREATER THAN 140 AND RECORDED AT 135 AND THEIR PLATELET COUNT WAS 25 AT 14.26P.M ON (B)(6). THERE FIBRINOGEN WAS LESS THAN 0.6 AT 19.48P.M ON (B)(6). MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT WHEN THE PATIENT ARRIVED AT ST (B)(6) HOSPITAL ON (B)(6), THE PATIENT WAS ON THE MAQUET QUADROX OXYGENATOR. THE PATIENT WAS RECEIVED FROM RCH. THE FIRST NAUTILUS OXYGENATOR WAS USED ON (B)(6) 2024. TWO NAUTILUS OXYGENATORS WERE USED. ON (B)(6) 2024, THE PATIENT WAS PLACED ON VA ECMO WITH A QUADROX OXYGENATOR FOR 10 DAYS. THERE WAS NO OXYGENATOR CHANGE INDICATED ON TRANSFER SHEET AT THIS POINT. ON (B)(6) 2024, THE PATIENT WAS PLACED ON VA ECMO WITH A NEW CIRCUIT AND NAUTILUS OXYGENATOR FOR 3 DAYS/70 HOURS. ON (B)(6) 2024, THE PATIENT WAS ON CARDIOPULMONARY BYPASS (CPB) WHILE USING AN AFFINITY FUSION OXYGENATOR FOR 123 MINUTES. ON (B)(6) 2024, THE PATIENT WAS ON RVAD WITH A NEW CIRCUIT AND NO OXYGENATOR FOR 54 MINUTES. ON (B)(6) 2024, THE PATIENT WAS ON RVAD WITH A NEW CIRCUIT AND NAUTILUS OXYGENATOR FOR 1 DAY/21 HOURS. THE LAST OXYGENATOR USED WAS THE COMPLAINT DEVICE. THERE WAS NO ANTICOAGULANTS, OTHER THAN HEPARIN, USED FOR THE PATIENT. HEPARIN WAS DISCONTINUED ON (B)(6) 2024 AT 8:40A.M. THE RVAD PLUS OXYGENATOR INITIATED ON (B)(6) 2024 AT 7:04 P.M. THERE WAS NO CLOTS VISUALIZED ON THE COMPLAINT DEVICE. THERE WAS CLOTS VISUALIZED ON THE PRE-MEMBRANE OXYGENATOR AND ¼" DISTAL LIMB PERFUSION LINE ON THE VA ECMO CIRCUIT WHICH WAS INITIATED ON (B)(6) 2024. THERE WAS NO CIRCUIT COMPONENTS CHANGED OUT ON (B)(6) 2024 DUE TO PATIENT FUTILITY AS PER SURGEON. THE PATIENT'S HEMATOCRIT AT INITIATION OF ECMO AND AT THE ONSET OF THE TIME THE EVENT WAS FIRST SUSPECTED ON (B)(6) 2024 AT 7.04P.M. WAS 22.9L/L (HCT). HEMOGLOBIN (HGB) WAS 75G/DL AT 6.28P.M. THE PATIENT'S HEMATOCRIT WAS 28L/L AND HGB WAS 95G/DL AT 7.48P.M. ONSET OF EVENT FIRST SUSPECTED, ON (B)(6) 2024 AT 4.05-4.10 P.M, THE PATIENT'S HEMATOCRIT WAS 34L/L AND HGB WAS 115G/DL. THERE WAS NO MENTION OF A DELTA P INCREASE DURING HANDOVER FOR BREAK RELIEF AT 3.52P.M. THE LAST ACT VALUE WAS RECORDED AT 162S ON (B)(6) 2024 AT 3.46P.M. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT (B)(6) HOSPITAL.
CRESCENT RA DUAL LUMEN CATHETER 13FR | K203409 | MEDTRONIC RECEIVED INFORMATION INFORMATION VIA LITERATURE REGARDING THE USE OF A CRESCENT CATHETER. IT WAS REPORTED THAT THE CATHETER HAD A FRACTURE, CRACK, OR BREAK, AND THE PATIENT SUFFERED MASSIVE AIR ENTRAINMENT INTO THE EXTRACORPOREAL CIRCUIT, RESULTING IN PATIENT HARM AND/OR DEATH. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. ONE OF THE CASES IDENTIFIED IN THE LITERATURE WAS PREVIOUSLY REPORTED TO MEDTRONIC, 3011468686-2024-00014. THE 2ND CASE IDENTIFIED IN THE LITERATURE WAS NOT FOUND TO HAVE BEEN REPORTED TO MEDTRONIC PREVIOUSLY. THIS REPORT IS FOR THE 2ND CASE. ADDITIONAL INFORMATION COULD NOT BE OBTAINED FROM THE HOSPITAL AND THE DETAILS OF THE 2ND CASE ARE UNCONFIRMED. | THE REPORT EXPLAINS THAT THERE IS NO SUSPICION OF MANUFACTURING DEFECT PRESENT AND THAT THE LENGTH OF USE AND DEVICE POSITION MAY CONTRIBUTE TO THE DEVICE FAILURE. THE MITIGATION STRATEGIES AND POTENTIAL EARLY WARNING SIGNS OUTLINED IN THE REPORT SHOULD BE CONSIDERED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND CHRONIC LYMPHOCYTIC LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION BIPAP AVAP | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, PANCREATIC AND LIVER AND KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION BIPAP ST30 | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, REDUCED CARDIOPULMONARY RESERVE, STROKE, HEART ATTACK AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, NEUROENDOCRINE (LIVER AND SPINE). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, ANAPLASTIC CANCER AND MULTIPLE MYELOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING AND CANCER CUTANEOUS T-CELL LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, CANCER, DEATH, LUNG CANCER AND BRAIN CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING) DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, REDUCED CARDIOPULMONARY RESERVE, HEART FAILURE/SUDDEN EMERGENCY SITUATION, RESULTING IN VFIB AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, LUNG CANCER PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RENAL CELL CARCINOMA, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER (DIFFUSE LARGE B-CELL LYMPHOMA) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | IN THE PERVIOUS REPORT DATE OF DEATH (RFB) WAS CAPTURE INCORRECTLY, IT IS CORRECTED IN THIS REPORT DATE OF DEATH (RFB) - NO INFORMATION.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND SURGERY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ZIO AT | K163512 | THE PATIENT EXPIRED DURING THEIR PRESCRIBED ZIO AT WEAR PERIOD. IRHYTHM ATTEMPTED TO GATHER MORE INFORMATION ABOUT THE CAUSE OF DEATH FROM THE ACCOUNT, BUT NO FURTHER DETAILS WERE PROVIDED. | A REVIEW OF THE COMPLAINT INDICATED THAT ON (B)(6) 2024, THE ¿BASELINE REPORT¿ (INITIAL TRANSMISSION REPORT) WAS RECEIVED, AND THE DEVICE TRIGGERED A SERVICE MONITORING TICKET SHOWING NO CONNECTION FOR A 4-HOUR PERIOD. THE DEVICE RECONNECTED BUT CONTINUED TO EXPERIENCE INTERMITTENT CONNECTIVITY ISSUES FROM (B)(6); HOWEVER, THE PATIENT¿S DAILY REPORT SHOWS RECORDED TRANSMISSIONS DURING THIS TIME. AN ACTIONABLE ARRHYTHMIA WAS IDENTIFIED ON (B)(6) 2024, AND PER STANDARD PROTOCOL, IRHYTHM NOTIFIED THE ACCOUNT APPROPRIATELY AND WAS INFORMED THAT THE ON-CALL ELECTROPHYSIOLOGIST WOULD BE PAGED AND THAT THE PATIENT HAD EXPIRED. AVAILABLE CLINICAL DATA SHOWED THE PATIENT WORE THE DEVICE FOR APPROXIMATELY 4 DAYS OF THE 14-DAY PRESCRIBED WEAR PERIOD. A REVIEW OF THE PATIENT¿S DAILY REPORTS SHOWED TRANSMISSION FROM (B)(6). THE ZIO AT WAS NOT RETURNED TO IRHYTHM FOR EVALUATION; THEREFORE, IRHYTHM CANNOT CONFIRM IF A DEVICE MALFUNCTION OCCURRED. HOWEVER, BASED ON THE AVAILABLE INFORMATION, NO KNOWN MISSED OR MISCLASSIFIED MDNS LIKELY ATTRIBUTED TO OR CAUSED THE EVENT. THIS REPORT IS BEING SUBMITTED OUT OF AN ABUNDANCE OF CAUTION. THIS EVENT IS BEING REPORTED PER 21CFR803 AS A SERIOUS INJURY. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY RHYTHM THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED. THESE TERMS ARE INCLUDED IN FORM FDA 3500A AND ARE FIXED TERMS FOR SELECTION CREATED BY THE FDA TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REGULATORY REPORTING.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAD DIED WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE DEVICE IS NOT PART OF RECALL.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LIVER DISEASE/TOXICITY AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. IN BOX H: TYPE OF REPORTED COMPLAINT IS UPDATED IN THIS FOLLOW-UP.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. UPON FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN), IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6)2024 WHILE CONNECTED TO HER LIBERTY SELECT CYCLER. THE PATIENT¿S CAUSE OF DEATH WAS FROM A MYOCARDIAL INFARCTION (MI) THAT OCCURRED DURING A PD TREATMENT ON THE EVENING OF (B)(6)2024. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS PRONOUNCED DECEASED IN HER HOME WITH NO TRANSPORT TO HOSPITAL. THE PATIENT DID NOT HAVE ANY KNOWN HISTORY OF CARDIAC DISEASE AND DID NOT EXPERIENCE ANY PRIOR MIS. THE PDRN STATED THE PATIENT¿S PHYSICIAN RULED OUT ANY INVOLVEMENT WITH PD THERAPY IN RELATION (OTHER THAN TEMPORAL) TO THE PATIENT¿S MI AND SUBSEQUENT DEATH. IT WAS CONFIRMED, THOUGH THE EXACT CAUSE OF THE PATIENT¿S MI WAS NOT REPORTED, THERE WAS NO INDICATION THE PATIENT¿S MI LEADING TO HER DEATH WAS RELATED TO PD THERAPY OR DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CAN BE ATTRIBUTED TO AN MI WITH NO INDICATION THE MI AND SUBSEQUENT DEATH WERE DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALSO REPORTED "DEATH". AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
STAR¿ DRIVE MA | K092562 | ON OCTOBER 30, 2024 FHC WAS CONTACTED BY NEUROSURGERY PROGRAM COORDINATOR AT EVERGREEN HEALTHCARE REPORTING THAT THEY HAD 3 PATIENTS RECENTLY WHO UNDERWENT DBS SURGERIES THAT DEVELOPED INFECTIONS. IN THE COURSE OF INVESTIGATING THESE ISSUES, THEY FOUND THAT ONE OF OUR STAR DRIVES WAS USED IN EACH SURGERY. THE FACILITY IS REVIEWING ALL THE POTENTIAL FACTORS AND THAT INCLUDES INSPECTION OF ALL EQUIPMENT. AS SUCH, THE HOSPITAL REQUESTED A REVIEW OF PROCESSING OF THE PRODUCT FROM FHC FIELD TECHNICIANS WITH THEIR STAFF. ADDITIONALLY, THEY NOTED THAT THEY WOULD BE SENDING THE PRODUCT TO FHC FOR PREVENTIVE MAINTENANCE BECAUSE IT WAS OVERDUE. ON 11/5/2024 FHC RECEIVED THE FOLLOWING ADDITIONAL INFORMATION FROM THE HOSPITALS NEUROSURGERY PROGRAM COORDINATOR: 1. (B)(6) 2024 LEAD PLACEMENT, PRESENTED TO ED WITH S/S OF INFECTION ON (B)(6), REMOVAL OF LEAD ON (B)(6). STAPH AUREUS, SERRATIA MARSESCENS. PT UNDERWENT COURSE OF IV ANTIBIOTICS 2. ON (B)(6) 2024 LEAD PLACEMENT, (B)(6) REPORTS S/S OF INFECTION, (B)(6) REMOVAL OF LEAD. COAG NEG STAPH AND NORMAL SKIN FLORA, YEAST. COURSE OF IV ANTIBIOTICS 3. ON (B)(6) 2024 LEAD PLACEMENT, (B)(6) S/S OF INFECTION, SERRATIA MARSESCENS. LEAD REMOVAL (UNKNOWN DATE), PT DECEASED. THERE WAS NO INDICATION OF ANY FAILURE OR DEFECT RELATED TO FHC PRODUCT. THIS PRODUCT IS NON-PATIENT CONTACTING. | FHC IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY FHC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. FHC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, FHC, OR ITS EMPLOYEES THAT THE DEVICE, FHC OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. FHC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
CVAC ASPIRATION SYSTEM | K233472 | ON (B)(6)2024, A MORBIDLY OBESE PATIENT WITH SUSPECTED SEPSIS AND SIGNIFICANT CO-MORBIDITIES (PRIOR HISTORY OF URINARY TRACT AND KIDNEY INFECTIONS, LYMPHEDEMA, ANEMIA), WHO HAD UNDERGONE A PREVIOUS RIGHT PCNL UNDERWENT A RIGHT SIDE URETEROSCOPY STONE REMOVAL PROCEDURE. INITIAL INSPECTION OF THE COLLECTING SYSTEM REVEALED DEBRIS, BLOOD CLOTS, AND VISCOUS FLUID IN ADDITION TO KIDNEY STONES. THE CVAC SYSTEM WAS USED TO CLEAR BLOOD CLOTS AND DEBRIS PRIOR TO INITIATING LASER LITHOTRIPSY. LASER LITHOTRIPSY AND STONE REMOVAL WAS THEN PERFORMED WITH THE CVAC SYSTEM AND UNKNOWN LASER FIBER. AN UNKNOWN STONE RETRIEVAL BASKET WAS ALSO USED FOR STONE FRAGMENT REMOVAL. THE MAJORITY OF THE PROCEDURE WAS CONDUCTED IN THE RIGHT KIDNEY UPPER POLE. APPROXIMATELY 2 CM OF STONE BURDEN WAS REMOVED FROM THE KIDNEY, WITH INITIAL PROCEDURAL DIFFICULTY CAUSED BY VISIBILITY CHALLENGES DUE TO CLOUDY URINE, INFLAMMATORY DEBRIS, AND A BLOOD CLOT FROM A PREVIOUS PCNL PROCEDURE (LOWER POLE), THAT MAY HAVE REDUCED IRRIGANT OUTFLOW RATE FROM THE KIDNEY. THE CVAC SYSTEM WAS REMOVED FROM THE PATIENT BETWEEN 5 AND 7 TIMES DURING THE PROCEDURE TO CHECK FOR CLOGGING, AND NO CLOGS OR OTHER DEVICE ISSUES WERE OBSERVED OR REPORTED. AT THE END OF THE PROCEDURE, THE CVAC ASPIRATION SYSTEM AND OTHER ACCESSORIES WERE REMOVED WITHOUT INCIDENT AND AN UNKNOWN URETERAL STENT WAS PLACED ON THE RIGHT SIDE. A URETERAL STENT WAS THEN INSERTED ON THE PATIENT'S LEFT SIDE. A FOLEY CATHETER WAS INSERTED AND THE PROCEDURE WAS CONCLUDED. THE ANESTHESIOLOGIST NOTED A DECREASE IN BLOOD PRESSURE. WHILE IN RECOVERY THE PATIENT EXPERIENCED FLANK PAIN, A HYPOTENSIVE EPISODE, HEMATURIA AND WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). ULTRASOUND REVEALED FLUID AROUND THE KIDNEY AND A BLEEDING CONCERN. AS REPORTED, THE PATIENT DEVELOPED HEMORRHAGIC SHOCK AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC). IN ADDITION, THERE WAS CONCERN FOR UROSEPSIS. THE PATIENT WAS TRANSFERRED TO AN OPERATING ROOM AND INTERVENTIONAL RADIOLOGY ATTEMPTED RENAL ARTERY EMBOLIZATION. THE PATIENT'S CONDITION WORSENED. AN EXPLORATORY LAPAROTOMY WITH NEPHRECTOMY WAS INITIATED AND A WOUND WAS OBSERVED IN THE UPPER POLE OF THE RIGHT KIDNEY. THE PATIENT RECEIVED 15 UNITS OF BLOOD; HOWEVER, THE PATIENT DEVELOPED CARDIOVASCULAR ARREST AND EFFORTS TO RESUSCITATE THE PATIENT WERE UNSUCCESSFUL. ON 02-DEC-2024 ADDITIONAL INFORMATION WAS RECEIVED FROM THE HOSPITAL THAT ACCORDING TO A REVIEW OF THE PHYSICIAN'S DOCUMENTATION IN THIS PATIENT'S MEDICAL RECORD, THE CAUSE OF DEATH WAS HEMORRHAGIC SHOCK AND DISSEMINATED INTRAVASCULAR COAGULATION (DIC). THERE WAS NO POST-MORTEM EXAMINATION PERFORMED. ON (B)(6)2024 FDA NOTIFIED CALYXO INC. OF THE HOSPITAL'S MEDSUN REPORT (UF# (B)(4)). | CORRECTION A2: PATIENT AGE WAS UPDATED. CORRECTION B5: UPDATED INFORMATION RELATED TO THE EVENT. CORRECTION E4: CHANGED TO YES AS HOSPITAL FILED A MEDSUN REPORT. (UF# (B)(4)). A REVIEW OF THE LOT HISTORY RECORD FOR LOT# P02084 WAS CONDUCTED, WHICH CONFIRMED THAT THERE WERE NO NON-CONFORMANCES DURING THE MANUFACTURING PROCESS THAT COULD BE RELATED TO THE REPORTED EVENT. THE REVIEW INDICATED THAT THE DEVICE MET ALL DESIGN AND MANUFACTURING SPECIFICATIONS WHEN RELEASED FOR DISTRIBUTION.
REM STAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER AND NON HODGKINS LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, KIDNEY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX E: HEALTH PROFESSIONAL (RFB) WAS CORRECTED. BOX H: MANUFACTURE DATE (RFB) WAS CORRECTED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, INFLAMMATORY RESPONSE, ASTHMA (NEW OR WORSENING), DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | K211575 | IT HAS BEEN REPORTED THAT ON (B)(6) 2024 THE PATIENT WAS HOSPITALIZED AT THE EMERGENCY ROOM (ER) OF (B)(6) HOSPITAL AFTER A FIGHT THAT CAUSED A SEVERE NOSEBLEED. FOR ISSUES UNRELATED TO DIABETES, THE PATIENT WAS TRANSFERRED AND HOSPITALIZED TO (B)(6) HOSPITAL IN (B)(6). ON (B)(6), THE PATIENT¿S CEREBRAL DEATH WAS DECLARED AND ON (B)(6), THE PATIENT WAS DECEASED. ACCORDING TO THE PATIENT'S HEALTH CARE PROFESSIONAL (HCP) FROM (B)(6) HOSPITAL, A BOLUS DOSE WAS ADMINISTERED USING THE PATIENT'S DASH PERSONAL DIABETES MANAGER (PDM) BY THE PATIENT¿S FAMILY OR FRIENDS, PRIOR TO THE PATIENT¿S ADMISSION TO ER. THE HCP STATED TO INSULET DURING A FOLLOW-UP CALL ON (B)(6) 2024 THAT THE PATIENT¿S BLOOD GLUCOSE VALUES AT THE TIME OF THE BOLUS ADMINISTRATION ARE UNKNOWN. REPORTEDLY, THE PATIENT WAS NOT WEARING HIS CONTINUOUS GLUCOSE MONITORING SENSOR AT THE TIME OF ADMISSION. THE HCP STATED TO INSULET THAT IT IS UNKNOWN HOW THE PDM WAS USED AND UNCONFIRMED THAT IT WAS USED CORRECTLY. INSULET IS REQUESTING FOR THE PDM TO BE RETURNED FOR INVESTIGATION. | FURTHER INFORMATION ON THE EVENT IS CURRENTLY BEING SOLICITED. IF ADDITIONAL INFORMATION IS PROVIDED IT WILL BE SUBMITTED IN A FOLLOW UP REPORT. THE DEVICE HAS BEEN REQUESTED TO BE RETURNED FOR EVALUATION. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR PRO SERIES | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION LIVER DISEASE/TOXICITY, CANCER, DEATH THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND CHRONIC BRONCHITIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RECEIVED BY THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND THROAT CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
INFINITY CENTRAL STATION | K061379 | ON OCTOBER 30TH, THE FSE RECEIVED A SERVICE CALL TO CHECK THE ALARMS AT THE INFINITY CONTROL CENTER (ICS). WHILE ON SITE THE CUSTOMER INFORMED THE FSE THAT A PATIENT HAD SUFFERED A CARDIAC INCIDENT ON (B)(6) AROUND 5:30. THE PATIENT WAS FOUND IN THEIR ROOM WITH THE M300 ALARMING, AND WAS TRANSFERRED TO THE CARDIOLOGY INTENSIVE CARE UNIT, THERE THE PATIENT PASSED AWAY. THE CUSTOMER REPORTED THAT THE M300 DID NOT ALARM ACOUSTICALLY AT THE ICS, AND SUSPECTED THAT SOME OF THE ALARMS WERE NOT WORKING PROPERLY. | THE INVESTIGATION HAS STARTED. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION.
BIPAP AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
UNKNOWN_VOCERA_PRODUCT | K130339 | IT WAS REPORTED BY THE CUSTOMER THAT AN ADVERSE EVENT WHERE A PATIENT ENDED UP PASSING AWAY WAS REPORTED. IT WAS REPORTED THAT ALERTS WERE POTENTIALLY NOT RECEIVED. TECHNICAL SUPPORT WILL BE PERFORMING A FULL RCA OF WORKFLOW AND CLIENT LEVEL EVENTS.
INTELLIVUE MX40 2.4GHZ | K113125 | IT WAS REPORTED THE PATIENT PASSED AWAY DUE TO POTENTIAL TELEMETRY ISSUES. THE USER ROUNDED ON THE PATIENT WITH THE PHYSICIAN AT 0950 AND FOUND THE PATIENT UNRESPONSIVE. THE PATIENT'S RHYTHM WAS ASYSTOLE AND A CODE TEAM WAS CALLED TO THE BEDSIDE, BUT RESUSCITATION WAS UNSUCCESSFUL. A NURSE SAW THE PATIENT ALIVE AROUND 0830 AND TELEMETRY STOPPED DISPLAYING AT 0901. ALL ECG WIRES WERE ATTACHED TO THE PATIENT AND THE BATTERY WAS IN PLACE, WITH NO LOW BATTERY ALARM BEING NOTED BY NURSING STAFF. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS HAVE BEEN INITIATED. THE DEVICE WAS IN USE ON A PATIENT. THERE WAS A REPORT OF PATIENT OR USER HARM. | D4: THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. (B)(6). A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
PATIENT INFORMATION CENTER IX | K211900 | THE CUSTOMER REPORTED THAT THE STAFF DID NOT HEAR THE ALARM FROM THE EMERGENCY ROOM SURVEILLANCE AND THE PATIENT PASSED AWAY. THE ENGINEER TESTED THE SYSTEM AND FOUND IT TO BE IN WORKING ORDER. THE UNIT MANAGER REQUESTED THAT THE VOLUME OF ALARMS BE INCREASED, THE ENGINEER INCREASED THE VOLUME FROM 4 TO 6. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | A PHILIPS FIELD SERVICE ENGINEER WAS DISPATCHED TO THE CUSTOMER SITE. THE MONITOR AND THE PATIENT INFORMATION CENTER WERE BOTH SUCCESSFULLY TESTED FOR ALARMING AND FOR SOUND. AUDIT LOGS WERE RETRIEVED AND SHARED WITH THE HOSPITAL ADMINISTRATION. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
INTELLIVUE MX500 PATIENT MONITOR | K131872 | THE CUSTOMER REPORTED THAT THE STAFF DID NOT HEAR THE ALARM FROM THE BEDSIDE MONITOR AND THE PATIENT PASSED AWAY. THE ENGINEER TESTED THE SYSTEM AND FOUND IT TO BE IN WORKING ORDER. IN ADDITION, THE FACILITY BIOMED REPORTED THERE WAS TAPE AND GAUZE WRAPPED AROUND THE SPEAKER OF THE OVERVIEW FOUND NEXT TO THE PATIENT ROOM, WHICH WAS REMOVED PRIOR TO THE ENGINEER'S ARRIVAL. THE UNIT MANAGER REQUESTED THAT THE VOLUME OF ALARMS BE INCREASED, THE ENGINEER INCREASED THE VOLUME FROM 4 TO 6. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | A PHILIPS FIELD SERVICE ENGINEER WAS DISPATCHED TO THE CUSTOMER SITE. THE MONITOR AND THE PATIENT INFORMATION CENTER WERE BOTH SUCCESSFULLY TESTED FOR ALARMING AND FOR SOUND. AUDIT LOGS WERE RETRIEVED AND SHARED WITH THE HOSPITAL ADMINISTRATION. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH STAGE 4 METASTATIC CANCER OF THE LIVER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A NURSE REPORTED A PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2024. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING OF (B)(6) 2024. THE PATIENT¿S DAUGHTER DISCOVERED THE PATIENT HAD EXPIRED WHEN ARRIVING TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA INDICATED THE PATIENT WAS IN THE FINAL DWELL CYCLE WHEN IT STOPPED, AND NO PRIOR ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO PATIENT'S ADVANCED AGE AND SIGNIFICANT CARDIAC HISTORY. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. THE PDRN REPORTED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, DUE TO HER POOR CARDIAC HEALTH AND ADVANCED AGE. THE SERIOUS ADVERSE EVENTS WERE REPORTEDLY UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
FLOWTRIEVER CATHETER, XL | K211013 | A 73-YEAR-OLD FEMALE ARRIVED AT THE HOSPITAL 2 WEEKS PRIOR DUE TO DEEP VEIN THROMBOSIS (DVT). THE PATIENT HAS A HISTORY OF DVT, ATRIAL FIBRILLATION, AND HAD A RECENT HYSTERECTOMY. THE HYSTERECTOMY LED TO BLEEDING COMPLICATIONS WHICH RESULTED IN THE PATIENT DISCONTINUING THEIR BLOOD THINNERS. AFTER THE PATIENT WAS ADMITTED TO THE HOSPITAL, THEY WERE PLACED ON HEPARIN, THIS LED TO HER BLOOD COUNT DROPPING. THE HEPARIN WAS DISCONTINUED, AND AN INFERIOR VENA CAVA (IVC) FILTER WAS PLACED. THE PATIENT RETURNED TO THE HOSPITAL DUE TO CONTINUED SYMPTOMS. A COMPUTED TOMOGRAPHY (CT) SCAN REVEALED MINIMAL PULMONARY EMBOLISM (PE) AND EXTENSIVE CLOT FROM THE LEFT ILIAC INTO THE IVC WHERE THE FILTER WAS PRESENT. THE DECISION WAS MADE TO PERFORM A DVT THROMBECTOMY AFTER REMOVING THE IVC FILTER. THE PATIENT WAS FLIPPED SUPINE, AND ACCESS WAS ACHIEVED. THE IVC FILTER WAS REMOVED THROUGH THE INTERNAL JUGULAR AND THE FLOWTRIEVER CATHETER, XL (FTXL) WITH 2 DISCS DEPLOYED. THE CLOTTRIEVER XL (CTXL) CATHETER WAS USED TO CLEAR THE IVC BUT, THE DISTAL RADIOPAQUE TIP FROM THE FTXL GOT CAUGHT ON THE CTXL. TO ADDRESS THIS, THE FTXL DISCS WERE BRIEFLY DROPPED TO PULL THE CTXL DOWN AND THE PHYSICIAN REDEPLOYED THE DISCS. THE THROMBECTOMY WAS CONTINUED USING THE CTXL AND CLOTTRIEVER BOLD CATHETER (CT BOLD). THE PATIENT BEGAN TO COMPLAIN OF SHORTNESS OF BREATH AND BRIEFLY BECAME UNRESPONSIVE. IMAGING WAS PERFORMED WHICH REVEALED A LACK OF CLOT DESPITE ANY INTERVENTION. AT THIS POINT, THE PATIENT BEGAN TO DECOMPENSATE, AND IMAGING REVEALED AN ENLARGED RIGHT VENTRICLE WITH A CLOT-IN-TRANSIT (CIT). ATTEMPTS WERE MADE TO GO ASPIRATE THE CIT, BUT THE PATIENT CODED. THE PHYSICIAN WENT AFTER THE PATIENT'S PE FIRST AND 2 ASPIRATIONS WERE PERFORMED. UNFORTUNATELY, CATHETER POSITIONING WAS LOST WHEN CARDIOPULMONARY RESUSCITATION (CPR) BEGAN. TPA WAS ADMINISTERED AND CPR CONTINUED FOR 20 MINUTES. THE PATIENT WAS UNABLE TO BE RESUSCITATED AND PASSED AWAY. | THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE NOT PROVIDED, THE LOT HISTORY RECORDS OF ALL POTENTIAL LOTS SHIPPED TO THE FACILITY WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS LIKELY THE RESULT OF CLOT EMBOLISM THAT OCCURRED WHEN THE FTXL DISKS WERE TEMPORARILY DROPPED. DISTAL EMBOLISM IS LISTED IN THE DEVICE LABELING AS A KNOWN POTENTIAL COMPLICATION/ADVERSE EVENT. MANUFACTURER REFERENCE #: (B)(4).
LINEAR 7.5FR. 40CC IAB | K041281 | IT WAS REPORTED THAT DURING INTRA-AORTIC BALLOON (IAB) THERAPY, (B)(6), THE CHIEF PERFUSIONIST REPORTED THAT THE PATIENT WENT INTO CARDIAC ARREST IN THE FIELD AND ARRIVED TO THE CCL AS A STEMI RECEIVING CPR. ACLS CONTINUED THROUGHOUT THE ENTIRETY OF THE CASE, NEVER ACHIEVING ROSC. HE SAID THAT THE PHYSICIAN PLACED THE IAB AS USUAL, THE TEAM HOOKED UP THE CONSOLE, PRESSED START TO INITIATE THERAPY AND IMMEDIATELY GOT AN AUTOFILL FAILURE ALARM. (B)(6) SAID THEY COMPLETED THE FOLLOWING TROUBLESHOOTING STEPS, ATTEMPTING TO INITIATE THERAPY AFTER EACH INTERVENTION: VERIFIED ALL LINES WERE APPROPRIATELY CONNECTED AND SECURE, PHYSICIAN REPOSITIONED THE CATHETER, CHANGED OUT ECG AND PRESSURE CABLES AND CHANGED THE CONSOLE. AFTER THESE INTERVENTIONS WERE UNSUCCESSFUL AND THE AUTOFILL FAILURE ALARM CONTINUED, RICK PERFORMED A SYRINGE FILL WHICH WAS SUCCESSFUL AS THEY COULD SEE THE IAB INFLATING UNDER FLUORO. THE GETINGE REPRESENTATIVE REVIEWED THE DIFFERENCE BETWEEN THE CARDIOSAVE USING IT¿S PRESSURE ALGORITHM TO FILL THE IAB VERSUS A SYRINGE FILL PERFORMED WITH NO RESISTANCE. (B)(6) SAID THEY REMOVED THE IAB AFTER THE PATIENT PASSED AND HE WAS ABLE TO GET IT TO PUMP OUTSIDE OF THE BODY. HE ALSO TESTED THE IAB ON THE FIRST CONSOLE THEY USED WHICH ALSO WORKED. HE THEN SWITCHED OVER TO HIS TRAINER IAB WHICH ALSO WORKED ON BOTH CONSOLES. THIS VERIFIED THE CONSOLES WERE NOT CAUSING THE AUTOFILL FAILURE TO TRIGGER. BASED ON THE TROUBLESHOOTING STEPS TAKEN AND THE PATIENT¿S CLINICAL PICTURE, THE GETINGE REPRESENTATIVE EXPLAINED THAT THE ALARM COULD HAVE BEEN TRIGGERED BY SEVERAL DIFFERENT FACTORS INCLUDING AN ANATOMICAL ABNORMALITY OR A SMALL KINK AT THE INSERTION SITE. | ADDITIONAL REPORTER: ALEXANDER COLON, CHIEF PERFUSIONIST EVENT SITE FULL NAME: (B)(6). THE DEVICE WAS NOT RETURNED AND COULD NOT BE EVALUATED. IT WAS DISCARDED BY THE USER. WE ARE UNABLE TO CONFIRM THE REPORTED EVENT. IF NEW INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. COMPLAINT RECORD ID # (B)(4).
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, FATIGUE AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING ACUTE RESPIRATORY FAILURE WITH HYPERCAPNIA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
| K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, DIZZINESS, HEADACHE, NAUSEA, VOMITING, CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD), ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LIVER DISEASE /TOXICITY, RESPIRATORY FAILURE, BREATHING PROBLEMS, LUNG DAMAGE, COUGHING UP PHLEGM, RESPIRATORY TRACT IRRITATION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
| K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH AND LUNG THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
| K083526 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH MUSCULAR DYSTROPHY PASSED AWAY WHILE USING A TRILOGY 100 VENTILATOR. IT WAS REPORTED THAT THE AIDE LEFT AROUND 0130 AND FAMILY WAS ASLEEP. IT APPEARS THE VENTILATOR STOPPED FUNCTIONING AROUND 0214 DUE TO INTERNAL BATTERY BEING DEPLETED. AROUND 0600 THE FAMILY DISCOVERED THE VENT WAS NOT ON, AND THE POWER CORD WAS IN THE VENT STAND BASKET. THE PATIENT WAS DECEASED. EMERGENCY MEDICAL SERVICES WAS CALLED. THE DEVICES SETTINGS AT THE TIME OF EVENT WERE REPORTED TO BE A MODE OF SIMV, RESPIRATORY RATE OF 16, INSPIRATORY TIME OF 1.0, PEEP OF 5, TIDAL VOLUME OF 500, PRESSURE SUPPORT OF 20, TRIGGER SET TO AUTO, RISE TIME OF 3. CIRCUIT DISCONNECT ALARM OF 30, AND LOW MINUTE VENTILATION ALARM SET TO 0.2. THE CAUSE OF DEATH WAS REPORTED TO BE THE VENTILATOR SHUT OFF DUE TO THE DEVICE BEING UNPLUGGED, WITH BATTERIES DEPLETED, AND PATIENT WAS NOT ABLE TO BREATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. UPON EVALUATION THE DEVICES APNEA ALARMS WERE NOTED TO BE SET TO OFF. IN REVIEW OF THE VENTILATORS' ERROR LOGS, THE LOG CONFIRMED THE OCCURRENCE OF A LOW EXHALED TIDAL VOLUME ALARM, LOW BATTERY ALARMS, AND A LOW MINUTE VOLUME ALARM. DURING THE EVALUATION IT WAS NOTED THAT BOTH THE DETACHABLE AND INTERNAL BATTERIES WERE PRESENT AND RECOGNIZED BY THE UNIT. IT WAS VERIFIED THAT BOTH THE INTERNAL AND DETACHABLE BATTERY WERE COMMUNICATING WITH THE UNIT, IN GOOD CONDITION, AND CHARGING PROPERLY WITH THE POWER CORD THAT WAS RECEIVED WITH THE UNIT. DURING EVALUATION IT WAS ALSO NOTED THAT THERE WERE NO SIGNS OF DAMAGE OR EVIDENCE OF DEFECTIVE OPERATION. THE TECH DID NOT DETECT ANY PINCHED OR BLOCKED TUBING. TECH ALSO FOUND THE AIRPATH FOAM TO BE FIRMLY SECURED, WITHOUT ANY SIGNS OF DELAMINATION OR DEGRADATION. THE TECH DID OBSERVE FOREIGN MATERIAL FROM EXTERNAL SOURCES INDICATIVE OF CONTAMINATION THROUGHOUT THE SUBASSEMBLIES, INCLUDING INSIDE THE TRANSITION TUBING AND FLOW SENSOR ASSEMBLY COMPONENTS. TECH ALSO NOTED AN INDETERMINATE ODOR INDICATIVE OF NICOTINE THAT ACCOMPANIED THE SUBASSEMBLIES. IN CONCLUSION, IT HAS BEEN DETERMINED THAT TRILOGY 100 VENTILATOR (TV117101005) OPERATES AS DESIGNED AND FUNCTIONS AS EXPECTED. THE SENSOR BOARD OF THE UNIT DRIFTED OUT OF SPEC DUE TO CONTAMINATION ISSUES, AND THE INTERNAL AND EXTERNAL BATTERIES CHARGED AND FUNCTIONED APPROPRIATELY.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE, TOXICITY, REDUCED CARDIOPULMONARY RESERVE, ENT, BONE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED DUE TO COPD. THE FAMILY STATED THAT HER CONDITION GOT WORSE WHILE USING THE DEVICE. THERE ARE OTHER ALLEGATIONS THAT THE PATIENT PASSED OUT DUE TO LACK OF OXYGEN AND ALSO HAD A PULMONARY OBSTRUCTION. THE PATIENT SPENT TIME IN THE HOSPITAL AND UNDERWENT TEST. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND STOMACH CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF SERIOUS PATIENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN B5 VERBIAGE WHICH WAS CORRECTED IN THIS REPORT.
ENDO ANCHOR SYSTEM - HELI-FX AAA | K140036 | AN ENDURANT IIS STENT GRAFT WAS IMPLANTED DURING THE ENDOVASCULAR TREATMENT OF A RUPTURED 45MM AAA. IT WAS REPORTED THAT THE PATIENT HAD A TYPE I ENDOLEAK AT THE END OF THE PROCEDURE. IT WAS REPORTED APPROXIMATELY 2 YEARS POST THE INDEX PROCEDURE THE PATIENT WAS TREATED FOR A RUPTURE DUE TO THE TYPE IA. 10 ENDOANCHORS WERE PLACED, HOWEVER ON THE POST-ANGIO THE TYPE IA ENDOLEAK REMAINED. PER THE PHYSICIAN THE CAUSE OF THE ENDOLEAK WAS DUE TO UNDER-SIZING. NO ADDITIONAL CLINICAL SEQUALAE WERE PROVIDED AND THE PATIENT WILL BE MONITORED. | ADDITIONAL INFORMATION RECEIVED: IT WAS REPORTED THAT THE MAIN BODY AND LIMBS WERE EXPLANTED. AS THE PATIENT WAS MOVED TO THE HOSPITAL BED, THE PATIENT HAD A HEART ATTACK AND EXPIRED IN THE OR ON (B)(6) 2024. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS VALVULAR HEART DISEASE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, LARGE B CELL NON-HODGKIN'S LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
OMNIPOD 5 POD | K203768 | AN EMAIL WAS RECEIVED FROM THE PATIENT'S HEALTH CARE PROFESSIONAL (HCP), REPORTING THAT THE PATIENT PASSED AWAY ON (B)(6)2024. REPORTEDLY THE PATIENT PASSED AWAY FROM HYPERGLYCEMIA AFTER HE 'FOLLOWED THE STEPS FOR TROUBLE SHOOTING WITH CHANGING A POD, BUT MAY HAVE BEEN TOO LATE AND DID NOT SEEK MEDICAL CARE'. | ACCORDING TO THE REPORT A PATIENT DEATH OCCURRED DUE TO HYPERGLYCEMIA. THE HCP CONFIRMED THE PATIENT SUCCESSFULLY PERFORMED TROUBLESHOOTING STEPS AND THE POD WAS REPLACED. THE PATIENT DID NOT SEEK MEDICAL ATTENTION WHEN TROUBLESHOOTING. THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. IF ADDITIONAL INFORMATION IS RECEIVED A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION FROM CUSTOMER IN REFERENCE TO A DS2ADV AUTO CPAP WITH AN ALLEGATION PATIENT PASSED AWAY DUE TO THE DEVICE. THERE WAS NO REPORT OF SERIOUS INJURY OR HARM. THERE IS NO MEDICAL INTERVENTION WAS SPECIFIED. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, AND CONCLUSION CODE GRID WAS UPDATED.
TRICLIP G4 SYSTEM | K241538 | SUBSEQUENT TO THE PREVIOUS REPORT, THE ADDITIONAL INFORMATION AS RECEIVED: THE CAUSE OF DEATH WAS SEVERE ANEMIA. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS AND A REVIEW OF THE LOT HISTORY RECORD REVIEW COULD NOT BE PERFORMED AS THE PART AND LOT INFORMATION REGARDING THE COMPLAINT DEVICE WERE NOT PROVIDED. BASED ON THE INFORMATION REVIEWED, A CAUSE FOR THE REPORTED PAIN, FATIGUE, DYSPNEA, HEMORRHAGE, HEMATOMA, AND SUBSEQUENT ANEMIA RESULTING IN DEATH CANNOT BE DETERMINED. HEMORRHAGE, HEMATOMA, PAIN, DYSPNEA AND DEATH ARE LISTED IN THE INSTRUCTIONS FOR USE (IFU) AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH TRICLIP PROCEDURES. UNEXPECTED MEDICAL INTERVENTION AND HOSPITALIZATION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
MITRACLIP | K190167 | IT WAS REPORTED THAT ON (B)(6) 2024, A MITRACLIP PROCEDURE WAS PERFORMED TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) GRADE 4+. XTW (LOT: 40429R1067) WAS CHOSEN FOR IMPLANTATION UTILIZING STEERABLE GUIDE CATHETER (SGC LOT: 40509R1001). WHEN INSERTING THE CLIP INTO THE SGC, PERICARDIAL EFFUSION WAS OBSERVED. WHEN THE CLIP WAS ADVANCED TO THE TIP OF THE SGC, TAMPONADE WAS OBSERVED. THE MITRACLIP PROCEDURE WAS ABANDONED AND PERICARDIOCENTESIS ALONG WITH SURGICAL INTERVENTION WAS PERFORMED. PATIENT WAS VERY SICK WITH MULTIPLE COMORBIDITIES AND INTER ATRIAL SEPTUM WAS FRAIL. DEATH RESULTED ON (B)(6) 2024. IT WAS THOUGHT THE PERICARDIAL EFFUSION AND SUBSEQUENT PE WAS LIKELY DUE TO THE TRANSEPTAL PUNCTURE. | IN THIS CASE, THERE WAS NO FUNCTIONAL TESTING THAT WAS NEEDED FOR THE RETURNED DEVICE AS THERE WAS NO REPORTED DEVICE MALFUNCTION ASSOCIATED WITH THE SGC. THE SGC HOWEVER, WAS INSPECTED AND THERE WERE NO ISSUES NOTED. A REVIEW OF THE LOT HISTORY RECORD IDENTIFIED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THE REPORTED EVENT. BASED ON THE INFORMATION REVIEWED AND PER THE PHYSICIAN, THE REPORTED PERICARDIAL EFFUSION AND CARDIAC TAMPONADE RESULTING IN DEATH APPEAR TO BE RELATED TO PROCEDURAL CONDITIONS AND DUE TO TRANSEPTAL PUNCTURE. IT WAS CONFIRMED BY THE ACCOUNT THAT THE DEVICE USED FOR THE TRANSEPTAL PUNCTURE WAS A NON-ABBOTT DEVICE AND THAT THE MITRACLIP AND THE SGC DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED DEATH. DEATH, PERICARDIAL EFFUSION AND CARDIAC TAMPONADE ARE LISTED IN THE INSTRUCTIONS FOR USE (IFU) AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH MITRACLIP PROCEDURES. UNEXPECTED MEDICAL INTERVENTION AND SURGICAL INTERVENTION WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE, TOXICITY, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
BIPAP AUTO BIFLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LIVER DISEASE, TOXICITY, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, NAUSEA, VOMITING, LIVER DISEASE, TOXICITY, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DAVINCI XI | K131861 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS THAT COULD HAVE CONTRIBUTED TO THE REPORTED EVENT OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, LUNG CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY IN 2023 AND THE PATIENT'S WIFE IS REQUESTING TO BE REMOVED FROM FURTHER CORRESPONDENCE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY DUE TO THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED AND DUST LIKE CONTAMINATION WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS SEVEN ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION, BUT DUST-LIKE CONTAMINATION WAS OBSERVED AND UNIT GOT CORRECTED. DEVICE PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID WAS UPDATED.
ALLURA XPER FD | K162859 | IT HAS BEEN REPORTED TO PHILIPS THAT DURING PREPARATION FOR A DIAGNOSTIC PROCEDURE, THE X-RAY COULD NOT BE TRIGGERED VIA THE WIRED AND WIRELESS FOOTSWITCHES. LOGFILE ANALYSIS, PERFORMED BY THE PHILIPS REMOTE SERVICE ENGINEER (RSE) AND ONSITE ANALYSIS PERFORMED BY THE FIELD SERVICE ENGINEER (FSE) CONFIRMED THE SYSTEM INTERFACE BOX (SIB BOX) WAS NOT WORKING AS INTENDED. NO ISSUE WAS FOUND WITH THE WIRED OR WIRELESS FOOTSWITCHES. REPLACEMENT OF THE SIB BOX RESOLVED THE ISSUE AND THE SYSTEM WAS RETURNED TO USE IN GOOD WORKING ORDER. THE PATIENT WAS MOVED TO ANOTHER ROOM FOR THE PROCEDURE. THE REPORT ALLEGED THAT THE PATIENT DIED; HOWEVER, NO FURTHER INFORMATION REGARDING THE DEATH HAS BEEN PROVIDED TO DATE. PHILIPS WILL CONTINUE ITS INVESTIGATION OF THIS COMPLAINT. | PHILIPS IS INVESTIGATING THIS COMPLAINT. ACCORDING TO THE ADDITIONAL INFORMATION COLLECTED, THE ISSUE OCCURRED WHILE THE PATIENT WAS BEING PREPARED FOR A NEUROVASCULAR TREATMENT FOR A CEREBRAL HEMORRHAGE. THE PATIENT¿S CAUSE OF DEATH WAS REPORTED TO BE CEREBRAL HEMORRHAGE. PHILIPS WILL CONTINUE ITS INVESTIGATION OF THIS COMPLAINT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH, CONSTANT RESPIRATORY INFECTIONS, SINUS INFECTIONS, AGGRAVATED ASTHMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PRLNE BLU 24IN 6-0 D/A BV-1 | K133356 | IT WAS REPORTED THAT A PATIENT WITH A VENTRICULAR SEPTAL DEFECT UNDERWENT SURGERY ON (B)(6) 2024 AND SUTURE WAS USED. ON (B)(6) 2024, AN ULTRASOUND/ ECHOCARDIOGRAPHY WAS PERFORMED. ON (B)(6) 2024, A REPEAT SURGERY WAS PERFORMED BECAUSE THE THREADS HAD BROKEN. ONE SUTURE ON THE TRICUSPID VALVE IN THE AREA OF THE ANTERIOR VALVE WAS UNRAVELED. THE PATIENT'S CONDITION IS SERIOUS. THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT OF THE CHILDREN'S HOSPITAL FOR APPROPRIATE INTENSIVE CARE IN THE POSTOPERATIVE PERIOD. ADDITIONAL INFORMATION WAS REQUESTED. | PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H6 COMPONENT CODE: G07002 - DEVICE NOT RETURNED. H6 INVESTIGATION FINDINGS: C22 - PHOTO REVIEW. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO RELATED NON-CONFORMANCES WERE IDENTIFIED. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. PLEASE PROVIDE THE CHILD'S DEMOGRAPHIC INFORMATION INCLUDING AGE, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE. ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION (NORMAL, THIN, CALCIFIED, FRAGILE, DISEASED)? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE TIED (SQUARE KNOT OR MULTIPLE KNOTS ONE END)? PLEASE DESCRIBE THE APPEARANCE OF THE SUTURE DURING THE SECOND PROCEDURE. WERE THERE ANY PATIENT STRESS FACTORS THAT PRECIPITATED THE EVENT OF SUTURE BREAKAGE POST OP? WHAT SYMPTOMS DID THE PATIENT EXPERIENCE FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? WHAT IS THE PATIENT'S CURRENT STATUS? WILL PRODUCT BE RETURNED? IF SO, PLEASE PROVIDE THE RETURN DATE AND TRACKING INFORMATION. SURGEON¿S NAME? INVESTIGATION SUMMARY : THIS IS AN ANALYSIS FOR A PHOTO SUBMITTED TO ETHICON FOR EVALUATION. DURING THE VISUAL ANALYSIS, THE FOLLOWING WAS OBSERVED: THE PHOTO SHOWS A BOX LABELED AS W8712 OPENED. THE IMAGE IS NOT CLEAR TO DETERMINE THE FAILURE MODE OR THE REPORTED CONDITION. BASED ON THE PHOTO REVIEW, NO CONCLUSION OR ROOT CAUSE COULD BE DETERMINED. HANDS ON DEVICE ANALYSIS MAY PROVIDE THE ADDITIONAL EVIDENCE NECESSARY TO CONFIRM THE ROOT CAUSE OF THE REPORTED EVENT. BECAUSE THE DEVICE WAS NOT RETURNED OUR EVALUATION IS LIMITED. WE VALUE THE OPPORTUNITY TO FULLY ANALYZE THE DEVICE UPON ITS RETURN. AS PART OF ETHICON QUALITY PROCESS ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS WILL BE CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF POST MARKET SURVEILLANCE.
SAFARI2¿ | K151244 | MEDICAL REVIEW SUMMARY RECEIVED FROM THE DISTRIBUTOR 02DEC2024: RECEIVED IMAGES: ANGIOGRAMS OF THE PROCEDURE. PRE-PROCEDURE CT SCAN NOT AVAILABLE. FEMORAL ANGIOGRAPHY PERFORMED, SHOWING PUNCTURE AT AN APPROPRIATE LEVEL, WITH NO EVIDENCE OF CONTRAST EXTRAVASATION. AORTIC ARCH ANGIOGRAM BEFORE THE INSTALLATION OF THE CEREBRAL EMBOLIC PROTECTION DEVICE. CONTRAST INJECTION AT THE AORTIC VALVE LEVEL, PIGTAIL IN THE NON-CORONARY SINUS, SHOWING A TRICUSPID AORTIC VALVE AND PATENT CORONARY ARTERIES. IN THE FOLLOWING SEQUENCE, GUIDEWIRE MANIPULATION IS OBSERVED WITH THE AID OF A PIGTAIL CATHETER. LEFT VENTRICULOGRAPHY WITH PIGTAIL CATHETER, ADEQUATE MYOCARDIAL CONTRACTILITY. PROGRESSION OF THE ISLEEVE TO THE ABDOMINAL AORTA, WITHOUT APPARENT DIFFICULTY. AORTIC VALVULOPLASTY WITH BALLOON UNDER RAPID PACING, POSSIBLY PERFORMED WITH DIRECT STIMULATION UNDER SAFARI2 GUIDEWIRE, WITHOUT PACEMAKER IMAGE IN THE RIGHT VENTRICULAR CAVITY. SAFARI IN SUBOPTIMAL POSITION DUE TO CURVATURE NOT FACING UP. PRE-DILATION APPARENTLY EFFECTIVE WITH AORTIC DISPLACEMENT OF THE STILL-INFLATED BALLOON. IN THE NEXT IMAGE, WE SEE THE DS ADVANCING TOWARDS THE AORTIC ARCH. NEXT, APPARENT LEFT VENTRICULAR HYPOCONTRACTILITY. IN THE IMAGE, THE VALVE IS STILL CLOSED, WITH CONTRAST REFLUX INTO THE LEFT VENTRICLE AND AN IMAGE SUGGESTIVE OF CONTRAST EXTRAVASATION THROUGH THE AORTIC VALVE APPARATUS. IN THE FOLLOWING IMAGES, A CLEAR CONTRAST EXTRAVASATION IS NOTED, POSSIBLY THROUGH THE LVOT, AND AN INCREASE IN CARDIAC RADIOPACITY, POSSIBLY DUE TO CONTRAST ACCUMULATION IN THE PERICARDIAL SAC TOPOGRAPHY. SUBSEQUENTLY, PERICARDIAL DRAINAGE PROCEDURE IS OBSERVED, AND THE ACCURATE NEO2 VALVE ALREADY DEPLOYED (THERE IS NO RECORD OF THE PROSTHESIS RELEASE PROCESS). CARDIOPULMONARY RESUSCITATION MANEUVERS FOLLOW, AND PACEMAKER POSITIONING IN THE RIGHT VENTRICLE IS NOTED. PERICARDIAL DRAINAGE APPARENTLY SUCCESSFUL, BUT SEVERE LEFT VENTRICULAR HYPOCONTRACTILITY REMAINS. FINAL AORTOGRAPHY SHOWING ADEQUATE VALVE IMPLANTATION HEIGHT, WITH PIGTAIL IN THE PERICARDIAL SAC, PATENT CORONARY ARTERIES, AND LEFT VENTRICULAR HYPOCONTRACTILITY. RESUMPTION OF CARDIOPULMONARY RESUSCITATION MANEUVERS. CONCLUSION: TAVR PROCEDURE COMPLICATED BY POSSIBLE LEFT VENTRICULAR PERFORATION, POSSIBLY DUE TO PRE-DILATION, FOLLOWED BY CARDIAC TAMPONADE. I DO NOT SEE ELEMENTS TO ATTRIBUTE ITS OCCURRENCE TO SAFARI2, BUT RATHER TO THE ACT OF PRE-DILATION ITSELF. PRE-PROCEDURE CT SCAN COULD HELP VERIFY ANNULAR MEASUREMENTS AND THE SUITABILITY OF THE BALLOON USED IN AORTIC VALVULOPLASTY." | SECTION B5 INCLUDES THE ADDITIONAL INFORMATION. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED.
EDWARDS ESHEATH+ INTRODUCER SET | K200258 | PER THE FIELD CLINICAL SPECIALIST (FCS), DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE USING A 29MM SAPIEN 3 ULTRA RESILIA VALVE A A PRE-BAV WAS PERFORMED THROUGH ESHEATH W/ A 20MM TRUE BALLOON. THEN LOADED THE 29MM SAPIEN3 ULTRA RESILIA, PREPPED PER IFU TO NOMINAL SPECS, INTO THE 16F ESHEATH, AND SOME SLIGHT RESISTANCE NEAR THE LATERAL CIRCUMFLEX FEMORAL ARTERY (LCFA) AREA, BUT THE VALVE PASSED. THE VALVE WAS TRACKING THROUGH AND ALL OF SUDDEN STOPPED. ON A CLOSER LOOK AT FLUORO, THE VALVE SEEMED TO BE OUTSIDE THE ESHEATH AND WHEN PUSHING THE DELIVERY SYSTEM, THE FLEX CATHETER SEEMED TO BE BUCKLING AND PROLAPSING, POSSIBLY OUTSIDE THE SHEATH. THE IMPLANTER OPTED TO TRY AND TAKE THE WHOLE SYSTEM OUT AS ONE UNIT AND VALVE WAS STUCK AND WOULD NOT COME OUT. A DECISION WAS MADE TO PLACE AN AORTIC OCCLUSION BALLOON THROUGH CONTRALATERAL ACCESS. THE PATIENT'S BLOOD PRESSURE WAS DROPPING RAPIDLY, AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. THE SURGEON NEEDED TO GET ACCESS TO GO ON CARDIOPULMONARY BYPASS (CPB) AND A STERNOTOMY WAS PERFORMED, AND THE PATIENT WAS CANNULATED. THE HEART TEAM WAS ABLE TO REMOVE THE VALVE AND DELIVERY SYSTEM EVENTUALLY WITH ADDITIONAL INCISIONS DOWN THE ABDOMEN AND LEFT GROIN AREA. THE PATIENT WAS STABILIZED, AND THE SURGICAL TEAM OPTED TO PLACE A NEW EDWARDS SURGICAL VALVE WHILE STERNOTOMY WAS AVAILABLE. THE PATIENT WAS TRANSFERRED OUT OF THE HYBRID OR HEMODYNAMICALLY STABLE. FOLLOW-UP INDICATED THAT THE IMPLANTER FELT THE ANATOMY, SEAM ON THE SHEATH, CALCIUM, AND VALVE ALL LINED UP PERFECTLY WHICH MAY HAVE ALLOWED THE DELIVERY SYSTEM (DS) TO PROTRUDE FROM THE SHEATH. THE DELIVERY SYSTEM MAY HAVE CAUGHT ON CALCIUM ON ONE TURN IN THE VESSEL AND THEN MAY HAVE JUMPED FORWARD AFTER. IT APPEARED THAT AFTER THEY TRIED TO REMOVE THE VALVE AND SYSTEM THERE WAS A BENT STRUT ON THE OUTFLOW SIDE. THE BENT STRUT WAS NOTICED ON FLUORO AND MAY HAVE HAPPENED WHEN PULLING BACK. THE PATIENT EXPIRED 11-12 DAYS AFTER THE PROCEDURE. THE DEVICES ARE NOT AVAILABLE FOR RETURN. | THE INVESTIGATION IS ONGOING.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE DEATH OF A PATIENT WHILE USING THE DEVICE DUE TO RESPIRATORY FAILURE WITH HYPOXIA, AND COPD. THE DEVICE SETTING WAS WEANING MODE PASSIVE, PEEP 5, PRESSURE SUPPORT 15, AND RISE TIME 2. THE CUSTOMER INQUIRED TO, ¿WHEN THEY DID A DOWNLOAD REPORT THE DATA DID NOT MATCH THE VENTILATOR DATA ON THE SCREEN AND THEREFORE WOULD LIKE THIS UNIT EVALUATED TO DETERMINE IF THERE WERE ANY ISSUES WITH THE UNIT AS A NORMAL PROCESS. THEY ARE NOT CLAIMING THAT THIS PATIENT DIED AS A RESULT OF A FAILURE OF THIS DEVICE.¿ ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH" THE COMPLAINT HAS BEEN REVIEWED BY A CLINICAL EXPERT AND CONFIRMED, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS BEEN RETURNED TO A THIRD-PARTY SERVICE CENTER, BUT EVALUATION HAS NOT YET BEGUN. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE SERVICE CENTER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED ALLEGING THE DEATH OF A PATIENT WHILE USING THE DEVICE DUE TO RESPIRATORY FAILURE WITH HYPOXIA, AND COPD. THE DEVICE SETTING WAS WEANING MODE PASSIVE, PEEP 5, PRESSURE SUPPORT 15, AND RISE TIME 2. THE CUSTOMER INQUIRED TO, ¿WHEN THEY DID A DOWNLOAD REPORT THE DATA DID NOT MATCH THE VENTILATOR DATA ON THE SCREEN AND THEREFORE WOULD LIKE THIS UNIT EVALUATED TO DETERMINE IF THERE WERE ANY ISSUES WITH THE UNIT AS A NORMAL PROCESS. THEY ARE NOT CLAIMING THAT THIS PATIENT DIED AS A RESULT OF A FAILURE OF THIS DEVICE.¿ ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH" THE COMPLAINT HAS BEEN REVIEWED BY A CLINICAL EXPERT AND CONFIRMED, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE WAS EVALUATED AT THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY (PIL). THE INVESTIGATION FINDINGS DID NOT UNCOVER ANY DETAILS THAT WOULD CHANGE OR MODIFY THE RISK OF THE DEVICE. THE PIL POWERED ON THE TRILOGY EVO, USA DEVICE UTILIZING THE EXISTING DEVICE SETTINGS AND RAN THERAPY FOR APPROXIMATELY 16 HOURS. THE DEVICE OPERATED WITHOUT ISSUE. PIL THEN POST TESTED THE DEVICE ON THE PIL TRILOGY EVO MULTIFUNCTIONAL TEST STATION AND PASSED ALL TESTING. PIL CONCLUDED THAT THE TRILOGY EVO, USA DEVICE OPERATES AS DESIGNED. ROOT CAUSE NOT APPLICABLE.
PRESSIO ICP AND ICT MONITORING KIT, PARENCHYMAL WITH BOLT | K162108 | IMPLANTATION ON (B)(6) 2024 ON A HYPERTENSIVE PATIENT WITH RIGHT DEEP INTRACEREBRAL HEMATOMA. THE POSITION OF THE ICP CATHETER WAS CONFIRMED BY CT SCAN ON (B)(6) 2024. AGGRAVATION OF INTRACRANIAL HYPERTENSION AT DAY 4 DIAGNOSED 1 DAY LATE, DUE TO CATHETER MIGRATION IN THE BOLT. THE CATHETER SLIPPED OUT OF THE BRAIN AND NO LONGER CORRECTLY MEASURED ICP (ICP AT 5MMHG BUT NOT PULSATILE). THE PATIENT EXPERIENCED ARTERIAL HYPOTENSION WITH AN AVERAGE OF 50 MMHG, BRADYCARDIA AT 35 BPM AND BILATERAL AREACTIVE MYDRIASIS. CATHETER REPLACEMENT ON (B)(6) 2024 SHOWING ICP AT 30MMHG. DEATH OF THE PATIENT OCCURS ON (B)(6) 2024. | NO CORRECTIVE OR PREVENTIVE ACTION HAS BEEN TAKEN SINCE WE DON'T KNOW AT WHICH MOMENT THE PROBLEM OCCURS. BUT A SENSIBILIZATION OF ALL THE OPERATOR HAS BEEN MADE. RÜTHER ANALYSIS MUST BE PERFORMED TO DETERMINE WHEN THE BOLT HAS BEEN BROKEN (AT THE MANUFACTURING SITE OR AT THE HOSPITAL). THE MANUFACTURER WILL CONTACT THE REPORTER TO INVESTIGATE ON THIS PART. ON THE SAME TIME, THE QUALITY TEAM WILL CONTROL BOLT OF PRESSIO KITS THAT ARE READY TO BE SEND TO CUSTOMER TO SEE IF THERE IS ANY BOLT BROKEN.
DREAMSTATION CPAP | K131982 | IT WAS REPORTED THAT THE DREAMSTATION CPAP DEVICE WAS THROWN AWAY A LONG TIME AGO AND THE USER HAD PASSED AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
PACEL BIPOLAR PACING CATHETER | K152784 | DURING A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE, A CARDIAC PERFORATION OCCURRED, AND THE PATIENT EXPIRED ONE DAY POST PROCEDURE. ON (B)(6) 2024, A PACEL CATHETER WAS PLACED DURING THE VALVE IMPLANT PROCEDURE AND CONNECTED TO AN EXTERNAL PACEMAKER GENERATOR. THE CATHETER WAS PLACED WITHOUT ANY DIFFICULTY. FOLLOWING THE PROCEDURE, THE PATIENT WAS TRANSFERRED TO THE ICU. THE FOLLOWING DAY, THE PATIENT EXPERIENCED HEMODYNAMIC DECOMPENSATION AND WAS DIAGNOSED WITH CARDIAC TAMPONADE SECONDARY TO A PERFORATION AT THE SITE OF THE PACEL (ANTERIOR SURFACE OF THE RIGHT VENTRICLE). THE PATIENT WAS TRANSFERRED TO SURGERY FOR A PERICARDIAL WINDOW PROCEDURE, HOWEVER THE PATIENT EXPIRED. THERE WERE NO PRODUCT PERFORMANCE ISSUES WITH ANY ABBOTT DEVICE. | THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED CARDIAC PERFORATION AND SUBSEQUENT DEATH REMAINS UNKNOWN. PER THE IFU, CARDIAC PERFORATION IS A KNOWN RISK DURING THE USE OF THIS DEVICE.
UNKNOWN_VOCERA_PRODUCT | K130339 | IT WAS REPORTED BY THE CUSTOMER THAT AN ADVERSE EVENT WHERE A PATIENT ENDED UP PASSING AWAY WAS REPORTED. LOGS ARE BEING REQUESTED TO GET FURTHER INFORMATION. | A FUNCTIONAL INSPECTION WAS PERFORMED BY STAFF TECHNICAL SUPPORT ENGINEER. AFTER REVIEWING DEVICE LOGS, IT WAS FOUND THAT THE ALARM NOTIFICATIONS WERE DELIVERED APPROPRIATELY, AND THAT THE ESCALATION PATHWAY FUNCTIONED AS EXPECTED. THE INVESTIGATION CONCLUDED THAT THE PRODUCT WAS BEHAVING AS EXPECTED, AND DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT¿S DEMISE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DECEASED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6) 2024 AT 0255. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 1119." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6)2024 AT 0255. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 1119." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6) 2024 AT 02:55. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 11"19." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THE DEVICE ALARMED "COMMUNICATION ERROR" AND "NO LONGER WORKING" DURING INFUSION OF LEVOPHED (8MG/250ML), VASOPRESSIN (50 UNITS/50ML), AND EPINEPHRINE (4MG/250ML) ON A PATIENT ADMITTED FOR ACUTE SEPTIC SHOCK. WHILE THE CLINICIAN WAS WAITING TO RECEIVE NEW DEVICES, THE PATIENT REPORTEDLY "LOST PULSE, AND CPR (CARDIOPULMONARY RESUSCITATION) WAS STARTED." PER REPORT, THE "PATIENT DID CODE DURING REPORTED INCIDENT REQUIRING INTUBATION, THEN ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS ACHIEVED." THE EVENT OCCURRED ON (B)(6) 2024 AT 0255. THE MEDICATIONS HAD REPORTEDLY BEEN INFUSING "WITHOUT INCIDENT" PRIOR TO THE REPORTED EVENT. THE PATIENT WAS THEN "MADE COMFORT CARE LATER IN THE DAY AND EXPIRED ON (B)(6) 2024 (AT) 1119." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DREAMSTATION AUTO CPAP | K131982 | A DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. DURING THE EVALUATION, SECONDARY FINDINGS WERE OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION RECEIVED ON 02/05/2025. H11 UPDATED AS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DS2ADV AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H6 HAS BEEN UPDATED.
DREAMSTATION CPAP PRO | K131982 | IT WAS REPORTED THAT THE DREAMSTATION CPAP PRO DEVICE WAS RETURNED TO THE PATIENT'S DME, AND THE USER HAD PASSED AWAY A YEAR AND A HALF AGO. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT'S WIFE REQUESTED THAT NO FURTHER CORRESPONDENCE BE SENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | K163000 | IT WAS REPORTED THAT ON (B)(6) 2024, AN UNKNOWN SIZE OCCLUDER WAS CHOSEN FOR IMPLANT USING A 12F AMPLATZER AMULET DELIVERY SHEATH. THE SHEATH WAS PREPARED AS PER INSTRUCTIONS FOR USE (IFU). DURING PROCEDURE, TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) REVELED AIR EMBOLISM OCCURRED DURING WIRE/INTRODUCER TO OCCLUDER EXCHANGE. BACK THE PATIENT COULD NOT BE SEDATED SUFFICIENTLY AND HIS BREATHING WAS IRREGULAR. DESPITE INSTRUCTIONS WITHIN THE INSTRUCTIONS FOR USE (IFU), THE SHEATH WAS NOT PULLED BACK INTO THE RIGHT ATRIUM (RA). AIR WAS DETECTED IN THE LEFT VENTRICLE (LV). TWO EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) SUPPORT WERE GIVEN, TEMPORARY PACEMAKER WAS FLOATED IN AND A DEFIBRILLATOR WAS USED, MEDICATIONS SUCH AS ADRENALINE, NOREPINEPHRINE, DOBUTAMINE, SALINE, AND AMIODARONE WERE ADMINISTERED DURING RESUSCITATION. THE SEDATION WAS INTERRUPTED AND GENERAL ANESTHESIA WAS IMPLEMENTED DURING REANIMATION. AN ASPIRATION WAS PERFORMED TO SUCK THE AIR OUT OF THE VENTRICLE USING A PIGTAIL CATHETER. THE PATIENT HAD VENTRICULAR FIBRILLATION, ELECTRONIC DECOUPLING, DILATED PUPILS, AND VERY LOW BLOOD PRESSURE AND THE CASE WAS ABORDED. THE PHYSICIAN MENTIONED THE CAUSE OF AIR EMBOLUS WAS THE PATIENT'S DEEP, SPONTANEOUS INHALATION PROBABLY RESULTED IN NEGATIVE VENOUS PRESSURE AT THE SHEATH. THE SHEATH WAS PROBABLY NOT FAR ENOUGH BELOW THE HEART LEVEL FOR THIS NEGATIVE PRESSURE. THE PATIENT WAS TRANSFERRED TO TO THE INTENSIVE CARE UNIT UNDER ALL LIFE SUPPORT MEASURES. | INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
LUCAS¿ CPR CHEST COMPRESSION SYSTEM | K090422 | THE CUSTOMER CONTACTED STRYKER, TO REPORT THAT THEIR LUCAS DEVICE WAS NOT LATCHING ONTO THE BACK OF THE COMPRESSION BOARD. THE CUSTOMER WAS UNABLE TO USE THE DEVICE. THE PATIENT INVOLVED IN THE REPORTED EVENT IS DECEASED. | STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. STRYKER CONTACTED THE CUSTOMER, TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS, PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED STRYKER WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. STRYKER PERFORMED A CLINICAL REVIEW OF THE REPORTED EVENT AND DETERMINED, THAT THERE WAS INSUFFICIENT INFORMATION PROVIDED, THAT THE DEVICE DID NOT CONTRIBUTE TO THE PATIENT OUTCOME.
FMC CASSETTE | K173718 | IT WAS REPORTED TO FRESENIUS THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER AND LIBERTY CYCLER SET FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE HOSPITALIZED ON (B)(6) 2024 DUE TO PERITONITIS. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) REVEALED THE PATIENT CONTACTED THE OUTPATIENT HOME DIALYSIS CLINIC ON (B)(6) 2024 WITH COMPLAINTS OF DIZZINESS, HYPOTENSION, DEHYDRATION, NAUSEA AND VOMITING. THE PATIENT REPORTED THEY WERE DEHYDRATED AFTER USING 2.5% DIALYSATE FROM RUNNING OUT OF 1.5% DIALYSATE AND WAS ADVISED BY THE PDRN TO GO TO THE EMERGENCY ROOM (ER). UPON ARRIVAL, A PERITONEAL EFFLUENT FLUID CULTURE AND CELL COUNT (WBC = 6,321 U/L) WERE OBTAINED, AND THE PATIENT WAS FORMALLY ADMITTED. THE PATIENT WAS PROPHYLACTICALLY TREATED WITH VANCOMYCIN AND ZOSYN (DOSAGES, ROUTE, FREQUENCY, DURATION NOT PROVIDED) WHILE IN THE ER, IN ADDITION TO POTASSIUM, ZOFRAN, NORMAL SALINE (500 ML), AND ASPIRIN. THE SPECIFICS REGARDING THE HOSPITALIZATION ARE LARGELY UNKNOWN, HOWEVER THE PDRN REPORTED PERITONITIS WAS RULED OUT AND THE PATIENT WAS DISCHARGED HOME ON (B)(6) 2024. ON (B)(6) 2024, THE PERITONEAL EFFLUENT FLUID CULTURES FROM (B)(6) 2024 RETURNED POSITIVE FOR GRAM-NEGATIVE PANTOEA AND CANDIDA ALBICANS (FUNGUS). UPON RECEIPT OF THE CULTURE RESULTS, THE NEPHROLOGIST ORDERED A REPEAT PERITONEAL EFFLUENT FLUID CULTURE AND CELL COUNT, AND THE PATIENT WAS PRESCRIBED DIFLUCAN (DOSE, ROUTE, FREQUENCY, DURATION NOT PROVIDED) THE SAME DAY. THE PATIENT OBTAINED THE DIFLUCAN ON (B)(6) 2024, HOWEVER IT IS UNKNOWN IF THE PATIENT RECEIVED A DOSE. ON (B)(6) 2024, THE PDRN RECEIVED THE PERITONEAL CELL COUNT RESULTS, AND THE PATIENT¿S WBC COUNT WAS CRITICALLY HIGH (8,491 U/L). THE PATIENT WAS ADVISED TO GO TO THE ER, HOWEVER THE PATIENT DECLINED UNTIL (B)(6) 2024 DESPITE THE SEVERITY OF THEIR SYMPTOMS AND EDUCATION. UPON READMISSION, THE PATIENT¿S PRELIMINARY PERITONEAL EFFLUENT FLUID CULTURES RETURNED POSITIVE FOR MOLD AND CANDIDA PARAPSILOSIS, AND THE PATIENT¿S PD CATHETER (NOT A FRESENIUS PRODUCT) WAS SURGICALLY REMOVED. ALTHOUGH THE TIMELINE AND SPECIFICS WERE NOT PROVIDED, IT APPEARS THE PATIENT WAS SEPTIC UPON READMISSION AND SUFFERED AN UNSPECIFIED CARDIAC EVENT WHICH LED TO THEIR DEATH ON (B)(6) 2024. IT WAS REPORTED THAT THE PATIENT¿S LAST CCPD TREATMENT OCCURRED ON (B)(6) 2024, APPROXIMATELY 10 DAYS PRIOR TO THEIR PASSING. THE PDRN ATTRIBUTED THE CAUSE OF THE PATIENT¿S PERITONITIS TO A PROBABLE TOUCH CONTAMINATION EVENT, IN WHICH THE PATIENT UTILIZED A DIRTY TOWEL INSTEAD OF PAPER TOWELS TO DRY THEIR HANDS. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). ALTHOUGH THE LIKELY CAUSE OF THE PATIENT¿S DEATH WAS A CARDIAC EVENT BROUGHT ON BY SEPTICEMIA, THE PDRN REPORTED THE CAUSE OF DEATH WAS PERITONITIS. THE PATIENT¿S PERSONAL LIBERTY SELECT CYCLER IS IN THE PROCESS OF BEING RETURNED AS OF 18/NOV/2024. | PLANT INVESTIGATION: THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER AND THE LOT NUMBER WAS NOT PROVIDED. A MANUFACTURING REVIEW WAS PERFORMED ON THE PRODUCTS SHIPPED TO THE PATIENT FOR THE THREE (3) MONTH TIME FRAME WHICH IMMEDIATELY PRECEDED THE EVENT OCCURRENCE DATE. THIS REVIEW INCLUDED THE LOT NUMBERS FOR ALL FRESENIUS LIBERTY CYCLER SETS SHIPPED TO THIS ACCOUNT WITHIN THE SELECTED TIME FRAME. THE ENTIRE SET OF LOTS HAVE BEEN SOLD AND DISTRIBUTED. THERE WERE NO NON-CONFORMANCES OR ABNORMALITIES IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE ASSOCIATED WITH THE REPORTED EVENT. AN INVESTIGATION OF THE DEVICE HISTORY RECORDS (DHR) WAS CONDUCTED AND CONFIRMED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE PRODUCT LOTS INVOLVED MET ALL SPECIFICATIONS FOR RELEASE. A REVIEW OF THE DHR DID NOT REVEAL A PROBABLE CAUSE FOR THE CUSTOMER COMPLAINT. AS A PHYSICAL EVALUATION COULD NOT BE PERFORMED, A DEFINITIVE CONCLUSION REGARDING THE REPORTED INCIDENT COULD NOT BE REACHED AND A CAUSE COULD NOT BE CONFIRMED.
DAVINCI XI | K131861 | A LITERATURE ARTICLE DESCRIBED A PROSPECTIVE SINGLE-CENTER STUDY OF PATIENTS WITH AORTOILIAC OCCLUSIVE DISEASE (AIOD) THAT UNDERWENT ROBOTIC-ASSISTED LAPAROSCOPIC (RAL) REVASCULARIZATION PROCEDURES. THE STUDY WAS CONDUCTED FROM FEBRUARY 2014 TO FEBRUARY 2019 AND INCLUDED 70 PATIENTS. THE AIM OF THE STUDY WAS TO ASSESS THE SURGEON LEARNING CURVE AND THE FEASIBILITY, SAFETY, AND MIDTERM OUTCOMES. THE ARTICLE DESCRIBED ONE PATIENT DEATH IN THE ROBOTIC COHORT. THE PATIENT HAD MULTIPLE COMORBIDITIES INCLUDING CHRONIC ISCHEMIC HEART DISEASE, HYPERTENSION, CHRONIC OBSTRUCTIVE PULMONARY DISEASE WITH CENTRO-LOBULAR EMPHYSEMA, A HISTORY OF ACUTE PANCREATITIS, ACTIVE SMOKING, AND TASC D AORTOILIAC LESIONS (REQUIRES SURGICAL INTERVENTION). THE PATIENT PRESENTED WITH A RETROPERITONEAL HEMATOMA ON POSTOPERATIVE DAY (POD) 1, WHICH REQUIRED SURGICAL DRAINAGE; ON POD 2, ILEAL AND COLONIC ISCHEMIA DEVELOPED, LEADING TO AN ILEAL AND RIGHT COLONIC RESECTION, FOLLOWED BY A LEFT COLECTOMY ON POD 13, ULTIMATELY SUCCUMBING TO SEPTIC SHOCK IN INTENSIVE CARE ON POD 51. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS OR ISSUES STATED, NOR DID THE AUTHORS SUGGEST THAT ANY INTUITIVE DA VINCI PRODUCTS CAUSED OR CONTRIBUTED TO THE ADVERSE EVENTS. THE DESIGNATED CONTACT AUTHOR HAS NOT RESPONDED TO REQUESTS FOR ADDITIONAL INFORMATION. THE ARTICLE DOCUMENTS THAT THE STUDY RESULTS ARE COMPARABLE TO STANDARD SURGICAL CONVENTIONAL OPEN OR LAPAROSCOPIC SURGERY. | NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM OR INSTRUMENT IDENTIFIERS. NO SPECIFIC INFORMATION REGARDING THE PROCEDURE SITE OR DATE WAS PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE LITERATURE ARTICLE DESCRIBES THE DEATH OF A PATIENT WITH A LONG LIST OF SIGNIFICANT COMORBIDITIES. THE DEATH OCCURRED A PROLONGED PERIOD AFTER AORTOILIAC ROBOTIC SURGERY AND NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS ARE MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SUTTER W, ALSAC JM, BEN ABDALLAH I, MICHEL C, JULIA P, EMPANA JP, EL BATTI S. TREATMENT OF AORTOILIAC OCCLUSIVE LESIONS BY AORTIC ROBOTIC SURGERY: LEARNING CURVE AND MIDTERM OUTCOME. ANN VASC SURG. 2024 JUL; 104:258-267. DOI: 10.1016/J.AVSG.2024.02.018. EPUB 2024 APR 7. PMID: 38593921. SECTION B2 DATE OF DEATH: THE PATIENT DATE OF DEATH WAS NOT PROVIDED; THE STUDY WAS CONDUCTED BETWEEN FEBRUARY 2014 AND FEBRUARY 2019. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
BIPAP SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE PATIENT'S SON ALLEGES THE PATIENT'S DEATH WAS DUE TO THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT PASSED AWAY, AND HIS WIFE WANTED TO RETURN IT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION AUTO CPAP DEVICE. THE SPOUSE OF A PATIENT REPORTED THE PATIENT HAS PASSED. THE REPORTER STATED THE PATIENT NEVER RECEIVED A RETURN LABEL AND DID NOT RECEIVE A REPLACEMENT. VETERANS AFFAIRS WOULD NOT TAKE THE DEVICE, THEREFORE THE REPORTER STATED THE DEVICE WAS "THROWN OUT". THE PATIENT'S CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
EFFICIA DFM100 | K110825 | PHILIPS RECEIVED A COMPLAINT ON THE DEFIBRILLATOR/MONITOR INDICATING THAT CUSTOMER WANTS TO KNOW WHY SHOCK IS NO LONGER NEEDED IN AED MODE WHILE IN USE. THE PATIENT EXPIRED. THE EFFICIA DFM100 DEFIBRILLATOR, MODEL# 866199, IS SUBSTANTIALLY SIMILAR TO THE HEARTSTART XL+ DEFIBRILLATOR (MODEL # 861290) AND WILL BE REPORTED IN THE UNITED STATES UNDER DEVICE MODEL # 861290. | THE PHILIPS TECHNICAL SUPPORT SPECIALIST CONTACTED THE DISTRIBUTOR WHO EVALUATED THE DEVICE ON SITE AND DID NOT FIND ANY ISSUES WITH THE DEVICE. THE DEVICE FUNCTIONS WITHIN SPECIFICATIONS. THE DEVICE REMAINS AT THE CUSTOMER SITE AND NO FURTHER EVALUATION IS WARRANTED AT THIS TIME. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS NOT DETERMINED. THE REPORTED PROBLEM WAS NOT CONFIRMED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION AUTO CPAP DEVICE. THE CLINIC HAS REPORTED THAT THE PATIENT HAS PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO A DREAMSTATION AUTO CPAP DEVICE. THE CLINIC HAS REPORTED THAT THE PATIENT HAS PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS NOT REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. UPDATED: EVALUATION METHOD CODE UPDATED. EVALUATION RESULTS CODE UPDATED. CONCLUSION CODE UPDATED.
DAVINCI XI | K131861 | A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE COHORT STUDY TO EVALUATE THE SAFETY AND EFFECTIVENESS OF THE USE OF THE DA VINCI SYNCHROSEAL INSTRUMENT COMPARED WITH THE DA VINCI VESSEL SEALER EXTEND INSTRUMENT FOR ROBOTIC PARENCHYMAL TRANSECTION OF THE LIVER. THE STUDY INCLUDED A TOTAL OF 155 ROBOTIC LIVER RESECTION (RLR) PATIENTS USING THE SYNCHROSEAL FROM JUNE 2021 TO MARCH 2023, AND 161 RLR PATIENTS USING THE VESSEL SEALER FROM FEBRUARY 2020 TO MARCH 2023 PERFORMED AT TWO INSTITUTIONS. DURING THE STUDY, ONE PATIENT IN THE VESSEL SEALER COHORT DIED AS A RESULT OF A COMPLICATED POSTOPERATIVE COURSE WITH A LUNG EMBOLUS AND BLEEDING FROM THE INFERIOR EPIGASTRIC ARTERY THAT REQUIRED A REOPERATION DURING WHICH IATROGENIC DAMAGE OF THE BLADDER OCCURRED. LATER, THE PATIENT DIED OF HEMORRHAGIC SHOCK STEMMING FROM BLEEDING AT THE SITE OF THE NEPHROSTOMY CATHETER. THERE WERE NO DEVICE ISSUES MENTIONED IN EITHER OF THE TWO GROUPS IN THE ARTICLE. THE ARTICLE CONCLUSION STATED THAT THE SYNCHROSEAL IS A SAFE AND EFFECTIVE DEVICE FOR ROBOTIC LIVER PARENCHYMAL TRANSECTION. | NO INVESTIGATION COULD BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE OR DA VINCI SYSTEM IDENTIFIERS OR INSTRUMENT LOT NUMBERS. THERE WAS NO REPORT OF, OR INDICATION OF, ANY DA VINCI PRODUCT ISSUES. THERE IS NO INDICATION THAT ANY DA VINCI PRODUCTS CONTRIBUTED TO THE REPORTED PATIENT COMPLICATIONS. THIS ARTICLE REPRESENTS THE PATIENT DEATH REPORT. SEE SECTION H10, RELATED REPORT NUMBER 1 FOR ADDITIONAL EVENTS FROM THIS ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THIS REPORT IS FROM A LITERATURE ARTICLE IN WHICH ONE PATIENT DIED AFTER A SERIES OF POST-OPERATIVE COMPLICATIONS INCLUDING A PULMONARY EMBOLISM, BLEEDING FROM THE INFERIOR EPIGASTRIC ARTERY, AND AN IATROGENIC BLADDER INJURY ON A RE-EXPLORATION. NO ALLEGATION HAS BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. CITATION: PILZ DA CUNHA G, DE MEYERE C, D¿HONDT M, ET AL. ROBOTIC LIVER PARENCHYMAL. TRANSECTION USING THE SYNCHROSEAL. SURG ENDOSC 38, 4947¿4955 (2024). HTTPS://DOI.ORG/10.1007/S00464-024-11005-4. B2: PATIENT DATE OF DEATH - THE DATE OF DEATH IS UNKNOWN; THE STUDY WAS CONDUCTED BETWEEN FEBRUARY 2020 TO MARCH 2023. SECTION B3: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE EVENT DATE ASSOCIATED WITH THE ADVERSE EVENT, AN ALTERNATE DATE, PUBLISHED DATE HAS BEEN USED. SECTION D: SUSPECT MEDICAL DEVICE: DUE TO THE LACK OF SPECIFIC INFORMATION REGARDING THE DA VINCI SYSTEM ASSOCIATED WITH THE ADVERSE EVENT, A GENERIC SYSTEM MATERIAL NUMBER WAS USED. SECTION E: THE PROCEDURES WERE PERFORMED AT TWO INSTITUTIONS: (B)(6) HOSPITAL.
REMSTAR PLUS C-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
OMNIPOD 5 POD | K203768 | INSULET BECAME AWARE OF A PATIENT EXPIRING WHILE USING THE OMNIPOD 5 SYSTEM WITH A DEXCOM CGM (CONTINUOUS GLUCOSE MONITORING). | DURING A CALL WITH INSULET AND A DEXCOM REPRESENTATION, THE ATTENDING HCP PROVIDED THE FOLLOWING DETAILS: PATIENT WAS A WELL CONTROLLED TYPE 1 DIABETIC WHO WAS ON OMNIPOD 5 SINCE 2022. PATIENT WAS USING OP5 WITH A DEXCOM G6 CGM SENSOR. PATIENT WAS A 20 YR OLD FEMALE WHO SHARED THE DEXCOM CLARITY APP WITH FAMILY FOR VISIBILITY IN HER CARE. (B)(6) 2024, THE PATIENT CONTACTED HER PARENTS TO DISCUSS BRINGING MORE SUPPLIES TO HER AT COLLEGE. (B)(6) 2024 AT 22:46 THE CGM STOPPED PROVIDING READINGS IN THE CLARITY APP. LAST KNOWN READING PER THE PARENTS- 76. (B)(6) 2024, PARENTS TEXT MESSAGES TO THE PATIENT WENT UNANSWERED. (B)(6) 2024, THE FAMILY CONTACTED LOCAL AUTHORITIES TO DO A WELLNESS CHECK AND THE PATIENT WAS FOUND EXPIRED. AUTOPSY IS PENDING. INSPECTION OF THE POD MANAGER HISTORY AND PATIENT HISTORY BUFFER FILES SHOWED THAT ON (B)(6) 2024 AT 22:51 THERE WAS NO LONGER AN ACTIVE CGM SENSOR, AND THE POD WAS NOT RECEIVING ANY ESTIMATED GLUCOSE VALUES. THE SYSTEM CORRECTLY SWITCHED INTO AUTOMATED MODE: LIMITED AFTER 20 MINUTES OF MISSING CGM SENSOR VALUES AND DELIVERED INSULIN AT THE LOWER OF THE USER¿S ADAPTIVE BASAL RATE AND MANUAL MODE BASAL RATE, AS EXPECTED. THE CONTROLLER WAS SHOWN TO GIVE MISSING CGM SENSOR ALERTS EVERY HOUR BEFORE THE POD WAS DEACTIVATED AT 04:43 ON (B)(6) 2024 DUE TO AN EXPIRATION ALARM AFTER 80 HOURS OF USE, WHICH IS AN EXPECTED SAFETY FEATURE OF THE SYSTEM. BASED ON THE AVAILABLE INFORMATION, THE OMNIPOD 5 OPERATED AS INTENDED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI X | K171294 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RIGHT HEMICOLECTOMY WITH INTRA-CORPOREAL ANASTOMOSIS, THE PATIENT EXPIRED ON POST-OPERATIVE DAY THREE. THERE WERE NO INTRAOPERATIVE COMPLICATIONS AND NO REPORT OF ANY DA VINCI SYSTEM, INSTRUMENT OR ACCESSORIES ISSUES. THE SURGEON INFORMED THE INTUITIVE CLINICAL SALES REPRESENTATIVE THAT WAS PRESENT DURING THE PROCEDURE THAT THE PATIENT DIED DUE TO RENAL FAILURE UNRELATED TO THE DA VINCI PRODUCTS. THE PATIENT WAS NOTED TO BE ELDERLY AND VERY UNHEALTHY AND COULD NOT BE SCANNED BEFORE SHE DIED. ADDITIONAL INFORMATION WAS REQUESTED BUT NO FURTHER DETAILS WERE PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, LARGE CLIP APPLIER, LARGE NEEDLE DRIVER, CADIERE FORCEPS, FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM STAPLER 60. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE ELDERLY PATIENT IN THIS REPORT WAS IN POOR HEALTH AND DIED FROM RENAL FAILURE 3 DAYS AFTER A HEMICOLECTOMY. THE PROCEDURE WAS REPORTEDLY UNEVENTFUL WITH NO INTRAOPERATIVE COMPLICATIONS AND NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION HUMIDIFIER | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION HUMIDIFIER DEVICE. THERE WAS A REPORT OF PATIENT DEATH. THE DEVICE WAS RETURNED FOR EVALUATION AND NO DEVICE PROBLEM WAS DETECTED. THE DEVICE WAS SCRAPPED.
FLEXCATH ADVANCE STEERABLE SHEATH | K123591 | IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING THE USE OF A SHEATH, A NON-MEDTRONIC INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER WAS PLACED, AND THE PHYSICIAN PERFORMED A TRANSSEPTAL PUNCTURE USING A NON-MEDTRONIC NEEDLE AND SHEATH. A NON-MEDTRONIC MAPPING CATHETER WAS INSERTED INTO THE LEFT ATRIUM AND LATER SEEN IN THE APEX OF THE LEFT VENTRICLE. THE NON-MEDTRONIC MAPPING CATHETER WAS PULLED BACK TO THE LEFT ATRIUM, AND MAPPING BEGAN. THE PATIENT EXPERIENCED HYPOTENSION AND A SIGNIFICANT DROP IN BLOOD PRESSURE. A PERICARDIAL EFFUSION WAS OBSERVED VIA ICE AND THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS, REMOVING 700 CUBIC CENTIMETERS (CC) OF BLOOD, WHICH WAS AUTO TRANSFUSED TO THE PATIENT. THE PATIENT APPEARED STABLE, AND THE PROCEDURE CONTINUED WITH THE PLACEMENT OF THE SHEATH IN THE LEFT ATRIUM. HOWEVER, BEFORE ADVANCING THE BALLOON CATHETER INTO THE PATIENT, THE EFFUSION PERSISTED, PROMPTING ANOTHER PER ICARDIOCENTESIS TO REMOVE 600-700CC OF BLOOD, WHICH WAS AGAIN AUTO TRANSFUSED. DESPITE SEEMING STABLE, THE CASE WAS ABORTED. CARDIAC SURGERY WAS SUBSEQUENTLY PERFORMED IN THE ELECTROPHYSIOLOGY (EP) LAB. THE PATIENT LATER DIED IN THE INTENSIVE CARE UNIT (ICU) DUE TO COMPLICATIONS FROM THE PERICARDIAL EFFUSION. IT WAS NOTED THAT A CARDIAC PERFORATION OCCURRED IN THE LEFT VENTRICLE. | CONTINUATION OF D10: PRODUCT ID: NON-MEDTRONIC UNKNOWN; PRODUCT TYPE: INTRACARDIAC ECHOCARDIOGRAPHY CATHETER PRODUCT ID: NON-MEDTRONIC UNKNOWN; PRODUCT TYPE: TRANSSEPTAL NEEDLE PRODUCT ID: NON-MEDTRONIC UNKNOWN; PRODUCT TYPE: SHEATH PRODUCT ID: NON-MEDTRONIC UNKNOWN; PRODUCT TYPE: MAPPING CATHETER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CAPIOX FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | K151791 | THE USER FACILITY REPORTED TO TERUMO CARDIOVASCULAR THAT DURING CARDIOPULMONARY BYPASS, THE OXYGENATOR FAILED AROUND 3.5 HOURS INTO BYPASS RUN. THERE WAS MINIMAL TO NO GAS TRANSFER OCCURRING. PRODUCT WAS NOT CHANGED OUT. NO BLOOD LOSS. PROCEDURE COMPLETED SUCCESSFULLY. PATIENT EXPIRED. | TERUMO HAS NOT RECEIVED THE DEVICE FOR EVALUATION; THEREFORE, THE INVESTIGATION HAS YET TO BE COMPLETED. TERUMO PLANS ON SUBMITTING A FOLLOW-UP REPORT WHEN THE INVESTIGATION IS COMPLETE AND WHEN MORE INFORMATION BECOMES AVAILABLE. FOR THIS REASON, TERUMO REFERENCES EVALUATION CONCLUSION CODE 11.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A NURSE REPORTED A PERITONEAL DIALYSIS (PD) PATIENT SAT UP TO DISCONNECT FROM TREATMENT ON THE LIBERTY SELECT CYCLER AND SLUMPED OVER. THE PATIENT PASSED AWAY AND CARDIAC ARREST WAS SUSPECTED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). ON (B)(6) 2024, THE PATIENT WAS NOT IN ACTIVE TREATMENT AND FELL WHILE OUTSIDE THE HOME. THE PATIENT WAS EXPERIENCING CHEST AND SHOULDER PAIN. THE PATIENT WAS BROUGHT TO THE EMERGENCY ROOM (ER). AN X-RAY REVEALED THE PATIENT HAD BROKEN HIS COLLAR BONE. NO ADDITIONAL INTERVENTIONS OR TESTING WERE PERFORMED. NO LABS WERE DRAWN. THE PATIENT WAS PLACED IN A SLING AND PRESCRIBED PAIN MEDICATION (UNSPECIFIED). THE PATIENT WAS DISCHARGED. LATER IN THE EVENING, THE PATIENT CONNECTED TO THE LIBERTY SELECT CYCLER FOR TREATMENT. IN THE MORNING ON (B)(6) 2024 THE PD TREATMENT WAS COMPLETE. THE PATIENT WENT TO STAND UP OUT OF THE BED TO DISCONNECT FROM THE LIBERTY SELECT CYCLER AND BEGAN SHAKING. THE PATIENT SUBSEQUENTLY FELL TO THE FLOOR. THE PATIENT COULD NOT BE REVIVED AND WAS PRONOUNCED DECEASED. THE SUSPECTED CAUSE OF DEATH IS CARDIAC ARREST; HOWEVER, THE CORONER HAS NOT COMPLETED THE REPORT. THE PATIENT WAS STILL CONNECTED TO THE CYCLER AT THE TIME OF DEATH. THE PDRN STATED THERE WERE NO ISSUES WITH THE PATIENT¿S PD TREATMENT AND THE PATIENT¿S SPOUSE AND DAUGHTER STATED THE PATIENT DID NOT HAVE ANY ISSUES THROUGHOUT THE NIGHT WITH THE TREATMENT. NO FURTHER INFORMATION WAS PROVIDED. | ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR INVESTIGATION A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. THE TEACH PUMPS TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. THERE WERE VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY AND THE BOTTOM COVER. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED. THERE IS NO ALLEGATION OF DEVICE CONTRIBUTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. REPORTEDLY, THE CAUSE OF DEATH WAS UNKNOWN. THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | THE DEVICE IS EXPECTED TO BE RETURNED; HOWEVER, THE DEVICE HAS NOT YET BEEN RECEIVED. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF THE DEVICE IS RECEIVED, AND EVALUATION IS PERFORMED.
ROTAREX | K211738 | IT WAS REPORTED THAT DURING A THROMBECTOMY AND ATHERECTOMY PROCEDURE FOR ENDOVASCULAR REDUCTION OF THE ILIAC ARTERY AND OPEN THE VESSEL THROUGH THE LOWER EXTREMITY ARTERIAL AREA, ILIAC ARTERIES AND DOWN TO THE SUPERFICIAL FEMORAL ARTERIES, RIGHT FEMORAL ARTERY CONTRALATERAL APPROACH, AS THE DISTAL END OF THE GUIDEWIRE WAS 15CM FROM THE TIP OF THE CATHETER A SUDDEN SPIN OF THE GUIDEWIRE OCCURRED STOPPING THE DEVICE. THEREFORE, THE CATHETER WAS RETRIEVED AS ONE UNIT, AND THE TIP OF THE GUIDEWIRE WAS FOUND ALLEGEDLY DISINTEGRATED. THE DETACHED SEGMENT WAS REPORTEDLY REMOVED FROM THE PATIENT WITH A GRABBER DEVICE, ANOTHER SAME DEVICE WAS USED FOR CONTRALATERAL SUCTION, CANNULATION, AND THROMBOLYSIS WAS PERFORMED SUCCESSFULLY COMPLETING THE PROCEDURE. IT WAS FURTHER REPORTED THAT SOMETIME POST PROCEDURE, THE PATIENT EXPIRED DUE TO HEART FAILURE, AND IT IS UNCLEAR WHETHER IT WAS DIRECTLY CAUSED BY AN ADVERSE EVENT OR THE DEVICE. THE RELATIONSHIP BETWEEN THE PATIENT DEATH AND THE DEVICE IS UNKNOWN. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: THE CATALOG NUMBER IDENTIFIED IN SECTION D4 HAS NOT BEEN CLEARED IN THE US BUT IS SIMILAR TO THE ROTAREX THAT ARE CLEARED IN THE US. THE PRO CODE AND 510 K NUMBER FOR THE ROTAREX ARE IDENTIFIED IN D2 AND G4. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS IS CURRENTLY BEING PERFORMED. THE RETURN OF THE SAMPLE IS PENDING. PHOTOS AND IMAGES WERE PROVIDED AND PENDING REVIEW. THE INVESTIGATION OF THE REPORTED EVENT IS CURRENTLY UNDERWAY. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
REMSTAR PLUS SERIES ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTO BI-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DECEASED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY AND THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
IPV-2C | K905236 | ON NOVEMBER 1ST, A SENTEC REPRESENTATIVE WAS MADE AWARE OF A PATIENT DEATH AFTER USE OF AN IPV-2C DEVICE. THE RESPIRATORY THERAPY STAFF COMMUNICATED THAT THE PATIENT WAS A CRITICALLY ILL ADULT WITH MANY COMORBIDITIES, AND THUS WAS COMPLEX TO TREAT, AND EXTREMELY DIFFICULT TO VENTILATE. ADDITIONALLY IT WAS STATED DUE TO THE PATIENT'S COMORBIDITIES, THEY LIKELY WOULD HAVE SUCCUMBED TO THEIR CONDITIONS WITH OR WITHOUT ADMINISTERING IPV THERAPY IN ANY SETUP (IN-LINE OR DIRECT TO ET TUBE). THE PATIENT WAS STATED TO BE ON A VENTILATOR BUT REMOVED FROM THE VENTILATOR TO PROVIDE THE IPV TREATMENT. THE IPV TREATMENT WAS ADMINISTERED THROUGH THE PATIENTS ET TUBE DIRECTLY. THROUGH ADDITIONAL COMMUNICATIONS FROM THE SENTEC REPRESENTATIVE AND HOSPITAL STAFF, IT WAS STATED THAT THE EXPIRATORY PORT OF THE PHASITRON WAS OCCLUDED WITH THE IN-LINE VALVE BLUE CAP. THE PATIENT EXPERIENCED A PNEUMOTHORAX; MULTIPLE ATTEMPTS WERE MADE TO REVIVE THE PATIENT, HOWEVER, THE PATIENT ULTIMATELY PASSED AWAY. THE AMOUNT OF TIME THE PATIENT WAS SET UP WITH THE IPV TREATMENT AND SETTINGS USED WERE UNABLE TO BE CONFIRMED. | THIS EVENT IS BEING REPORTED DUE TO PATIENT DEATH. BASED ON THE DETAILS PROVIDED FROM THE SENTEC REPRESENTATIVE'S COMMUNICATIONS WITH THE HOSPITAL, THE ROOT CAUSE OF THIS EVENT IS ASSUMED TO BE RELATED TO THE CRITICAL STATE OF THE PATIENT'S CONDITION AND THEIR COMORBIDITIES, IN COMBINATION WITH THE FAILURE TO REMOVE THE IN-LINE VALVE BLUE CAP FROM THE PHASITRON CIRCUIT EXPIRATORY PORT PRIOR TO PROVIDING THERAPY DIRECTLY TO THE PATIENT'S AIRWAY. THE AMOUNT OF TIME THE PATIENT UNDERWENT THE IPV THERAPY IS UNCLEAR, HOWEVER, IT WAS STATED THAT THE PATIENT EXPERIENCED A PNEUMOTHORAX PRIOR TO PASSING AWAY. PERCUSSIONAIRE INITIATED A FIELD SAFETY CORRECTIVE ACTION (FSCA) FOR THE IN-LINE VALVE BLUE CAP MISUSE IN SEPTEMBER 2024 (1000524541-08/21/2024-001-C). ALL IMPACTED CUSTOMERS WERE PROVIDED NOTIFICATION OF THIS ISSUE, UPDATED INSTRUCTIONS FOR USE WITH UPDATED WARNINGS AND IMAGES FOR PROPER BLUE CAP USE, AS WELL AS CAUTIONARY LABELS TO BE PLACED ON CUSTOMER'S IN-LINE VALVE INVENTORY. A WARNING WAS ALREADY INCLUDED IN THE PREVIOUS REVISION OF THE INSTRUCTIONS FOR USE TO ENSURE BLUE CAP WAS REMOVED PRIOR TO DIRECT AIRWAY USE; THIS WARNING WAS UPDATED FOR CLARITY TO STATE THE BREATHING CIRCUIT EXPIRATORY PORT SHOULD REMAIN OPEN AT ALL TIMES. THE CAUTION LABELS PROVIDED TO CUSTOMERS DETAILED PROPER USE OF THE IN-LIN VALVE, WITH A PROMPT TO KEEP THE PHASITRON EXPIRATORY PORT OPEN AT ALL TIMES. ALL HOSPITALS RECEIVED LABELS AND A HARD COPY OF THE FSCA NOTIFICATION. PERCUSSIONAIRE COMMUNICATED TO REACH OUT IF ADDITIONAL CAUTIONARY LABELS WERE NEEDED. IT WAS CONFIRMED THAT ELLIOT HOSPITAL RECEIVED MULTIPLE EMAILS ABOUT THIS FSCA, AS WELL AS THE SHIPPED LABELS AND A HARD COPY OF THE FSCA NOTIFICATION. PERCUSSIONAIRE IS ATTEMPTING TO GATHER ANY ADDITIONAL INFORMATION. IF ADDITIONAL INFORMATION IS OBTAINED, A FOLLOW-UP MDR WILL BE SUBMITTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, ASTHMA (NEW OR WORSENING),INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS ONE ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, DEATH AND ACUTE RESPIRATORY DISTRESS SYNDROME. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
VIVO 45 LS | K193586 | ON (B)(6) 2024 AT 8:13PM, BREAS RECEIVED THE FOLLOWING REQUEST FROM THE HEALTHCARE PROVIDER: DESCRIPTION: PATIENT EXPIRED ON (B)(6) 2024. PER WIFE, SHE FOUND HIM DISCONNECTED. WE ARE WANTING VERIFICATION THAT THE VENT WAS FUNCTIONING PROPERLY. PER DOWNLOAD ALARMS APPEAR TO HAVE FUNCTIONED AS INTENDED CAREGIVER ATTENDANT: YES, WHEN WAS THE FAILURE FIRST DETECTED? DURING THE ACTIVE TREATMENT PATIENT OUTCOME: DEAD THERE IS NO ALLEGATION AT THIS TIME THAT THE DEVICE HAS MALFUNCTIONED OR CAUSED OR CONTRIBUTED TO THE EVENT. THE CUSTOMER SENT THE DEVICE TO BREAS MEDICAL INC. FOR INVESTIGATION ON (B)(6) 2024 AND THE DEVICE ARRIVED ON (B)(6) 2024. | INVESTIGATION OF THE DEVICE STARTED ON NOV 26, 2024. A FLOW TEST FAILURE NOTED PRIOR TO CALIBRATION. THIS DOES NOT INDICATE AN ALARM FAILURE. DEVICE PASSED FUNCTIONAL TESTING AFTER FLOW RECALIBRATION. PER COMPLAINANT, DOWNLOAD ALARMS APPEAR TO HAVE FUNCTIONED AS INTENDED. THE INVESTIGATION COULD AT THIS POINT NOT CONFIRM THAT THERE HAD BEEN MALFUNCTION. A COMPLETE ANALYSIS OF THE LOG FILE IS STILL ONGOING.
TURNPIKE SPIRAL 135CM | K151981 | IT WAS REPORTED BY THE SALES REPRESENTATIVE WHO WAS CALLED INTO THE CASE THAT: "TURNPIKE SPIRAL 135CM USED TO GAIN ACCESS TO DISTAL LAD THROUGH 7FR GUIDE. TURNPIKE REACHED DESTINATION POINT. UPON REMOVAL, THE TIP OF TURNPIKE SPIRAL BROKE OFF IN THE PATIENT'S DISTAL LAD. PATIENT CODED AND NEEDED TO BE RESUSCITATED. PATIENT DEFIBRILLATED TWICE AND RECOVERED. TIP COULD NOT BE RETRIEVED AND HAD TO BE LEFT IN THE LAD. PERIPHERAL VENTRICULAR ASSIST DEVICE PVAD WAS INSERTED FOR LV SUPPORT. PATIENT WAS ON A VASOPRESSOR INFUSION AND STABLE. PATIENT NOT A CANDIDATE FOR SURGERY, BASED ON THE AMOUNT OF DISEASE IN THE VESSEL AND NO LANDMARKS." ADDITIONAL INFORMATION RECEIVED FROM SALES REP WHO WAS IN CONTACT WITH THE INTERVENTIONAL CARDIOLOGIST ON (B)(6) 2024 REPORTED THAT: "LAD WAS CALCIFIED AND THERE WERE ATTEMPTS TO SWAP GUIDE WIRE OUT FOR ANOTHER WIRE. TURNPIKE SPIRAL WAS TORQUED TO DISTAL LAD. UPON BACKING OUT SPIRAL, WHILE TURNING IN COUNTER-CLOCKWISE DIRECTION THE TIP BROKE OFF. THERE WAS NOT RESISTANCE PRIOR TO TIP DISLODGMENT. TIP OF TURNPIKE SPIRAL OCCLUDED THE DISTAL LAD, CAUSING NO FLOW TO APEX OF THE HEART. PATIENT HAD ARRHYTHMIAS AND NEEDED TO BE CARDIOVERTED, CPR INITIATED AND WAS INTUBATED. PVAD CARDIAC PUMP WAS DEPLOYED FOR SUPPORT AND CASE WAS ABORTED. PATIENT WAS TAKEN TO ICU ON A VENTILATOR. THE PATIENT PASSED AWAY LESS THAN 12 HOURS LATER ON THE (B)(6)." | QN# (B)(4).
TRILOGY 100 | K083526 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT WAS DEATH AND OTHERS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WERE OBSERVED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WERE 22 ERRORS FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX D: PRODUCT UDI WAS CORRECTED IN THIS REPORT. BOX E: REPORTER COUNTRY WAS CORRECTED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER, AND CHOLANGIOCARCINOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WERE OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY A YEAR AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAS PASSED AWAY COUPLE OF YEARS AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAS PASSED AWAY COUPLE OF YEARS AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER PREVIOUSLY INCORRECTLY REPORTED HEALTH IMPACT GRID CODE AS C50675. IN THIS FOLLOW UP REPORT HEALTH IMPACT GRID IS UPDATED AS C28554. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, NAUSEA, VOMITING AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER BECAME AWARE OF A DREAMSTATION AUTO BIPAP STATING THAT HUSBAND PASSED AWAY AND THAT THERE IS NO ALLEGATION THAT THE DEVICE HAS CONTRIBUTED TO THE DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED. THERE IS NO ALLEGATION OF DEVICE CONTRIBUTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER BECAME AWARE OF A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION BIPAP AUTOSV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING UPPER COLON CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ACUTE MYELOID LEUKEMIA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, DIZZINESS, HEADACHE, CANCER AND METASTATIC BREAST CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER, KIDNEY CANCER AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER, LYMPHOMA, LUNG DISEASE, DIZZINESS AND HEADACHE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, DEATH, RESPIRATORY ISSUES AND PULMONARY ISSUES. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
REACT CATHETER | K180715 | MEDTRONIC RECEIVED INFORMATION REGARDING A PATIENT WHO HAD COMPLICATIONS AFTER UNDERGOING A THROMBECTOMY PROCEDURE IN WHICH A REACT-68 ASPIRATION CATHETER AND SOLITAIRE X STENT RETRIEVER WERE USED. IT WAS REPORTED THAT THE PATIENT'S BASELINE TICI WAS 0, MRS WAS 0, AND NIHSS WAS 18. ON (B)(6) 2024, THE PATIENT UNDERWENT THE THROMBECTOMY PROCEDURE. NO DEVICE MALFUNCTION OR DEFICIENCY WAS REPORTED AND TICI 3 WAS ACHIEVED. POST-OPERATIVE COMPUTED TOMOGRAPHY (CT) SHOWED RIGHT FRONTOPARIETAL HEMORRHAGIC TRANSFORMATION WITH SUBFALCORAL AND MEDIAL TEMPORAL INVOLVEMENT WITH SYMPTOMS INCLUDING: HICCUPS, VIGILANCE DISORDERS, AND ANISOCORIA. THE DECISION WAS MADE TO LIMIT AND/OR DISCONTINUE OTHER INTERVENTIONAL THERAPIES AND PATIENT DIED THE FOLLOWING DAY, ON (B)(6) 2024. SITE ASSESSMENT CONCLUDED THE EVENTS WERE CAUSED BY THE PATIENT DISEASE UNDER STUDY/INDEX STROKE AND NOT RELATED TO THE DEVICES OR THE PROCEDURE. HOWEVER THE CLINICAL STUDY SPONSOR ASSESSMENT CONCLUDED THE EVENTS WERE ASSOCIATED WITH THE PROCEDURE. ADDITIONAL INFORMATION RECEIVED REPORTED THAT AUTOPSY RECORDS WERE NOT AVAILABLE, HOWEVER THE SOURCE DOCUMENTS MENTIONED THE DEATH OF THE PATIENT ON (B)(6) 2024 DUE TO MASSIVE RIGHT HEMISPHERIC HEMORRHAGIC SOFTENING FOLLOWING SUCCESSFUL REPERFUSION OF THE MIDDLE CEREBRAL ARTERY (MCA)/RIGHT CAROTID TERMINUS. 1 PASS HAD BEEN MADE IN THE THROMBECTOMY PROCEDURE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
SOLITAIRE FR4 | K183022 | MEDTRONIC RECEIVED INFORMATION REGARDING A PATIENT WHO HAD COMPLICATIONS AFTER UNDERGOING A THROMBECTOMY PROCEDURE IN WHICH A REACT-68 ASPIRATION CATHETER AND SOLITAIRE X STENT RETRIEVER WERE USED. IT WAS REPORTED THAT THE PATIENT'S BASELINE TICI WAS 0, MRS WAS 0, AND NIHSS WAS 18. ON (B)(6) 2024, THE PATIENT UNDERWENT THE THROMBECTOMY PROCEDURE. NO DEVICE MALFUNCTION OR DEFICIENCY WAS REPORTED AND TICI 3 WAS ACHIEVED. POST-OPERATIVE COMPUTED TOMOGRAPHY (CT) SHOWED RIGHT FRONTOPARIETAL HEMORRHAGIC TRANSFORMATION WITH SUBFALCORAL AND MEDIAL TEMPORAL INVOLVEMENT WITH SYMPTOMS INCLUDING: HICCUPS, VIGILANCE DISORDERS, AND ANISOCORIA. THE DECISION WAS MADE TO LIMIT AND/OR DISCONTINUE OTHER INTERVENTIONAL THERAPIES AND PATIENT DIED THE FOLLOWING DAY, ON (B)(6) 2024. SITE ASSESSMENT CONCLUDED THE EVENTS WERE CAUSED BY THE PATIENT DISEASE UNDER STUDY/INDEX STROKE AND NOT RELATED TO THE DEVICES OR THE PROCEDURE. HOWEVER THE CLINICAL STUDY SPONSOR ASSESSMENT CONCLUDED THE EVENTS WERE ASSOCIATED WITH THE PROCEDURE. ADDITIONAL INFORMATION RECEIVED REPORTED THAT AUTOPSY RECORDS WERE NOT AVAILABLE, HOWEVER THE SOURCE DOCUMENTS MENTIONED THE DEATH OF THE PATIENT ON (B)(6) 2024 DUE TO MASSIVE RIGHT HEMISPHERIC HEMORRHAGIC SOFTENING FOLLOWING SUCCESSFUL REPERFUSION OF THE MIDDLE CEREBRAL ARTERY (MCA)/RIGHT CAROTID TERMINUS. 1 PASS HAD BEEN MADE IN THE THROMBECTOMY PROCEDURE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DEATH, HEART FAILURE AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, | K131982 | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE OF A DREAMSTATION AUTO CPAP STATING THAT HUSBAND PASSED AWAY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
DREAMSTATION AUTO BIPAP | K131982 | IT WAS REPORTED THAT THE DREAMSTATION AUTO BIPAP DEVICE WAS NO LONGER AVAILABLE AND THE USER PASSED AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION BIPAP PRO | K131982 | IT WAS REPORTED THAT THE DREAMSTATION BIPAP PRO DEVICE WAS DISCARDED, AND THE USER PASSED AWAY ON (B)(6) 2024. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
FLEXCATH ADVANCE STEERABLE SHEATH | K123591 | LITERATURE WAS REVIEWED REGARDING THE OPTIMAL TIMING OF CRYOBALLOON ABLATION AFTER THE FIRST CLINICAL DIAGNOSIS OF ATRIAL FIBRILLATION (AF) AND ITS PROGNOSIS FOR PATIENTS WITH HEART FAILURE (HF). PATIENTS UNDERGOING AN ABLATION WERE DIVIDED INTO TWO GROUPS; THOSE WITH AN ABLATION LESS THAN SIX MONTHS AFTER THE DIAGNOSIS OF AF (EARLY-ABLATION GROUP) AND THOSE GREATER THAN SIX MONTHS POST DIAGNOSIS (DELAYED-ABLATION GROUP). THE AUTHORS DESCRIBED PATIENT DEATHS WHICH OCCURRED IN BOTH GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND CATEGORIZED BY ALL CAUSE OR CARDIOVASCULAR DEATHS. THERE IS NO ALLEGATION OF DEVICE OR PROCEDURE DEATH RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE CAUSE OF DEATH AND DEVICE-RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THERE WERE PATIENTS WHO EXPERIENCED STROKE, CARDIAC TAMPONADE, PHRENIC NERVE INJURY (PNI), PSEUDOANEURYSM, ARTERIOVENOUS FISTULA, GROIN HEMATOMA AND PROLONGED HOSPITALIZATION DUE TO WORSENED HF. THE GROIN HEMATOMA REQUIRED A BLOOD TRANSFUSION, AND ALL PHRENIC NERVE INJURIES RECOVERED DURING THE FOLLOW-UP PERIOD WITHOUT SERIOUS SYMPTOMS. THE STATUS OF THE CATHETERS AND SHEATHS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. | THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EARLY CRYOABLATION AFTER FIRST DIAGNOSIS OF ATRIAL FIBRILLATION REDUCES ARRHYTHMIA RECURRENCE IN HEART FAILURE PATIENTS. JACC ASIA. 2024;4:857¿871 DOI: 10.1016/J.JACASI.2024.08.005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
ACHIEVE MAPPING CATHETER - 20 MM | K102588 | LITERATURE WAS REVIEWED REGARDING THE OPTIMAL TIMING OF CRYOBALLOON ABLATION AFTER THE FIRST CLINICAL DIAGNOSIS OF ATRIAL FIBRILLATION (AF) AND ITS PROGNOSIS FOR PATIENTS WITH HEART FAILURE (HF). PATIENTS UNDERGOING AN ABLATION WERE DIVIDED INTO TWO GROUPS; THOSE WITH AN ABLATION LESS THAN SIX MONTHS AFTER THE DIAGNOSIS OF AF (EARLY-ABLATION GROUP) AND THOSE GREATER THAN SIX MONTHS POST DIAGNOSIS (DELAYED-ABLATION GROUP). THE AUTHORS DESCRIBED PATIENT DEATHS WHICH OCCURRED IN BOTH GROUPS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND CATEGORIZED BY ALL CAUSE OR CARDIOVASCULAR DEATHS. THERE IS NO ALLEGATION OF DEVICE OR PROCEDURE DEATH RELATEDNESS INDICATED IN THE ARTICLE; HOWEVER, THE CAUSE OF DEATH AND DEVICE-RELATEDNESS HAS BEEN REQUESTED, BUT NOT YET RECEIVED. THERE WERE PATIENTS WHO EXPERIENCED STROKE, CARDIAC TAMPONADE, PHRENIC NERVE INJURY (PNI), PSEUDOANEURYSM, ARTERIOVENOUS FISTULA, GROIN HEMATOMA AND PROLONGED HOSPITALIZATION DUE TO WORSENED HF. THE GROIN HEMATOMA REQUIRED A BLOOD TRANSFUSION, AND ALL PHRENIC NERVE INJURIES RECOVERED DURING THE FOLLOW-UP PERIOD WITHOUT SERIOUS SYMPTOMS. THE STATUS OF THE CATHETERS AND SHEATHS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. | THIS INFORMATION IS BASED ENTIRELY ON JOURNAL LITERATURE. MEDTRONIC WAS MADE AWARE OF THIS EVENT THROUGH A SEARCH OF LITERATURE PUBLICATIONS. THIS EVENT OCCURRED OUTSIDE THE US. PATIENT INFORMATION IS LIMITED DUE TO CONFIDENTIALITY CONCERNS. OF NOTE, MULTIPLE PATIENTS WERE NOTED IN THE ARTICLE; HOWEVER, A ONE-TO-ONE CORRELATION COULD NOT BE MADE WITH UNIQUE PRODUCT SERIAL/LOT NUMBERS. THE DATE OF DEATH IS NOT AVAILABLE AT THE TIME OF THIS REPORT AS THERE IS NO INDICATION OF SPECIFIC SERIAL NUMBER/PATIENT INFORMATION. WITHOUT A LOT NUMBER OR DEVICE SERIAL NUMBER, THE MANUFACTURING DATE CANNOT BE DETERMINED. SINCE NO DEVICE ID WAS PROVIDED, IT IS UNKNOWN IF THIS EVENT HAS BEEN PREVIOUSLY REPORTED. A REQUEST FOR ADDITIONAL INFORMATION WAS MADE. IF ADDITIONAL INFORMATION IS OBTAINED REGARDING THIS EVENT, IT WILL BE ADDED, AND A SUPPLEMENTAL REPORT WILL BE SUBMITTED ACCORDINGLY. REFERENCED ARTICLE: EARLY CRYOABLATION AFTER FIRST DIAGNOSIS OF ATRIAL FIBRILLATION REDUCES ARRHYTHMIA RECURRENCE IN HEART FAILURE PATIENTS. JACC ASIA. 2024; 4:857¿871 DOI: 10.1016/J.JACASI.2024.08.005. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH OF THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | IT WAS PREVIOUSLY REPORTED THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH OF THE PATIENT. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER WHERE IT WAS INSPECTED. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS AND NO COSMETIC DAMAGE FOUND. THE DEVICE PASSED ALL TESTS.
IMPEDE-FX EMBOLIZATION PLUG | K182390 | ON (B)(6) 2024, DR. (B)(6) PERFORMED A STENT ON A HEPATIC ARTERY AND A SAC FILL OF A GASTRODUODENAL ARTERY (GDA) ANEURYSM WHERE 70 IMP-FX-12 PLUGS WERE PLACED USING A 7F MEDTRONIC LAUNCHER GUIDE CATHETER AND 180CM AMPLATZ S. S. GUIDEWIRE. THE PHYSICIAN ACCESSED THE HEPATIC ARTERY FIRST USING AN 8F X 90CM TERUMO DESTINATION SHEATH WHICH THE 7F MEDTRONIC LAUNCHER WAS DELIVERED INTO/THROUGH. THE PHYSICIAN ALSO USED 3 GORE VIABAHN VBX STENTS TO ELIMINATE PERFUSION OF THE GDA ANEURYSM. AT THE END OF THE PROCEDURE, EMBOLIZATION WAS COMPLETE, THERE WAS NO FLOW IN THE GDA ANEURYSM, AND NO PERFUSION OF THE SAC FROM THE HEPATIC ANEURYSM. THE PROCEDURE WAS COMPLETED, AND THE PATIENT WAS IN STABLE CONDITION. THE PATIENT REMAINED IN THE HOSPITAL PER NORMAL PROCEDURE, BUT 3 DAYS LATER ON (B)(6) 2024 EXPERIENCED A RUPTURE OF THE ABDOMINAL ARTERY AND THE PATIENT EXPIRED. UNDER COMPUTED TOMOGRAPHY (CT) SCAN, THE PHYSICIAN COULD NOT IDENTIFY WHICH ARTERY WAS RUPTURED. IT WAS REPORTED THAT "DR. (B)(6) DOES NOT BELIEVE IT WAS THE GDA THAT RUPTURED, GIVEN THAT IT HAD BEEN EMBOLIZED AND EXCLUDED, I.E., NO LONGER PERFUSED AND PRESSURIZED BY THE HEPATIC ARTERY." | BASED ON THE INFORMATION RECEIVED, THE PATIENT HAD UNDERGONE PROCEDURE ON (B)(6) 2024 OF A STENT ON A HEPATIC ARTERY AND A SAC FILL OF A GASTRODUODENAL ARTERY (GDA) ANEURYSM USING IMP-FX-12 DEVICES. 70 DEVICES WERE DEPLOYED TO THE TARGET SITE WITHOUT ANY INCIDENT USING A 7F MEDTRONIC LAUNCHER GUIDE CATHETER AND 180CM AMPLATZ S. S. GUIDEWIRE. THE PHYSICIAN ACCESSED THE HEPATIC ARTERY FIRST USING AN 8F X 90CM TERUMO DESTINATION SHEATH WHICH THE 7F MEDTRONIC LAUNCHER WAS DELIVERED INTO/THROUGH. THE PHYSICIAN ALSO USED 3 GORE VIABAHN VBX STENTS TO ELIMINATE PERFUSION OF THE GDA ANEURYSM. AT THE END OF THE PROCEDURE, EMBOLIZATION WAS COMPLETE, THERE WAS NO FLOW IN THE GDA ANEURYSM, AND NO PERFUSION OF THE SAC FROM THE HEPATIC ANEURYSM. THE PROCEDURE WAS COMPLETED, AND THE PATIENT WAS IN STABLE CONDITION. THE PATIENT HAD EXPIRED ON (B)(6) 2024 DUE TO A RUPTURE OF AN ABDOMINAL ARTERY. UNDER COMPUTED TOMOGRAPHY (CT) SCAN, THE PHYSICIAN COULD NOT MAKE A DETERMINATION (OF WHICH ABDOMINAL ARTERY RUPTURED OR WHY IT RUPTURED) BASED ON THE INFORMATION GAINED FROM THE IMAGING STUDY. IT WAS REPORTED THAT "(B)(6) DOES NOT BELIEVE IT WAS THE GDA THAT RUPTURED, GIVEN THAT IT HAD BEEN EMBOLIZED AND EXCLUDED, I.E., NO LONGER PERFUSED AND PRESSURIZED BY THE HEPATIC ARTERY." NO ADDITIONAL EVALUATION WAS PERFORMED AT THE TIME OF THE EVENT AS WELL AS NO AUTOPSY OR FURTHER EXAMINATIONS WERE PERFORMED TO CONFIRM THE CAUSE OF DEATH. A REVIEW OF THE LOT HISTORY RECORDS AND ACCOMPANYING LOT RELEASE REPORTS IDENTIFIED NO QUALITY ISSUES RELATED TO THE DEVICE PERFORMANCE OR INDICATIVE OF A DEVICE MALFUNCTION WHICH COULD BE ATTRIBUTABLE TO THE EVENT. HOWEVER, WITHOUT ADDITIONAL INFORMATION THE SPECIFIC ROOT CAUSE ASSOCIATED WITH THE EVENT CANNOT BE DEFINITIVELY DETERMINED. NOTE: TWO LOTS WERE USED DURING TREATMENT AND THERE IS ONLY SPACE ON THE 3500A FORM TO LIST EXPIRATION AND MANUFACTURING DATES FOR ONE LOT. BELOW IS THE COMPLETE INFORMATION FOR BOTH LOTS. F24020501U (67 IMP-FX-12 DEVICES), MANUFACTURE DATE: 02/15/2024, EXPIRATION DATE: 05/15/2027. F22122104U (3 IMP-FX-12 DEVICES), MANUFACTURE DATE: 01/18/2023, EXPIRATION DATE: 01/18/2025. A SUPPLEMENTAL REPORT WILL BE FILED SHOULD ADDITIONAL INFORMATION BE RECEIVED. SHAPE MEMORY MEDICAL INC. COMPLAINT REFERENCE NUMBER: (B)(4)
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE CUSTOMER REPORTED THE PATIENT PASSED AWAY AND THE DREAMSTATION AUTO CPAP DEVICE WILL BE RETURNED. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S WIFE STATING THAT HER HUSBAND PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | REPAIR EVALUATION INFORMATION WAS RECEIVED ON 12/4/2024 AND H11 IS UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS ONE ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. SECTION G AND H IS UPDATED IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
MICRA | K132030 | IT WAS REPORTED THAT DURING IMPLANT OF A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG), THE DEVICE WAS PASSED THROUGH THE TRICUSPID VALVE IN THE AP VIEW, AND WHILE FEELING SLIGHT RESISTANCE, IT WAS INSERTED INTO THE VENTRICLE WITH A BEHAVIOR THAT SEEMED TO BE CAUGHT. DURING CONTRAST IMAGING, THE HINGE PART WAS IMAGED, SO IT WAS NOT PLACED THERE, AND A REQUEST WAS MADE TO CHANGE THE LOCATION, AND THE LOCATION WAS SEARCHED TWO OR THREE TIMES. HOWEVER, IT ONLY WENT TO A SIMILAR LOCATION AND WHILE CONSIDERING THIS, THE PATIENT SUDDENLY DETERIORATED, COMPLAINED OF BACK PAIN, BECAME AGITATED, AND LOST CONSCIOUSNESS. ST ELEVATION WAS SEEN ON THE ELECTROCARDIOGRAM (ECG) AND BREATHING HAD STOPPED SO INTUBATION WAS PERFORMED TO SECURE THE AIRWAY. AFTER THAT, PERICARDIAL EFFUSION WAS CONFIRMED, AND DRAINAGE WAS PERFORMED. ABOUT 600ML TO 800ML WAS DRAINED, AND WHILE PERFORMING BLOOD TRANSFUSION, CARDIAC MASSAGE WAS PERFORMED. THE PATIENT ALSO EXPERIENCED CARDIAC PERFORATION AND TAMPONADE. THIS SITUATION CONTINUED FOR ABOUT AN HOUR, AND SINCE THE BLOOD PRESSURE DID NOT RISE, DEATH WAS CONFIRMED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION | K131982 | THE MANUFACTURER BECAME AWARE OF A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, CANCER AND RESPIRATORY FAILURE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ENDO ANCHOR SYSTEM - HELI-FX AAA | K140036 | LITERATURE WAS REVIEWED REGARDING 'LONG-TERM OUTCOMES IN PATIENTS MANAGED WITH THE ENDURANTTM ENDOGRAFT UNDER ELECTIVE SETTING'. THE TIME FRAME OF THIS STUDY WAS '2008 TO 2024' THE FOLLOWING MEDTRONIC DEVICES WERE USED: 'ENDURANT II/IIS ENDOGRAFTS AND HELI-FX ENDOANCHORS AMONG PATIENT ADVERSE EVENTS INCLUDED: DEATHS OCCURRED IN THE STUDY POPULATION. THE AUTHOR CONFIRMED THAT NO EVENT WAS DIRECTLY RELATED TO MEDTRONIC DEVICES. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | CONTINUATION OF D10: PRODUCT ID: UNK-CV-SR-ENDUR-II (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0180-AORTIC STENT GRAFT; IMPLANT DATE ; EXPLANT DATE; PRODUCT ID UNK-CV-SR-ENDUR-II (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0180-AORTIC STENT GRAFT; IMPLANT DATE ; EXPLANT DATE PRODUCT ID UNK-CV-SR-ENDUR-II (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0180-AORTIC STENT GRAFT; IMPLANT DATE ; EXPLANT DATE PRODUCT ID SA-85 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0001-ACCESSORY; IMPLANT DATE ; EXPLANT DATE PRODUCT ID SA-85 (UNKNOWN SERIAL/LOT); PRODUCT TYPE: 0001-ACCESSORY; IMPLANT DATE ; EXPLANT DATE G2: CITATION: AUTHORS: KONSTANTINOS SPANOS 1, PETROULA NANA 2, GEORGE VOLAKAKIS 1, GEORGE KOUVELOS 1, KONSTANTINOS DAKIS 1, CHRISTOS KARATHANOS 1, ELENI ARNAOUTOGLOU 3, MILTIADIS MATSAGKAS 1, ATHANASIOS GIANNOUKAS. LONG-TERM OUTCOMES IN PATIENTS MANAGED WITH THE ENDURANTTM ENDOGRAFT UNDER ELECTIVE SETTING. JOURNAL OF CLINICAL MEDICINE 2024. DOI: 10.3390/JCM13185601 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
LEAD EXTRACTION BYRD POLYPROPYLENE DILATOR SHEATH SET | K893480 | ON 25NOV2024, PHILIPS JAPAN PROVIDED THE FOLLOWING INFORMATION REGARDING ADVERSE EVENTS OCCURRING DURING LEAD REMOVAL CASES. DEVICES USED: LR-PPLBES-10.0 (LOT UNKNOWN). LR-PPLBES-11.5 (LOT UNKNOWN). PHILIPS/GLIDELIGHT 14FR. PHILIPS/GLIDELIGHT 16FR. THIS WAS A LEAD EXTRACTION PROCEDURE FOR INFECTION. RV LEAD EXTRACTION WAS PROCEEDED USING A PHILIPS/GLIDELIGHT 14FR, A BYRD 10F (LR-PPLBES-10.0), AND A BYRD 11.5F (LR-PPLBES-11.5). PHILIPS/GLIDELIGHT 14FR AND THE OUTER HAD BEEN ADVANCED FROM NEAR THE TRICUSPID VALVE, BUT THE OUTER DID NOT ADVANCE PROPERLY A FEW CENTIMETERS BEFORE THE TIP OF THE LEAD, SO GLIDELIGHT 14FR WAS CHANGED TO A GLIDELIGHT 16FR. THE GLIDELIGHT 16FR WAS ADVANCED, AND THE LASER WAS IRRADIATED FOR 1 SECOND A FEW CENTIMETERS BEFORE THE TIP OF THE LEAD. THE GLIDELIGHT 16FR WAS REMOVED WITHOUT FURTHER LASER IRRADIATION, BUT AT THE MOMENT OF REMOVAL, FLUOROSCOPY SHOWED THAT THE GLIDELIGHT HAD ADVANCED SEVERAL CM TOWARD THE CARDIAC APEX. IMMEDIATELY AFTER REMOVAL OF THE GLIDELIGHT 16FR, PERICARDIAL EFFUSION AND A DROP IN BLOOD PRESSURE WERE CONFIRMED. THE CARDIOVASCULAR SURGERY DEPARTMENT WAS CONTACTED, AND EMERGENCY ECMO, THORACOTOMY, AND RIGHT VENTRICULAR REPAIR WERE PERFORMED. THE GLIDELIGHT PROTRUDED SEVERAL CENTIMETERS FROM THE CARDIAC APEX. ALTHOUGH THE RA LEAD HAD NOT BEEN COMPLETELY REMOVED, IT HAD ALREADY DETACHED FROM THE ATRIAL WALL DURING THE REPAIR SURGERY, AND IT WAS DECIDED THAT IT COULD BE EASILY REMOVED ONCE THE REPAIR HAD SETTLED, SO REMOVAL WAS SCHEDULED TO BE PERFORMED PENDING THE REPAIR. ON 27NOV2024: THE FOLLOWING INFORMATION WAS ADDED BY THE SALES REP. PATIENT ANATOMY BEFORE SURGERY AND AT THE TIME OF FAILURE: PACEMAKER LEAD HAD BEEN IMPLANTED FROM LEFT CHEST TO RIGHT ATRIUM AND RIGHT VENTRICLE RESPECTIVELY. PRIMARY DISEASE: INFECTION THE RIGHT VENTRICLE WAS APPROACHED VIA THE LEFT SUBCLAVIAN VEIN FOR LEAD EXTRACTION. THE USER STARTED WITH A PHILIPS/GLIDELIGHT 14FR AND CONTINUED TO USE IT UP TO THE TRICUSPID VALVE, AND USED A PPLBES 10FR AND A PPLBES 11.5FR IN AREAS WHERE THE GLIDELIGHT 14FR COULD NOT ADVANCE. THE REASON FOR USING TWO SIZES OF PPLBES WAS THAT A 10FR WAS FIRST USED, AND THEN A LARGER SIZE WAS USED AFTER IT WAS DETERMINED THAT THERE WAS SIGNIFICANT ADHESION. THE USER USED PPLBES UP UNTIL SVC, THEN SWITCHED BACK TO GLIDELIGHT 14FR. HOWEVER, IT DID NOT ADVANCE AT THE TRICUSPID VALVE, SO IT WAS CHANGED TO A PHILIPS/GLIDELIGHT 16FR. THERE WAS STRONG ADHESION AT THE CARDIAC APEX, SO THE USER USED THE LASER TO PERFORM COUNTERTRACTION (KEEPING THE SHEATH IN PLACE AND PULLING THE LEAD), AND THE BLOOD PRESSURE DROPPED. PCPS WAS ADMINISTERED AND THEN THORACOTOMY WAS PERFORMED. <SALES REP'S COMMENT> THE RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND COOK PRODUCTS CANNOT BE CONFIRMED (IS UNKNOWN). IT SEEMS THE PATIENT HAD AN INFECTION AND IT WAS NECESSARY TO REMOVE THE LEAD COMPLETELY. THE PATIENT EXPERIENCED RIGHT VENTRICULAR PERFORATION WITH A GLIDELIGHT 16FR, CARDIAC TAMPONADE, AND THE PATIENT WAS OPERATED ON TO SURGICALLY REPAIR THE AFFECTED AREA, BUT DIED IN THE END. HOWEVER, IT IS UNCLEAR WHETHER THE DIRECT CAUSE OF THE PATIENT'S DEATH WAS THE RIGHT VENTRICULAR PERFORATION. ON 28NOV2024: THE FOLLOWING INFORMATION WAS ADDED BY THE SALES REP. ACCORDING TO OUR SALES REP, A PHILIPS REPRESENTATIVE WHO WAS PRESENT DURING THE SURGERY CONFIRMED THAT THE GLIDELIGHT 16FR HAD CLEARLY PROTRUDED FROM THE RV (RIGHT VENTRICLE) WHEN THE CHEST WAS OPENED. AT THAT POINT, THE PHILIPS REPRESENTATIVE REPORTEDLY DECIDED THAT HE COULD NOT BE PRESENT ANY LONGER AND LEFT THE SITE. OUR SALES REPRESENTATIVE ATTEMPTED TO CONFIRM THE CIRCUMSTANCES LEADING UP TO THE PATIENT'S DEATH AND THE CAUSAL RELATIONSHIP/CONNECTION OF THE COOK PRODUCT TO WHAT THE PHYSICIAN BELIEVED WAS THE CAUSE OF DEATH, BUT WAS UNABLE TO OBTAIN ANY INFORMATION FROM THE FACILITY (PHILIPS APPARENTLY DID NOT HAVE THIS INFORMATION EITHER, AND WAS UNABLE TO OBTAIN ANY INFORMATION FROM PHILIPS EITHER). | THE UDI INFORMATION PER 21CFR 803.52(C)(4) AND 21 CFR 803.50(B) THAT IS REQUIRED IS REASONABLY UNKNOWN DUE TO THE DEVICE ITSELF WAS DISCARDED/NOT RETURNED FOR INVESTIGATION AFTER MANUFACTURER REQUESTED FOR THE RETURN AND THE LOT OF THE DEVICE WAS KNOWN. E3 - OCCUPATION: DOCTOR. THE DEVICE WAS NOT RETURNED FOR THE COMPLAINT; THEREFORE, A PHYSICAL INVESTIGATION COULD NOT BE PERFORMED, AND THE CUSTOMER'S COMPLAINT IS ACKNOWLEDGED, BUT COULD NOT BE CONFIRMED, OTHER THAN BY THE CUSTOMER'S TESTIMONY. THE COMPLAINT/EVENT THAT WAS ENTERED AND REPORTED WITHIN TRACKWISE: "PATIENT DIED AFTER PERFORATION OF THE CARDIA APEX AND CONVERSION TO OPEN CHEST SURGERY." THE DEVICE HISTORY RECORD (DHR) WAS REVIEWED. THERE WERE NO SIGNS THAT THE DEVICE WAS NONCONFORMING OR THAT THE DEVICE WAS NOT MANUFACTURED TO CURRENT SPECIFICATIONS. THIS COMPLAINT MODE WILL BE TRACKED, TRENDED AND MONITORED IN CVI COMPLAINT HANDLING AND POST MARKET SURVEILLANCE PROCEDURES. A RISK ASSESSMENT WILL BE PERFORMED WITHIN THE COMPLAINT SUMMARY TAB OF TRACKWISE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED. THERE IS NO ALLEGATION OF DEVICE CONTRIBUTION. THE PATIENT'S WIFE HAS RETURNED DEVICE TO THE DME ALL CARE HOME HEALTH YEARS AGO. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
LEVEEN SUPERSLIM | K113090 | IT WAS REPORTED THAT AN INTRAUTERINE FETAL DEATH OCCURRED. A 2.0CM X 25CM LEVEEN SUPERSLIM NEEDLE ELECTRODE WAS SELECTED FOR USE WITH AN RF 3000 GENERATOR. RADIOFREQUENCY ABLATION (RFA) WAS PERFORMED IN A TWIN REVERSED ARTERIAL PERFUSION (TRAP) PROCEDURE. THE PROCEDURE WAS SUCCESSFULLY COMPLETED UNDER ULTRASOUND, AND AFTER 25 CUMULATIVE MINUTES, THERE WAS ROLL OFF WITH AN IMPEDANCE OF 270OHM. PROCEDURALLY, THE PHYSICIAN DEEMED THE CASE A SUCCESS. THERE WERE NO DEVICE ISSUES. EVERYTHING PROCEDURALLY WAS AS EXPECTED. HOWEVER, A SLIGHTLY DECREASED HEARTBEAT OF THE PUMP TWIN (INTENDED SURVIVING TWIN) WAS NOTED BY THE END OF THE PROCEDURE. THE DECREASED HEARTBEAT WAS A CONSEQUENCE OF HEMODYNAMIC CHANGES IN RESPONSE TO RFA. IT WAS REPORTED THAT THE POSSIBILITY OF THIS OCCURRING WAS DISCUSSED WITH THE PATIENT. SUBSEQUENTLY, THE PUMP TWIN WAS UNWELL TO THE POINT OF BEING NON-VIABLE/DECEASED. THE MOTHER WAS WELL. | CORRECTION TO H6: EVALUATION CONCLUSION CODES.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER WAS CONTACTED WITH INFORMATION THAT THE PATIENT HAD PASSED AWAY. THE PATIENT'S SON CONTACTED THE MANUFACTURER ABOUT A CHECK HE RECEIVED THAT IS IN HIS FATHER'S NAME, SEEKING TO CHANGE THE NAME ON THE CHECK. THAT IS HOW THE MANUFACTURER WAS INFORMED OF THE PATIENT'S DEATH. PER THE PMS CLINICAL EXPERT, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. ADDITIONAL INFORMATION HAS BEEN REQUESTED TO VERIFY IF AN EVENT RELATED TO DEVICE USAGE MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DUODOPA_DUOPA | K142793 | THE REPORTED EVENT INVOLVES A PATIENT IN ITALY WHO UNDERWENT A PROCEDURE FOR THE PLACEMENT OF A PERCUTANEOUS ENDOSCOPIC GASTROSTOMY (PEG) TUBE WITH JEJUNAL (PEG-J) TUBE ON (B)(6) 2024. ON (B)(6) 2024, THE PATIENT DIED OF CARDIAC COMPLICATIONS. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. DESPITE MULTIPLE QUERIES, THERE IS NO CLEAR EVIDENCE LINKING THE DEVICE TO THE REPORTED EVENT, AND NO ADDITIONAL INFORMATION REGARDING DEVICE-RELATED ALLEGATIONS, MALFUNCTIONS, COMPLICATIONS, DIAGNOSTIC TESTING, OR TREATMENT IS AVAILABLE. THEREFORE, OUT OF ABUNDANCE OF CAUTION, ABBVIE IS CONSERVATIVELY REPORTING THE EVENT WITHOUT ATTRIBUTING CAUSALITY TO THE DEVICE. | REFERENCE NUMBER (B)(6). CATALOG NUMBER IN D4 IS THE INTERNATIONAL LIST NUMBER WHICH IS SIMILAR TO US LIST NUMBER OF 062910 (062945). THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. POTENTIAL POST PROCEDURAL COMPLICATION IS A KNOWN COMPLICATION OF A PEG- J TUBE PLACEMENT. H6 CODE OF E2402 WAS CHOSEN TO CAPTURE THE EVENT OF POSSIBLE POST PROCEDURAL COMPLICATION. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
DS2ADV AUTO CPAP W/HUMID CELL/BT | K200480 | THIS REPORT IS BEING SUBMITTED TO INCLUDE THE DEVICE EVALUATION AND THE PRODUCT UDI. IT WAS REPORTED THAT THE DREAMSTATION 2 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE PRODUCT INVESTIGATION LAB (PIL) FOR EVALUATION. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGESTS THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. PIL IS NOT ABLE TO ADDRESS ANY SYMPTOMS SPECIFIED IN THE COMPLAINT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | THIS REPORT IS BEING SUBMITTED TO INCLUDE THE DEVICE EVALUATION AND THE PRODUCT UDI. IT WAS REPORTED THAT THE DREAMSTATION 2 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE WAS RETURNED TO THE PRODUCT INVESTIGATION LAB (PIL) FOR EVALUATION. THERE IS NO VISIBLE DAMAGE OR FUNCTIONALITY FAILURES OF THE DEVICE, WHICH SUGGESTS THAT THE SOURCE OF CONTAMINATION WAS EXTERNAL TO THE DEVICE. PIL IS NOT ABLE TO ADDRESS ANY SYMPTOMS SPECIFIED IN THE COMPLAINT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION CPAP | K131982 | IT WAS REPORTED THAT THE DREAMSTATION CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | IT WAS REPORTED THAT THE REP DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT'S WIFE REQUESTED THAT NO FURTHER CORRESPONDENCE BE SENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER WAS CONTACTED WITH INFORMATION THAT THE PATIENT HAD PASSED AWAY - NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED TO VERIFY IF AN EVENT RELATED TO DEVICE USAGE MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED REGARDING BEING CONTACTED WITH INFORMATION THAT THE PATIENT HAD PASSED AWAY - NO FURTHER INFORMATION WAS PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED TO VERIFY IF AN EVENT RELATED TO DEVICE USAGE MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE ATTEMPTS WERE UNSUCCESSFUL. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER, WHICH ASSESSED IT AND FOUND NO ERRORS. ADDITIONALLY, NO FAULTS WERE IDENTIFIED, AND THE DEVICE SUCCESSFULLY PASSED ALL TESTS BEFORE BEING RETURNED TO LOAN. NO FURTHER INVESTIGATION WILL BE CONDUCTED CURRENTLY. HOWEVER, IF NEW INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE SUBMITTED.
REP DREAMSTATION AUTO CPAP, DOM | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S WIFE STATING THAT HER HUSBAND PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER WHERE IT WAS INSPECTED. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS LOCATED IN THE LOGS, AND NO FAULTS WERE FOUND. THE DEVICE'S FILTERS, TUBING, AND MANUAL WERE REPLACED AND THE DEVICE SENT TO BE PUT BACK INTO LOAN STOCK. SOFTWARE V1.0.6.620. THE DEVICE PASSED ALL TESTS.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH." DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, "NO FAULTS WERE FOUND. ZERO ERRORS LOCATED IN THE LOGS. THE DEVICE'S FILTERS, TUBING AND MANUAL WERE REPLACED. THE DEVICE TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO LOAN STOCK."
BIPAP AUTOSV ADVANCED SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION FROM PATIENT'S WIFE, STATING THAT THE PATIENT IS DECEASED, AND SHE WOULD LIKE TO SEND BACK THE RECALLED DEVICE. SHE NEEDS A SHIPPING LABEL TO RETURN THE DEVICE. THE CALLER WAS ASKED IF THE LABEL COULD BE SENT VIA EMAIL, BUT CALLER WOULD PREFER THE LABEL SENT VIA MAIL. SHIPPING LABEL HAS BEEN SENT THE CALLERS 'ADDRESS. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE PASSED AWAY AND THE DEVICE WILL BE RETURNED TO THE MANUFACTURER. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE PASSED AWAY AND THE DEVICE WILL BE RETURNED TO THE MANUFACTURER. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND ERROR CODE E-370 (ERR_FR_ERROR_MISC) A CONTINUE LEVEL-ERROR NOT DUE TO ANY COMPONENT FAILURE. THERE WAS NO VALID COMPLAINT TO BE REPRODUCED. THE DEVICE WAS IN GOOD CONDITION AND WAS WORKING FINE. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER UPDATED BOX H IN THIS REPORT.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED CANCER, DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PLUS C-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT HAD DIED BECAUSE OF CANCER BUT SHE DOESN'T KNOW OF THE MACHINE WAS THE CAUSE OF HIS DEATH. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND VISIBLE FOAM PARTICLES WERE NOT OBSERVED. ADDITIONALLY, THE PATIENT¿S COMPLAINT WAS NOT CONFIRMED, AND THE DEVICE WAS SCRAPPED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
AZURION | K200917 | IT HAS BEEN REPORTED TO PHILIPS THAT THE C-ARM STOPPED MOVING DURING A PROCEDURE. A SYSTEM RESTART DID NOT RESOLVE THE ISSUE, AND THE PATIENT WAS MOVED TO ANOTHER ROOM. THE REPORT INDICATED THAT THE PATIENT PASSED AWAY BEFORE THE PROCEDURE COULD BE CONTINUED IN THE OTHER ROOM. PHILIPS HAS STARTED AN INVESTIGATION OF THIS COMPLAINT.
75MM SELECTABLE NEW TLC | K092577 | IT WAS REPORTED THAT THE DOCTOR COMMENTED THAT ON THE SECOND POSTOPERATIVE DAY THE PATIENT HAD A STAPLE LINE OPEN AND DIED. | (B)(4). DATE SENT: 12/4/2024. D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WAS THE PATIENT'S ORIGINAL DIAGNOSIS? PATIENT SCHEDULED FOR A WHIPPLE PROCEDURE CAN YOU PROVIDE DETAILS OF THE ORIGINAL PROCEDURE? WHIPPLE WHAT STAPLERS AND STAPLES (CARTRIDGES) WERE USED? I USE NTLC 75. WHAT STRUCTURES WAS THE STAPLER USED ON? DUODENUM WHICH STAPLE LINE OPENED? DUODENAL FOR THE STAPLE LINE THAT OPENED, WAS IT A VASCULAR FIRING OR A GI FIRING? G.I. HOW WAS THE OPEN STAPLE LINE IDENTIFIED? REINTERVENTION 2 DAYS AFTER THE PROCEDURE. HOW WAS THE OPEN STAPLE LINE ADDRESSED? I DO NOT HAVE INFORMATION WERE THE ETHICON DEVICES USED IN THIS PROCEDURE REPROCESSED? NO WAS AN AUTOPSY PERFORMED? NO INFORMATION IS AVAILABLE IS THE CAUSE OF DEATH KNOWN? THE SPECIALIST STATES: MULTIFACTORIAL IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL? WE EXPRESSED OUR INTEREST IN CLARIFYING THE SITUATION, IT WAS EXPLAINED TO HIM THAT A PRODUCT EVENT REPORT WAS ESTABLISHED WITH THE LITTLE DATA THAT IS AVAILABLE AND WE SHARED THAT WE COULD BE CONTACTING HIM TO FOLLOW UP. HE STATED THAT HE HAS NO PROBLEM WITH IT, BUT DURING THAT CONVERSATION THERE WAS NO EVIDENCE OF ANY INTENTION TO CONTACT ETHICON MEDICAL. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
FMC CASSETTE | K173718 | ON 19/NOV/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE HOSPITALIZED ON (B)(6) 2024 DUE TO SEPSIS CAUSED BY PERITONITIS. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. MEDICAL RECORDS REVEALED THE PATIENT PRESENTED TO THE EMERGENCY ROOM (ER) ON (B)(6) 2024 WITH SYMPTOMS OF SEPTIC SHOCK INCLUDING: FEVER, DIAPHORESIS, ABDOMINAL PROTUBERANCE/TENDERNESS, TACHYCARDIA (HEART RATE = 124), HYPOTENSION (BLOOD PRESSURE = 98/58), TACHYPNEIC (RATE = 35), ABDOMINAL PAIN (ONSET - ON (B)(6) 2024), AND CLOUDY EFFLUENT (ONSET - ON (B)(6) 2024). THE PATIENTS¿ POINT OF CONTACT (POC) REPORTED THE PATIENT RECENTLY UNDERWENT A DIAGNOSTIC LAPAROSCOPY (DATE NOT PROVIDED) FOR THE EVALUATION OF HIS PD CATHETER (NOT A FRESENIUS PRODUCT) AND HAS NOT ¿ENTIRELY BEEN HIMSELF EVER SINCE.¿ HEMATOLOGICAL STUDIES REVEALED THE PATIENT WAS SUFFERING FROM SEVERE METABOLIC ACIDOSIS (LACTIC ACID = 8.6 MMOL/L), ELEVATED LIVER FUNCTION TESTS, AND LEUKOCYTOSIS. THE PATIENT WAS INITIALLY TREATED WITH 2.0 LITERS OF INTRAVENOUS (IV) NORMAL SALINE, WHICH IMPROVED THE LACTIC ACID LEVEL (5.6 MMOL/L) AND BLOOD PRESSURE, HOWEVER THE PATIENT¿S HYPOTENSION RETURNED AND REQUIRED THE ADMINISTRATION OF NOREPINEPHRINE. A PERITONEAL EFFLUENT CELL COUNT WAS ALSO PERFORMED AND WAS POSITIVE FOR GRAM-POSITIVE COCCI, AND A COMPUTED TOMOGRAPHY (CT) SCAN DISCOVERED A BOWEL PERFORATION, SMALL BOWEL ISCHEMIA, ASCITES, PORTAL VENOUS GAS, PNEUMATOSIS OF THE SMALL INTESTINE, AND ALTHOUGH THE TIMELINE FOR COLLECTION IS UNKNOWN, PERITONEAL CULTURES WERE POSITIVE FOR MULTIPLE BACTERIA (ORGANISMS NOT PROVIDED). THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) AND WAS INTUBATED SHORTLY THEREAFTER. THE PATIENT RECEIVED IV FLUIDS, VANCOMYCIN, ZOSYN, AND 3 DIFFERENT VASOPRESSORS (VASOPRESSIN, ANGIOTENSIN II, NOREPINEPHRINE) DUE TO SEVERE HYPOTENSION. A SURGICAL CONSULTATION WAS PLACED FOR THE BOWEL PERFORATION, AND AFTER CONSULTING WITH THE GENERAL SURGEON, THE PATIENT/FAMILY CHOSE TO WITHDRAW CARE AND PROVIDE COMFORT MEASURES ONLY (HOSPICE CARE) DUE TO HIS POOR PROGNOSIS (MULTISYSTEM ORGAN FAILURE, ACUTE RESPIRATORY FAILURE, SEPTIC SHOCK, BOWEL PERFORATION/ISCHEMIA, METABOLIC ENCEPHALOPATHY). LIFE SUPPORT WAS REMOVED LATER THE SAME DAY, AND THE PATIENT EXPIRED AT 2122 ON (B)(6) 2024 WITH FAMILY PRESENT. THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED APPROXIMATELY 24 HOURS PRIOR TO HIS PASSING. PER THE MEDICAL EXAMINER, THE PATIENT¿S DEATH WAS DUE TO NATURAL CAUSES AND NO AUTOPSY WAS PERFORMED. THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS SEPTICEMIA, SECONDARY TO A BOWEL PERFORATION AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. THE PATIENT¿S LIBERTY SELECT CYCLER HAS NOT BEEN RETURNED TO THE MANUFACTURER AT THE TIME OF THIS INVESTIGATION. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXIST BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF BOWEL PERFORATION/ISCHEMIA, SEPTIC PERITONITIS, MULTISYSTEM ORGAN FAILURE, ACUTE RESPIRATORY FAILURE, SEPTIC SHOCK, AND METABOLIC ENCEPHALOPATHY (CHARACTERIZED BY MENTAL STATUS CHANGES, FEVER, ABDOMINAL PROTUBERANCE/TENDERNESS, DIAPHORESIS, TACHYCARDIA, HYPOTENSION, TACHYPNEA, CLOUDY PERITONEAL EFFLUENT FLUID) AND DEATH, WHICH REQUIRED HOSPITALIZATION, ICU ADMISSION, AND INTUBATION. THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED APPROXIMATELY 24 HOURS PRIOR TO HIS PASSING, AND PER THE MEDICAL EXAMINER, THE PATIENT¿S DEATH WAS DUE TO NATURAL CAUSES. THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS SEPTICEMIA, SECONDARY TO A BOWEL PERFORATION, AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. SEPTICEMIA IS A LIFE-THREATENING CONDITION WITH A MORTALITY RATING OF 12-22% AND IS PRIMARILY CAUSED BY SEVERE BACTERIAL INFECTIONS. MEDICAL DOCUMENTATION CONFIRMED THE DECISION WAS MADE TO TRANSITION THE PATIENT INTO HOSPICE CARE ON (B)(6) 2024 AFTER A COMPLICATED HOSPITAL COURSE, AND THE PATIENT EXPIRED AS A RESULT LATER THE SAME DAY. HOSPICE CARE IS CONSIDERED MEDICAL INTERVENTION, WITH THE EXPECTATION BEING THE PATIENT WILL EXPIRE AS A RESULT. THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION PATENT WAS DECEASED. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PARKER | K984528 | IT WAS STATED BY CUSTOMER THAT "WHILE INTUBATING THE PATIENT, THE INFLATING TUBE WAS DISLODGED AND THE CUFF ATROPHIED.THE PATIENT LATER DIED. THE HOSPITAL WAS OF THE OPINION THAT THERE WAS NO CAUSAL RELATIONSHIP BETWEEN THE DISLODGED INFLATING TUBE AND THE DEATH. WE DO NOT KNOW THE CAUSE-AND-EFFECT RELATIONSHIP". | H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): ENDOTRACHEAL TUBE. THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 05 DEC 2024 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE . AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT - (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. COMPLAINT WAS CONSIDERED CONFIRMED BASED ON CUSTOMER NARRATIVE OF EVENTS. A 24 MONTH REVIEW WAS CONDUCTED AND 1 ADDITIONAL COMPLAINTS WERE IDENTIFIED RELATED TO THIS ISSUE HOWEVER NO TREND WAS OBSERVED. THE LOT NUMBER WAS NOT PROVIDED, HOWEVER THE VENDOR WAS NOTIFIED OF THE INCIDENT. THIS INCIDENT WAS NOT IDENTIFIED AS AN INCREASING TREND. NO FURTHER ACTION TO BE TAKEN.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | ON 19/NOV/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE HOSPITALIZED ON (B)(6) 2024 DUE TO SEPSIS CAUSED BY PERITONITIS. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. MEDICAL RECORDS REVEALED THE PATIENT PRESENTED TO THE EMERGENCY ROOM (ER) ON (B)(6) 2024 WITH SYMPTOMS OF SEPTIC SHOCK INCLUDING: FEVER, DIAPHORESIS, ABDOMINAL PROTUBERANCE/TENDERNESS, TACHYCARDIA (HEART RATE = 124), HYPOTENSION (BLOOD PRESSURE = 98/58), TACHYPNEIC (RATE = 35), ABDOMINAL PAIN (ONSET - 17/NOV/2024), AND CLOUDY EFFLUENT (ONSET - 18/NOV/2024). THE PATIENTS¿ POINT OF CONTACT (POC) REPORTED THE PATIENT RECENTLY UNDERWENT A DIAGNOSTIC LAPAROSCOPY (DATE NOT PROVIDED) FOR THE EVALUATION OF HIS PD CATHETER (NOT A FRESENIUS PRODUCT) AND HAS NOT ¿ENTIRELY BEEN HIMSELF EVER SINCE.¿ HEMATOLOGICAL STUDIES REVEALED THE PATIENT WAS SUFFERING FROM SEVERE METABOLIC ACIDOSIS (LACTIC ACID = 8.6 MMOL/L), ELEVATED LIVER FUNCTION TESTS, AND LEUKOCYTOSIS. THE PATIENT WAS INITIALLY TREATED WITH 2.0 LITERS OF INTRAVENOUS (IV) NORMAL SALINE, WHICH IMPROVED THE LACTIC ACID LEVEL (5.6 MMOL/L) AND BLOOD PRESSURE, HOWEVER THE PATIENT¿S HYPOTENSION RETURNED AND REQUIRED THE ADMINISTRATION OF NOREPINEPHRINE. A PERITONEAL EFFLUENT CELL COUNT WAS ALSO PERFORMED AND WAS POSITIVE FOR GRAM-POSITIVE COCCI, AND A COMPUTED TOMOGRAPHY (CT) SCAN DISCOVERED A BOWEL PERFORATION, SMALL BOWEL ISCHEMIA, ASCITES, PORTAL VENOUS GAS, PNEUMATOSIS OF THE SMALL INTESTINE, AND ALTHOUGH THE TIMELINE FOR COLLECTION IS UNKNOWN, PERITONEAL CULTURES WERE POSITIVE FOR MULTIPLE BACTERIA (ORGANISMS NOT PROVIDED). THE PATIENT WAS ADMITTED TO THE INTENSIVE CARE UNIT (ICU) AND WAS INTUBATED SHORTLY THEREAFTER. THE PATIENT RECEIVED IV FLUIDS, VANCOMYCIN, ZOSYN, AND 3 DIFFERENT VASOPRESSORS (VASOPRESSIN, ANGIOTENSIN II, NOREPINEPHRINE) DUE TO SEVERE HYPOTENSION. A SURGICAL CONSULTATION WAS PLACED FOR THE BOWEL PERFORATION, AND AFTER CONSULTING WITH THE GENERAL SURGEON, THE PATIENT/FAMILY CHOSE TO WITHDRAW CARE AND PROVIDE COMFORT MEASURES ONLY (HOSPICE CARE) DUE TO HIS POOR PROGNOSIS (MULTISYSTEM ORGAN FAILURE, ACUTE RESPIRATORY FAILURE, SEPTIC SHOCK, BOWEL PERFORATION/ISCHEMIA, METABOLIC ENCEPHALOPATHY). LIFE SUPPORT WAS REMOVED LATER THE SAME DAY, AND THE PATIENT EXPIRED AT 2122 ON (B)(6) 2024 WITH FAMILY PRESENT. THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED APPROXIMATELY 24 HOURS PRIOR TO HIS PASSING. PER THE MEDICAL EXAMINER, THE PATIENT¿S DEATH WAS DUE TO NATURAL CAUSES AND NO AUTOPSY WAS PERFORMED. THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS SEPTICEMIA, SECONDARY TO A BOWEL PERFORATION AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. THE PATIENT¿S LIBERTY SELECT CYCLER HAS NOT BEEN RETURNED TO THE MANUFACTURER AT THE TIME OF THIS INVESTIGATION. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXIST BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF BOWEL PERFORATION/ISCHEMIA, SEPTIC PERITONITIS, MULTISYSTEM ORGAN FAILURE, ACUTE RESPIRATORY FAILURE, SEPTIC SHOCK, AND METABOLIC ENCEPHALOPATHY (CHARACTERIZED BY MENTAL STATUS CHANGES, FEVER, ABDOMINAL PROTUBERANCE/TENDERNESS, DIAPHORESIS, TACHYCARDIA, HYPOTENSION, TACHYPNEA, CLOUDY PERITONEAL EFFLUENT FLUID) AND DEATH, WHICH REQUIRED HOSPITALIZATION, ICU ADMISSION, AND INTUBATION. THE PATIENT¿S LAST CCPD TREATMENT WAS COMPLETED APPROXIMATELY 24 HOURS PRIOR TO HIS PASSING, AND PER THE MEDICAL EXAMINER, THE PATIENT¿S DEATH WAS DUE TO NATURAL CAUSES. THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS SEPTICEMIA, SECONDARY TO A BOWEL PERFORATION, AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. SEPTICEMIA IS A LIFE-THREATENING CONDITION WITH A MORTALITY RATING OF 12-22% AND IS PRIMARILY CAUSED BY SEVERE BACTERIAL INFECTIONS. MEDICAL DOCUMENTATION CONFIRMED THE DECISION WAS MADE TO TRANSITION THE PATIENT INTO HOSPICE CARE ON (B)(6) 2024 AFTER A COMPLICATED HOSPITAL COURSE, AND THE PATIENT EXPIRED AS A RESULT LATER THE SAME DAY. HOSPICE CARE IS CONSIDERED MEDICAL INTERVENTION, WITH THE EXPECTATION BEING THE PATIENT WILL EXPIRE AS A RESULT. THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING INFORMATION WAS REPORTED: ON (B)(6) 2024, THE PATIENT WAS BEING TREATED FOR A LARGE SYMPTOMATIC LEFT COMMON ILIAC ARTERY ANEURYSM. IT WAS REPORTED THAT DURING ACCESS THE PHYSICIAN WAS HAVING ISSUES WITH THE WIRE BECOMING COILED. DURING THIS TIME THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA) HAD NOTED THAT THE PATIENTS BLOOD PRESSURE WAS LOW. REPORTEDLY THE BLOOD PRESSURE WAS LOW DURING ACCESS. THE PHYSICIAN OBTAINED BILATERAL (LEFT AND RIGHT) COMMON FEMORAL ARTERY (CFA) ACCESS AND PLACED A MEDTRONIC MVP 9Q INTO THE LEFT INTERNAL ILIAC ARTERY VIA THE RIGHT CFA ACCESS. NEXT, HE PLACED AN AMPLATZ WIRE INTO THE AORTA VIA THE LEFT CFA ACCESS WITH A 5FR 65CM GLIDECATH. THIS WIRE AND CATHETER BECAME COILED IN THE 4.6CM COMMON ILIAC ARTERY. IT HAD BEEN DETERMINED THAT THE LEFT COMMON ILIAC ANEURYSM HAD RUPTURED. THE PHYSICIAN INSERTED AN OCCLUSION BALLOON IN THE DISTAL AORTA THROUGH THE RIGHT CFA ACCESS. THE PATIENT DID NOT STABILIZE, AND BLOOD WAS ORDERED. A GORE® DRYSEAL SHEATH WAS ADVANCED AND A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG DEVICE WAS IMPLANTED BUT WAS NOT LONG ENOUGH TO COVER THE RUPTURED LEFT COMMON ILIAC ARTERY. WIRE ACCESS WAS LOST AND THE PHYSICIAN CHOSE TO PERFORM A CUT DOWN REPAIR AND LIGATE THE ILIAC ARTERY. A FEM-FEM BYPASS WAS PERFORMED. TEN UNITS OF BLOOD WAS GIVEN AND THE PATIENT CODED ONCE THE OCCLUSION BALLOON WAS REMOVED AND HAD A MASSIVE MYOCARDIAL INFARCTION. THE PATIENT EXPIRED IN THE OPERATING ROOM. PATIENT HAD VERY TORTUOUS ANATOMY, AND A HISTORY OF CARDIAC DISEASE. IT IS UNKNOWN IF THE ANEURYSM RUPTURED BEFORE OR AFTER THE INSERTION OF THE DRYSEAL SHEATH. | THE GORE® DRYSEAL SHEATH INSTRUCTIONS FOR USE (IFU) STATES THAT ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: VASCULAR TRAUMA (RUPTURE) AND DEATH.
HEATER/COOLER SYSTEM | K220635 | LIVANOVA RECEIVED REPORT OF PATIENT INFECTION. THE SURGERY WAS AN AORTIC VALVE REPLACEMENT DONE ON (B)(6) 2015. PATIENT DECEASED (B)(6) 2022. | A.1.-A.5. PATIENT INFORMATION WAS NOT PROVIDED. D.4. SERIAL NUMBER IS UNKNOWN. SINCE THE SN IS UNKNOWV, ALSO UDI IS UNKNOWN. THIS INFORMATION WILL BE PROVIDED IN A SUPPLEMENTAL REPORT IF MADE AVAILABLE. H.4. AS THE SERIAL NUMBER IS UNKNOWN, THE DEVICE MANUFACTURE DATE COULD NOT BE DETERMINED. THIS INFORMATION WILL BE PROVIDED IN A SUPPLEMENTAL REPORT IF MADE AVAILABLE. H.9. LIVANOVA DEUTSCHLAND IMPLEMENTED A FIELD SAFETY NOTICE FOR DISINFECTION AND CLEANING OF HEATER-COOLER DEVICES. THE Z NUMBER IS Z-2076/2081-2015. H11: LIVANOVA DEUTSCHLAND MANUFACTURES THE HEATER-COOLER SYSTEM 3T. A DHR REVIEW COULD NOT BE PERFORMED SINCE POSSIBLE SERIAL NUMBERS INVOLVED WERE NOT PROVIDED. SOURCE OF PATIENT CONTAMINATION REMAINS UNKNOWN AND THE LIVANOVA DEVICE INVOLVEMENT AS WELL. INVOLVED DEVICE SERIAL NUMBER REMAINS UNKNOWN AND WAS NOT EQUIPPED WITH VACUUM AND SEALING KIT AT THE TIME OF SURGERY (2015) SINCE UPGRADE ACTIVITY STARTED IN 2017. LIVANOVA ALREADY IMPLEMENTED A STRATEGY TO DECREASE THE PROBABILITY OF BACTERIA GROW IN THE HC DEVICE BY APPLYING MULTIPLE MEASURES IMPLEMENTED OVER THE PAST FEW YEARS THROUGH DEDICATED CAPA AND FIELD ACTION. LIVANOVA BECAME AWARE OF THESE CASES THROUGH LEGAL ACTIONS, WE DO NOT EXPECT TO RECEIVE ADDITIONAL INFORMATION IN THE NEAR FUTURE. THE RISK IS IN THE ACCEPTABLE REGION. NO SPECIFIC ACTION WAS CURRENTLY DEEMED NECESSARY. LIVANOVA WILL KEEP MONITORING THE MARKET.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR SURGERY OR INTERVENTION - UNPLANNED TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 77 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 4.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR COMPLICATION - MAJOR TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 6 VASCULAR COMPLICATION - MAJOR AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 75 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 2 FEMALE(S) AND 4 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 6 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.33. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR SURGERY OR INTERVENTION - UNPLANNED TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 11 VASCULAR SURGERY OR INTERVENTION - UNPLANNED AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 94 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 6 FEMALE(S) AND 5 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 11 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 4.73. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR COMPLICATION - MAJOR TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 18 VASCULAR COMPLICATION - MAJOR AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 98 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 10 FEMALE(S) AND 8 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 18 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 42.94. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - HEMATOMA AT ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 2 BLEEDING - HEMATOMA AT ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 58 - 89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 2 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 4 BLEEDING - ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 79 - 89 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 FEMALE(S) AND 3 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES (B)(4) BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.5. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
CONFIDA BRECKER CURVE GUIDEWIRE | K132623 | MEDTRONIC RECEIVED INFORMATION THAT PRIOR TO THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A PRE-IMPLANT BALLOON AORTIC VAL VULVOPLASTY WAS PERFORMED DUE TO CALCIFICATION. SUBSEQUENTLY, THE VALVE AND DELIVERY CATHETER SYSTEM WERE INSERTED INTO THE PATIENT, AND THE VALVE WAS IMPLANTED. FAILURE OF THE ACCESS SITE CLOSURE DEVICE WAS OBSERVED, AND A LEFT VENTRICLE PERFORATION WAS NOTED. SUBSEQUENTLY, THE PATIENT'S CONDITION BEGAN TO DECOMPENSATE. A STERNOTOMY WAS PERFORMED AND CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED; HOWEVER, THE PATIENT PASSED AWAY. PER THE PHYSICIAN, THE CAUSE OF DEATH WAS THE LEFT VENTRICLE PERFORATION AND THE VALVE CONTRIBUTED TO THE PATIENT'S SYMPTOMS. THE PROCEDURE WAS DELAYED BY APPROXIMATELY 7 TO 8 HOURS. ADDITIONAL INFORMATION WAS RECEIVED THAT A RETROPERITONEAL BLEED WAS ALSO OBSERVED AND CONTRIBUTED TO THE DEATH. PER THE PHYSICIAN, THE POSSIBLE CAUSE OF THE LEFT VENTRICLE PERFORATION WAS AN UNCONFIRMED GUIDEWIRE. THE PERFORATION WAS FIRST OBSERVED DURING THE STERNOTOMY AND TREATED DURING THE OPEN HEART SURGERY. ADDITIONALLY, THE RIGHT FEMORAL ARTERY, WITH A MINIMUM DIAMETER OF 6.5 MM, WAS USED AS THE ACCESS SITE FOR THE DCS. SURGICAL CUTDOWN WAS PROVIDED FOR THE ACCESS SITE INJURY. THE CPR WAS INITIATED POST-PROCEDURE. NO ADDITIONAL ADVERSE PATIENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION AS RECEIVED THAT THE GUIDEWIRE USED WAS A MEDTRONIC GUIDEWIRE. | CONTINUATION OF D10: PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012338858); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012328696); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID EVFXPLUS-23 (SERIAL: (B)(6); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE: (B)(6) 2024; EXPLANT DATE N/A. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR COMPLICATION - MAJOR TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 27 VASCULAR COMPLICATION - MAJOR AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 46 - 96 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 13 FEMALE(S) AND 14 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 27 VASCULAR COMPLICATION - MAJOR DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 41.78. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE THV CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - RETROPERITONEAL TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - RETROPERITONEAL AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 87 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - RETROPERITONEAL DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
ESHEATH PLUS INTRODUCER SHEATH | K200258 | AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT UNDERWENT A TRANSFEMORAL TAVR PROCEDURE WITH A 29MM SAPIEN 3 ULTRA RESILIA VALVE IN AORTIC POSITION. DURING THE PROCEDURE, AN ILIAC RUPTURE/DISSECTION OCCURRED WHILE ADVANCING THE VALVE THROUGH THE COMMON ILIAC. THE VALVE AND DELIVERY SYSTEM EXITED THE SHEATH AT THE COMMON ILIAC. THE PATIENT IMMEDIATELY LOST BLOOD PRESSURE AND NEVER RECOVERED. THE SURGICAL HISTORY OF THE PATIENT SHOWS THE PATIENT APPEARS TO HAVE A TUBE GRAFT IN THE DISTAL ABDOMINAL AORTA THAT MAY HAVE CREATED SOME ARTIFICIAL, ADDITIONAL UNEXPECTED TORTUOSITY GOING FROM THE ILIACS TO THE AORTA. AS PER FOLLOW-UP WITH THE FCS, THERE WAS RESISTANCE DURING THE INSERTION OF THE ESHEATH AND DELIVERY SYSTEM. THE ESHEATH HAD SPLIT (MID SHEATH) AND THE VALVE AND DELIVERY SYSTEM BECAME EXTRAVASCULAR. THE DELIVERY SYSTEM AND VALVE WERE NOT ABLE TO BE RETRIEVED AND REMAINED IN THE PATIENT WHO PASSED. THE ROOT CAUSE OF THE VASCULAR COMPLICATION IS DUE TO AN ESHEATH SPLIT. THE PERCEIVED ROOT CAUSE FOR THE ESHEATH SPLIT IS DUE TO SEVERE TORTUOSITY AND VALVE/DS WOULD NOT PASS THROUGH THE ANASTOMOSIS OF THE COMMON ILIAC AND A 35 YEAR OLD ABDOMINAL AORTA GRAFT. | INVESTIGATION IS ONGOING.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR VASCULAR SURGERY OR INTERVENTION - UNPLANNED TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 19 VASCULAR SURGERY OR INTERVENTION - UNPLANNED AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 59 - 96 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 10 FEMALE(S) AND 9 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 19 VASCULAR SURGERY OR INTERVENTION - UNPLANNED DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.26. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). ACCORDING TO THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR COMPLICATIONS, INCLUDING PERFORATION OR DISSECTION OF VESSELS WHICH MAY REQUIRE INTERVENTION, ARE POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT PROCEDURE. ACCORDING TO LITERATURE REVIEW, AND AS DOCUMENTED IN A CLINICAL TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, VASCULAR COMPLICATIONS ARE A WELL-RECOGNIZED COMPLICATION OF THE TRANSFEMORAL TRANSCATHETER VALVE REPLACEMENT PROCEDURE IN THIS ELDERLY POPULATION WITH MULTIPLE CO-MORBIDITIES. EDWARDS HAS REVIEWED MANY REPORTS, INCLUDING SCREENING DATA RECORDS AND SOURCE DOCUMENTATION OF VASCULAR COMPLICATIONS, AND HAS FOUND THAT THE ROOT CAUSE IS TYPICALLY RELATED TO A COMBINATION OF VESSEL SIZE, TORTUOSITY, AND CALCIFICATIONS. ALTHOUGH THE INCIDENCE IS DECREASING WITH SMALLER SHEATH/DELIVERY SYSTEM SIZES AND PHYSICIAN EXPERIENCE, THERE WILL CONTINUE TO BE CASES IN WHICH VASCULAR COMPLICATIONS WILL OCCUR. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IT ALSO NOTES THAT CALCIFICATION MAY REDUCE LUMEN DIAMETER AND LIMIT OR PREVENT TRANSFEMORAL PASSAGE OF THE DEVICES. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE SAPIEN VALVE CAN BE DELIVERED TRANSFEMORALLY. ASSESSMENT OF LOCATION AND AMOUNT OF CIRCUMFERENTIAL CALCIUM WILL AID IN DETERMINING AREAS OF REDUCED VESSEL DIAMETERS. THE OPERATORS ARE TRAINED TO MEASURE MINIMUM VESSEL DIAMETER TAKING CALCIUM INTO ACCOUNT. THE PHYSICIAN TRAINING MANUAL ALSO LISTS THE MINIMUM RECOMMENDED VESSEL SIZE FOR EACH SIZE DEVICE. DESPITE THE BEST SCREENING TOOLS, A SMALL PERCENTAGE OF PATIENTS WILL HAVE FEMORAL/ILIAC VESSELS THAT ARE NOT AMENABLE TO THE TRANS-FEMORAL APPROACH OR WHERE INCREASED RESISTANCE IS ENCOUNTERED DURING INSERTION OF DEVICES. IN MANY CASES, THE VESSEL MINIMUM LUMINAL DIAMETER (MLD) MAY BE BORDERLINE OR BELOW THE INDICATED SIZE. IN ADDITION, SIGNIFICANT CALCIFICATION AND/OR TORTUOSITY, NOT ALWAYS APPRECIABLE ON IMAGING, COULD BE CONTRIBUTING FACTORS TO THE EVENT. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - HEMATOMA AT ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 53 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 1.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION THEREFORE, THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR AORTIC DEATH EVENTS FOR BLEEDING - ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 8 BLEEDING - ACCESS SITE AORTIC DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 59 - (B)(6) YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 4 FEMALE(S) AND 4 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER (B)(4) AND SUMMARIZES 8 BLEEDING - ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE AORTIC POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 0.5. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4) . THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION, THEREFORE THE EVENT IS BEING REPORTED. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
EDWARDS TRANSFEMORAL ESHEATH UNKNOWN | K200258 | THV/TVT REGISTRY SUMMARY REPORTING FOR ADVERSE EVENT SUBMISSION FOR QUARTER 4 2024 DATA EXTRACT FOR MITRAL DEATH EVENTS FOR BLEEDING - HEMATOMA AT ACCESS SITE TRANSCATHETER HEART VALVE. THIS REPORT SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE MITRAL DEATH EVENT(S) FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE. THE AGE RANGE FOR THESE EVENT(S) IS 82 YEARS. THE BREAKDOWN FOR GENDER IS AS FOLLOWS: 1 MALE(S). | THIS REPORT PROVIDES DATA FROM THV/TVT REGISTRY EXEMPTION NUMBER E2016006 AND SUMMARIZES 1 BLEEDING - HEMATOMA AT ACCESS SITE DEATH EVENTS FOR THE SAPIEN 3 ULTRA RESILIA TRANSCATHETER HEART VALVE IN THE MITRAL POSITION. THE TIME TO EVENT (TTE, IN DAYS) FOR THIS EVENT IS 6.0. DUE TO THE LIMITED INFORMATION AND THE DATA RECEIVED, EDWARDS CANNOT CONFIRM WHICH ESHEATH WAS USED IN THIS CASE, THE ESHEATH+ OR ESHEATH INTRODUCER SET. THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS EXPANDABLE INTRODUCER SHEATH + SET ARE (B)(4). THE DEVICE IDENTIFICATION (DI) NUMBERS FOR EDWARDS ESHEATH INTRODUCER SET ARE (B)(4). THE INSTRUCTIONS FOR USE (IFU) LIST HEMORRHAGE REQUIRING TRANSFUSION OR INTERVENTION AS A POTENTIAL RISK ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT PROCEDURE, INCLUDING POTENTIAL ACCESS COMPLICATIONS ASSOCIATED WITH STANDARD CARDIAC CATHETERIZATION. THE MAJORITY OF TRANSFEMORAL ACCESS RELATED BLEEDING COMPLICATIONS ARE RELATED TO DISEASED ILEO-FEMORAL VESSELS AND/OR PROCEDURAL TECHNIQUE DURING THE INSERTION OR REMOVAL OF THE SHEATH AND/OR DILATORS. THERE ARE CASES WHERE THE BLEEDING IS SIGNIFICANT AND MORE COMPLEX INTERVENTION OR TRANSFUSION IS REQUIRED TO TREAT/PREVENT A PERMANENT INJURY FROM OCCURRING. THE THV PHYSICIAN TRAINING MANUALS INSTRUCT ON PROCEDURAL CONSIDERATIONS FOR SHEATH INSERTION WITH REGARDS TO PROPER SCREENING CRITICAL TO REDUCING VASCULAR COMPLICATIONS. THE IFU CONTRAINDICATES PATIENTS WITH ILIO-FEMORAL VESSEL CHARACTERISTICS THAT WOULD PRECLUDE SAFE PLACEMENT OF SHEATHS SUCH AS SEVERE OBSTRUCTIVE CALCIFICATION OR SEVERE TORTUOSITY. PRE-PROCEDURE SCREENING AND ASSESSMENT OF THE FEMORAL/ILIAC ARTERY INTERNAL DIAMETERS WILL ENABLE THE CLINICIAN TO DETERMINE IF THE VALVE CAN BE DELIVERED TRANSFEMORALLY. IN THIS CASE, SPECIFIC PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE; HOWEVER, A TRANSFUSION OF BLOOD WAS REPORTED IN THE REGISTRY. BY VARC DEFINITION, THE MEET CRITERIA FOR MAJOR VASCULAR COMPLICATION; THEREFORE, THE EVENT IS BEING REPORTED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. THE REPORTED EVENT IS A KNOWN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. SPECIFIC CLINICAL AND PROCEDURAL DETAILS ARE NOT AVAILABLE TO DETERMINE POTENTIAL CONTRIBUTING FACTORS TO THE EVENT, OR IF THE EVENT IS RELATED TO AN EDWARDS DEVICE. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED.
MITRACLIP | K190167 | IT WAS REPORTED THAT ON (B)(6) 2024, A PATIENT PRESENTED WITH GRADE 4 DEGENERATIVE MITRAL REGURGITATION (MR) AND HEART FAILURE FOR A MITRACLIP PROCEDURE. TWO CLIPS WERE IMPLANTED. THE XT CLIP WAS DIFFICULT TO MAINTAIN CAPTURE OF THE POSTERIOR LEAFLET, BUT THE CLIP WAS MOVED LATERAL TO THE FIRST CLIP. THE CLIP WAS IMPLANTED AND STABLE. THE MR WAS REDUCED TO GRADE <1. ON (B)(6) 2024, A SINGLE LEAFLET DEVICE ATTACHMENT (SLDA) WAS OBSERVED DURING THE PRE-DISCHARGE ECHOCARDIOGRAM. THE POSTERIOR LEAFLET WAS DETACHED. THE MR WAS RECURRENT AT GRADE 3. ON (B)(6) 2024, A SECOND CLIP INTERVENTION WAS PERFORMED. THE SAME TRANSSEPTAL PUNCTURE WAS USED FROM THE PREVIOUS CASE. AN NT CLIP WAS ATTEMPTED TO PLACE LATERAL TO THE SLDA CLIP. IT TOOK APPROXIMATELY 30 MINUTES TO GRASP THE LEAFLETS DUE TO MOBILITY FORM THE SLDA, AND IMAGING WAS NOTED TO BE CHALLENGING DUE SHADOWING FROM THE PREVIOUS CLIP. THE CLIP WAS UNABLE TO GRASP BOTH LEAFLETS. BLOOD PRESSURE STARTED TO DROP, AND A PERICARDIAL EFFUSION WAS NOTED ON THE RIGHT SIDE OF THE HEART. A PERICARDIAL TAP WAS PERFORMED, AND 500 CCS OF BLOOD WAS DRAINED, REQUIRING A BLOOD TRANSFUSION. A PLAN TO ATTEMPT AN XT WAS ABORTED DUE TO HEMODYNAMIC INSTABILITY. A TEMPORARY PACEMAKER WAS PLACED. THE PATIENT COULD NOT RECOVER AND PASSED AWAY ON THE TABLE. IT WAS NOTED THAT THE ACT WAS GREATER THAN 500 DURING THE PROCEDURE. PER THE PHYSICIAN, THE CLIPS DID NOT CONTRIBUTE TO THE DEATH. IT IS UNKNOWN IF THE SGC CAUSED OR CONTRIBUTED TO THE DEATH. | INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
NXSTAGE SYSTEM ONE | K170469 | A REPORT WAS RECEIVED ON (B)(6) 2024 FROM THE HOME THERAPY NURSE (HTN) OF A 63 YEAR OLD MALE PATIENT APPROVED FOR PERFORMING SOLO HEMODIALYSIS WITH A MEDICAL HISTORY INCLUDING END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON (B)(6) 2024 FROM THE HOME THERAPY NURSE (HTN) WHO STATED THE PATIENT WAS FOUND UNRESPONSIVE AND STILL CONNECTED TO THE MACHINE, SURROUNDED BY AN UNSPECIFIED AMOUNT OF BLOOD. PER THE NEPHROLOGIST THE SUSPECTED CAUSE OF DEATH WAS CARDIAC EVENT, WITH SECONDARY POSSIBLE HEMORRHAGIC SHOCK. | THERE WAS NO DEVICE MALFUNCTION. THE CYCLER WAS RECEIVED FOR EVALUATION AND SUCCESSFULLY PASSED TESTING. EVALUATIONS OF THE INVOLVED LOG FILES, CARTRIDGE AND DIALYSATE SAK WERE UNREMARKABLE WITH NO DEVICE PROBLEM FOUND. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, CANCER, DEATH AND BRAIN GLIOBLASTOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | IN PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN BOX E. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. IN BOX E: REPORTER COUNTRY HAS BEEN CORRECTED.
DREAMSTATION AUTO | K131982 | THE MANUFACTURER RECEIVED INFORMATION AS PART OF THE UNO RECALL EU PATIENTS CLASS ACTION IN REFERENCE TO A DREAMSTATION AUTO DEVICE WITH AN ALLEGATION OF THE USER BECAME ILL WHILE USING THE DEVICE. THERE WERE NO FURTHER DETAILS PROVIDED EXCEPT PLAINTIFF'S HAVE SUFFERED PERSONAL INJURY (VARYING PER PATIENT, BUT FOR EXAMPLE CHRONIC INFLAMMATION AND CANCER). THE USER/LAWYER IS CLAIMING DAMAGES BASED ON EXPOSURE TO THE RISK/FEAR OF GETTING SICK/DYING, PERSONAL INJURY DAMAGES, AND MORAL DAMAGES OR RELATIVES OF PATIENTS THAT ALLEGEDLY HAVE DECEASED BECAUSE OF THE USE OF THEIR PHILIP'S DEVICE. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND NO EVIDENCE OF FOAM DEGRADATION OR PARTICLES. THE DEVICE WAS SCRAPPED PER THE CUSTOMER REQUEST. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
COLONIC FTRD SET | K170867 | THE PROCEDURE WAS DONE IN AN ELDERLY PATIENT WITH SIGNIFICANT CARDIAC COMORBIDITY AND CORRESPONDING ANTICOAGULATION DRUG THERAPY. THE PATIENT SUFFERED FROM A 3.5 CM LESION IN THE TRANSVERSE COLON. CLIP APPLICATION WAS NOT VERIFIED BEFORE RESECTION. SURGICAL REPAIR WAS NECESSARY. AFTER SURGERY THE PATIENT WAS IN GOOD MEDICAL CONDITION. THE FOLLOWING DAY THE PATIENT WENT INTO CARDIOLOGIST ARREST AND RESUSCITATION BECAME NECESSARY. ACLS REMAINED UNSUCCESSFUL, ALSO GIVEN THE BAD PROGNOSIS DUE TO THE SEVERE PRE-EXISTING UNDERLYING CARDIAC DISEASE. | AS THE DEVICE IN QUESTION WAS DISCARDED THE ASSESSMENT BASED ON THE INFORMATION AVAILABLE. THE PHYSICIAN DESCRIBED THE PROCEDURE COULD BE PERFORMED WITHOUT REALIZING ANY STRANGE. WHEN RETRIEVING THE ENDOSCOPE THE WHITE APPLICATION RING WAS COMPLETELY MOVED FORWARD AND THE CLIP WAS STILL SITTING ON THE CAP. AS RESECTION WAS CONDUCTED WITHOUT VERIFYING CLIP APPLICATION, PERFORATION WAS CAUSED. THIS SITUATION WAS TRIED TO REPLICATE WITH A DEMO FTRD PRODUCT TO SIMULATE THE DESCRIBED CIRCUMSTANCES. RESULTING INTO FOLLOWING CONCLUSIONS: AFTER SEVERAL TRIALS IT WAS NOT POSSIBLE TO REPLICATE THE SEQUENCE OF ISSUES THE USER DESCRIBED. DUE TO THE DIMENSIONS OF THE CLIP AND APPLICATION RING, WHICH ARE DEFINED BY THE TOOLS USED DURING MANUFACTURE, IT IS TECHNICALLY IMPOSSIBLE FOR THE APPLICATION RING TO SLIDE OVER THE CLIP DURING APPLICATION. THE REVIEW OF THE LOT BASED QUALITY RECORDS REVEALED NO ABNORMALITIES. NO INDICATION OF A TECHNICAL DEFECT OF THE FTRD USED. THE DEPLOYMENT PROBLEM APPEARS AS MULTIFACTORIAL. NOTE: THIS REPORT IS PART OF THE ADDITIONAL INFORMATION REQUESTED BY THE FDA AND THE SUBSEQUENT NOTIFICATION, AS COMMUNICATED BY OVESCO ON 24.10.2024 AND REFERS TO: 'FOLLOW UP QUESTIONS FTRD SYSTEM PERFORATION-CLIP DETACHED FAILURE_10-03-2024 ANNEX 1'.
COLONIC FTRD SET | K170867 | INITIAL FTRD TREATMENT WAS WITHOUT COMPLICATIONS. AFTER 1 WEEK, A DELAYED PERFORATION OCCURRED. PATIENT HAD BEEN REGULARLY TREATED WITH PERITONEAL DIALYSIS, WHICH WAS NOT PAUSED AFTER THE PROCEDURE. THE PERFORATION WAS CLOSED BY SURGERY BUT THE COMORBID PATIENT HAD MORE COMPLICATIONS IN THE FURTHER CLINICAL COURSE AND DIED 4 WEEKS AFTER THE PROCEDURE. | INITIAL FTRD TREATMENT WAS WITHOUT COMPLICATIONS. AFTER 1 WEEK, A DELAYED PERFORATION OCCURRED. PATIENT HAD BEEN REGULARLY TREATED WITH PERITONEAL DIALYSIS, WHICH WAS NOT PAUSED AFTER THE PROCEDURE. THE PERFORATION WAS CLOSED BY SURGERY BUT THE COMORBID PATIENT HAD MORE COMPLICATIONS IN THE FURTHER CLINICAL COURSE AND DIED 4 WEEKS AFTER THE PROCEDURE. PATIENTS ON PERITONEAL DIALYSIS ARE COMMONLY SWITCHED TO HEMODIALYSIS BEFORE AND AFTER SURGICAL OR TRANSMURAL INTERVENTIONAL COLORECTAL PROCEDURES. IT IS KNOWN FROM THE LITERATURE THAT PATIENTS OTHERWISE HAVE A SIGNIFICANTLY INCREASED RISK OF ABDOMINAL COMPLICATIONS. THIS WAS NOT DONE IN THIS CASE. THE DEVICE WAS DISCARDED AND NOT AVAILABLE FOR TECHNICAL ANALYSIS BY US. BASED ON THE AVAILABLE INFORMATION, THERE IS NO INDICATION OF A TECHNICAL FAILURE OF THE DEVICE. THIS REPORT IS PART OF THE ADDITIONAL INFORMATION REQUESTED BY THE FDA AND THE SUBSEQUENT NOTIFICATION, AS COMMUNICATED BY OVESCO ON 24.10.2024 AND REFERS TO: "FOLLOW UP QUESTIONS FTRD SYSTEM PERFORATION-CLIP DETACHED FAILURE_10-03-2024 ANNEX 1".
STREAMLINE TEMPORARY PACING LEADS | K012454 | MEDTRONIC RECEIVED INFORMATION THAT ON THE SAME DAY OF IMPLANT OF THIS TEMPORARY PACING WIRE AFTER AN AORTIC VALVE REPLACEMENT PROCEDURE, IT WAS INDICATED THE PATIENT PRESENTED WITH COMPLETE HEART BLOCK WITH SLOW ESCAPE AT 30 BEATS PER MINUTE (BPM) AND REQUIRED ELECTRO-TRAINING. IT WAS STATED THAT AN EXTERNAL PACEMAKER WAS PUT ON SENTINEL. SIGNIFICANT VOMITING WITH INHALATION WAS OBSERVED. TWO DAYS FOLLOWING THE PROCEDURE, THE PATIENT PRESENTED WITH A BRIEF CARDIORESPIRATORY ARREST ON COMPLETE BLOCK. THE PATIENT RECOVERED FOLLOWING A FEW CHEST COMPRESSIONS AND REACTIVATION OF THE EXTERNAL PACEMAKER. IT WAS REPORTED THAT 3 DAYS POST IMPLANT OF THIS UNIPOLAR TEMPORARY PACING WIRE, THE PACING SYSTEM STOPPED WORKING WHEN THE STERNAL WOUND DRESSING WAS REMOVED AND CHANGED. IT WAS CONFIRMED THAT THE PACING WIRES WERE NOT ACCIDENTLY REMOVED DURING WOUND RE-DRESSING. IT WAS FURTHER REPORTED THAT THE EXTERNAL PACEMAKER BOX WAS SWITCHED ON, BUT NO PACING WAS COMMENCED. THE PATIENT WENT INTO CARDIORESPIRATORY ARREST. A EXTERNAL DEFIBRILLATOR WAS USED TO PACE AT 80 BEATS PER MINUTE (BPM). THE PATIENT THEN WENT INTO ARREST AGAIN AND CARDIOPULMONARY RESUSCITATION AND THERAPY WERE PROVIDED AND THE PATIENT RAPIDLY RECOVERED. IT WAS MENTIONED A FEMORAL SEES WAS INSTALLED. SUBSEQUENTLY, 2 DAYS LATER, THE PATIENT DIED. THE CAUSE OF DEATH WAS RECEIVED AS MULTIPLE ORGAN FAILURE. ADDITIONAL INFORMATION WAS RECEIVED WHICH INDICATED THAT TWO TEMPORARY PACING WIRES WERE USED ON THE PATIENT. IT WAS CLARIFIED THAT DURING THE IMPLANT PROCEDURE OF THE TEMPORARY PACING WIRES POST AORTIC VALVE REPLACEMENT, PACING WAS OBSERVED WHILE TESTING. IT WAS FURTHER INDICATED THAT UPON THE PATIENT PRESENTING WITH COMPLETE HEART BLOCK 2 DAYS POST PROCEDURE, PACING WAS PERFORMED USING A EXTERNAL PACEMAKER AT DDD MODE "80/MIN." IT WAS NOTED THAT THERE WERE NO ISSUES OBSERVED WITH THE EXTERNAL PACEMAKER AND ELECTRODES AT THIS TIME. IT WAS FURTHER INDICATED THAT 3 DAYS POST PROCEDURE, JUST AFTER THE STERNAL DRESSING WAS ATTEMPTED TO BE CHANGED AND THE PATIENT HAD GONE INTO CARDIOPULMONARY ARREST, THE EXTERNAL PACEMAKER WAS ON BUT NO LONGER PACING AND THE ELECTRODES DID NOT APPEAR TO HAVE BEEN MOBILIZED. IT WAS MENTIONED THAT THE CABLES AND PACEMAKER WERE NOT CHANGED. IT WAS REPORTED THAT AT THE SAME TIME RESUSCITATION WAS STARTED, THE PATIENT INHALED MASSIVELY BEFORE INTUBATION. IT WAS STATED THAT PACING DID NOT OCCUR WHEN REQUIRED, LEADING TO "PROLONGED CARDIAC ARREST AND FURTHER COMPLICATIONS OF INHALATION, PNEUMOPATHY AND LOW FLOW COMPLICATIONS". IT WAS INDICATED THAT A ELECTRO-DRIVE PROBE (SEES) WAS PLACED BY THE CARDIOLOGY TEAM FOR CONTINUED PACING. IT WAS STATED THAT THE PATIENT WAS TREATED WITH ANTIBIOTICS FOR TREATMENT OF THE INHALATION PNEUMONITIS. IT WAS ALSO STATED THAT EPINEPHRINE AND A NON-SELECTIVE BETA-ADRENERGIC RECEPTOR AGONIST WAS ALSO USED TO TREAT THE PATIENT. LATER THE SAME DAY, IT WAS OBSERVED THAT THE EXTERNAL PACEMAKER AND THESE SAME ELECTRODES WERE WORKING AGAIN WITHOUT ANY ADDITIONAL MANIPULATION. THE SEES WAS LEFT AS SENTINEL AND THE PATIENT WAS EFFECTIVELY "ELECTRO-TRAINED" BY THE EXTERNAL PACEMAKER UNTIL THE PATIENT DIED 2 DAYS LATER. IT WAS REPORTED THAT IT IS THOUGHT THAT THE ELECTRO-DRIVE DEFECT AT THE ORIGIN OF THE CARDIORESPIRATORY ARREST IS LINKED TO THE PATIENT'S DEATH. IT WAS NOTED THAT THE PULMONARY INHALATION IS RESPONSIBLE FOR SEPTIC SHOCK WITH MULTI-ORGAN FAILURE IN THE COURSE, LEADING TO THE PATIENT'S DEATH. IT WAS REPORTED THAT THE LOW FLOW AND PNEUMOPATHY RESULTING FROM COMPLICATIONS OF CARDIAC ARREST, SECONDARY TO LACK OF STIMULATION AND ATRIOVENTRICULAR (AV) BLOCKAGE CONTRIBUTED TO THE PATIENT'S DEATH. ADDITIONAL INFORMATION WAS RECEIVED WHICH STATED THAT A DIFFERENT SET OF TWO TEMPORARY PACING WIRES HAD BEEN IMPLANTED IN THE PATIENT ON THE SAME DAY POST AORTIC VALVE REPLACEMENT, HOWEVER, IT WAS REPORTED THAT THESE WIRES HAD VERY BAD STIMULATION THRESHOLD SO THEY WERE THEN SWITCHED TO A SECOND SET OF WIRES WHICH WERE FOUND TO WORK WELL AND HAVE VERY GOOD STIMULATION THRESHOLD AT THE TIME OF IMPLANT. THE SECOND SET OF WIRES THAT WERE IMPLANTED IN THE PATIENT REMAINED IN THE PATIENT UNTIL THEIR DEATH 5 DAYS FOLLOWING IMPLANT AND WERE SUBSEQUENTLY EXPLANTED AT THAT TIME. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
| K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE PATIENT'S DAUGHTER REPORTS THE PATIENT DID NOT PASS DUE TO USE OF CPAP DEVICE. THE CAUSE OF DEATH IS UNKNOWN. THE DEVICE WILL BE RETURNED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K083526 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND ACUTE RESPIRATORY DISTRESS SYNDROME. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. DESPITE THE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, ESOPHAGEAL CANCER, LUNG CANCER, LYMPHATIC CANCER AND NON-HODGKIN'S LYMPHOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER, DEATH AND PANCREATIC. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | BOX B WAS CORRECTED AND UPDATED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. A REPORT WILL BE FILED WITH ALL APPLICABLE REGULATORY AGENCIES. THE MANUFACTURER'S INVESTIGATION IS ONGOING. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW-UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATS, SINUS INFECTIONS, PNEUMONIA, AND NOSE BLEEDS. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
| K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE TESTING OF THE DEVICE, NO ERRORS OR FAULTS WERE NOTED. THE DEVICE WAS TESTED AND PASSED. THE DEVICE WAS RETURNED TO LOAN STOCK. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, KIDNEY DISEASE/ TOXICITY AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DIZZINESS, HEADACHE, HYPERSENSITIVITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BTK TURBOHAWK | K170191 | LITERATURE WAS REVIEWED REGARDING 'DIRECTIONAL ATHERECTOMY FOR CRITICAL LOWER LIMB ISCHEMIA: A RETROSPECTIVE OBSERVATIONAL STUDY. A TOTAL OF 42 ATHERECTOMIES WERE PERFORMED IN 41 PATIENTS. A TOTAL OF 42 DIRECTIONAL ATHERECTOMY PROCEDURES WITH TURBOHAWK WERE PERFORMED IN 41 PATIENTS. NITREX OR SPIDER FX GUIDEWIRES WERE USED AS PART OF THE PROCEDURE, ALONG WITH OTHER MANUFACTURERS DEVICES. THE DISTAL EMBOLIC PROTECTION FILTER SPIDERFX WAS USED IN 57% OF CASES. ADJUVANT PROCEDURES, SUCH AS TRANSLUMINAL BALLOON ANGIOPLASTY, WERE PERFORMED IN THE PRESENCE OF RESIDUAL STENOSES GREATER THAN 30% AFTER THE PASSAGE OF THE ATHERECTOMY WITH A CHOICE BETWEEN DRUG-ELUTING BALLOONS (NON-MEDTRONIC) OR CONVENTIONAL BALLOONS (NON MEDTRONIC, EVERCROSS OR NANOCROSS) BEING DETERMINED BASED ON PATIENT COMORBIDITIES, LESION CHARACTERISTICS, AND THE AVAILABILITY OF DRUG-ELUTING BALLOONS IN THE INSTITUTION. TECHNICAL SUCCESS WAS ACHIEVED IN 78.6% OF CASES AND PROCEDURAL SUCCESS IN 97.6%. THE FAILURE WAS OBSERVED IN ONE PATIENT WITH A 90% STENOSIS OF THE SUPERFICIAL FEMORAL ARTERY, WHICH, AFTER DIRECTIONAL ATHERECTOMY AND DRUG-ELUTING BALLOON ANGIOPLASTY PERSISTED WITH A 30% STENOSIS. NO FURTHER INTERVENTION WAS PERFORMED, AND LIMB SALVAGE WAS ACHIEVED AFTER SIX MONTHS OF FOLLOW-UP. THE NEED FOR ADJUNCTIVE CONVENTIONAL BALLOON ANGIOPLASTY WAS 21.4%, WHILE DRUG-COATED BALLOON ANGIOPLASTIES WERE PERFORMED IN 50% OF LIMB. PRE-DILATION ANGIOPLASTY WAS REQUIRED IN 42.9% OF PROCEDURES. THERE WERE NINE DEATHS (21%) DURING THE STUDY, WHICH INCLUDED ONE CASE OF SOFT TISSUE SEPSIS IN THE LEG WHERE THE ATHERECTOMY WAS PERFORMED, WITHIN THE FIRST 30 DAYS (N = 1). DEATHS OCCURRING BETWEEN 31 AND 90 DAYS WERE DUE TO DECOMPENSATED HEART FAILURE (N = 1), SEVERE TRAUMATIC BRAIN INJURY ASSOCIATED WITH STATUS EPILEPTICUS (N = 1), SOFT TISSUE SEPSIS DUE TO PROGRESSION OF ARTERIAL DISEASE WITH A LACK OF ACCEPTANCE OF AMPUTATION UPON READMISSION (N = 1), SACRAL AND TROCHANTERIC PRESSURE SORE INFECTION (N = 1), MULTI-LOBAR KPC PNEUMONIA (N = 1), AND URINARY SEPSIS (N = 1). BEYOND 90 DAYS AND UNTIL THE END OF THE FOLLOW-UP, DEATHS WERE DUE TO A STROKE (N = 1) AND A PATIENT WITH A HIP FRACTURE AND END-STAGE RENAL FAILURE (N = 1). COMPLICATIONS OCCURRED IN 14.3% OF ALL PROCEDURES AND ADVERSE EVENTS INCLUDED DISTAL EMBOLISM IN 4.8%, DISSECTION N IN 4.8%, ARTERIOVENOUS FISTULAS OCCURRED IN 4.8% OF THE INTERVENED LIMBS, ONE CASE OF SUPERFICIAL FEMORAL ARTERY PERFORATION (2.4%) AND PUNCTURE SITE COMPLICATIONS WITH AN INFECTED INGUINAL HEMATOMA. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | LITERATURE REF: DOI 10.7759/CUREUS.70840. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
NITREX | K040345 | LITERATURE WAS REVIEWED REGARDING 'DIRECTIONAL ATHERECTOMY FOR CRITICAL LOWER LIMB ISCHEMIA: A RETROSPECTIVE OBSERVATIONAL STUDY. A TOTAL OF 42 ATHERECTOMIES WERE PERFORMED IN 41 PATIENTS. A TOTAL OF 42 DIRECTIONAL ATHERECTOMY PROCEDURES WITH TURBOHAWK WERE PERFORMED IN 41 PATIENTS. NITREX OR SPIDER FX GUIDEWIRES WERE USED AS PART OF THE PROCEDURE, ALONG WITH OTHER MANUFACTURERS DEVICES. THE DISTAL EMBOLIC PROTECTION FILTER SPIDERFX WAS USED IN 57% OF CASES. ADJUVANT PROCEDURES, SUCH AS TRANSLUMINAL BALLOON ANGIOPLASTY, WERE PERFORMED IN THE PRESENCE OF RESIDUAL STENOSES GREATER THAN 30% AFTER THE PASSAGE OF THE ATHERECTOMY WITH A CHOICE BETWEEN DRUG-ELUTING BALLOONS (NON-MEDTRONIC) OR CONVENTIONAL BALLOONS (NON-MEDTRONIC, EVERCROSS OR NANOCROSS) BEING DETERMINED BASED ON PATIENT COMORBIDITIES, LESION CHARACTERISTICS, AND THE AVAILABILITY OF DRUG-ELUTING BALLOONS IN THE INSTITUTION. TECHNICAL SUCCESS WAS ACHIEVED IN 78.6% OF CASES AND PROCEDURAL SUCCESS IN 97.6%. THE FAILURE WAS OBSERVED IN ONE PATIENT WITH A 90% STENOSIS OF THE SUPERFICIAL FEMORAL ARTERY, WHICH, AFTER DIRECTIONAL ATHERECTOMY AND DRUG-ELUTING BALLOON ANGIOPLASTY PERSISTED WITH A 30% STENOSIS. NO FURTHER INTERVENTION WAS PERFORMED, AND LIMB SALVAGE WAS ACHIEVED AFTER SIX MONTHS OF FOLLOW-UP. THE NEED FOR ADJUNCTIVE CONVENTIONAL BALLOON ANGIOPLASTY WAS 21.4%, WHILE DRUG-COATED BALLOON ANGIOPLASTIES WERE PERFORMED IN 50% OF LIMB. PRE-DILATION ANGIOPLASTY WAS REQUIRED IN 42.9% OF PROCEDURES. THERE WERE NINE DEATHS (21%) DURING THE STUDY, WHICH INCLUDED ONE CASE OF SOFT TISSUE SEPSIS IN THE LEG WHERE THE ATHERECTOMY WAS PERFORMED, WITHIN THE FIRST 30 DAYS (N = 1). DEATHS OCCURRING BETWEEN 31 AND 90 DAYS WERE DUE TO DECOMPENSATED HEART FAILURE (N = 1), SEVERE TRAUMATIC BRAIN INJURY ASSOCIATED WITH STATUS EPILEPTICUS (N = 1), SOFT TISSUE SEPSIS DUE TO PROGRESSION OF ARTERIAL DISEASE WITH A LACK OF ACCEPTANCE OF AMPUTATION UPON READMISSION (N = 1), SACRAL AND TROCHANTERIC PRESSURE SORE INFECTION (N = 1), MULTI-LOBAR KPC PNEUMONIA (N = 1), AND URINARY SEPSIS (N = 1). BEYOND 90 DAYS AND UNTIL THE END OF THE FOLLOW-UP, DEATHS WERE DUE TO A STROKE (N = 1) AND A PATIENT WITH A HIP FRACTURE AND END-STAGE RENAL FAILURE (N = 1). COMPLICATIONS OCCURRED IN 14.3% OF ALL PROCEDURES AND ADVERSE EVENTS INCLUDED DISTAL EMBOLISM IN 4.8%, DISSECTION N IN 4.8%, ARTERIOVENOUS FISTULAS OCCURRED IN 4.8% OF THE INTERVENED LIMBS, ONE CASE OF SUPERFICIAL FEMORAL ARTERY PERFORATION (2.4%) AND PUNCTURE SITE COMPLICATIONS WITH AN INFECTED INGUINAL HEMATOMA. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | LITERATURE REF: DOI 10.7759/CUREUS.70840 A2: AVERAGE AGE A3A: MAJORITY SEX. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
NANOCROSS 014 | K082854 | LITERATURE WAS REVIEWED REGARDING 'DIRECTIONAL ATHERECTOMY FOR CRITICAL LOWER LIMB ISCHEMIA: A RETROSPECTIVE OBSERVATIONAL STUDY. A TOTAL OF 42 ATHERECTOMIES WERE PERFORMED IN 41 PATIENTS. A TOTAL OF 42 DIRECTIONAL ATHERECTOMY PROCEDURES WITH TURBOHAWK WERE PERFORMED IN 41 PATIENTS. NITREX OR SPIDER FX GUIDEWIRES WERE USED AS PART OF THE PROCEDURE, ALONG WITH OTHER MANUFACTURERS DEVICES. THE DISTAL EMBOLIC PROTECTION FILTER SPIDERFX WAS USED IN 57% OF CASES. ADJUVANT PROCEDURES, SUCH AS TRANSLUMINAL BALLOON ANGIOPLASTY, WERE PERFORMED IN THE PRESENCE OF RESIDUAL STENOSES GREATER THAN 30% AFTER THE PASSAGE OF THE ATHERECTOMY WITH A CHOICE BETWEEN DRUG-ELUTING BALLOONS (NON-MEDTRONIC) OR CONVENTIONAL BALLOONS (NON MEDTRONIC, EVERCROSS OR NANOCROSS) BEING DETERMINED BASED ON PATIENT COMORBIDITIES, LESION CHARACTERISTICS, AND THE AVAILABILITY OF DRUG-ELUTING BALLOONS IN THE INSTITUTION. TECHNICAL SUCCESS WAS ACHIEVED IN 78.6% OF CASES AND PROCEDURAL SUCCESS IN 97.6%. THE FAILURE WAS OBSERVED IN ONE PATIENT WITH A 90% STENOSIS OF THE SUPERFICIAL FEMORAL ARTERY, WHICH, AFTER DIRECTIONAL ATHERECTOMY AND DRUG-ELUTING BALLOON ANGIOPLASTY PERSISTED WITH A 30% STENOSIS. NO FURTHER INTERVENTION WAS PERFORMED, AND LIMB SALVAGE WAS ACHIEVED AFTER SIX MONTHS OF FOLLOW-UP. THE NEED FOR ADJUNCTIVE CONVENTIONAL BALLOON ANGIOPLASTY WAS 21.4%, WHILE DRUG-COATED BALLOON ANGIOPLASTIES WERE PERFORMED IN 50% OF LIMB. PRE-DILATION ANGIOPLASTY WAS REQUIRED IN 42.9% OF PROCEDURES. THERE WERE NINE DEATHS (21%) DURING THE STUDY, WHICH INCLUDED ONE CASE OF SOFT TISSUE SEPSIS IN THE LEG WHERE THE ATHERECTOMY WAS PERFORMED, WITHIN THE FIRST 30 DAYS (N = 1). DEATHS OCCURRING BETWEEN 31 AND 90 DAYS WERE DUE TO DECOMPENSATED HEART FAILURE (N = 1), SEVERE TRAUMATIC BRAIN INJURY ASSOCIATED WITH STATUS EPILEPTICUS (N = 1), SOFT TISSUE SEPSIS DUE TO PROGRESSION OF ARTERIAL DISEASE WITH A LACK OF ACCEPTANCE OF AMPUTATION UPON READMISSION (N = 1), SACRAL AND TROCHANTERIC PRESSURE SORE INFECTION (N = 1), MULTI-LOBAR KPC PNEUMONIA (N = 1), AND URINARY SEPSIS (N = 1). BEYOND 90 DAYS AND UNTIL THE END OF THE FOLLOW-UP, DEATHS WERE DUE TO A STROKE (N = 1) AND A PATIENT WITH A HIP FRACTURE AND END-STAGE RENAL FAILURE (N = 1). COMPLICATIONS OCCURRED IN 14.3% OF ALL PROCEDURES AND ADVERSE EVENTS INCLUDED DISTAL EMBOLISM IN 4.8%, DISSECTION N IN 4.8%, ARTERIOVENOUS FISTULAS OCCURRED IN 4.8% OF THE INTERVENED LIMBS, ONE CASE OF SUPERFICIAL FEMORAL ARTERY PERFORATION (2.4%) AND PUNCTURE SITE COMPLICATIONS WITH AN INFECTED INGUINAL HEMATOMA. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | LITERATURE REF: DOI 10.7759/CUREUS.70840. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
SPIDER FX | K111010 | LITERATURE WAS REVIEWED REGARDING 'DIRECTIONAL ATHERECTOMY FOR CRITICAL LOWER LIMB ISCHEMIA: A RETROSPECTIVE OBSERVATIONAL STUDY. A TOTAL OF 42 ATHERECTOMIES WERE PERFORMED IN 41 PATIENTS. A TOTAL OF 42 DIRECTIONAL ATHERECTOMY PROCEDURES WITH TURBOHAWK WERE PERFORMED IN 41 PATIENTS. NITREX OR SPIDER FX GUIDEWIRES WERE USED AS PART OF THE PROCEDURE, ALONG WITH OTHER MANUFACTURERS DEVICES. THE DISTAL EMBOLIC PROTECTION FILTER SPIDERFX WAS USED IN 57% OF CASES. ADJUVANT PROCEDURES, SUCH AS TRANSLUMINAL BALLOON ANGIOPLASTY, WERE PERFORMED IN THE PRESENCE OF RESIDUAL STENOSES GREATER THAN 30% AFTER THE PASSAGE OF THE ATHERECTOMY WITH A CHOICE BETWEEN DRUG-ELUTING BALLOONS (NON-MEDTRONIC) OR CONVENTIONAL BALLOONS (NON MEDTRONIC, EVERCROSS OR NANOCROSS) BEING DETERMINED BASED ON PATIENT COMORBIDITIES, LESION CHARACTERISTICS, AND THE AVAILABILITY OF DRUG-ELUTING BALLOONS IN THE INSTITUTION. TECHNICAL SUCCESS WAS ACHIEVED IN 78.6% OF CASES AND PROCEDURAL SUCCESS IN 97.6%. THE FAILURE WAS OBSERVED IN ONE PATIENT WITH A 90% STENOSIS OF THE SUPERFICIAL FEMORAL ARTERY, WHICH, AFTER DIRECTIONAL ATHERECTOMY AND DRUG-ELUTING BALLOON ANGIOPLASTY PERSISTED WITH A 30% STENOSIS. NO FURTHER INTERVENTION WAS PERFORMED, AND LIMB SALVAGE WAS ACHIEVED AFTER SIX MONTHS OF FOLLOW-UP. THE NEED FOR ADJUNCTIVE CONVENTIONAL BALLOON ANGIOPLASTY WAS 21.4%, WHILE DRUG-COATED BALLOON ANGIOPLASTIES WERE PERFORMED IN 50% OF LIMB. PRE-DILATION ANGIOPLASTY WAS REQUIRED IN 42.9% OF PROCEDURES. THERE WERE NINE DEATHS (21%) DURING THE STUDY, WHICH INCLUDED ONE CASE OF SOFT TISSUE SEPSIS IN THE LEG WHERE THE ATHERECTOMY WAS PERFORMED, WITHIN THE FIRST 30 DAYS (N = 1). DEATHS OCCURRING BETWEEN 31 AND 90 DAYS WERE DUE TO DECOMPENSATED HEART FAILURE (N = 1), SEVERE TRAUMATIC BRAIN INJURY ASSOCIATED WITH STATUS EPILEPTICUS (N = 1), SOFT TISSUE SEPSIS DUE TO PROGRESSION OF ARTERIAL DISEASE WITH A LACK OF ACCEPTANCE OF AMPUTATION UPON READMISSION (N = 1), SACRAL AND TROCHANTERIC PRESSURE SORE INFECTION (N = 1), MULTI-LOBAR KPC PNEUMONIA (N = 1), AND URINARY SEPSIS (N = 1). BEYOND 90 DAYS AND UNTIL THE END OF THE FOLLOW-UP, DEATHS WERE DUE TO A STROKE (N = 1) AND A PATIENT WITH A HIP FRACTURE AND END-STAGE RENAL FAILURE (N = 1). COMPLICATIONS OCCURRED IN 14.3% OF ALL PROCEDURES AND ADVERSE EVENTS INCLUDED DISTAL EMBOLISM IN 4.8%, DISSECTION N IN 4.8%, ARTERIOVENOUS FISTULAS OCCURRED IN 4.8% OF THE INTERVENED LIMBS, ONE CASE OF SUPERFICIAL FEMORAL ARTERY PERFORATION (2.4%) AND PUNCTURE SITE COMPLICATIONS WITH AN INFECTED INGUINAL HEMATOMA. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | LITERATURE REF: DOI 10.7759/CUREUS.70840. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
EVERCROSS 035 | K082579 | LITERATURE WAS REVIEWED REGARDING 'DIRECTIONAL ATHERECTOMY FOR CRITICAL LOWER LIMB ISCHEMIA: A RETROSPECTIVE OBSERVATIONAL STUDY. A TOTAL OF 42 ATHERECTOMIES WERE PERFORMED IN 41 PATIENTS. A TOTAL OF 42 DIRECTIONAL ATHERECTOMY PROCEDURES WITH TURBOHAWK WERE PERFORMED IN 41 PATIENTS. NITREX OR SPIDER FX GUIDEWIRES WERE USED AS PART OF THE PROCEDURE, ALONG WITH OTHER MANUFACTURERS DEVICES. THE DISTAL EMBOLIC PROTECTION FILTER SPIDERFX WAS USED IN 57% OF CASES. ADJUVANT PROCEDURES, SUCH AS TRANSLUMINAL BALLOON ANGIOPLASTY, WERE PERFORMED IN THE PRESENCE OF RESIDUAL STENOSES GREATER THAN 30% AFTER THE PASSAGE OF THE ATHERECTOMY WITH A CHOICE BETWEEN DRUG-ELUTING BALLOONS (NON-MEDTRONIC) OR CONVENTIONAL BALLOONS (NON MEDTRONIC, EVERCROSS OR NANOCROSS) BEING DETERMINED BASED ON PATIENT COMORBIDITIES, LESION CHARACTERISTICS, AND THE AVAILABILITY OF DRUG-ELUTING BALLOONS IN THE INSTITUTION. TECHNICAL SUCCESS WAS ACHIEVED IN 78.6% OF CASES AND PROCEDURAL SUCCESS IN 97.6%. THE FAILURE WAS OBSERVED IN ONE PATIENT WITH A 90% STENOSIS OF THE SUPERFICIAL FEMORAL ARTERY, WHICH, AFTER DIRECTIONAL ATHERECTOMY AND DRUG-ELUTING BALLOON ANGIOPLASTY PERSISTED WITH A 30% STENOSIS. NO FURTHER INTERVENTION WAS PERFORMED, AND LIMB SALVAGE WAS ACHIEVED AFTER SIX MONTHS OF FOLLOW-UP. THE NEED FOR ADJUNCTIVE CONVENTIONAL BALLOON ANGIOPLASTY WAS 21.4%, WHILE DRUG-COATED BALLOON ANGIOPLASTIES WERE PERFORMED IN 50% OF LIMB. PRE-DILATION ANGIOPLASTY WAS REQUIRED IN 42.9% OF PROCEDURES. THERE WERE NINE DEATHS (21%) DURING THE STUDY, WHICH INCLUDED ONE CASE OF SOFT TISSUE SEPSIS IN THE LEG WHERE THE ATHERECTOMY WAS PERFORMED, WITHIN THE FIRST 30 DAYS (N = 1). DEATHS OCCURRING BETWEEN 31 AND 90 DAYS WERE DUE TO DECOMPENSATED HEART FAILURE (N = 1), SEVERE TRAUMATIC BRAIN INJURY ASSOCIATED WITH STATUS EPILEPTICUS (N = 1), SOFT TISSUE SEPSIS DUE TO PROGRESSION OF ARTERIAL DISEASE WITH A LACK OF ACCEPTANCE OF AMPUTATION UPON READMISSION (N = 1), SACRAL AND TROCHANTERIC PRESSURE SORE INFECTION (N = 1), MULTI-LOBAR KPC PNEUMONIA (N = 1), AND URINARY SEPSIS (N = 1). BEYOND 90 DAYS AND UNTIL THE END OF THE FOLLOW-UP, DEATHS WERE DUE TO A STROKE (N = 1) AND A PATIENT WITH A HIP FRACTURE AND END-STAGE RENAL FAILURE (N = 1). COMPLICATIONS OCCURRED IN 14.3% OF ALL PROCEDURES AND ADVERSE EVENTS INCLUDED DISTAL EMBOLISM IN 4.8%, DISSECTION N IN 4.8%, ARTERIOVENOUS FISTULAS OCCURRED IN 4.8% OF THE INTERVENED LIMBS, ONE CASE OF SUPERFICIAL FEMORAL ARTERY PERFORATION (2.4%) AND PUNCTURE SITE COMPLICATIONS WITH AN INFECTED INGUINAL HEMATOMA. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | LITERATURE REF: DOI 10.7759/CUREUS.70840 A2:AVERAGE AGE A3A: MAJORITY SEX MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
EMBOTRAP | K193063 | IT WAS REPORTED, VIA A PERSONAL INTERACTION, THAT AN EMBOTRAP III 5 MM X 37 MM (ET309537/ 23L089AV) WAS USED FOR A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT AN OCCLUSION FROM THE M1 TO M2 SEGMENTS OF THE MIDDLE CEREBRAL ARTERY (MCA). THERE WAS CALCIFICATION AND STENOSIS AREA AT THE OCCLUDED SITE. DURING THE PROCEDURE, THE USER REPORTED WITHDRAWAL DIFFICULTY OF THE EMBOTRAP III DEVICE FROM THE VESSEL AND INTO THE MICROCATHETER. DURING THE SECOND PASS, THE STRUT OF THE DEVICE BECAME STUCK WITHIN A STENOTIC AREA. THE DEVICE WAS WITHDRAWN WHILE IN DEPLOYED STATE, AND A HEMORRHAGE OCCURRED. THE PARENT ARTERY WAS OCCLUDED BY COILING FOR HEMOSTASIS. NO ADDITIONAL TREATMENT WAS PERFORMED. AFTER SEVERAL DAYS OF MONITORING THE PATIENT, THE PATIENT ULTIMATELY EXPIRED. THE PHYSICIAN JUDGED THAT THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE HEALTH HAZARD AND THE PRODUCT. THE PHYSICIAN FURTHER COMMENTED, ¿THE PATIENT DIED DUE TO BLEEDING WHEN THE EMBOTRAP III WAS PULLED. HOWEVER, THE PRODUCT WAS NOT THE CAUSE.¿ THE EVENT WAS REPORTED AS SUCH, ¿THE PROCEDURE WAS A MECHANICAL THROMBECTOMY FOR AN ACUTE ISCHEMIC STROKE OF M1D AND M2 MIDDLE CEREBRAL ARTERY. THE OCCLUDED LESION WAS BETWEEN M1D AND M2 AND THE EMBOTRAP III WAS USED FOR THE PROCEDURE BY COMBINED TECHNIQUE. THE PROCEDURE WAS PERFORMED WITHOUT ANY ISSUES UNTIL THE STENT DEPLOYMENT. THERE WAS CALCIFICATION AND STENOSIS AREA AT THE OCCLUDED SITE. RETRIEVAL RESISTANCE OCCURRED WHEN RETRACTING THE EMBOTRAP III DURING THE FIRST PASS. RETRIEVAL RESISTANCE AGAIN OCCURRED DURING THE SECOND PASS, SO THE MICROCATHETER WAS ATTEMPTED TO BE ADVANCED OVER THE EMBOTRAP III, BUT THE MICROCATHETER COULD NOT BE ADVANCED BECAUSE THE STENT STRUT WAS STUCK IN THE STENOTIC AREA. THE EMBOTRAP WAS PULLED AS IT WAS, AND HEMORRHAGE OCCURRED. THE PARENT ARTERY WAS OCCLUDED BY COILING FOR HEMOSTASIS AS THE PRIMARY TREATMENT. NO ADDITIONAL TREATMENT WAS PERFORMED. POST-PROCEDURE CHANGES IN THE PATIENT: THE PATIENT EXPIRED TWO DAYS AGO FOLLOWING SEVERAL DAYS OF PATIENT MONITORING. THE PHYSICIAN¿S ASSESSMENT ON HEALTH HAZARD: SERIOUS. THE PHYSICIAN JUDGED THAT THERE WAS A CAUSAL RELATIONSHIP BETWEEN THE HEALTH HAZARD AND THE PRODUCT. THE PHYSICIAN¿S COMMENT REGARDING THE RELATIONSHIP BETWEEN THE EVENT AND THE PRODUCT: THE PATIENT DIED DUE TO BLEEDING WHEN THE EMBOTRAP III WAS PULLED. HOWEVER, THE PRODUCT WAS NOT THE CAUSE. THE DEVICES WERE USED ACCORDING TO IFU. CONTINUOUS FLUSH WAS DONE. THE LESION WAS MIDDLE CEREBRAL ARTERY. OTHER CONCOMITANT DEVICE WAS TRAK MICROCATHETER (STRYKER).¿ NO FURTHER INFORMATION WAS MADE AVAILABLE AT THE TIME OF THIS REVIEW. | PRODUCT COMPLAINT # (B)(4). UPDATED SECTIONS ON THIS MED WATCH REPORT: A2, A5, A6, B4, G3, G6, H2, AND H11. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT.
ADMIRAL XTREME OTW | K173515 | LITERATURE WAS REVIEWED TITLED 'NEWLY DESIGNED, SELF-EXPANDING LARGE-BORE NITINOL STENTS FOR SYMPTOMATIC CENTRAL VENOUS STENOSIS: TECHNICAL AND LONGTERM CLINICAL OUTCOME'. MULTIPLE MANUFACTURES DEVICES WERE USED IN THE PROCEDURE. MEDTRONIC DEVICES INCLUDED ADMIRAL EXTREME BALLOON USED FOR PRE-DILATION OF STENOSIS PRIOR TO NON-MEDTRONIC STENT IMPLANTATION. THERE WAS ONE (2%) PROCEDURE-RELATED COMPLICATION IN A PATIENT WITH A MALIGNANT SVC STENOSIS IN WHOM SMALL PULMONARY EMBOLI WERE FOUND ON FOLLOW-UP CT. THIS PATIENT DEVELOPED RIGHT HEART FAILURE AND WAS SUCCESSFULLY MANAGED WITH CONSERVATIVE MEASURES. IN ANOTHER THREE PATIENTS (5%) WITH A MALIGNANT SVC STENOSIS, COMPLICATIONS POTENTIALLY RELATED TO THE PROCEDURE WERE ENCOUNTERED, INCLUDING SINUS BRADYCARDIA 2 DAYS AFTER THE PROCEDURE (N = 1) AND SEPSIS (N = 2). THERE WAS ONE PATIENT WITH AN EARLY SYMPTOMATIC STENT OCCLUSION (1.9%) AND 3 PATIENTS WITH A LATE SYMPTOMATIC RESTENOSIS (5.6%). THE EARLY, THROMBOTIC STENT OCCLUSION OCCURRED 10 DAYS AFTER THE PROCEDURE IN A NSCLC PATIENT. THE LATE SYMPTOMATIC RESTENOSIS OCCURRED 105, 249 AND 420 DAYS AFTER THE INDEX STENT PLACEMENT FOR MALIGNANT (N = 2) OR BENIGN (N = 1) SVC OCCLUSIVE DISEASE. THE OVERALL PRIMARY PATENCY RATE, PRIMARY-ASSISTED PATENCY RATE AND SECONDARY PATENCY RATE WERE 92.6%, 98.2% AND 100%, RESPECTIVELY. ONE OF THE 2 PATIENTS WITH POSTPROCEDURAL SEPSIS DIED 1 DAY AFTER THE STENT PROCEDURE. | LITERATURE REF: DOI.ORG/10.1007/S00270-024-03820-Y. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
WECK VISTA | K152663 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RADICAL PROSTATECTOMY (WITHOUT LYMPHADENECTOMY), A SMALL BOWEL INJURY OCCURRED WHILE PORTS WERE BEING PLACED DUE TO BAD ADHESIONS. THE INJURY OCCURRED WHEN PLACING A DA VINCI 8MM, SHORT, BLADELESS OBTURATOR. THE INJURY WENT UNDIAGNOSED AND RESULTED IN POST-OPERATIVE COMPLICATIONS. ON POST-OPERATIVE DAYS 1 AND 2, FOLLOWING THE ONSET OF SEPSIS, THE PATIENT UNDERWENT TWO ADDITIONAL, NON-ROBOTIC SURGERIES TO RESOLVE THE BOWEL INJURY; IT IS UNKNOWN WHAT REPAIRS WERE PERFORMED DURING THESE PROCEDURES. THE PATIENT EXPIRED ON POST-OPERATIVE DAY 3. IT WAS REPORTED THAT THE DA VINCI SYSTEM AND INSTRUMENTS USED DURING THE PROCEDURE WORKED FINE WITHOUT ANY ISSUES. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE TWO SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, 0 DEGREE ENDOSCOPE PLUS, TWO LARGE NEEDLE DRIVERS, MARYLAND BIPOLAR FORCEPS, PROGRASP FORCEPS, AND MONOPOLAR CURVED SCISSORS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN UNRECOGNIZED TROCAR INJURY TO THE BOWEL WHICH OCCURRED DURING A ROBOTIC PROSTATECTOMY. THIS INJURY WAS NOT DISCOVERED AT THE TIME OF THE INITIAL SURGERY AND THE PATIENT BECAME SEPTIC REQUIRING ADDITIONAL PROCEDURES. THEY ULTIMATELY DIED AS A RESULT OF THIS COMPLICATION. HOW THE INJURY OCCURRED IS NOT REPORTED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
BIPAP SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER BECAME DUE TO A USER OF BIPAP AUTO BIFLEX. DAUGHTER CALLED IN STATING PATIENT PASSED AWAY. SHE WANTS TO STOP COMMUNICATION. THE DEVICE IS NO LONGER IN POSSESSION OF PATIENT¿S FAMILY AND WILL NOT BE RETURNING FOR EVALUATION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
REMSTAR SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A LETTER RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP W/HUMID CELL/BT | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A LETTER RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERATOR | K800742 | AS REPORTED: "ON (B)(6) 2024, WE [(B)(4), DISTRIBUTOR] RECEIVED INFORMATION ABOUT THE DEATH OF A PATIENT WHO WAS BEING SUPPLIED WITH OXYGEN BY US [(B)(4), DISTRIBUTOR]. THE REASON FOR THE CAUSE OF DEATH: APARTMENT FIRE DUE TO SMOKING IN BED. OUR [(B)(4), DISTRIBUTOR] INITIAL INFORMATION (FORWARDED TO THE POLICE BY OUR SERVICE PARTNER) WAS THAT OXYGEN AND THUS THE LIBERATOR [BU, BASE UNIT] WERE SUPPOSEDLY NOT INVOLVED / OR WERE NOT IN USE. WE ALSO RECEIVED PHOTOS OF THE LIBERATOR. THE CONDITION OF THE BU IN THE PHOTOS SUGGESTS THAT THE BU WAS NOT DIRECTLY INVOLVED IN THE FIRE. 29TH OF NOVEMBER 2024: AFTER CONSULTING THE POLICE, WE RECEIVED FURTHER INFORMATION THAT THE USE OF OXYGEN CAN NEITHER BE CONFIRMED NOR RULED OUT. THIS MEANS THAT WE CANNOT RULE OUT A DIRECT CONNECTION BETWEEN OXYGEN THERAPY WITH THE LIBERATOR AND THE PATIENT'S DEATH. THE PATIENT ALSO HAD A STROLLER (3740545292C). ACCORDING TO INFORMATION FROM THE POLICE, IT WAS IN A ROOM THAT WAS FILLED WITH SMOKE, BUT NOT WHERE THE FIRE ACTUALLY BROKE OUT. WE THEREFORE ASSUME THAT THE STROLLER WAS NOT INVOLVED IN THE INCIDENT. WE WILL CONFISCATE BOTH DEVICES AND SEND THEM BACK TO OUR TECHNICAL CENTER. WE WILL ALSO REQUEST ACCESS TO THE FILE FROM THE POLICE." | PURSUANT TO TITLE 21 - FOOD AND DRUGS, CHAPTER I - FOOD AND DRUG ADMINISTRATION DEPARTMENT OF HEALTH AND HUMAN SERVICES, SUBCHAPTER H -0 MEDICAL DEVICE, PART 803 - MEDICAL DEVICE REPORTING, SUBPART A - GENERAL PROVISIONS, SECTION 803.16, NEITHER THIS REPORT NOR ANY INFORMATION SUBMITTED HEREIN CONSTITUTES AN ADMISSION BY CAIRE INC. THAT THE DEVICE STATED IN THIS REPORT, CAIRE INC., OR CAIRE INC.'S EMPLOYEES, CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT STATED HEREIN. SECTION D4 PRIMARY UNIQUE DEVICE IDENTIFIER (UDIS) # IS INTENTIONALLY BLANK. THE DEVICE WAS INTENDED FOR EXPORT ONLY. CAIRE DETERMINED THAT THE EVENT MEETS REPORTING CRITERIA OUT OF AN ABUNDANCE OF CAUTION AND DUE TO THE DEATH OF THE PATIENT INVOLVED IN THIS INCIDENT. CAIRE CONCURS WITH THE INITIAL REPORTER (B)(4) THAT CAIRE'S EQUIPMENT DOES NOT APPEAR TO HAVE BEEN INVOLVED IN THE FIRE CAUSED BY PATIENT SMOKING, BUT CANNOT RULE OUT THE CONNECTION BASED ON AVAILABLE EVIDENCE. THE INFORMATION PROVIDED REPRESENTS THE BEST INFORMATION AVAILABLE TO CAIRE. THE AVAILABLE EVIDENCE SUPPORTS PATIENT SMOKING A CIGARETTE WHILE USING THE DEVICE. NO SMOKING WARNINGS ON THE EQUIPMENT AND IN THE IFU WERE REVIEWED AND DEEMED ADEQUATE IN THE COURSE OF THE INVESTIGATION. "NO SMOKING" SYMBOLS (REG# P002) ARE AFFIXED TO THE FACE OF THE UNITS. THROUGHOUT THE UNIT IFU, WARNINGS STATE NOT TO ALLOW SMOKING OR OPEN FLAME NEAR THE DEVICE AND TO KEEP ALL FLAMMABLE MATERIALS AWAY FROM THE EQUIPMENT. ADDITIONAL WARNINGS STATE THAT SMOKING WHILE WEARING AN OXYGEN CANNULA CAN CAUSE FACIAL BURNS AND POSSIBLY RESULT IN DEATH. WARNINGS ALSO ADDRESS THE RISK OF OXYGEN SATURATION OF FABRICS WHEN REMOVING A CANNULA WHILE THE UNIT IS TURNED ON, LEADING TO INCREASED RISK OF FLASH FIRE. THIS CASE REPRESENTS PATIENT MISUSE OF THE DEVICE AND FAILURE TO HEED THE WARNINGS COMMUNICATED TO THE USER. LIBERATOR RISK ASSESSMENT SLOX-RA-001 REV T WAS REVIEWED AND DETERMINED TO BE ADEQUATE WITHOUT REVISION. THE USABILITY ASSESSMENT ADEQUATELY ADDRESSES USABILITY ERROR "PATIENT SMOKES OR USES AN OPEN FLAME NEAR THE DEVICE." MITIGATION INCLUDES VARIOUS WARNINGS ON THE EQUIPMENT AND IN THE INSTRUCTIONS FOR USE. "NO SMOKING" SYMBOLS ARE AFFIXED TO THE FACE OF THE UNITS. THROUGHOUT THE UNIT IFU, WARNINGS STATE NOT TO ALLOW SMOKING OR OPEN FLAME NEAR THE DEVICE AND TO KEEP ALL FLAMMABLE MATERIALS AWAY FROM THE EQUIPMENT. ADDITIONAL WARNINGS STATE THAT SMOKING WHILE WEARING AN OXYGEN CANNULA CAN CAUSE FACIAL BURNS AND POSSIBLY RESULT IN DEATH.
DREAMSTATION | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. EVALUATION OF THE DEVICE INDICATED THE DEVICE PASSED AND ALL TESTING WITHOUT ANY FAULTS FOUND. BASED ON THE SERVICE EVALUATION OF THE DEVICE, THERE IS NO EVIDENCE INDICATING A REPORTABLE MALFUNCTION OCCURRED OR THAT THERAPY WAS AFFECTED. THE INVESTIGATION FINDINGS DID NOT UNCOVER ANY DETAILS THAT WOULD CHANGE OR MODIFY THE RISK OF THE DEVICE.
EPSTAR FIXED ELECTROPHYSIOLOGY CATHETER | K183649 | IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A PREMATURE VENTRICULAR CONTRACTION (PVC) ABLATION PROCEDURE, AN EPSTAR CATHETER WAS SELECTED FOR USE. THE PATIENT WAS HAVING INFREQUENT PVC. ISUPREL WAS GIVEN TO INCREASE THE FREQUENCY TO MAP WITH A NON-BOSTON SCIENTIFIC MAPPING SYSTEM. THE SYSTEM WAS MAPPED WITH NON-BOSTON SCIENTIFIC CATHETERS. THE PHYSICIAN WAS NOT HAPPY WITH THE ENDOCARDIAL LOCATIONS IN THE LEFT VENTRICULAR OUTFLOW TRACT (LVOT AND RIGHT VENTRICULAR OUTFLOW TRACT (RVOT), HENCE DECIDED TO USE THE CATHETER THROUGH A NON-BOSTON SCIENTIFIC LAMP SHEATH TO CANULATE THE CORONARY SINUS (CS) WITH SOME CHALLENGING ANATOMY. THE PHYSICIAN THEN PROCEEDED TO EXCHANGE THE CATHETER FOR A 4FR SHEATH. THE PHYSICIAN THEN PUT THE 2FR EPSTAR CATHETER THROUGH THE SHEATH TO MAP THE SEPTAL BRANCHES. THE PHYSICIAN CANNULATED SEVERAL BRANCHES BUT NONE OF INTEREST TO THE PVC FOCUS. THE EPSTAR WAS REMOVED. THE PHYSICIAN PROCEEDED TO TRY AND MANIPULATE THE LAMP SHEATH TO AIM AT DIFFERENT BRANCHES OF THE ANTERIOR INTER VENTRICULAR VEINS. THE LAMP SHEATH CAME OUT. THE PHYSICIAN THEN REMOVED THE 4FR SHEATH. THEY PROCEEDED TO LOOK AT THE RVOT FOR PVCS AGAIN TO SEE IF ENDOCARDIAL ABLATION IS POSSIBLE FROM THE RVOT. THE PHYSICIAN THEN PROCEEDED TO PUT THE NON-BOSTON SCIENTIFIC CATHETER THROUGH THE LAMP SHEATH WITHOUT FLUOROSCOPY AND WAS GOING TO ATTEMPT TO CANULATE THE CS USING THE MAPPING SYSTEM. THE PATIENT WAS HEPARINIZED DUE TO MAPPING THE LEFT VENTRICULAR OUTFLOW TRACT. THE PHYSICAL STRUGGLED TO CANULATE THE CS SIMILAR TO THE FIRST TIME WITH SOME MOVEMENT ON THE CATHETER TOWARD THE RIGHT ATRIAL APPENDAGE. THE PATIENTS HEART RATE DROPPED SUDDENLY, AND PRESSURES DROPPED. THE PHYSICIAN SUSPECTS A CORONARY SINUS DISSECTION. AN EFFUSION WAS SEEN ON INTRACARDIAC ECHOCARDIOGRAPHY (ICE). HEPARIN REVERSAL AND PERICARDIOCENTESIS WAS CALLED FOR IMMEDIATELY. A THORACOTOMY WAS PERFORMED, AND A PATCH WAS PLACED ON THE RIGHT ATRIAL APPENDAGE AT THE SITE OF THE PERFORATION. THE PATIENT WAS STABILIZED. THE PHYSICIAN BELIEVES THAT THE PERFORATION OCCURRED WHEN THE NON-BOSTON SCIENTIFIC CATHETER WAS INTRODUCED BLINDLY WITH THE MAPPING SYSTEM AS HE DID NOT KNOW THE LOCATION OF THE SHEATH AT THAT TIME. THE COMPLICATION OCCURRED WHILE NO BOSTON SCIENTIFIC EQUIPMENT WAS IN THE BODY. HE DOES NOT BELIEVE THE EPSTAR CAUSED ANY ISSUES DURING THE PROCEDURE. THE PHYSICIAN BELIEVES THAT HIGH DOSE ADRENERGIC AGONISTS MASKED EARLY SIGNS OF AN EFFUSION AND LEAD TO A LATE RESPONSE AND A NECESSARY PATCH OF THE RIGHT ATRIAL APPENDAGE. THE PATIENT PASSED AWAY ON (B)(6) 2024. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A LETTER RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS FOUND DECEASED BY A SECURITY GUARD IN THE PATIENT'S APARTMENT COMPLEX. THE GUARD HEARD THE CYCLER ALARMING AND CONDUCTED A WELFARE CHECK ON THE PATIENT. THE PATIENT WAS ALREADY DECEASED WHEN FOUND. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER. THE CAUSE OF DEATH IS UNKNOWN. IT IS UNKNOWN IF AN AUTOPSY WILL BE PERFORMED. THE PDRN STATED THAT THE PATIENT¿S DAUGHTER HAS NOT ANSWERED THEIR INQUIRIES. NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER. IT WAS REPORTED THAT PRIOR TO THE PATIENT¿S DEATH, THE PATIENT WAS FEELING WEAK AND FATIGUED. THE PATIENT HAD AN EVENT OF COVID AND PERITONITIS IN OCTOBER AND NEVER FULLY RECOVERED FROM FEELING SICK. NO ADDITIONAL INFORMATION WAS AVAILABLE. | THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. THE TEACH PUMPS TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. THERE WERE VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY AND THE BOTTOM COVER. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S). UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, HYPERSENSITIVITY, LUNG DISEASE, DEATH, OTHER AND LOST SENSE OF TASTE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
MICRA | K132030 | IT WAS NOTED THAT THE LEADLESS IPG WAS REPROGRAMMED POST HYPOTENSION AND SHOCK. IT WAS FURTHER NOTED THAT THE PATIENT EXPERIENCED PACING-INDUCED CARDIOMYOPATHY (PICM) DUE TO VVI PACING AND CONGESTIVE HEART FAILURE, THE CAUSE OF DEATH WAS NOTED TO BE SICK SINUS SYNDROME. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT DIED IN (B)(6) 2024. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER WHERE IT WAS INSPECTED. EVALUATION RESULTS DETERMINED NO ERRORS WERE FOUND. THE TUBING, MANUAL, AND FILTERS WERE REPLACED. ADDITIONALLY, THE DEVICE WAS CLEANED AND PASSED ALL TESTS.
AVALON FETAL MONITOR FM50 | K140535 | IT WAS REPORTED THAT ON (B)(6) 2024, THE MATERNAL AND FETAL HEART RATES WERE TRACING TOGETHER. THE MOTHER WAS AT 38 WEEKS 6 DAYS GESTATION AT THE TIME OF THE EVENT, AND THE NEWBORN PASSED AWAY SECONDARY TO HYPOXIA. | THE FOLLOWING FUNCTIONAL TESTS WERE PERFORMED: THE AVALON FM50 WAS RETURNED FOR BENCH REPAIR FOR EVALUATION. THE UNIT WAS POWERED ON, AND IT BOOTED UP WITH NORMAL DISPLAY GRAPHICS. THE BENCH COULD NOT DUPLICATE THE FAILURE AS DESCRIBED BY THE CUSTOMER. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION, AND A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED THE STRIPS AND PROVIDED DATA FROM THE EVENT, AND THIS ISSUE SHOWS INSTANCES OF COINCIDENCE AS THE MEASURED VALUES FOLLOW A SIMILAR TREND. WHEN THE CROSS-CHANNEL VERIFICATION DETECTS THAT THE SIGNAL OF THE MATERNAL HEART RATE COINCIDES WITH THE FETAL HEART RATE (FHR), THE COINCIDENCE INOP IS ISSUED. SINCE THE PROVIDED STRIPS ARE NOT FROM THE AVALON FETAL MONITOR, WE CAN ONLY ASSUME THAT IT WAS ISSUED CORRECTLY BY THE DEVICE. BELOW ARE EXCERPTS FROM THE AVALON FETAL MONITOR INSTRUCTIONS FOR USE INSTRUCTIONS FOR USE (IFU) FM20/30, FM40/50, AVALON CL 453564898421 RELEASE L.3 REV. A.40.00 PAGES 165-167: COINCIDENCE OF MATERNAL PULSE AND FHR. WHEN THE MATERNAL PULSE AND FHR ARE BEING MONITORED, AND THE MEASURED VALUES ARE VERY SIMILAR OR THE SAME, THE COINCIDENCE QUESTION MARK IS DISPLAYED ON THE MONITOR¿S SCREEN ABOVE BOTH OF THE CORRESPONDING NUMERICS (IN THIS CASE MATERNAL PULSE AND FHR). OFTEN THE SIGNAL LOSS OR COINCIDENCE HAPPENS BECAUSE THE FETAL OR MATERNAL MOVEMENT DISPLACED THE ULTRASOUND TRANSDUCER, AND A REPOSITIONING OF THE TRANSDUCER IS NECESSARY. RECOMMENDED ACTIONS FOR COINCIDENCE INOP. 1. CONFIRM FETAL LIFE BY PALPATION OF FETAL MOVEMENT OR AUSCULTATION OF FETAL HEART SOUNDS USING A FETOSCOPE, STETHOSCOPE, OR PINARD STETHOSCOPE. 2. MANUAL DETERMINATION OF THE MATERNAL PULSE AND COMPARISON WITH THE FETAL HEART RATE SOUND SIGNALS FROM THE LOUDSPEAKER. 3. REPOSITION THE TRANSDUCER, OR ENSURE THAT THE FETAL SCALP ELECTRODE IS PLACED CORRECTLY, UNTIL YOU RECEIVE A CLEAR SIGNAL AND THE MONITOR IS NO LONGER ISSUING THE COINCIDENCE INOP. 4. IN CASE OF DIFFICULTIES DERIVING A STABLE MATERNAL PULSE READING USING THE TOCO MP OR CL TOCO+MP TRANSDUCER, USE SPO2 OR THE CL SPO2 POD INSTEAD. IN CASE OF SIMILAR PROBLEMS WITH THE PULSE MEASUREMENT FROM SPO2, USE MECG INSTEAD. REASONS TO SWITCH THE METHOD FOR DERIVING A MATERNAL PULSE OR HEART RATE INCLUDE: MOTION ARTIFACTS, ARRHYTHMIA, AND INDIVIDUAL DIFFERENCES IN PULSE SIGNAL QUALITY ON THE ABDOMINAL SKIN (VIA TOCO MP). 5. IF YOU CANNOT HEAR THE FETAL HEART SOUNDS, AND YOU CANNOT CONFIRM FETAL MOVEMENT BY PALPATION, CONFIRM FETAL LIFE USING OBSTETRIC ULTRASONOGRAPHY. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE FETAL MONITOR WAS WORKING AS INTENDED. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE MONITOR WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEWED BY THE CLINICAL EXPERT AND THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ILET BIONIC PANCREAS | K231485 | ON 11/13/2024 A BETA BIONICS CLINICAL DIABETES SPECIALIST (CDS) WAS MADE AWARE THAT AN ILET USER PASSED AWAY, WITH THE ESTIMATED DATE OF THE INCIDENT BEING (B)(6) 2024. IT IS UNCLEAR WHETHER THE USER WAS WEARING OR USING THE ILET SYSTEM AT THE TIME. THE CLINIC EDUCATORS NOTED THE USER LIKELY EXPERIENCED DIABETIC KETOACIDOSIS (DKA) AT THE TIME OF DEATH BUT LACKED SPECIFIC DETAILS. MULTIPLE FOLLOW-UP ATTEMPTS BY BETA BIONICS CUSTOMER CARE TO GATHER INFORMATION FROM THE USER'S HUSBAND HAVE BEEN UNSUCCESSFUL. THE CDS VISITED THE USER'S CLINIC TO OBTAIN ADDITIONAL DETAILS BUT WAS UNABLE TO ACQUIRE FURTHER INFORMATION. | NO PRODUCT WAS RETURNED FOR EVALUATION. THE ILET LOGS WERE REVIEWED BY BETA BIONICS FAILURE INVESTIGATION DEPARTMENT. NO EVIDENCE OF DEVICE MALFUNCTION WAS FOUND IN THE ENGINEERING LOGS. THE ILET WAS BEHAVING AS INTENDED AND CAN BE SEEN REACTING APPROPRIATELY TO CGM GLUCOSE VALUES. THE ILET WAS APPROPRIATELY TRIGGERING DEVICE AND CGM GLUCOSE ALERTS AS OF (B)(6) 2024 WHICH IS THE LAST AVAILABLE DAY OF DEVICE DATA. NO PRODUCT PERFORMANCE ISSUES WERE IDENTIFIED. IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE COMPLAINT WILL BE REOPENED AND INVESTIGATED ACCORDINGLY. NO ANOMALIES WERE OBSERVED. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. BASED ON CASE NOTES AND DEVICE DATA, THE ILET OPERATED AS INTENDED. WITHOUT DEVICE DATA FROM THE REPORTED EVENT DATE (9/30/24), THE PERFORMANCE OF THE DEVICE CANNOT BE EVALUATED, AND THE CAUSE FOR THE EVENT CANNOT BE DETERMINE. MULTIPLE ATTEMPTS HAVE BEEN MADE TO REACH THE HUSBAND FOR ADDITIONAL INFORMATION WITHOUT SUCCESS.
PRLNE BLU 24IN 7-0 D/A BV-1 | K133356 | IT WAS REPORTED THAT A PATIENT UNDERWENT A TCRAT MIS CABG PROCEDURE AND SUTURE WAS USED. THE SUTURE SNAPPED DURING PROXIMAL ANASTOMOSES WITH THE KNOT PUSHER EVEN WITH A STATIC KNOT PUSHER. CUSTOMER OPENED A NEW SUTURE PACK. SURGEON HAD TO REDO THE ANASTOMOSES WHICH IN RETURN WILL TRAUMATIZE THE SAPHENOUS VEIN GRAFT USED IN CABG. THE PATIENT PASSED AWAY ON (B)(6)2024. ADDITIONAL INFORAMTION WAS REQUESTED. | PRODUCT COMPLAINT # ==>(B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H6 COMPONENT CODE: G07002 NO DEVICE PROBLEM FOUND ADDITIONAL INFORMATION: D9, H3, H6 INVESTIGATION SUMMARY ==> THE PRODUCT WAS RETURNED TO ETHICON FOR EVALUATION. THREE UNOPENED SAMPLES WERE RECEIVED FOR ANALYSIS. PRODUCT CODE 8702H. IN ORDER TO EVALUATE THE CONDITION OF THE RETURNED SAMPLES, THE PACKETS WERE OPENED. THE SWAGE AND ATTACHMENT AREA WERE NOTED AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO ISSUES RELATED TO BREAKAGE SUTURES OR ANOMALIES WERE OBSERVED DURING THE EVALUATION. ALSO, A FUNCTIONAL TEST WAS PERFORMED USING INSTRON EQUIPMENT AND TENSILE FORCE WAS ABOVE THE MINIMUM REQUIREMENTS. AS PART OF THE ETHICON QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. THE DEVICE PERFORMED WITHOUT ANY DIFFICULTIES NOTED. ALTHOUGH NO PRODUCT DEFECT WAS IDENTIFIED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE LABORATORY ANALYSIS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS IS CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF POST-MARKET SURVEILLANCE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DIZZINESS AND/OR HEADACHE, CANCER. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION INDICATING A PATIENT USING A TRILOGY EVO VENTILATOR HAS DIED. BASED ON INFORMATION PROVIDED AND AVAILABLE, THE REPORTED EVENT OF THE PASSING AWAY DUR TO RESPIRATORY FAILURE, DESPITE RESUSCITATION MEASURES AND IT WAS REPORTED THAT THEY SUCTIONED THE PATIENT WITH THE DEVICE IN STANDBY MODE WHILE THE DEVICE WAS BEING USED WITH NON-INVASIVE POSITIVE PRESSURE VENTILATION AND FORGOT TO PRESS THE START BUTTON AFTER THAT. ALTHOUGH THE ALARM WAS OBSERVED ON ANOTHER VITAL SIGNS MONITOR, IT WAS TOO LATE, RESULTING IN THE PATIENT¿S DEATH, LIKELY AN UNINTENDED USE ERROR. IT WAS REPORTED THAT IT IS BELIEVED THAT NO MALFUNCTIONS WERE ON THE DEVICE. IT WAS ALSO LATER REPORTED THAT IT WAS CONFIRMED ON THE VITAL SIGNS MONITOR THAT THE PATIENT EXPERIENCED BRADYCARDIA AND RESUSCITATION MEASURES WERE PERFORMED. HOWEVER, THE PATIENT UNFORTUNATELY PASSED AWAY. DEVICE EVALUATION AND LOG ATTACHMENTS ARE CURRENTLY PENDING. THEREFORE, DEVICE CAUSE OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH WHILE ON THE DEVICE CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE EVIDENCE PRESENTED. FURTHER GOOD FAITH EFFORTS AND INFORMATION GATHERING ARE REQUIRED TO DETERMINE IF THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT¿S EVENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER. THEREFORE, DEVICE FUNCTIONALITY IS NOT ABLE TO BE CONFIRMED AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION HAS BEEN COMPLETED.
PRLNE BLU 24IN 7-0 D/A BV-1 | K133356 | IT WAS REPORTED THAT A PATIENT UNDERWENT A TRACT MIS CABG PROCEDURE AND SUTURE WAS USED. THE SUTURE SNAPPED DURING PROXIMAL ANASTOMOSES WITH THE KNOT PUSHER EVEN WITH A STATIC KNOT PUSHER. CUSTOMER OPENED A NEW SUTURE PACK. SURGEON HAD TO REDO THE ANASTOMOSES WHICH IN RETURN WILL TRAUMATIZE THE SAPHENOUS VEIN GRAFT USED IN CABG. THE PATIENT PASSED AWAY ON (B)(6) 2024. ADDITIONAL INFORMATION WAS REQUESTED. | PRODUCT COMPLAINT # ==> (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H6 COMPONENT CODE: G07002 NO DEVICE PROBLEM FOUND ADDITIONAL INFORMATION: D9, H3, H6 INVESTIGATION SUMMARY ==> THE PRODUCT WAS RETURNED TO ETHICON FOR EVALUATION. THREE UNOPENED SAMPLES WERE RECEIVED FOR ANALYSIS. PRODUCT CODE 8702H. IN ORDER TO EVALUATE THE CONDITION OF THE RETURNED SAMPLES, THE PACKETS WERE OPENED. THE SWAGE AND ATTACHMENT AREA WERE NOTED AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO ISSUES RELATED TO BREAKAGE SUTURES OR ANOMALIES WERE OBSERVED DURING THE EVALUATION. ALSO, A FUNCTIONAL TEST WAS PERFORMED USING INSTRON EQUIPMENT AND TENSILE FORCE WAS ABOVE THE MINIMUM REQUIREMENTS. AS PART OF THE ETHICON QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. THE DEVICE PERFORMED WITHOUT ANY DIFFICULTIES NOTED. ALTHOUGH NO PRODUCT DEFECT WAS IDENTIFIED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE LABORATORY ANALYSIS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS IS CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF POST-MARKET SURVEILLANCE.
PRLNE BLU 24IN 7-0 D/A BV-1 | K133356 | IT WAS REPORTED THAT A PATIENT UNDERWENT A TCRAT MIS CABG PROCEDURE AND SUTURE WAS USED. THE SUTURE SNAPPED DURING PROXIMAL ANASTOMOSES WITH THE KNOT PUSHER EVEN WITH A STATIC KNOT PUSHER. CUSTOMER OPENED A NEW SUTURE PACK. SURGEON HAD TO REDO THE ANASTOMOSES WHICH IN RETURN WILL TRAUMATIZE THE SAPHENOUS VEIN GRAFT USED IN CABG. THE PATIENT PASSED AWAY ON (B)(6) 2024. ADDITIONAL INFORAMTION WAS REQUESTED. | PRODUCT COMPLAINT #(B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H6 COMPONENT CODE: G07002 NO DEVICE PROBLEM FOUND. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON FOR EVALUATION. THREE UNOPENED SAMPLES WERE RECEIVED FOR ANALYSIS. PRODUCT CODE 8702H. IN ORDER TO EVALUATE THE CONDITION OF THE RETURNED SAMPLES, THE PACKETS WERE OPENED. THE SWAGE AND ATTACHMENT AREA WERE NOTED AS EXPECTED. THE SUTURES WERE DISPENSED WITHOUT PROBLEMS AND EXAMINED ALONG OF THE STRAND AND NO ISSUES RELATED TO BREAKAGE SUTURES OR ANOMALIES WERE OBSERVED DURING THE EVALUATION. ALSO, A FUNCTIONAL TEST WAS PERFORMED USING INSTRON EQUIPMENT AND TENSILE FORCE WAS ABOVE THE MINIMUM REQUIREMENTS. AS PART OF THE ETHICON QUALITY PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND RELEASED TO APPROVED SPECIFICATIONS. THE PERFORMED WITHOUT ANY DIFFICULTIES NOTED. ALTHOUGH NO PRODUCT DEFECT WAS IDENTIFIED, THERE MAY HAVE BEEN OTHER CIRCUMSTANCES OR ISSUES THAT OCCURRED DURING THE USE OF THE DEVICE THAT COULD NOT BE REPLICATED DURING THE LABORATORY ANALYSIS. ADDITIONAL COMPLAINT INFORMATION MONITORING FOR POTENTIAL SAFETY SIGNALS IS CONDUCTED THROUGH COMPLAINT TRENDING AS PART OF POST-MARKET SURVEILLANCE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY AND A RETURN LABEL WAS BEING REQUESTED. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT YET BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY AND A RETURN LABEL WAS BEING REQUESTED. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND NO ERRORS PRESENT. THERE WAS NO EVIDENCE OF FOAM PARTICLES OR DEGRADATION. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT FOR THE DEVICE EVALUATION.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WERE 24 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REP DREAMSTATION AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED AND SECTION H11 SHOULD BE REPORTED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT PASSED AWAY. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WHICH DUST CONTAMINATION WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS THIRTY-TWO ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. ADVERSE EVENT/PRODUCT PROBLEM (UPDATED FROM ADVERSE EVENT TO BOTH), PROBLEM CODE GRID, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID WERE UPDATED.
WATCHMAN FXD CURVE ACCESS SYSTEM | K212228 | IT WAS REPORTED THAT PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A 30MM WATCHMAN FLX LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN SUCCESSFULLY PERFORMED THE TRANSSEPTAL PUNCTURE AND POSITIONED THE WAS IN THE LAA OF THE PATIENT. THE PIGTAIL WIRE AND VERSACROSS CONNECT DILATOR WERE REMOVED FROM THE PATIENT. THE PATIENT TOOK A DEEP BREATH IN, AND THE PHYSICIAN FELT THEY SUCKED AIR IN THROUGH THE WAS. THE PHYSICIAN ASSESSED THE POSITION OF THE WAS AND CONFIRMED IT WAS STILL IN THE LEFT ATRIUM. AT THIS TIME THE PATIENT BECAME BRADYCARDIC AND WENT INTO HEART BLOCK. THE PATIENT WAS INTUBATED, AND CHEST COMPRESSIONS WERE PERFORMED. THE PHYSICIAN PLACED A NON-BOSTON SCIENTIFIC HEART PUMP TO HELP THE PATIENT. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS DUE TO A MYOCARDIAL INFARCTION.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
STEALTHSTATION¿ S8 SYSTEM | K162309 | MEDTRONIC RECEIVED INFORMATION REGARDING A NAVIGATION SYSTEM USED DURING A STEREOELECTROENCEPHALOGRAPHY (SEEG) PROCEDURE. IT WAS REPORTED THAT THE PATIENT HAD DIED. | CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID 29631 (SERIAL NUMBER: (B)(6)). H3, H6) NO PARTS HAVE BEEN RECEIVED BY THE MANUFACTURER FOR EVALUATION. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
AVALON CABLELESS ECG/IUP TRANSDUCER | K140535 | IT WAS REPORTED A UTERINE RUPTURE OCCURRED DURING ACTIVE LABOR AND THE INTRAUTERINE PRESSURE CATHETER WAS DISPLAYING INADEQUATE CERVICAL CHANGE. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | PHILIPS RECEIVED A COMPLAINT ON THE AVALON CABLELESS ECG/IUP TRANSDUCER INDICATING THAT A UTERINE RUPTURE OCCURRED DURING ACTIVE LABOR AND THE INTRAUTERINE PRESSURE CATHETER WAS DISPLAYING INADEQUATE CERVICAL CHANGE. A REQUEST FOR ADDITIONAL INFORMATION WAS SENT AND NEW INFORMATION WAS PROVIDED. THE PATIENT WAS NOT DIAGNOSED WITH PLACENTA PREVIA, AND NO BLEEDING WAS PRESENT DURING LABOR ON THIS PATIENT WITH NO PRIOR C-SECTIONS. THE PATIENT RECEIVED A BLOOD TRANSFUSION, THE RUPTURE WAS REPAIRED AND WAS DISCHARGED TO HOME IN STABLE CONDITION. THE OPERATIVE NOTE STATED, 'UTERUS WITH STELLATE RUPTURE ON THE LOWER LEFT SIDE THAT INCORPORATED THE POSTERIOR WALL AND CARRIED AROUND THROUGH THE SIDE WALL AND THE LOWER UTERINE SEGMENT'. THE PRESSURE CATHETER WAS NOT SPECIFICALLY THE CAUSE OF THE RUPTURE, BUT THE CUSTOMER WAS CONCERNED WITH THE ACCURACY OF THE AS A DRIVING FORCE FOR MANIPULATION OF OXYTOCIN DOSING. BASED ON THIS INFORMATION, THE PRESSURE CATHETER WAS NOT THE DIRECT PHYSICAL CAUSE OF THE UTERINE RUPTURE, BUT THE CUSTOMER WAS CONCERNED WITH THE VALUES PROVIDED MAY HAVE DRIVEN OXYTOCIN DOSING, WHICH MAY HAVE BEEN A FACTOR IN THE EVENT; HOWEVER, THE TRUE CAUSE REMAINS UNKNOWN. THROUGH COMMUNICATION WITH THE ACCOUNT MANAGER, IT WAS CONFIRMED THAT THE IUPC CATHETER BEING USED FOR MEASUREMENT WAS A KOALA IUPC, WHICH IS NOT MANUFACTURED BY PHILIPS. THE MANUFACTURER OF THE CATHETER PART ID PRODUCT NUMBER IPC-5000 WAS IDENTIFIED TO BE CLINICAL INNOVATIONS WHICH IS PART OF LABORIE MEDICAL TECHNOLOGIES. THE CUSTOMER INDICATED THEY HAVE BEEN IN CONTACT WITH THEIR KOALA IUPC CATHETER REPRESENTATIVE REGARDING THE ISSUE AS WELL. BASED ON THE INFORMATION AVAILABLE, THE REPORTED PROBLEM WAS RELATED TO THE KOALA IUPC CATHETER AND WAS UNRELATED TO THE PHILIPS TRANSDUCER. NO FURTHER INVESTIGATION BY PHILIPS IS NECESSARY AS THE PRODUCT IS NOT MANUFACTURED BY PHILIPS. LABORIE MEDICAL TECHNOLOGIES HAS BEEN NOTIFIED OF THE ISSUE FOR FURTHER INVESTIGATION. IT HAS BEEN CONCLUDED THAT NO FURTHER ACTION BY PHILIPS IS REQUIRED AT THIS TIME.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973, AND Z-1974.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY WHILE WAITING FOR THE REPLACEMENT DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972, Z-1973, AND Z-1974.
DS2ADV AUTO CPAP W/HUMID CELL/BT | K200480 | THE MANUFACTURER BECAME AWARE OF A USER OF BIPAP AUTO BIFLEX STATING DAUGHTER IN LAW CALLED IN STATING PATIENT PASSED AWAY. SHE WANTS TO RETURN THE DEVICE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION IN REFERENCE TO A USER OF DS2ADV AUTO CPAP STATING DAUGHTER IN LAW CALLED IN STATING PATIENT PASSED AWAY THE DEVICE WAS RETURNED TO A THIRD PARTY SERVICE CENTER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. DURING THE EVALUATION OF THE DEVICE AT THE THIRD PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED FOR PARTICLES AND SCRAPPED. THERE WERE NO REPORTABLE PRODUCT MALFUNCTIONS REPORTED IN THE EVALUATION.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY AND PATIENT'S WIFE WANTS TO RETURN OLD DEVICE. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT IS DECEASED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. ADDITIONALLY, CONNECTOR IS CORRODED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT TO ENROLL RECALLED DEVICE FOR THE CLASS ACTION LAWSUIT SETTLEMENT CLAIM. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AVAPS30 | K131982 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE REP DREAMSTATION AVAPS30 HAD STATES DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE REP DREAMSTATION AUTO CPAP HAD STATES DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT HER HUSBAND'S NAME TO BE REMOVED FROM EMAIL AND RECALL LIST. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT HER HUSBAND'S NAME TO BE REMOVED FROM EMAIL AND RECALL LIST. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE CORRECT B5 SHOULD BE- THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANT HER HUSBAND'S NAME TO BE REMOVED FROM EMAIL AND RECALL LIST. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. OUTCOMES ATTRIBUTED TO AE SECTION HAS BEEN CORRECTED FROM OTHER TO DEATH. APPROPRIATE CODE UPDATED/CORRECTED. IN THE INITIAL REPORT IT WAS WRONGLY REPORTED UNDER RECALL, CORRECTED IN THIS FOLLOW-UP REPORT. THIS REPORT IS NOT A PART OF RECALL.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY 2 YEARS AGO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR W/SD CARD | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAD PASSED AWAY, SO PATIENT'S WIFE WOULD LIKE TO RETURN THE DEVICE SINCE IT WAS NOT USED BY RECEIVING SHIPPING LABEL VIA MAIL. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAD PASSED AWAY, SO PATIENT'S WIFE WOULD LIKE TO RETURN THE DEVICE SINCE IT WAS NOT USED BY RECEIVING SHIPPING LABEL VIA MAIL.. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WERE 32 ERROR CODES FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT WAS SCRAPPED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CONGESTED HEART FAILURE, BREATHING ISSUES WHICH LED TO THE DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
WATCHMAN FXD CURVE ACCESS SYSTEM | K212228 | IT WAS REPORTED THAT PATIENT DEATH OCCURRED. A PULSED FIELD ABLATION (PFA) PROCEDURE WAS SUCCESSFULLY COMPLETED WITH THE FARAPULSE PFA DEVICE. THE PHYSICIAN THEN STARTED A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT AND THEN INSERTED A PIGTAIL WIRE. CONTRAST WAS RELEASED IN THE LAA, AND THE PATIENT WENT INTO CARDIAC ARREST. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION FOR TEN (10) MINUTES BEFORE THERE WERE STABILIZED. THE PHYSICIAN THEN IMPLANTED A NON-BOSTON SCIENTIFIC HEART PUMP. THE PROCEDURE WAS ENDED, AND NO CLOSURE DEVICE WAS IMPLANTED IN THE PATIENT. TWO (2) DAYS LATER THE PATIENT DIED WITH AN EXACT CAUSE OF DEATH UNKNOWN.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. ACCORDING TO THE NURSE, THE CUSTOMER EXPERIENCED KETOACIDOSIS THAT WAS "NOT HANDLED CORRECTLY" RESULTING IN CARDIAC ARREST. THE NURSE CONFIRMED THAT THE PUMP WAS "WORKING ACCORDINGLY TO SPECIFICATION." NO ADDITIONAL INFORMATION COULD BE CONFIRMED AT THE TIME OF THIS REPORT, AS PUMP DATA IS NOT AVAILABLE FOR REVIEW. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
BIPAP AVAPS C SERIES | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REP DREAMSTATION AUTO BIPAP THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DREAMSTATION AUTO CPAP DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
CONFIDA BRECKER CURVE GUIDEWIRE | K132623 | ADDITIONAL INFORMATION WAS RECEIVED THAT PER THE PHYSICIAN, THE CAUSE OF DEATH WAS UNKNOWN BUT WAS LIKELY DUE TO THE GUIDEWIRE PERFORATION AND EXTENSIVE SURGERY TO REPAIR THE VENTRICLE. ADDITIONALLY, PER THE PHYSICIAN, THE VALVE CAN BE EXCLUDED AS A CONTRIBUTING CAUSE TO THE DEATH. | UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
VOLUMAT MC AGILIA US | K121613 | THE FOLLOWING HAS BEEN REPORTED: (B)(6) HAD A TRAVELING NURSE RECENTLY SELECT MCG/MIN INADVERTENTLY WHEN MCG/KG/MIN WAS INTENDED TO BE SELECTED. SPECIFICALLY THE ORDER WAS FOR 0.03MCG/KG/MIN, BUT THE NURSE INITIALLY SELECTED THE DOSING OPTION OF MCG/MIN ON THE PUMP WHICH ULTIMATELY MEANT THE DOSE THE PATIENT RECEIVED WAS 106X LESS THAN WHAT THE ORIGINAL ORDER WAS INTENDED TO BE. THE PATIENT IN QUESTION WEIGHED 106KG. NO SAMPLE AVAILABLE NO LETTER REQUESTED EMAILED CUSTOMER ON 12/9/2024 TO DETERMINE IF PATIENT SUFFERED ANY ADVERSE EVENTS RELATED TO THE UNDERDOSE. ADDITIONAL INFORMATION RECEIVED FROM THE CUSTOMER'S PHARMACIST " "PER EMAIL ON 12/10/2024 I MESSAGED ONE OF OUR EMERGENCY ROOM PROVIDERS AND HOSPITALIST TO GET THEIR TAKE ON IF THE PATIENT SUFFERED ANY ADVERSE EVENTS RELATED TO THE UNDERDOSE AND THEY WAS THEIR RESPONSE: "HARD TO SAY. I MEAN SHE ULTIMATELY PASSED FROM SEPTIC SHOCK THE NEXT DAY AND THAT IS WHAT WE WERE TREATING, BUT SHE WAS BEING TREATED WITH OTHER PRESSORS UNTIL THE ISSUE WAS RESOLVED." OVERALL THE NOREPINEPHRINE APPEARED TO NOT BE EFFECTIVE, AS IT WAS ACTUALLY UNDERDOSED, AND OTHER PRESSERS WERE USED INSTEAD WHICH IS APPROPRIATE COURSE OF TREATMENT IF FOR SOME REASON NOREPINEPHRINE WERE NOT EFFECTIVE AT THE TYPICAL DOSE." " REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REP DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | DEVICE EVALUATION HAS BEEN COMPLETED. THE FOLLOWING FIELDS HAS BEEN UPDATED IN THIS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS NO REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN CONFIRMED NO PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID HAS BEEN UPDATED.
EVERFLO | K061261 | THE MANUFACTURER WAS NOTIFIED THAT A PATIENT USING AN EVERFLO OXYGEN CONCENTRATOR HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DUAL CHAMBER TEMPORARY PACEMAKER | K150246 | IT WAS REPORTED THAT FOLLOWING A PROCEDURE FOR AORTIC VALVE REPLACEMENT AN EXTERNAL PULSE GENERATOR (EPG) WAS USED AT DDD MODE "80/MIN WITH A PACING WIRE TO TREAT A PATIENT WITH A COMPLETE HEART BLOCK. THE FOLLOWING DAY WHEN THE STERNAL DRESSING WAS CHANGED THE PATIENT SUFFERED A CARDIOPULMONARY ARREST, IT WAS CONFIRMED THE EXTERNAL PULSE GENERATOR (EPG) WAS ON BUT NO LONGER PACING AND THE ELECTRODES DID NOT APPEAR TO HAVE BEEN MOBILIZED. IT WAS REPORTED THAT AT THE SAME TIME RESUSCITATION WAS STARTED, THE PATIENT INHALED MASSIVELY BEFORE INTUBATION. AN ELECTRO-DRIVE PROBE (SEES) WAS PLACED. IT WAS STATED THAT PACING DID NOT OCCUR WHEN REQUIRED, LEADING TO "PROLONGED CARDIAC ARREST AND FURTHER COMPLICATIONS OF INHALATION, PNEUMOPATHY AND LOW FLOW COMPLICATIONS". THE PATIENT WAS TREATED WITH ANTIBIOTICS FOR TREATMENT OF THE INHALATION PNEUMONITIS. IT WAS ALSO STATED THAT EPINEPHRINE AND A NON-SELECTIVE BETA-ADRENERGIC RECEPTOR AGONIST WAS ALSO USED TO TREAT THE PATIENT. LATER THE SAME DAY, IT WAS OBSERVED THAT THE EXTERNAL PACEMAKER AND THE SAME ELECTRODES WERE WORKING AGAIN WITHOUT ANY ADDITIONAL MANIPULATION. THE SEES WAS LEFT AS SENTINEL AND THE PATIENT WAS EFFECTIVELY "ELECTRO-TRAINED" BY THE EXTERNAL PACEMAKER UNTIL THE PATIENT DIED 2 DAYS LATER DUE TO MULTI ORGAN FAILURE. IT WAS REPORTED THAT IT WAS CONSIDERED THAT THE ELECTRO-DRIVE DEFECT WAS AT THE ORIGIN OF THE CARDIORESPIRATORY ARREST IS LINKED TO THE PATIENT'S DEATH. IT WAS NOTED THAT THE PULMONARY INHALATION IS RESPONSIBLE FOR SEPTIC SHOCK WITH MULTI-ORGAN FAILURE IN THE COURSE, LEADING TO THE PATIENT'S DEATH. IT WAS REPORTED THAT THE LOW FLOW AND PNEUMOPATHY RESULTING FROM COMPLICATIONS OF CARDIAC ARREST, SECONDARY TO LACK OF STIMULATION AND ATRIOVENTRICULAR (AV) BLOCKAGE CONTRIBUTED TO THE PATIENT¿S DEATH. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. THE DEVICE WAS DISCARDED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LUCAS CPR CHEST COMPRESSION SYSTEM | K090422 | THE CUSTOMER CONTACTED STRYKER TO REPORT THAT DURING USE OF THE DEVICE ON A WOUNDED PATIENT, THE DEVICE ALARMED AND STOPPER COMPRESSION. IF A DEVICE MALFUNCTIONS, THE DEVICE OPERATING INSTRUCTIONS INDICATE THAT THE USER IS TO REMOVE THE DEVICE AND IMMEDIATELY START MANUAL CHEST COMPRESSIONS. THE CUSTOMER REPORTED THAT MANUAL CHEST COMPRESSION COULD NOT BE PERFORMED AS THE PATIENT WAS BEING LIFTED AND RESCUED FROM A CLIFF. THE PATIENT ASSOCIATED WITH THE REPORTED EVENT DID NOT SURVIVE. | STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56. STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED STRYKER WITH THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION IS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. STRYKER PERFORMED A CLINICAL ASSESSMENT OF THE EVENT AND DETERMINED THAT THE DEVICE USE MAY HAVE CONTRIBUTED TO THE PATIENT'S OUTCOME.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY (B)(6) 2022. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE IS NOT AVAILABLE TO BE RETURNED FOR INVESTIGATION. THE USER'S WIFE STATES SHE NO LONGER HAS CPAP DEVICE AS SOMEONE CAME AND PICKED IT UP BEFORE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE REP DREAMSTATION AUTO CPAP HAD STATES DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER BECAME AWARE THAT A USER OF THE REP DREAMSTATION AUTO CPAP HAD STATES DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN CONFIRMED NO PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY CUSTOMER COMPLAINT IS NOT REPLICATED, AND THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX G: DATE RECEIVED BY MFG HAS BEEN UPDATED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE PASSED AWAY AND THE PATIENT¿S WIFE WANTS TO MAKE SURE ANOTHER DEVICE IS NOT BEING SENT TO HIM. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE REP DREAMSTATION AUTO CPAP HAD STATES DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
NXSTAGE SYSTEM ONE | K122051 | A REPORT WAS RECEIVED ON 20 NOV 2024 FROM THE HOME THERAPY NURSE (HTN) OF A 67-YEAR-OLD MALE WITH A MEDICAL HISTORY INCLUDING CARDIAC COMORBIDITIES AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 21 NOV 2024 FROM THE HTN WHO STATED THE PATIENT WAS IN TREATMENT FOR APPROXIMATELY 3.5 HOURS BEFORE THE EVENT OCCURRED WITH NO ALARMS OR CAUTIONS NOTED. EMERGENCY MEDICAL SERVICES (EMS) WAS CALLED AND THE PATIENT WAS PRONOUNCED ON THE SCENE, CAUSE OF DEATH CARDIAC ARREST. | THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
OMNIPOD 5 CONTROLLER | K203768 | IT WAS REPORTED ON SOCIAL MEDIA BY THE PATIENT'S DAUGHTER-IN-LAW THAT PATIENT PASSED AWAY FROM A DEVICE MALFUNCTION OF THE OMNIPOD SYSTEM. THE DAUGHTER-IN-LAW STATED PATIENT PASSED AWAY IN THEIR SLEEP DUE TO AN OVERDOSE OF INSULIN FROM THE OMNIPOD SYSTEM. NO OTHER INFORMATION IS AVAILABLE DUE TO ISSUE BEING REPORTED ON SOCIAL MEDIA. | ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED DEATH. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED.
ALPHAVAC | K240397 | AN ANGIODYNAMICS TERRITORY MANAGER (TM) REPORTED A PATIENT DEATH OCCURRED AFTER A PROCEDURE, USING AN ALPHAVAC MULTIPURPOSE MECHANICAL ASPIRATOR F18 C85. DURING THE PROCEDURE, THE WHITE HANDLE AND FIN DETACHED FROM THE DEVICE, WHEN THE PHYSICIAN WAS TURNING IT. THE TM NOTED THAT HANDLE AND FIN ARE LOCATED OUTSIDE OF THE PATIENT, AT ALL TIMES DURING THE PROCEDURE; THEREFORE, THIS DID NOT IMPACT THE PATIENT. THE PROCEDURE WAS PERFORMED ON AN EXTREMELY ILL PATIENT, WHO WAS IN THE INTENSIVE CARE UNIT, IN AN ATTEMPT TO REMOVE A PULMONARY EMBOLISM FROM THEIR LUNG. DURING THE PROCEDURE, WHICH WAS PERFORMED IN ACCORDANCE WITH THE IFU, THE EMBOLISM SHIFTED TO A POSITION WHERE IT WAS CUTTING OFF BLOOD SUPPLY. WHEN THIS OCCURRED, THE CANNULA WAS REMOVED FROM THE PATIENT AND THE PATIENT WAS SUCCESSFULLY REVIVED; HOWEVER, WHILE THE PATIENT WAS BEING TRANSPORTED TO SURGERY, SHE CODED AND ULTIMATELY EXPIRED. THE PHYSICIAN REPORTED THAT THERE IS NO CASUAL RELATIONSHIP BETWEEN THE ALPHAVAC DEVICE AND THE PATIENT'S EXPIRATION AND THEIR EXPIRATION IS DIRECTLY RELATED TO THEIR POOR HEALTH STATUS. BOTH THE PHYSICIAN AND THE TM WHO WERE PRESENT FOR THE PROCEDURE, HAVE PERFORMED SEVERAL SUCCESSFUL ALPHAVAC PROCEDURES AND ARE WELL EDUCATED IN THE USAGE OF THE DEVICE. | THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF THE HANDLE/FIN ON THE ALPHAVAC CANNULA HUB - BECAME DETACHED, WAS CONFIRMED SINCE CUSTOMER PROVIDED A PICTURE, HOWEVER, THE COMPLAINT SAMPLE WAS NOT RETURNED. WITHOUT RECEIVING THE COMPLAINT SAMPLE FOR EVALUATION, WE ARE UNABLE TO DETERMINE A ROOT CAUSE FOR THE HANDLE/FIN DETACHMENT. THIS HANDLE FIN COMPONENT IS BONDED TO THE CANNULA HUB USING ADHESIVE. POTENTIAL ROOT CAUSE FOR THIS TYPE OF FAILURE MODE IS INSUFFICIENT AMOUNT OF ADHESIVE APPLIED AND/OR END USER APPLYING EXCESSIVE AMOUNT OF TORQUE ON THE HANDLE FIN. THE PURPOSE OF THE FIN IS TO PROVIDE ORIENTATION/STEERABILITY OF THE FUNNEL TIP. A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE, ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. DIRECTIONS FOR USE AND MANUALS FOR THE ALPHAVAC SYSTEM AND ALL RELATED ACCESSORIES SHOULD BE READ PRIOR TO USE AND DEVICES USED AS INDICATED. AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES AS WELL AS PHYSICIANS TRAINED AND EXPERIENCED IN PERCUTANEOUS, INTRAVASCULAR, DIAGNOSTIC AND INTERVENTIONAL TECHNIQUES REQUIRING FLUOROSCOPIC OR IMAGE GUIDANCE AND VISUALIZATION. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: DISTAL EMBOLIZATION OF THROMBUS, PULMONARY EMBOLISM, CARDIAC ARREST, DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
DS2ADV AUTO CPAP | K200480 | THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
AGILIA SP MC WIFI GB | K121613 | THE FOLLOWING HAS BEEN REPORTED: EDIT BY CA-C. THERE ARE 2 PARTS TO THIS COMPLAINT, THE DRUG ERROR RESULTING IN PATIENT DEATH, AND THE EVENT LOG NOT BEING CORRECT. PLEASE SEE BELOW FOR HOW THIS WAS REPORTED TO US: EMAIL (FORMAL) A DRUG ERROR HAD OCCURRED ON THE INTENSIVE CARE UNIT RESULTING IN THE DEATH OF A PATIENT. THIS WAS ON THE SP MC WIFI. THE DRUG, INTENDED DOSE AND DELIVERED DOSE INFORMATION HAS BEEN REQUESTED BY MYSELF. CUSTOMER IS TRYING TO DOWNLOAD THE EVENT LOG FOR THE AGILIA SP MC S.N (B)(6) (TYPO, THE CORRECT SN IS (B)(6)) WHICH IS INVOLVED IN AN INCIDENT ON (B)(6) 2024, HOWEVER THE DATE IS MISSING, THE MONTHS IS ALL OVER THE PLACE E.G WHEN READING THE MONTHS IT WILL STARTS (B)(6) THEN WENT BACK TO (B)(6) WHICH DOES NOT MAKE SENSE REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EVENT EFFECTS WERE REPORTED. ADDITIONAL INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.
SIMPLYGO | K111885 | THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
BIPAP AUTOSV ADV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THE PATIENT PASSED AWAY AND THE CUSTOMER REQUESTED ASSISTANCE IN REVIEWING THE LOGS. THE STAFF REPORTED THEY DID NOT SEE A YELLOW LEADS OFF ALARM AND THE PATIENT PASSED AWAY. THE SETTINGS AND AUDIT LOGS WERE REVIEWED WITH THE NURSE MANAGER, REVEALING THE LEADS OFF ALARM WAS SET TO YELLOW AND INOPS WERE SET TO RE-ALARM. A YELLOW ALARM FOR LEADS OFF WAS GENERATED AND ACKNOWLEDGED AND RE-ALARMED FIVE TIMES. NO DEVICES WERE TAKEN OUT OF SERVICE AND THE MANAGER DID NOT FEEL THE DEVICE WAS AT FAULT. | THE REMOTE CLINICAL SUPPORT SPECIALIST REVIEWED THE AUDIT LOGS AND DEVICE SETTINGS WITH THE NURSE MANAGER, AND CONFIRMED THE LEADS OFF ALARM WAS SET TO YELLOW AND INOPS WERE SET TO RE-ALARM. A YELLOW ALARM FOR LEADS OFF WAS NOTED IN THE AUDIT LOGS AND IT WAS ACKNOWLEDGED AND RE-ALARMED 5 TIMES. THE MANAGER DID NOT FEEL THE MONITORING SYSTEM WAS AT FAULT; THEREFORE, THE DEVICE REMAINS IN SERVICE. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED WE WERE UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DAVINCI XI | K131861 | IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNEVENTFUL DA VINCI-ASSISTED INGUINAL HERNIA REPAIR WITHOUT ANY INTRA-OPERATIVE COMPLICATIONS. THE PATIENT WAS NOTED TO HAVE EXTENSIVE ADHESIONS AND ADHESIOLYSIS WAS PERFORMED OR BEDSIDE BEFORE PLACING PORTS FOR THE ROBOTIC PROCEDURE. POST-OPERATIVELY, THE PATIENT WAS DISCHARGED HOME AS PLANNED. THE SURGEON INFORMED THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE (CSR) THAT SEVERAL DAYS AFTER BEING DISCHARGED, THE PATIENT FELL WHILE AT HOME. THE PATIENT WAS ADMITTED TO A DIFFERENT HOSPITAL THAN WHERE THE DA VINCI PROCEDURE WAS PERFORMED AND EXPIRED SEVERAL DAYS AFTER BEING HOSPITALIZED. THE SURGEON OF THE DA VINCI PROCEDURE DID NOT KNOW THE PATIENT'S CAUSE OF DEATH AND STATED THAT THE FALL EVENT WAS NOT RELATED TO THE DA VINCI ROBOT. THE SURGEON STATED THAT IT WAS POSSIBLE THE PATIENT SUSTAINED A LIVER LACERATION AS A RESULT OF THE FALL; HOWEVER, IT WAS NOT PROVIDED HOW THAT WAS THEORIZED. THE SURGEON DECLINED TO PROVIDE ANY ADDITIONAL INFORMATION, INCLUDING THE NAME OF THE TREATING PHYSICIAN AT THE OTHER HOSPITAL. | NO SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE WAS PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. ON 11/19/2024, THE SURGEON STATED THAT THE PROCEDURE HAD BEEN PERFORMED ¿WITHIN THE LAST 2 MONTHS¿. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT AN INGUINAL HERNIA REPAIR AND FELL AT HOME A FEW DAYS LATER, WAS ADMITTED TO THE HOSPITAL, AND EVENTUALLY DIED A FEW DAYS LATER FOR REASONS THAT ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
WECK VISTA | K152663 | IT WAS REPORTED THAT WHILE PLACING PORTS FOR A DA VINCI-ASSISTED HYSTERECTOMY, THE INITIAL TROCAR/OBTURATOR PLACEMENT INADVERTENTLY INJURED THE PATIENT'S AORTA RESULTING IN BLEEDING; THE PATIENT EXPIRED. IT IS UNKNOWN WHAT INTERVENTIONS WERE PERFORMED DUE TO THIS EVENT. MULTIPLE ATTEMPTS TO CONTACT THE SURGEON FOR ADDITIONAL INFORMATION HAVE BEEN MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | AS THIS EVENT OCCURRED DURING PORT PLACEMENT, THE DA VINCI SYSTEM WAS NOT DOCKED; THEREFORE, THERE ARE NO SYSTEM OR INSTRUMENT LOGS AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF BLEEDING FROM AN AORTA INJURY DURING THE INITIAL ATTEMPT TO GET ABDOMINAL ACCESS. HOW THIS OCCURRED WAS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS CONTRIBUTED TO THESE EVENTS.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | ON 4/DEC/2024, IT WAS REPORTED THAT THIS PATIENT ON PERITONEAL DIALYSIS (PD) EXPIRED WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE STATED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF END STAGE RENAL DISEASE (ESRD), ATHEROSCLEROTIC HEART DISEASE WITH PREVIOUS MYOCARDIAL INFARCTION (UNRELATED TO DIALYSIS), HYPERLIPIDEMIA, DIABETES MELLITUS AND ULCERATIVE COLITIS. ADDITIONALLY, THE NURSE REPORTED THAT THE PATIENT WAS IMMOBILE AND IN POOR HEALTH RECEIVING DIALYSIS TREATMENT IN A NURSING HOME. PER THE NURSE, THE PATIENT ALSO HAD AN INDWELLING FOLEY CATHETER AND THAT THE PATIENT HAD SYMPTOMS OF A URINARY TRACT INFECTION SINCE (B)(6) 2024. HOWEVER, IT WAS REPORTED THAT THE PATIENT WAS NOT ON ANTIBIOTIC THERAPY. THE PD NURSE STATED THAT ON (B)(6) 2024 THE PATIENT BEGAN PD TREATMENT WITH THE FRESENIUS CYCLER (SERIAL NUMBER (B)(6)) WITHOUT ANY ISSUES IN THE NURSING HOME. AT APPROXIMATELY 5:30AM (ON (B)(6) 2024), THE PATIENT WAS ALIVE RECEIVING HER DIALYSIS OBSERVED ON MORNING NURSING CHECK. SUBSEQUENTLY, AT APPROXIMATELY 8:30 AM, THE PATIENT WAS FOUND UNRESPONSIVE. THE NURSE STATED THE PATIENT WAS IN THE LAST DRAIN OF PD TREATMENT AND IT IS UNKNOWN IF THE TREATMENT WAS COMPLETED. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION (CPR). HOWEVER, CPR WAS NOT SUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED IN THE NURSING HOME (NOT BROUGHT TO THE HOSPITAL). THE NURSE REPORTS THAT THE CAUSE OF DEATH HAS NOT BEEN FIRMLY ESTABLISHED. THE PATIENT¿S ESRD DEATH FORM IS NOT AVAILABLE. HOWEVER, THE NURSE INDICATED THAT IT IS SUSPECTED THE PATIENT MAY HAVE BEEN SEPTIC DUE TO CONCURRENT URINARY TRACT INFECTION WITH THE INDWELLING FOLEY CATHETER. THE NURSE CONFIRMED THERE WERE NO ISSUES WITH PD TREATMENT OR PRODUCT IN RELATION TO THE DEATH. THE CYCLER IS IN THE PD CLINIC PENDING RETURN TO MANUFACTURER. | ADDITIONAL INFORMATION: B7.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THE PATIENT RECEIVED PCA (PATIENT-CONTROLLED ANALGESIA) DOSE "OUTSIDE THE GUARDRAILS LIMIT" AND THE EVENT ALLEGEDLY RESULTED TO DEATH. PER REPORT, THE PCA INFUSION SET UP WAS NOT CONNECTED TO AN END-TIDAL CO2 (ETCO2) MONITOR AT THE TIME OF THE EVENT THEREFORE THE PATIENT'S ETCO2 LEVEL WAS NOT MONITORED. THIS WAS REPORTED TO A BD FIELD SPECIALIST WHILE THEY WERE ON THE CUSTOMER'S FACILITY WORKING ON A PROJECT. PER REPORT, THE FACILITY HAS SINCE CHANGED THEIR POLICY REQUIRING USE OF ETCO2 MONITOR WHENEVER PCA MODULE IS IN USE. PER CUSTOMER (BIOMEDICAL MANGER), "I HAVE BEEN TOLD THAT THIS HAS BEEN RESOLVED, AND NO FURTHER INVESTIGATION IS NEEDED." MULTIPLE COMMUNICATIONS HAVE BEEN MADE REQUESTING FOR ADDITIONAL INFORMATION AND PRODUCT RETURN. NO RESPONSE HAS BEEN PROVIDED TO DATE. | OMIT: C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. CORRECTION: DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES.
ALARIS SYSTEM | K043299 | IT WAS REPORTED THAT THE PATIENT RECEIVED PCA (PATIENT-CONTROLLED ANALGESIA) DOSE "OUTSIDE THE GUARDRAILS LIMIT" AND THE EVENT ALLEGEDLY RESULTED TO DEATH. PER REPORT, THE PCA INFUSION SET UP WAS NOT CONNECTED TO AN END-TIDAL CO2 (ETCO2) MONITOR AT THE TIME OF THE EVENT THEREFORE THE PATIENT'S ETCO2 LEVEL WAS NOT MONITORED. THIS WAS REPORTED TO A BD FIELD SPECIALIST WHILE THEY WERE ON THE CUSTOMER'S FACILITY WORKING ON A PROJECT. PER REPORT, THE FACILITY HAS SINCE CHANGED THEIR POLICY REQUIRING USE OF ETCO2 MONITOR WHENEVER PCA MODULE IS IN USE. PER CUSTOMER (BIOMEDICAL MANGER), "I HAVE BEEN TOLD THAT THIS HAS BEEN RESOLVED, AND NO FURTHER INVESTIGATION IS NEEDED." MULTIPLE COMMUNICATIONS HAVE BEEN MADE REQUESTING FOR ADDITIONAL INFORMATION AND PRODUCT RETURN. NO RESPONSE HAS BEEN PROVIDED TO DATE. | OMIT: C20 - NO FINDINGS AVAILABLE, D15 - CAUSE NOT ESTABLISHED. CORRECTION: DESCRIBE EVENT OR PROBLEM. ADDITIONAL INFORMATION: IMDRF ANNEX A, B, C, D, G CODES.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 01/28/2025 AND H11 IS UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED BUT THERE WAS NO MENTION WHETHER THE DEVICE WILL BE AVAILABLE FOR EVALUATION, THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. AFTER DEVICE RETURN AND COMPLETION OF EVALUATION, A FOLLOW-UP REPORT WILL BE FILED. SECTION G AND H IS UPDATED IN THIS REPORT.
ESSENCE RENTAL | K042655 | THE MANUFACTURER RECEIVED INFORMATION REGARDING AN ESSENCE RENTAL. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION PATENT PASSED AWAY. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON, AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED. | THE MANUFACTURER HAS PREVIOUSLY FILED THIS REPORT AS REP DEVICE, BUT DEVICE INFORMATION HAS BEEN UPDATED. HENCE A CORRECTION REPORT HAS BEEN FILED. BOX D HAS BEEN UPDATED AND CORRECTED.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE PASSED AWAY AND THE PATIENT¿S WIFE WANTS TO RETURN IT. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP | K200480 | THE PATIENT HAS ALLEGED PATIENT DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
BIPAP AUTO BIFLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO BIPAP AUTO BIFLEX DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THERE WAS NO ALLEGATION OF THE DEVICE CONTRIBUTING TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DS2ADV AUTO CPAP | K200480 | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
EDWARDS ESHEATH PLUS INTRODUCER SHEATH | K200258 | AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 26MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION, MULTIPLE SHOCKWAVE PERCUTANEOUS TRANSLUMINAL ANGIOPLASTY AND STENTING (PTA'S) WERE DONE FOR INTRODUCTION OF THE 14F ESHEATH+. THE ESHEATH+ WAS ADVANCED INTO THE ABDOMINAL AORTA AND THE VALVE WAS ADVANCED INTO THE COMMON ILIAC. THE DELIVERY SYSTEM BUCKLED AND APPEARED TO TEAR THE VESSEL. THE REMOVAL OF THE DELIVERY SYSTEM AND VALVE APPEARED TO AVULSE THE ILIAC ARTERY. THE ENTIRE SYSTEM AND VALVE WERE REMOVED AND A NEW 14F ESHEATH+ WAS INTRODUCED. A CODA BALLOON WAS USED TO STABILIZE THE BLEEDING. MULTIPLE VIABAHN STENTS WERE DEPLOYED THE LENGTH OF THE ILIAC SYSTEM, HOWEVER THE PATIENT ULTIMATELY EXPIRED. IT WAS RECOGNIZED AND DISCUSSED PATIENT ANATOMY LED TO COMPLICATION. | THE INVESTIGATION IS ONGOING.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REP DREAMSTATION AUTO BIPAP. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
EDWARDS ESHEATH INTRODUCER SET | K200258 | AS REPORTED BY OUR AFFILIATES IN AUSTRALIA, THIS WAS A CASE WITH A 29MM SAPIEN 3 VALVE, IN AORTIC POSITION BY SUBCLAVIAN APPROACH WITH SURGICAL CUTDOWN. DURING THE PROCEDURE, THERE WAS RESISTANCE ADVANCING THE DELIVERY SYSTEM WITH THE CRIMPED VALVE THROUGH THE ESHEATH AND, SUBSEQUENTLY, THE SHEATH KINKED/BENT. AT THIS POINT, THE NOSE CONE EXITED THE DISTAL TIP OF ESHEATH, BUT THE VALVE DID NOT EXIT. IT WAS NOT POSSIBLE TO PUSH THE VALVE FORWARD OR PULL BACK. WITH WORSENING HEMODYNAMICS, THE ENTIRE SYSTEM WAS REMOVED AS A SINGLE UNIT. A MAJOR SUBCLAVIAN INJURY WAS NOTED. EMERGENCY STERNOTOMY WAS PERFORMED, AND A BYPASS PUMP WAS USED. ALL RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL. THE PATIENT PASSED AWAY. AS PER MEDICAL OPINION, THE PERCEIVED ROOT CAUSE OF THE SUBCLAVIAN INJURY WAS THE SPLIT IN THE ESHEATH. PER IMAGE EVALUATION, IT WAS OBSERVED THAT THE SHEATH LINER WAS TORN IN THE MIDDLE SHAFT WITH THE COMMANDER DELIVERY SYSTEM AND VALVE PROTRUDING THROUGH LINER. | UPDATED SECTION H6. THE EVENT REPORTED IS ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION AS IT WAS DISCARDED. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. THE PROVIDED IMAGERY WAS EVALUATED AND REVEALED THE FOLLOWING: COMMANDER DELIVERY SYSTEM WITH CRIMPED VALVE WAS OBSERVED INSERTED THROUGH THE ESHEATH. COMMANDER DELIVERY SYSTEM NOSE TIP EXITED THE ESHEATH TIP. THE ESHEATH TIP APPEARS UNOPENED. THE ESHEATH LINER WAS TORN IN THE MIDDLE SHAFT. WITH THE COMMANDER DELIVERY SYSTEM AND VALVE PROTRUDING THROUGH LINER. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH USING THE S3 VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, SHEATH SHAFT KINKS CAN BE A RESULT OF ANY EXCESSIVE DEVICE MANIPULATION APPLIED TO OVERCOME THE RESISTANCE. THE COMPLAINT FOR RESISTANCE BETWEEN SYSTEM COMPONENTS LEADING TO INABILITY TO ADVANCE THROUGH THE SHEATH WAS CONFIRMED BASED ON THE EVALUATION OF THE PROVIDED IMAGERY. IT IS POSSIBLE THAT ONE OR MORE PATIENT/PROCEDURAL FACTORS (ACCESS VESSEL CALCIFICATION, TORTUOSITY, UNDERSIZED VESSEL DIAMETERS, AND/OR STEEP ANGLE OF INSERTION) MAY HAVE BEEN PRESENT DURING THE PROCEDURE. HOWEVER, DUE TO INSUFFICIENT INFORMATION, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED. THE COMPLAINT FOR KINKED SHEATH WAS UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF APPLICABLE IMAGERY OR RETURNED DEVICE FOR EVALUATION. AVAILABLE INFORMATION SUGGESTS PROCEDURAL FACTORS (EXCESSIVE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENT AS IT WAS REPORTED THAT "IT WAS NOT POSSIBLE TO PUSH THE VALVE FORWARD OR PULL BACK". IF EXCESSIVE DEVICE MANIPULATION IS APPLIED IN AN ATTEMPT TO OVERCOME THE EXPERIENCED RESISTANCE, THE SHEATH SHAFT CAN KINK, ESPECIALLY IF COMPOUNDED WITH VESSEL CALCIFICATION AND/OR TORTUOSITY. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, SHEATH LINER TEAR EVENTS HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, AND/OR WITHDRAWAL OF AN ALTERED BALLOON PROFILE OF THE DELIVERY SYSTEM OR BAV. THE COMPLAINT FOR PUNCTURED LINER WAS CONFIRMED BASED ON THE EVALUATION OF THE PROVIDED IMAGERY. IT IS POSSIBLE THAT ONE OR MORE PATIENT/PROCEDURAL FACTORS (ACCESS VESSEL CALCIFICATION, TORTUOSITY) MAY HAVE BEEN PRESENT DURING THE PROCEDURE. HOWEVER, DUE TO INSUFFICIENT INFORMATION, A DEFINITIVE ROOT CAUSE IS UNABLE TO BE DETERMINED. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
NXSTAGE SYSTEM ONE | K122051 | A REPORT WAS RECEIVED ON (B)(6) 2024 FROM THE HOME THERAPY NURSE (HTN) OF A 76-YEAR-OLD MALE WITH A MEDICAL HISTORY INCLUDING A RECENT HOSPITALIZATION FOR ATRIAL FIBRILLATION AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 26 AND 27 NOV 2024 FROM THE HTN WHO STATED THE PATIENT LOST CONSCIOUSNESS APPROXIMATELY 2 HOURS INTO TREATMENT. EMERGENCY MEDICAL SERVICES (EMS) WAS CALLED AND THE PATIENT WAS TRANSPORTED TO HOSPITAL. RESUSCITATION EFFORTS (NOS) WERE UNSUCCESSFUL AND THE PATIENT WAS PRONOUNCED AT THE HOSPITAL, SUSPECTED CAUSE OF DEATH CARDIAC ARREST. | THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
DREAMSTATION BIPAP ST30 | K102465 | THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO CORRECT THE PREVIOUS REPORT AND PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE DREAMSTATION BIPAP ST30 DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. AS A RESULT OF AN INTERNAL INSPECTION, NO FAULTS WERE FOUND, AND THE FILTERS, TUBING, AND MANUAL WERE REPLACED. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO LOAN STOCK. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO CORRECT THE PREVIOUS REPORT AND PROVIDE THE MANUFACTURER'S COMPLETED INVESTIGATION. THE DREAMSTATION BIPAP ST30 DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. AS A RESULT OF AN INTERNAL INSPECTION, NO FAULTS WERE FOUND, AND THE FILTERS, TUBING, AND MANUAL WERE REPLACED. THE DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS SENT TO BE PUT BACK INTO LOAN STOCK. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
MULTI-FUNCTIONAL RESPIRATORY P | K945578 | IT WAS REPORTED THAT AN SPO2, CONTINUED TO DISPLAY A 100% READING TWO HOURS AFTER A PATIENT'S DEATH. THE SENSOR REMAINED ATTACHED TO THE PATIENT POST-MORTEM. THE SPO2 WAS 100% BECAUSE THE RATIO OF NEAR-INFRARED LIGHT ABSORPTION TO RED LIGHT ABSORPTION (I.E., R VALUE) CONTINUES TO BE CALCULATED AS A LOW VALUE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
2008T HEMODIALYSIS SYSTEM WITHOUT CDX | K150708 | ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM 29/DEC/2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER. | ADDITIONAL INFORMATION: A1, A2, A3, A4, B5, G2, H6, H10. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HD THERAPY UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, COMBI-SET BLOODLINES AND THE SERIOUS ADVERSE EVENTS OF HYPOTENSION, LOSS OF CONSCIOUSNESS, AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT¿S EXPIRATION OCCURRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON THE POST-EVENT TESTING, THERE IS EVIDENCE TO SUPPORT THE EXISTENCE OF INTERNAL DEVICE FAILURES, WHICH MAY HAVE CAUSED THE 2008T HEMODIALYSIS SYSTEM TO PERFORM OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR APPROXIMATELY 45% OF ALL ESRD DEATHS, AND APPROXIMATELY 25% OF THOSE DEATHS ARE DUE TO SUDDEN CARDIAC DEATH. BASED ON THE INFORMATION AVAILABLE, THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS ACTIVELY UNDERGOING HD THERAPY WHEN HE WAS FOUND UNRESPONSIVE, THE FACILITY¿S REQUEST FOR POST-EVENT FUNCTIONAL COMPLIANCE TESTING, THE UF PUMP FAILURE REQUIRING RECALIBRATION, THE PRESSURE HOLDING TEST FAILURE, AND A PENDING MANUFACTURER EVALUATION OF THE DIALYZER/BLOODLINES, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
DAVINCI X | K171294 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION B5 ADDITIONAL INFORMATION: THE DA VINCI-ASSISTED PROCEDURE WAS A RECTOSIGMOIDECTOMY DUE TO RECTAL CANCER. THE OFFICIAL CAUSE OF DEATH WAS DOCUMENTED AS "CHALLENGE NEOPLASIA". CULTURES OF THE REPORTED YELLOW SECRETIONS FOUND IN THE SECONDARY OPERATION IDENTIFIED ABUNDANT AEROMONAS HYDROPHILA/CAVIAE MICROORGANISMS. THE CUSTOMER STATED THAT NO DA VINCI PRODUCTS CAUSED OR CONTRIBUTED TO THIS EVENT.
BENEVISION N1 PATIENT MONITOR | K213799 | THE CUSTOMER REPORTED THAT A PATIENT WAS CONNECTED TO SPO2 AND THE ECG DISPLAYED "LEADS OFF" AS THE PATIENT WAS MOVING TOO MUCH. ECG ALARM WAS MISSED AND THE PATIENT WAS FOUND UNRESPONSIVE AND CPR WAS INITIATED. RESUSCITATION EFFORTS WERE NOT SUCCESSFUL AND THE PATIENT EXPIRED ON (B)(6) 2024 AT 2:30AM. | THE MINDRAY FIELD SERVICE REPRESENTATIVE VERIFIED THE PROPER OPERATION OF BENEVISION N1 (SERIAL NUMBER (B)(6)). SYSTEM LOGS OF PATIENT DATABASES, DEVICE IP ADDRESSES, AND ROOM NUMBERS WERE COLLECTED FOR ANALYSIS, AND NO MALFUNCTION WAS CONFIRMED. THE PATIENT MONITOR WAS GENERATING ECG TECHNICAL ALARMS, INDICATING THAT THE ECG LEADS WERE NOT CONNECTED TO THE PATIENT. IN ADDITION, IT ALARMED WHEN ECG WAVEFORMS WERE UNAVAILABLE. THESE ALARMS ARE CONSISTENT WITH THE EVENT DESCRIPTION, WHICH STATED THAT THE PATIENT WAS MOVING TOO MUCH. ACCORDING TO THE BENEVISION N SERIES PATIENT MONITOR OPERATOR'S MANUAL (PART NUMBER 046-011259-00) ALARMS TABLE, THE RECOMMENDATION IS THAT AFTER THE DEVICE ALARMS FOR ECG LEAD-OFF, THE USER NEEDS TO CHECK THE CONNECTION OF THE ELECTRODES AND LEAD WIRES. MINDRAY BENEVISION N1 PERFORMED PER SPECIFICATIONS.
KYPHON® HV-R® BONE CEMENT | K041584 | INFORMATION WAS RECEIVED FROM THE REGULATORY BODY VIA MANUFACTURER REPRESENTATIVE REGARDING AN EVENT WHICH OCCURRED DURING A SPINAL FUSION PROCEDURE FROM THE TENTH THORACIC VERTEBRA (T10) TO THE SECOND LUMBAR VERTEBRA (L2) IN A PATIENT DIAGNOSED WITH PATHOLOGIC FRACTURE OF THE TWELFTH THORACIC VERTEBRA (T12). IT WAS REPORTED THAT DURING SURGERY, THE SURGEON INJECTED BONE CEMENT INTO THE FENESTRATED SCREWS UNDER FLUOROSCOPIC IMAGING TO OBSERVE THE FILLING. AT THORACIC T10 LEVEL THE SURGEON NOTED THAT THE CEMENT WAS MINIMIZED DUE TO THE OBSERVATION OF SOME CEMENT EXTRAVASATION AT THAT LEVEL. NO OTHER ISSUES WERE OBSERVED WHILE FILLING AT T11, L1 OR L2 LEVELS. SURGERY CONTINUED AND THE SURGEON WAS INFORMED THAT THERE WAS A PULMONARY EVENT WITH CONCERN FOR PULMONARY EMBOLISM. SURGEON NOTED DUE TO TIMING OF EVENT, CONCERN FOR CEMENT MONOMER EMBOLIZATION LEADING TO REACTION. PATIENT QUICKLY FLIPPED, CARDIOPULMONARY RESUSCITATION WAS INITIATED BUT THE PATIENT EXPIRED. | H6 - NEITHER THE DEVICE NOR FILMS OF APPLICABLE IMAGING STUDIES WERE RETURNED TO THE MANUFACTURER FOR EVALUATION. THEREFORE, WE ARE UNABLE TO DETERMINE THE DEFINITIVE CAUSE OF THE REPORTED EVENT. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
ATRICLIP® PRO-V® LAA EXCLUSION SYSTEM ( | K233407 | A PATIENT UNDERWENT A ROBOTIC VIDEO-ASSISTED THORACOSCOPIC SURGICAL (VATS) CONVERGENT PROCEDURE WITH CONCOMITANT LEFT ATRIAL APPENDAGE EXCLUSION (LAAE) USING A PROV45 CLIP. THE ABLATION PORTION WAS COMPLETED AND THE PROV245 WAS PLACED ON THE APPENDAGE. APPROXIMATELY 3-5 MINUTES AFTER THE PROV245 CLIP WAS PLACED ON THE APPENDAGE, PATIENT WENT INTO VENTRICULAR FIBRILLATION (VFIB). ELECTRICAL DEFIBRILLATION WAS ATTEMPTED THREE TIMES WITHOUT SUCCESS. SURGEONS ATTEMPTED TO PERFORM CORONARY ARTERY BYPASS GRAFT (CABG) PROCEDURE BUT WERE UNSUCCESSFUL AND PATIENT EXPIRED. WHILE THERE IS NO EVIDENCE THAT AN ATRICURE DEVICE DIRECTLY CAUSED THE VENTRICULAR FIBRILLATION, CONTRIBUTION FROM THE LAAE PROCEDURE CANNOT BE RULED OUT AND SO THIS EVENT IS BEING REPORTED PER PART 803. THERE WAS NO REPORTED DEVICE MALFUNCTION, AND THE ADVERSE EVENT WAS THE RESULT OF A PROCEDURAL COMPLICATION. | THE DEVICE WAS NOT RETURNED FOR INVESTIGATION, AND A DEVICE HISTORY RECORD REVIEW WAS UNABLE TO BE CONDUCTED AS THE LOT NUMBER FOR THE DEVICE WAS NOT REPORTED OR ABLE TO BE ASCERTAINED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED REGARDING THE VOLUNTARY FIELD SAFETY NOTICE/RECALL CONCERNING THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER HAS BEEN INFORMED THAT A USER PASSED AWAY DUE TO A LACK OF OXYGEN TO THE BRAIN, WHICH WAS ALLEGEDLY CAUSED BY NOT USING THE AFFECTED DEVICE. THE USER HAD NOT BEEN USING THE DEVICE FOR TWO YEARS AND PASSED AWAY IN 2023. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
INTELLIVUE MX40 802.11A/B/G | K113125 | IT WAS REPORTED THE CUSTOMER WANTED TO KNOW IF THE EQUIPMENT HAS AN OPTIONAL FUNCTION FOR EGC ANALYSIS AND ALARMING DURING POSTOPERATIVE CARE AND CARDIAC ARREST. IT WAS INDICATED THAT A PATIENT DEATH OCCURRED WHILE THE PATIENT WAS CONNECTED TO THE MX40 WHICH WAS CONNECTED TO THE PIC IX. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K091319 | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
PATIENT INFORMATION CENTER IX | K211900 | IT WAS REPORTED THE CUSTOMER WANTED TO KNOW IF THE EQUIPMENT HAS AN OPTIONAL FUNCTION FOR EGC ANALYSIS AND ALARMING DURING POSTOPERATIVE CARE AND CARDIAC ARREST. IT WAS INDICATED THAT A PATIENT DEATH OCCURRED WHILE THE PATIENT WAS CONNECTED TO THE MX40 WHICH WAS CONNECTED TO THE PIC IX. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
BIPAP AUTOSV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. THERE WAS REPORT OF PATIENT DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE CONTACTED RECALL SUPPORT AND STATED THAT HER HUSBAND HAD A RECALLED PHILIPS DEVICE AND NEVER GOT A REPLACEMENT FOR IT. SHE SAID THE PATIENT HAS BEEN DECEASED FOR 2 YEARS AND DOES NOT NEED A REPLACEMENT ANY LONGER. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K220414 | IT WAS REPORTED THAT PATIENT DEATH OCCURRED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED, A VERSACROSS CONNECT FOR LAAC KIT WAS SELECTED FOR USE. THE PATIENT WAS UNDER CONSCIOUS SEDATION AND WAS NOT SNORING. THE PHYSICIAN SUCCESSFULLY PERFORMED THE TRANSSEPTAL PUNCTURE AT MID-MID UNDER INTRACARDIAC ECHOCARDIOGRAPHY (ICE) GUIDANCE. THE PHYSICIAN WENT LEFT SIDED WITH THE RF PIGTAIL WIRE, CONNECT DILATOR AND WATCHMAN SHEATH (WAS), THE PATIENT STARTED SNORING. THE RF WIRE AND DILATOR WERE REMOVED FROM THE PATIENT. THE PATIENT TOOK A DEEP BREATH IN, AND THE PHYSICIAN FELT THEY SUCKED AIR IN THROUGH THE WAS. THE PHYSICIAN ASSESSED THE POSITION OF THE WAS AND CONFIRMED IT WAS STILL IN THE LEFT ATRIUM. AT THIS TIME THE PATIENT BECAME BRADYCARDIC AND WENT INTO HEART BLOCK. THE PATIENT WAS INTUBATED, AND CHEST COMPRESSIONS WERE PERFORMED. ONCE INTUBATED THE TEE PROBE WAS DROPPED, 10% EF WAS NOTED. THE PHYSICIAN PLACED A NON-BOSTON SCIENTIFIC HEART PUMP TO HELP THE PATIENT. THE PATIENT DID NOT RECOVER FROM THIS EVENT AND THE PATIENT DIED. THE OFFICIAL CAUSE OF DEATH WAS DUE TO A MYOCARDIAL INFARCTION. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). THERE WAS NO DIFFICULTY NOTED WITH THE TSP WORKFLOW. AIR EMBOLISM WAS NOT RECOGNIZED ON IMAGING. THE SHEATH VALVE WAS OPEN WHEN THE VERSACROSS DILATOR WAS TAKEN OUT, AND THEN SUBSEQUENTLY CLOSED AFTER IT THE DILATOR WAS OUT. DURING ASPIRATION, THERE WERE NO CATHETER PRESENT IN THE SHEATH WHEN THE GURGLING WAS NOTED. THEN THE VERSACROSS RF WIRE WAS PUT BACK UP. THERE WAS NO VERSACROSS DEVICE MALFUNCTIONS REPORTED. IN THE PHYSICIAN OPINION, THE WATCHMAN SHEATH HEMOSTASIS VALVE SHOULD HAVE BEEN BETTER DESIGNED TO MITIGATE AIR IN CONSCIOUS SEDATION PATIENTS. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
GEM PREMIER 5000 PAK | K203790 | IT WAS REPORTED THAT A 14-MONTH-OLD MALE PATIENT WITH IDIOPATHIC CHRONIC PLEURAL EFFUSION UNDERWENT FOLLOW-UP FOR IDIOPATHIC CHRONIC PERICARDIAL EFFUSION, WITHOUT STRUCTURAL HEART DISEASE OR HEMODYNAMIC REPERCUSSIONS. THE PATIENT WAS ADMITTED FOR SURGERY: PLEUROPERICARDIAL WINDOW BY THORACOSCOPY AND OPERATED ON (B)(6) 2024. FOLLOWING SURGERY, THE PATIENT WAS TRANSFERRED TO THE ICU. A BLOOD GAS ANALYSIS WAS PERFORMED AND SHOWED A BLOOD GLUCOSE LEVEL OF 23 MG/DL. IT WAS REPORTED THAT HYPOGLYCEMIA WAS SEEMINGLY NOT DETECTED DURING SURGERY DUE TO REPEATED ERRORS DURING THE PROCESSING OF THREE ARTERIAL BLOOD GAS SAMPLES IN A TOTAL OF FIVE TESTS. ALL THREE SAMPLES WERE TAKEN FROM AN ARTERIAL CATHETER IN WHICH NO MEDICATION WAS BEING ADMINISTERED THROUGH. IN THE FIVE ANALYSES PERFORMED, THE RESULTS FOR GLUCOSE, LACTATE, PH, PCO2, PO2, IONS AND HEMATOCRIT APPEAR AS INCALCULABLE. ONLY THE COOXIMETRY PARAMETERS SHOWED A RESULT. THERE WERE NO INDICATIONS REGARDING THE POSSIBLE CAUSE. THE INCIDENT WAS REPORTED BY THE CUSTOMER TO DETERMINE THE CAUSE OF THE INCALCULABLE PARAMETERS. | AN INVESTIGATION WAS COMPLETED ON NOVEMBER 1, 2024 AND DETERMINED THE INSTRUMENT WAS OPERATING TO SPECIFICATION. THIS INCIDENT IS BEING REPORTED DUE TO ADDITIONAL INFORMATION RECEIVED ON DECEMBER 4, 2024, INDICATING PATIENT DEATH. FURTHER INFORMATION IS CURRENTLY BEING OBTAINED ON PATIENT OUTCOME. INVESTIGATION OF THE CARTRIDGE DATA SHOWED THAT FOR THE SAMPLES IN QUESTION, IQM2 CORRECTLY DETECTED SAMPLES THAT WERE POTENTIALLY AFFECTED WITH A PRE-ANALYTICAL OR ANALYTICAL ISSUE AND, THEREFORE, SAMPLES RESULTS WERE CORRECTLY BLOCKED. REVIEW OF MANUFACTURING RECORDS FOR THE CARTRIDGE IN QUESTION, GEM PREMIER 5000 PAK SERIAL # (B)(6), SHOWED THAT THE CARTRIDGE PASSED ALL MANUFACTURING AND QA SPECIFICATIONS. A REVIEW OF THE COMPLAINT DATABASE REVEALS THERE IS NO TREND EXCURSION PERTAINING TO THE REPORTED FAILURE MODE. A FOLLOW-UP REPORT WILL BE FILED WHEN MORE INFORMATION IS AVAILABLE.
AVALON US TRANSDUCER | K140535 | IT WAS REPORTED THAT WHEN USING THE DEVICE TO MONITOR THE HEARTBEATS ON TWINS THERE WAS A FALSE TRACING. THE SOURCE NOTES INDICATE THAT IT WAS "FOUND OUT THERE WAS A SECOND BABY, HOWEVER BABY EXPIRED". THE DEVICE WAS SHOWING NO ISSUE WITH THE HEARTBEAT, BUT THE TRACING WAS FALSE. | PHILIPS INVESTIGATED THE COMPLAINT AND CONFIRMED THAT THIS ISSUE WAS CAUSED BY CHANGES IMPLEMENTED UNDER AN ENGINEERING CHANGE, RELEASED ON MAY 25TH, 2023. THESE CHANGES UPDATED THE ULTRASOUND TRANSDUCER WITH A NEW CPU BOARD AND FIRMWARE (FW) IN RESPONSE TO A SHORTAGE OF MULTIPLE COMPONENTS. DUE TO A CHIP SHORTAGE SITUATION, A HARDWARE (HW)/SOFTWARE (SW) DESIGN CHANGE WAS REQUIRED, WHICH INTRODUCED A PERFORMANCE ISSUE TO THE FETAL HEART RATE (FHR) MEASUREMENT OF WIRED ULTRASOUND (US) TRANSDUCERS. THE NEED FOR SUBTLE ADJUSTMENTS OF THE TRANSMIT/RECEIVE TIME WINDOWS IN US MEASUREMENTS UNEXPECTEDLY INTRODUCED INTERFERENCE WITH ADJACENT TRANSDUCERS (TWIN/TRIPLET USE CASES) BY GENERATING MEASURABLE ARTIFICIAL RHYTHMIC SIGNALS THAT CAN BE INTERPRETED AS ACTUAL PHYSIOLOGICAL SIGNALS (FETAL HEART BEATS). THIS INVESTIGATION DETERMINED THAT THE NEW FIRMWARE (453564254621-S-FW-01, REV L.01.04) WHICH IS DESIGNED FOR ALL AVALON TRANSDUCERS, SHOWS AN UNEXPECTED ISSUE FOR THE ULTRASOUND TRANSDUCER (867246). WHEN MONITORING MULTIPLES (TWINS OR TRIPLETS), THE WIRED AVALON ULTRASOUND TRANSDUCERS (PRODUCT NO. 867246) HAVE A TENDENCY TO PRODUCE AN ARTIFICIAL FETAL HEART RATE (FHR) INSTEAD OF FHR GAPS, MOSTLY AT APPROXIMATELY 180 BPM, IN SITUATIONS WHERE THE PHYSIOLOGICAL SIGNAL (ECHO FROM PULSATING FETAL HEART) IS ABSENT OR VERY WEAK. TO ADDRESS THE ISSUE, AN URGENT FIELD SAFETY NOTICE (UFSN) WAS SENT MAY 22ND, 2024, PROVIDING NOTIFICATION OF THE POTENTIAL FOR INACCURATE FETAL HEART RATE MEASUREMENTS WHEN THE AFFECTED ULTRASOUND TRANSDUCERS WERE USED WHILE MONITORING MULTIPLES (TWINS OR TRIPLETS). ON DECEMBER 3, 2024, A FIELD CHANGE ORDER (FCO) WAS IMPLEMENTED ONSITE AT NATIVADAD MEDICAL CENTER CORRECTING THE TRANSDUCER ISSUE BY UPDATING FW VERSION L.01.04 WITH FW VERSION L.01.05. ALL, BUT THREE, TRANSDUCERS WERE UPDATED. TWO OF THE TRANSDUCERS ARE CURRENTLY SEQUESTERED PENDING SITE INVESTIGATION INTO THIS ISSUE. A THIRD TRANSDUCER IS MISSING. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
EVERFLO 120V OPI | K061261 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A EVERFLO 120V OPI DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE PASSED AWAY AND THE PATIENT¿S WIFE WOULD LIKE THE LETTERS TO STOP. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
MSERIES HEATED HUMIDIFIER | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO M SERIES HEATED HUMIDIFIER DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | DUE TO A DISCREPANCY WITH THE CUSTOMER REPORTED SERIAL NUMBER OF THE HUMIDIFIER COMPARED TO THE BASE DEVICE UNITS REGISTERED TO THE CUSTOMER, THE MODEL AND SERIAL NUMBER OF THE BASE DEVICE IS UNKNOWN. ADDITIONAL INFORMATION IS NOT AVAILABLE AT THIS TIME.
LIFELINE | K013896 | IT WAS REPORTED TO DEFIBTECH THAT AN AED DISPLAYED A LOW BATTERY WARNING DURING A RESCUE ATTEMPT, FOLLOWED BY A SERVICE CODE 200 AFTER THE RESCUE. SERVICE CODE 200 IS A CODE ASSOCIATED WITH THE BATTERY VOLTAGE BEING UNREASONABLY LOW. IT WAS REPORTED THAT THE PATIENT WAS NOT RESUSCITATED. LOG FILES FROM THE AED RELATED TO THIS EVENT ARE NOT AVAILABLE; THEREFORE, A COMPREHENSIVE REVIEW OF THE DEVICE'S PERFORMANCE AND MAINTENANCE COULD NOT BE CONDUCTED. HOWEVER, GIVEN THE NATURE OF THE COMPLAINT, THIS MDR IS BEING FILED OUT OF AN ABUNDANCE OF CAUTION. | ALTHOUGH REQUESTED, THE LOG FILES FROM THE DEVICE HAVE NOT BEEN PROVIDED AND THE CAUSE OF THE COMPLAINT IS NOT KNOWN. SHOULD ADDITIONAL INFORMATION BECOME AVAILABLE, A FOLLOW-UP MDR SHALL BE SUBMITTED.
REP DREAMSTATION GO AUTO W/BT, DOM | K131982 | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF REP DREAMSTATION GO AUTO, CALLER STATING THA THE PATIENT HAS PASSED AWAY AND IS ASKING WHAT TO DO WITH THE DEVICE. ADVISED THAT IT WAS A REPLACEMENT WHICH SHE CAN DONATE IF NO LONGER NEEDED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
INFINITY ACUTE CARE SYSTEM (M540) | K113798 | IT WAS REPORTED THAT THE INFINITY ACUTE CARE SYSTEM (IACS) M540 DID NOT ALARM, AND ELECTROCARDIOGRAM (ECG) COULD NOT BE DERIVED, AS THE DEVICE DID NOT RECOGNIZE ASYSTOLE, AND THE PATIENT PASSED AWAY. | THE INVESTIGATION HAS STARTED. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION.
RENTAL, DREAMSTATION ST30, GB | K102465 | IT WAS REPORTED THAT THE DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
CLOTTRIEVER XL CATHETER | K242557 | THE PATIENT WAS A 51-YEAR-OLD FEMALE WITH LOWER LEFT EXTREMITY DEEP VEIN THROMBOSIS (DVT). AN ULTRASOUND SHOWED ACUTE DVT. NO CHEST COMPUTED TOMOGRAPHY SCAN WAS PERFORMED TO CHECK FOR PRE-EXISTING PULMONARY EMBOLISM (PE). ACCESS AND WIRE POSITIONING WERE OBTAINED WITHOUT ISSUE. VENOGRAMS REVEALED OCCLUSIVE DVT IN THE LEFT POPLITEAL WITH NO OUTFLOW TO THE INFERIOR VENA CAVA (IVC). THE RIGHT POPLITEAL WAS OPEN EXCEPT FOR A SMALL AMOUNT OF THROMBUS AT THE PROXIMAL COMMON ILIAC. THE CLOTTRIEVER SHEATH, 16 FR WAS PLACED IN THE RIGHT POPLITEAL AND THE CLOTTRIEVER SHEATH, 13 FR WAS PLACED IN THE LEFT POPLITEAL. THE PLAN WAS TO USE THE CLOTTRIEVER XL CATHETER (CTXL) TO CLEAR THE IVC AND FOR THE CLOTTRIEVER BOLD CATHETER (CT BOLD) TO CLEAR THE ILIAC ON THE RIGHT. CT BOLD WOULD THEN BE USED TO CLEAN THE LEFT AT THE LOWER LEFT EXTREMITY. THE 1ST PASS WITH THE CTXL RESULTED IN A COLLAPSED CORING ELEMENT AT THE RIGHT COMMON ILIAC. THE PATIENT NOTICED SIGNIFICANT PAIN AND THE PHYSICIAN FELT RESISTANCE. THE CORING ELEMENT WAS SLIGHTLY SHEATHED INTO THE CATHETER AND THE UNIT WAS MOVED UNTIL RESISTANCE WAS MET AGAIN. THE CORING ELEMENT WAS PULLED FURTHER INTO THE OUTER CATHETER WHICH ALLOWED THE UNIT TO BE REMOVED. A 2ND PASS WAS PERFORMED WHICH LED TO A SIMILAR COLLAPSE AT THE RIGHT COMMON ILIAC. FORTUNATELY, THE DEVICE WAS REMOVED WITHOUT ISSUE. UPON THE 3RD PASS, ANOTHER COLLAPSE OCCURRED AT THE RIGHT COMMON ILIAC. THE PATIENT BECAME VERY ANXIOUS AND DECOMPENSATED. THE CTXL WAS REMOVED AND A VENOGRAM WAS PERFORMED TO CHECK FOR VEIN AVULSION. THE VENOGRAM SHOWED NOTHING OF CONCERN AND A PE WAS SUSPECTED. THE PATIENT WAS FLIPPED ONTO HER BACK AND LOST HER PULSE SHORTLY AFTER. THE PATIENT CODED FOR 48 MINUTES BEFORE PASSING AWAY. POST-PROCEDURAL DISCUSSION REVEALED THAT THE PATIENT HAD A HISTORY OF MULTIPLE PES. ADDITIONALLY, THE PATIENT WAS MORBIDLY OBESE WITH MULTIPLE COMORBIDITIES. A TRANSESOPHAGEAL ECHOCARDIOGRAM PERFORMED DURING THE CODE REVEALED A MASSIVE RIGHT VENTRICLE AND MASSIVELY DILATED PULMONARY ARTERIES. | THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE NOT PROVIDED, THE LOT HISTORY RECORDS OF ALL POTENTIAL LOTS SHIPPED TO THE FACILITY WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS THE RESULT OF INADVERTENT CLOT EMBOLISM. DISTAL EMBOLIZATION OF BLOOD CLOTS IS IDENTIFIED IN THE DEVICE LABELING AS POSSIBLE ADVERSE EVENTS/COMPLICATIONS. THE DEVICE INSTRUCTIONS FOR USE INCLUDE THE FOLLOWING WARNING STATEMENT: "AVOID USING EXCESSIVE FORCE TO ADVANCE OR RETRACT AGAINST RESISTANCE. IF EXCESSIVE RESISTANCE IS ENCOUNTERED, RETRACT, AND COLLAPSE THE COLLECTION BAG AND CORING ELEMENT INTO THE OUTER CATHETER AND REMOVE THE DEVICE. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE DEVICE OR VESSEL." MANUFACTURER REFERENCE #: (B)(4).
DREAMSTATION GO AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. THE DEVICE WAS EVALUATED. THERE WERE NO PARTS REPLACED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER. THERE WAS ONE ERROR FOUND. DEVICE WAS SCRAPPED.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REP DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THIS ISSUE HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THIS ISSUE HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AS PART OF THE COMPLAINT HANDLING PROCESS, ATTEMPT TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION HAS BEEN UNSUCCESSFUL BECAUSE THERE IS NO CUSTOMER CONTACT INFORMATION AVAILABLE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI ASSISTED URETERAL REIMPLANTATION PROCEDURE, AN UNSPECIFIED ARTERY WAS DAMAGED AND THE PROCEDURE WAS CONVERTED TO OPEN; THE PATIENT EXPIRED IN THE OPERATING ROOM. THE PATIENT HAD A HISTORY OF RETROPERITONEAL FIBROSIS. THE SITE ROBOTICS COORDINATOR INFORMED AN INTUITIVE CLINICAL REPRESENTATIVE THAT THE PROCEDURE WAS NEAR COMPLETION WHEN THE COMPLICATION OCCURRED. IT WAS REPORTED THAT THE ARTERY WAS DAMAGED WHEN THE SURGEON WAS ATTEMPTING TO DISSECT VERY DIFFICULT ADHESIONS THAT WERE STUCK TO THE URETER. A DA VINCI GRASPING INSTRUMENT WAS USED TO CLAMP THE VESSEL TO CONTROL THE BLEEDING WHILE THE PROCEDURE WAS CONVERTED TO OPEN. THE PATIENT ULTIMATELY EXPIRED IN THE OPERATING ROOM. THE ROBOTICS COORDINATOR STATED THAT THE COMPLICATIONS WERE NOT RELATED TO THE DA VINCI PRODUCTS IN USE. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, LARGE NEEDLE DRIVER, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A BLEEDING EVENT FROM AN ARTERY DURING A DIFFICULT DISSECTION NEAR THE URETER. WHILE THE SPECIFIC VESSEL IS NOT MENTIONED, THE MOST LIKELY MAJOR VESSELS IN THIS AREA INCLUDE THE COMMON AND EXTERNAL ILIACS. OTHER POSSIBLE LARGE VESSELS MAY INCLUDE THE INTERNAL ILIACS, AORTA, AND SACRAL VESSELS. THERE WAS DIFFICULTY DURING DISSECTION DUE TO PATIENT DISEASE AND ADHESIONS WHICH MAY HAVE CONTRIBUTED TO THIS INJURY. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | IT WAS REPORTED THAT A PATIENT EXPIRED POST-OPERATIVELY AFTER A 5.5 HOUR DA VINCI-ASSISTED PARTIAL HEPATECTOMY - WEDGE RESECTION. A SURGEON INFORMED THE INTUITIVE CLINICAL SALES REPRESENTATIVE THAT THE PATIENT OUTCOME AND POST-OPERATIVE COMPLICATIONS WERE NOT RELATED TO THE DA VINCI PRODUCTS; THERE WERE NO ISSUES WITH THE DA VINCI PRODUCTS USED DURING THE PROCEDURE. ONE ASSISTING PHYSICIAN REPORTED THAT AN UNSPECIFIED VEIN WAS DAMAGED WHILE REMOVING A PORTION OF THE LIVER. THE ROBOTIC PORTION OF THE PROCEDURE WAS NOT YET COMPLETE WHEN THE INJURY OCCURRED. THE ESTIMATED BLOOD LOSS WAS UNKNOWN; BLOOD TRANSFUSIONS WERE ADMINISTERED; IT WAS UNKNOWN WHAT OTHER INTERVENTIONS WERE PERFORMED. IT WAS REPORTED THAT THE PATIENT ALSO EXPERIENCED A GAS EMBOLISM; TREATMENT FOR THE GAS EMBOLISM WAS UNKNOWN. THE PHYSICIAN DID NOT KNOW IF THE PATIENT¿S DETERIORATING POST-OPERATIVE CONDITION WAS RELATED TO THE INTRA-OPERATIVE VASCULAR INJURY, OR THE GAS EMBOLISM. ANOTHER ASSISTING SURGEON REPORTED THAT THERE WAS NO UNEXPECTED VASCULAR DAMAGE DURING THE PROCEDURE, THE BLEEDING WAS WITHIN THE EXPECTED RANGE FOR THE PROCEDURE, AND THE BLOOD TRANSFUSIONS ADMINISTERED WERE NOT DUE TO UNEXPECTED BLEEDING. THE PATIENT DEVELOPED A POST-OPERATIVE GAS EMBOLISM, BUT THE DETAILS OF THE PATIENT¿S POST-OPERATIVE COURSE WAS UNKNOWN. NO ADDITIONAL INFORMATION WAS PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND THAT NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THIS REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MONOPOLAR CURVED SCISSORS, MEDIUM-LARGE CLIP APPLIER, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, LARGE SUTURE CUT NEEDLE DRIVER, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING A PARTIAL HEPATECTOMY. AN INJURY OCCURRED WHILE EXTRACTING THE SPECIMEN. THE BLEEDING THAT OCCURRED AND THE IMPACT IT MAY HAVE HAD IS UNCLEAR, AND THERE ARE SOMEWHAT CONFLICTING ACCOUNTS AS TO THE DETAILS OF THIS BLEED. IT IS UNCLEAR HOW OR IF THAT EVENT CONTRIBUTED TO THE EVENTUAL DEMISE OF THE PATIENT. WHILE THE CAUSE OF DEATH IS ALSO NOT CONFIRMED, THERE WAS SUGGESTION THAT A GAS EMBOLISM MAY HAVE OCCURRED. GAS EMBOLISM CAN OCCUR AT THE TIME OF MINIMALLY INVASIVE SURGERY DUE TO INSUFFLATION OF CO2. THERE ARE NO DETAILS AS TO HOW A GAS EMBOLISM OCCURRED NOR THE DETAILS OF HOW IT WAS DIAGNOSED. IT IS ALSO UNKNOWN IF OR HOW THIS MAY HAVE CONTRIBUTED TO THE EVENTUAL DEATH OF THE PATIENT. THERE IS NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
WAVELINQ 4FR P | K192239 | IT WAS REPORTED THROUGH THE RESULTS OF A CLINICAL TRIAL THAT APPROXIMATELY TWO WEEKS POST-ENDOVASCULAR ARTERIOVENOUS FISTULA CREATION, THE ENDOVASCULAR ARTERIOVENOUS FISTULA WAS ALLEGEDLY FOUND TO BE NOT MATURED. IT WAS FURTHER REPORTED THAT APPROXIMATELY TWO MONTHS AND FIVE DAYS POST INDEX PROCEDURE, THE SUBJECT UNDERWENT SECOND STAGE PROCEDURE OF BALLOON ASSISTED MATURATION TO SUPPORT CANNULATION OF THE ARTERIALIZED VEIN SEGMENTS. FURTHERMORE, APPROXIMATELY SIX MONTHS AND TWENTY FOUR DAYS POST ENDOVASCULAR FISTULA CREATION PROCEDURE, THE SUBJECT HAD AN ADVERSE EVENT OF CEPHALIC VEIN STENOSIS. REPORTEDLY, THE ADVERSE EVENT WAS NOT RELATED TO STUDY DEVICE AND PROCEDURE BUT DEFINITELY RELATED TO THE ARTERIOVENOUS ACCESS CIRCUIT. EVIDENTLY, APPROXIMATELY ONE YEAR, FOUR MONTHS, AND FOURTEEN DAYS AFTER INDEX PROCEDURE, THE SUBJECT WAS EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REP DREAMSTATION AUTO BIPAP, DOM-RECRT DOM DEVICE'S EVENT. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE DEVICE HAS HEART FAILURE AND PASSED AWAY. THERE WAS NO SERIOUS PATIENT HARM OR INJURY. THE PATIENT 'S REQUIRED NO MEDICAL INTERVENTION. AT THIS TIME, NO MEDICAL INTERVENTION HAS BEEN REPORTED, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, CANCER, AND DEATH. NO FURTHER CLINICAL DETAILS OR MEDICAL INTERVENTION WERE REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION REGARDING A REMSTAR PRO C-FLEX+ DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. THE MANUFACTURER RECEIVED NEW AND RELEVANT INFORMATION ON 02/14/2024. THE DEVICE WAS RETURNED TO THE SERVICE CENTER FOR EVALUATION. DURING THE EVALUATION OF THE DEVICE, THERE WAS NO VISIBLE FOAM PARTICLES FOUND. DUST CONTAMINATION FOUND DURING THE DEVICE EVALUATION. THE DEVICE WAS REPAIRED. SECTION G3 HAS BEEN UPDATED. IN ADDITION, APPROPRIATE CODE UPDATED/CORRECTED.
BIPAP AUTOSV ADV | K090539 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF BIPAP AUTOSV ADV. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF BIPAP AUTOSV ADV. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTRE FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF VISIBLE FOAM PARTICLES. THE TECHNICIAN CONFIRMED NO PARTICLES IN THE AIR PATH. DURING THE EVALUATION, SECONDARY FINDINGS WAS OBSERVED. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTRE CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND UNIT SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
REMSTAR PRO C-FLEX | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REMSTAR PRO C-FLEX. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF REMSTAR PRO C-FLEX. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DUODOPA_DUOPA | K142793 | ON AN UNKNOWN DATE A PATIENT IN SPAIN UNDERWENT A PROCEDURE FOR THE PLACEMENT OF A PERCUTANEOUS ENDOSCOPIC GASTROSTOMY(PEG) TUBE. ON (B)(6) 2024, IT WAS REPORTED THE PATIENT HAD PERITONITIS DUE TO THE TRANSLOCATED PEG TUBE THAT WAS COMPLICATED WITH PNEUMONIA AND PATIENT PASSED AWAY 2 YEARS AGO. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. DESPITE MULTIPLE QUERIES, THERE IS NO CLEAR EVIDENCE LINKING THE DEVICE TO THE REPORTED EVENT, AND NO ADDITIONAL INFORMATION REGARDING DEVICE-RELATED ALLEGATIONS, MALFUNCTIONS, COMPLICATIONS, DIAGNOSTIC TESTING, OR TREATMENT IS AVAILABLE. THEREFORE, OUT OF ABUNDANCE OF CAUTION, ABBVIE IS CONSERVATIVELY REPORTING THE EVENT WITHOUT ATTRIBUTING CAUSALITY TO THE DEVICE. | REFERENCE NUMBER (B)(4). THE DEVICE MANUFACTURER AND LOT NUMBER OF THE DEVICE INVOLVED IN THIS COMPLAINT WAS NOT PROVIDED. THEREFORE, IT IS UNKNOWN IF THE DEVICE INVOLVED WAS ABBVIE BRANDED TUBING. CONSERVATIVELY, ABBVIE HAS CHOSEN TO REPORT THIS COMPLAINT DUE TO THE POTENTIAL THAT THE DEVICE INVOLVED COULD HAVE BEEN ABBVIE BRANDED TUBING. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
DS2ADV AUTO CPAP W/HUMID CELL/BT | K200480 | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF DS2ADV AUTO CPAP, PATIENT'S WIFE CALLED IN STATING THAT THE PATIENT HAS PASSED AWAY AND IS ASKING FOR A RETURN LABEL AS SHE RECEIVED A REPLACEMENT DEVICE BUT DOES NOT NEED ONE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER INQUIRY ABOUT THE RETURN OF A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON A COMPENSATION CHECK ISSUED TO HER HUSBAND WHO IS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION CPAP PRO DEVICE. THERE WAS A REPORT OF PATIENT PASSING AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PANCREATIC AND PROSTATE CANCER, CORONARY ARTERY DISEASE, ACUTE HYPOXIC RESPIRATORY FAILURE WITH ASSOCIATED ALTERED MENTAL STATUS, SEPTIC SHOCK AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REACT CATHETER | K180715 | LITERATURE WAS REVIEWED REGARDING "EFFICACY AND SAFETY OF A NEW DELIVERY ASSIST CATHETER WITH A FLEXIBLE, SPINDLE-SHAPED SHAFT IN MECHANICAL THROMBECTOMY". THE PURPOSE OF THIS STUDY WAS TO EVALUATE WHETHER NAVIGATING AN ASPIRATION CATHETER TOWARD THE OCCLUSION SITE COULD BE DONE SAFELY, STEADILY, AND RAPIDLY USING A DELIVERY ASSIST INNER CATHETER. THE TIME FRAME OF THIS STUDY WAS JANUARY 2018 AND MAY 2022. DEVICE MANUFACTURERS: MULTIPLE MANUFACTURER¿S DEVICES WERE USED INTHE STUDY POPULATION. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 68 ASPIRATION CATHETER AND SOLITAIRE STENT RETRIEVER SIX DEATHS OCCURRED IN THE STUDY POPULATION. THE CAUSE OF DEATH WAS NOT NOTED AMONG PATIENT ADVERSE EVENTS INCLUDED: -1 ARTERY DISSECTION -8 SYMPTOMATIC INTRACRANIAL HEMORRHAGES -MEDIAN MRS AT 90 DAYS WAS 2, RANGE 1-4 NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CONTOUR 3D | K101212 | IT WAS REPORTED THAT A POSTOPERATIVE DEATH OCCURRED FOLLOWING THE IMPLANTATION OF A 30MM CONTOUR 3D TRICUSPID ANNULOPLASTY RING. THE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO ALLEGED PRODUCT ISSUES REPORTED. THE DEVICE REMAINS IN THE PATIENT. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CARDIOROOT WOVEN | K103347 | COMPLAINT # (B)(4). | CORRECTED DATA: IN BLOCK D4, THE SUSPECTED MEDICAL DEVICE INFORMATION HAS BEEN UPDATED AFTER A THOROUGH PRODUCT TRACEABILITY REVIEW TO IDENTIFY THE INVOLVED CARDIOROOT, AS DETAILED BELOW. ADDITIONAL MANUFACTURER NARRATIVE: (4115) THE INVOLVED DEVICE WAS EXPLANTED AND DISCARDED. (4112/213) THE MANUFACTURER CONDUCTED A THOROUGH PRODUCT TRACEABILITY INVESTIGATION TO IDENTIFY THE INVOLVED CARDIOROOT WITH MODEL HEWROOT0028. THE GETINGE SALES OPERATIONS MANAGER (FRENCH SUBSIDIARY) CONFIRMED THAT BETWEEN JAN 2022 AND MAY 2023, ONLY ONE CARDIOROOT WITH MODEL HEWROOT0030, SERIAL NUMBER (B)(6) AND LOT NUMBER 21F10 WAS IMPLANTED AT THE INVOLVED HOSPITAL ON (B)(6) 2022. NOTE THAT THERE WAS NO CARDIOROOT CHARGED FOR THIS HOSPITAL IN 2023; GRAFTS ARE CHARGED TO HOSPITALS ONCE THEY ARE IMPLANTED. THEREFORE, IT IS NOT EXCLUDED THAT THE INITIAL INFORMATION REGARDING CARDIOROOT REFERENCE (HEWROOT0028) PROVIDED BY THE HOSPITAL REGARDING THE COMPLAINT IS NOT ENTIRELY ACCURATE. (4121/213) THE DEVICE HISTORY REVIEW (DHR) OF THE HYPOTHETICALLY INVOLVED CARDIOROOT HEWROOT0030 (SN (B)(6) AND LOT# 21F10) WAS PERFORMED BY THE QUALITY ASSURANCE SUPERVISOR IN ORDER TO DETECT ANY RELEVANT INFORMATION THAT MAY HAVE IMPACTED THE PRODUCT AND BE RELATED TO THE COMPLAINT. IT WAS CONCLUDED THAT NO DEVIATION WAS IDENTIFIED IN RELATION WITH THE REPORTED EVENT. (4109/213) THE REVIEW OF HISTORICAL DATA INDICATED THAT NO OTHER SIMILAR COMPLAINT WAS REPORTED FOR THE SAME STERILIZATION LOT NUMBER 21F10. (4110/213) THE OCCURRENCE RATE WAS CALCULATED FOR THE SIMILAR EVENTS REPORTED (INCLUDING THIS COMPLAINT) FOR GRAFTS AND PATCHES ON THE SAME MANUFACTURING LINE (INTERGARD / HEMAGARD) FOR 1-YEAR PERIOD FROM 01-JAN-24 TO 31-JAN-25. THE OCCURRENCE RATE IS (B)(4), WHICH IS BELOW THE ANTICIPATED OCCURRENCE RATE (MAXIMUM) OF (B)(4) IN THE PRODUCT RISK ASSESSMENT. (4112/4315) BASED ON THE LIMITED PROVIDED INFORMATION AND THE MEDICAL REVIEW, NO CONCLUSION CAN BE DRAWN ON THE EXACT ORIGIN OF THE ADVERSE EVENT AND THE INVOLVEMENT OF THE DEVICE IN THE SERIOUS INCIDENT. AS CONCLUDED IN THE MEDICAL REVIEW, DUE TO THE LACK OF ANY INFORMATION REGARDING THE PATIENT, OR THE PROCEDURE, IT IS NOT POSSIBLE TO COMPLETE A MEDICAL ASSESSMENT OF THE EVENT. MOREOVER, THE CONDUCTED INVESTIGATION ON THE HYPOTHETICALLY INVOLVED CARDIOROOT HEWROOT0030 SUGGESTS THAT THE PRODUCT WAS NOT DEFECTIVE AT THE TIME OF MANUFACTURING.
EVIS EXERA II ULTRASOUND GASTROVIDEOSCOPE | K093395 | NO ADDITIONAL INFORMATION WAS RECEIVED. | THIS REPORT IS BEING SUPPLEMENTED TO PROVIDE ADDITIONAL INFORMATION BASED ON THE APPROVED FINAL INVESTIGATION. UPDATED FIELDS: D8, H6. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE IS NO REPORT OF MEDICAL INTERVENTION BEING REQUIRED. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE MANUFACTURER SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE MANUFACTURER SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.
TFNA FEM NAIL Ø12 R 130° L400 TIMO15 | K160167 | IT WAS REPORTED THAT ON (B)(6) 2024, THE PATIENT UNDERWENT AN UNKNOWN SURGERY WITH THE PRODUCT IN QUESTION. THE SURGERY WAS COMPLETED SUCCESSFULLY WITH NO SURGICAL DELAY. THE MORNING AFTER THE DAY OF THE SURGERY, THE PATIENT DIED. THE CAUSAL RELATIONSHIP WITH THE PRODUCT IN QUESTION IS UNKNOWN. | PRODUCT COMPLAINT # (B)(4). DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
SENTRANT INTRODUCER SHEATH | K123990 | A SENTRANT SHEATH (SENSH1428W) WAS INTENDED TO BE USED AS A CONDUIT DURING A TAVI PROCEDURE. IT WAS NOTED THAT THE SENTRANT SHEATH WAS FREE FROM ANY DEFECTS. THE PATIENT PRESENTED WITH HYPOTENSION AND PERICARDIAL EFFUSION IT WAS REPORTED DURING THE INDEX PROCEDURE, THAT THE MEDICAL TEAM INADVERTENTLY PERFORATED THE LEFT VENTRICLE WHILE POSITIONING THE GUIDE WITHIN THE VENTRICLE. DESPITE EFFORTS, INCLUDING CARDIORESPIRATORY RESUSCITATION AND PERICARDIAL DRAINAGE, THE INTERVENTIONS WERE UNSUCCESSFUL, AND THE PATIENT EXPIRED DURING THE INDEX PROCEDURE. PER THE PHYSICIAN THE CAUSE OF THE PERFORATION AND SUBSEQUENT DEATH WAS DUE TO USER ERROR. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
MITRACLIP | K190167 | IT WAS REPORTED THAT STEERABLE GUIDE CATHETER (SGC) WAS PREPPED AND INSERTED PER INSTRUCTIONS FOR USE (IFU). UPON REMOVAL OF THE DILATOR, IT WAS OBSERVED THAT THE PATIENT'S BLOOD PRESSURE WAS LOW. PATIENT WAS NOTED TO BE BRADYCARDIC AND HYPOTENSIVE. CPR WAS PERFORMED. THE SGC WAS REMOVED FROM THE LEFT ATRIUM AND BACK INTO THE INFERIOR VENA CAVA. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED DEATH AND HYPOTENSION. DEATH AND HYPOTENSION ARE LISTED IN THE INSTRUCTIONS FOR USE IS A KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION WAS A RESULT OF CASE SPECIFIC CIRCUMSTANCE AS CPR WAS PERFORMED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
VIVO 45LS | K193586 | 12/11/2024: BREAS RECEIVED FROM EASTERN MEDTECH NOTIFYING BREAS THAT A 'PT PASSED USING OUR VENTILATOR AND THE VENT MAY HAVE BEEN TAKEN BY A HAZMAT TEAM OR ANOTHER DME COMPANY. BREAS REACHED OUT FOR MORE INFORMATION AND WAS PROVIDED THE FOLLOWING ON 12/23/2024: WE HAD A PATIENT WHO PASSED AWAY. HE WAS USING THE VENT WHEN HE PASSED AWAY ON (B)(6) 2024 AND HIS BODY WAS NOT DISCOVERED UNTIL 2-3 DAYS LATER. WE WERE TOLD THAT THE HAZMAT TEAM TOOK THE VENT DUE TO IT BEING "HAZARDOUS" BECAUSE HE WAS HOOKED UP TO IT WHEN HE PASSED. WE HAVE CALLED MANY PLACES TRYING TO TRACK DOWN WHO TOOK THE VENT FROM THE HOUSE AND WHERE IT COULD POSSIBLY BE. THE PATIENT'S UNCLE IS IN CHARGE OF EVERYTHING AND THOUGHT THAT ANOTHER COMPANY TOOK IT THINKING IT WAS THEIRS. THEY SAID THEY LEFT IT BECAUSE IT HAD OUR NAME ON IT. WE WERE THEN TOLD THAT THE FUNERAL HOME TOOK IT. I CALLED THERE AND THEY SAID THAT THEY DO NOT TOUCH THAT TYPE OF THING AND IT WAS MORE LIKELY THAT THE HAZMAT TEAM TOOK IT. WE CALLED THE POLICE DEPARTMENT WHICH UNFORTUNATELY WAS ANOTHER DEAD END. AT THIS POINT, WE ARE NOT SURE WHERE IT IS OR WHAT HAPPENED TO IT. BREAS HAS NOT RECEIVED ANY ADDITIONAL INFORMATION TO DATE. THE VENTILATOR IS NO LONGER IN POSSESSION OF EASTERN MEDTECH. THE VENTILATOR HAS NOT BEEN RETURNED TO BREAS FOR INVESTIGATION. | DEVICE WAS TAKEN INTO POSSESSION BY A THIRD PARTY WHICH IS UNKNOWN AT THIS TIME. LOG FILES HAVE NOT BEEN MADE AVAILABLE TO THE MANUFACTURER AT THIS TIME. THE US AGENT HAS MADE MULTIPLE ATTEMPTS TO REQUEST ADDITIONAL INFORMATION AND TO OBTAIN POSSESSION OF THE DEVICE TO NO AVAIL.
DUODOPA_DUOPA | K142816 | ON (B)(6) 2024 A PATIENT IN SPAIN UNDERWENT A PROCEDURE FOR THE PLACEMENT OF PERCUTANEOUS ENDOSCOPIC GASTROSTOMY-JEJUNAL (PEG-J) TUBES, APPARENTLY WITHOUT COMPLICATIONS AND RETURNED HOME. AT HOME, THE PATIENT STARTED EXPERIENCING BLOCKAGES, THE INTESTINAL TUBE WAS REMOVED, PROMPTING A RETURN TO THE HOSPITAL ON (B)(6) 2024. IN THE HOSPITAL, AN ABDOMINAL CT SCAN REVEALED INTESTINAL PERFORATION, RESULTING IN PERITONITIS AND SEPSIS, BOTH WITH A POOR PROGNOSIS. BY (B)(6) 2024, THE PATIENT DEVELOPED A STOMACH INFECTION AND ABUNDANT DRAINAGE THROUGH THE STOMA. AFTER THE CONTINUOUS INFUSION PUMP WAS REMOVED, COMPLICATIONS PERSISTED, LEADING TO ALL PARKINSON'S TREATMENTS BEING WITHDRAWN AND REPLACED WITH MORPHIC CHLORIDE. THE PATIENT PASSED AWAY ON (B)(6) 2024. IT IS UNKNOWN IF AN AUTOPSY WAS PERFORMED. DESPITE MULTIPLE QUERIES, THERE IS NO CLEAR EVIDENCE LINKING THE DEVICE TO THE REPORTED EVENT, AND NO ADDITIONAL INFORMATION REGARDING DEVICE-RELATED ALLEGATIONS, MALFUNCTIONS, COMPLICATIONS, DIAGNOSTIC TESTING, OR TREATMENT IS AVAILABLE. THEREFORE, OUT OF ABUNDANCE OF CAUTION, ABBVIE IS CONSERVATIVELY REPORTING THE EVENT WITHOUT ATTRIBUTING CAUSALITY TO THE DEVICE. | REFERENCE NUMBER (B)(4). THE DEVICE MANUFACTURER AND LOT NUMBER OF THE DEVICE INVOLVED IN THIS COMPLAINT WAS NOT PROVIDED. THEREFORE, IT IS UNKNOWN IF THE DEVICE INVOLVED WAS ABBVIE BRANDED TUBING. CONSERVATIVELY, ABBVIE HAS CHOSEN TO REPORT THIS COMPLAINT DUE TO THE POTENTIAL THAT THE DEVICE INVOLVED COULD HAVE BEEN ABBVIE BRANDED TUBING. THE DEVICE INVOLVED IN THE EVENT WAS NOT RETURNED; THEREFORE, A RETURN SAMPLE EVALUATION IS UNABLE TO BE PERFORMED. POTENTIAL POST PROCEDURAL COMPLICATION IS A KNOWN COMPLICATION OF A PEG- J TUBE PLACEMENT. H6 CODE OF E2402 WAS CHOSEN TO CAPTURE THE EVENT OF POSSIBLE POST PROCEDURAL COMPLICATION. IF ANY FURTHER RELEVANT INFORMATION IS IDENTIFIED OR OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
GST BLUE RELOAD 45MM 6 ROW | K163454 | IT WAS REPORTED THAT DURING A BARIATRIC SURGERY, THE PATIENT HAD DEHISCENCE OF THE VERTICAL STAPLE LINE OF THE GASTRIC POUCH IN THE SECOND POST-OPERATIVE DAY. HE DIED. | (B)(4). DATE SENT 1/6/2025. D4 BATCH # UNK. B3 EXACT DATE IS UNK. ASSUMED FIRST DAY OF THE MONTH. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE LOT#722C37 HISTORY RECORDS WERE REVIEWED AND CERTIFIED BY EXTERNAL MANUFACTURING THAT THE MANUFACTURING CRITERIA WERE MET PRIOR TO THE RELEASE OF THE EQUIPMENT. THE CERTIFICATE RECORDS ARE ACCESSIBLE THROUGH EXTERNAL MANUFACTURING. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT IS THE TIMELINE OF THE EVENT? WHAT DEVICE WAS USED WITH THE RELOAD? ANY DIFFICULTY WITH THE DEVICE INTRAOPERATIVELY? ANY ISSUE NOTED WITH STAPLE FORMATION THROUGHOUT THE PROCEDURE? WHAT WAS THE SITE OF THE LEAK? WAS THE SITE OF THE DEHISCENCE FAR SUPERIOR ON THE STAPLE LINE? WHAT WAS THE CAUSE OF DEATH? IS IT RELATED TO THE INITIAL SURGERY? DOES THE SURGEON BELIEVE THAT THERE IS AN ALLEGED DEFICIENCY WITH THE ETHICON DEVICE OR WERE THERE OTHER CONTRIBUTING PATIENT FACTORS? IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON PERSONNEL REGARDING THE EVENT? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
BIPAP S/T C-SERIES | K102465 | THE MANUFACTURER RECEIVED INFORMATION FROM A FAMILY FRIEND INFORMING PHILIPS THAT THE PATIENT HAD A NEW BIPAP S/T C SERIES DEVICE SINCE THE ONE THEY HAD WAS SENT IN TO BE FIXED. THE FAMILY FRIEND WANTED TO KNOW IF THE NEW DEVICE WAS REAL OR FAKE BECAUSE THEY ALLEGED THE NEW DEVICE GIVEN TO THE PATIENT HAD CAUSED THE PATIENT TO BECOME "VERY SICK USING THE NEW DEVICE AND WAS IN AND OUT OF THE HOSPITAL UNTIL HE PASSED AWAY." THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION FROM A FAMILY FRIEND INFORMING PHILIPS THAT THE PATIENT HAD A NEW BIPAP S/T C SERIES DEVICE SINCE THE ONE THEY HAD WAS SENT IN TO BE FIXED. THE FAMILY FRIEND WANTED TO KNOW IF THE NEW DEVICE WAS REAL OR FAKE BECAUSE THEY ALLEGED THE NEW DEVICE GIVEN TO THE PATIENT HAD CAUSED THE PATIENT TO BECOME "VERY SICK USING THE NEW DEVICE AND WAS IN AND OUT OF THE HOSPITAL UNTIL HE PASSED AWAY." REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
CEREPAK FREEFORM | K220040 | IT WAS REPORTED, VIA A HEALTHCARE PROFESSIONAL, THAT AN UNKNOWN CEREPAK COIL (PRODUCT CODE/LOT # UNKNOWN) WAS USED FOR AN ENDOVASCULAR EMBOLIZATION PROCEDURE. DURING THE PROCEDURE, THE USER REPORTED THE COIL FAILED TO DETACH. THE EVENT RESULTED IN THE PHYSICIAN USING A BALLOON GUIDE CATHETER (UNKNOWN BRAND) TO ARREST BLOOD FLOW AND REMOVE THE COIL. COMPETITOR COILS WERE USED TO COMPLETE THE SURGERY. THE REPORTED BELIEVED THE ¿PATIENT DID END UP RE RUPTURING DURING THE CASE.¿ THE EVENT WAS REPORTED AS SUCH, ¿SURGEON CALLED ME ON SATURDAY 12-14-24 AT 4:22PM EST TO LET ME KNOW HE WOULD BE USING A CEREPAK FREEFORM MINI COIL TO TREAT A RUPTURED ACOM ANEURYSM. HE MENTIONED OVER THE PHONE HE DID NOT HAVING ANY DETACHERS AVAILABLE AND THAT HE WOULD NEED TO USE MANUAL BREAK AS A MECHANISM OF DETACHMENT. DR. CALLED ME BACK AT 5:22PM TO LET ME KNOW HE WAS HAVING AN ISSUE, AND HE COULDN¿T GET THE COIL DETACHED. ONCE THE COIL WAS INSIDE THE DOME OF THE ANEURYSM, THE MANUAL BREAK DETACHMENT FAILED. WITH THE COIL ALREADY IN THE ANEURYSM PROTECTING THE RUPTURE SITE HE WAS TRYING TO AVOID TAKING THE COIL OUT BECAUSE THE PATIENT COULD END UP RE RUPTURING. DR. OPTED TO PUT UP A BALLOON AND TAKE THE COIL OUT AND PUT UP ANOTHER COIL FROM ANOTHER VENDOR. I BELIEVE THE PATIENT DID END UP RE RUPTURING DURING THE CASE. WE ARE UNAWARE OF THE SKU#¿S/LOTS OF COILS USED ON SATURDAY NIGHT AND WE ARE UNSURE IF THE OPERATOR/STAFF SAVED THE DEVICE IN QUESTION. I WAS NOT PRESENT AT THE CASE.¿ NO FURTHER INFORMATION WAS MADE AVAILABLE AT THE TIME OF THIS REVIEW. | MANUFACTURER'S REF. NO: (B)(4). SECTION B5: PER THE ADDITIONAL INFORMATION RECEIVED ON 06-MAR-2025, IT WAS DISCLOSED THAT A LOOP OF THE COIL LANDED IN THE SUBARACHNOID SPACE, RUPTURED THE NECK OF THE ANEURYSM, AND RESULTED IN AN ACUTE LEAK AND SUBARACHNOID BLEEDING. THE PATIENT SUBSEQUENTLY DIED AFTER THE PROCEDURE DUE TO SUBARACHNOID HEMORRHAGE. ALTHOUGH THE PATIENT HAD SIGNIFICANT HEMORRHAGING PRIOR TO THE PROCEDURE WITH POOR PROGNOSIS (HUNT & HESS 4), THE CORRELATION BETWEEN THE ALLEGED DEVICE DEFICIENCY AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT. BASED ON THIS INFORMATION, THIS EVENT WILL BE CONSERVATIVELY REPORTED TO THE US FDA REPORTING UNDER CRITERIA 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH,¿ WITH AN AWARENESS DATE OF 26-FEB-2025. THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED TO IDENTIFY STATISTICAL SIGNALS FOR CONSIDERATION OF FURTHER CORRECTION ACTION. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT.
PHOENIX/INNOVA MACHINES | K070643 | IT WAS REPORTED THAT DURING HEMODIALYSIS THERAPY WITH A PHOENIX MACHINE, "WHILE DRAWING BLOOD", THE PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. NO ADDITIONAL INFORMATION IS AVAILABLE. | H11: THE DEVICE WAS NOT RETURNED AND THE LOT NUMBER IS UNKNOWN; THEREFORE, A DEVICE ANALYSIS COULD NOT BE COMPLETED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
INTELLIVUE MX450 PATIENT MONITOR | K130849 | IT WAS REPORTED THE ECG LEADS WERE OFF AND THE PATIENT PASSED AWAY. THE STAFF WANTED TO CHECK THE FUNCTIONALITY OF THE MONITOR. THE CLINICAL AUDIT TRAIL WAS REVIEWED WITH THE CUSTOMER, REVEALING THE FIRST INDICATION OF ECG LEADS OFF WAS AT 00:37. THE RED ECG LEADS OFF ENDED AND REPEATED NUMBERS TIMES UNTIL 03:27. AT THAT TIME, ECG LEADS OFF PERSISTED UNTIL 06:16 WHEN THE RED ECG LEADS OFF ALARM ENDED. OVER THE COURSE OF THAT SAME TIME PERIOD, THE RED ECG LEADS OFF REMINDERS WERE SILENCED NUMEROUS TIMES. THE PATIENT PASSED AWAY AT APPROXIMATELY 06:00. | A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED. THE FSE TESTED AND VERIFIED DEVICE FUNCTION. THE FSE FOUND THAT BOTH THE SURVEILLANCE AND THE OVERVIEW HAD AUDIO WHICH WAS SET AT VOLUME LEVEL OF 4. THE PATIENT PASSED AWAY AT APPROXIMATELY 06:00 ON (B)(6) 2024 IN BED M510. THE CLINICAL AUDIT TRAIL WAS REVIEWED WITH THE CUSTOMER, REVEALING THE FIRST INDICATION OF ECG LEADS OFF WAS AT 00:37. THE RED ECG LEADS OFF ENDED AND REPEATED NUMEROUS TIMES UNTIL 03:27. AT THAT TIME, ECG LEADS OFF PERSISTED UNTIL 06:16 WHEN THE RED ECG LEADS OFF ALARM ENDED. OVER THE COURSE OF THAT SAME TIME PERIOD, THE RED ECG LEADS OFF REMINDERS WERE SILENCED NUMEROUS TIMES FROM THE UNIT OVERVIEW HOST (B)(6). THE REPORTED PROBLEM WAS NOT CONFIRMED. THE INVESTIGATION VERIFIED THAT THE ECG LEADS WERE OFF PRIOR TO THE PATIENT EXPIRATION; ECG COULD NOT BE MEASURED ON THE PATIENT. THE DEVICE PROVIDED INOPS FOR THIS CONDITION. THERE WAS NO MALFUNCTION OF THE PHILIPS DEVICES. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY PROCEDURE, THE PATIENT EXPIRED. THE PROCEDURE DATE WAS UNKNOWN, AND THE CAUSE OF DEATH WAS NOT PROVIDED. THE DA VINCI PROCEDURE WAS CONVERTED TO OPEN; THE SITE OPERATING ROOM MANAGER REPORTED THAT THE CONVERSION WAS NOT RELATED TO A DEVICE ISSUE OR A DA VINCI PRODUCT MALFUNCTION. THE REASON FOR THE CONVERT TO OPEN WAS NOT PROVIDED. ADDITIONAL INFORMATION HAS BEEN REQUESTED; NO RESPONSE HAS BEEN RECEIVED. | NO SPECIFIC INFORMATION REGARDING THE PROCEDURE DATE WAS PROVIDED; THEREFORE, NO DA VINCI SYSTEM OR ACCESSORIES LOG FILES OR DEVICE HISTORY RECORD ARE AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A ROBOTIC CHOLECYSTECTOMY WHICH REQUIRED CONVERSION TO AN OPEN PROCEDURE. NO INFORMATION AS TO THE REASONS FOR THE CONVERSION WERE PROVIDED. THE PATIENT SUBSEQUENTLY DIED BUT NO DETAILS AS TO THE EVENTS THAT LED TO THEIR DEATH OR HOW THEY DIED WERE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI SI | K081137 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED RIGHT PULMONARY LOBECTOMY PROCEDURE, THE PATIENT EXPERIENCED BLEEDING AND THE PROCEDURE WAS CONVERTED TO OPEN. THE PATIENT EXPIRED THE SAME DAY AS THE PROCEDURE. IT IS UNKNOWN IF THE PATIENT DIED INTRA-OPERATIVELY, OR POST-OPERATIVELY. IT WAS REPORTED THAT VASCULAR RECONSTRUCTION WAS PERFORMED TO CONTROL THE INTRA-OPERATIVE BLEEDING. THE CAUSE AND SOURCE OF THE BLEEDING WAS NOT KNOWN TO THE REPORTER. THE AMOUNT OF BLOOD LOSS WAS LESS THAN 500ML; THE PATIENT WAS ADMINISTERED UNSPECIFIED BLOOD PRODUCT(S). IT WAS REPORTED THAT THERE WAS NO DA VINCI MALFUNCTION PRIOR TO THE CONVERT TO OPEN. ATTEMPTS WERE MADE TO SPEAK WITH THE SURGEON AND HAVE BEEN SO FAR UNSUCCESSFUL. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD LIKELY HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) INVOLVED WITH THIS REPORTED EVENT AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A ROBOTIC PULMONARY LOBECTOMY AND ULTIMATELY DIED THE SAME DAY. BLEEDING OCCURRED DURING THE CASE. THE DETAILS OF HOW THAT BLEEDING OCCURRED, THE FACTORS THAT MAY HAVE CONTRIBUTED TO THAT EVENT AND HOW THE BLEEDING WAS ADDRESSED, AND HOW THIS EVENT MAY HAVE CONTRIBUTED TO THE DEATH OF THE PATIENT WERE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), CANCER, OTHER (UNSPECIFIED EVENT) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
CARESCAPE | K210384 | THE HOSPITAL REPORTED A PATIENT CONNECTED TO A CARESCAPE R860 WHEN IT WAS ALLEGED THAT THE UNIT ALARMED FOR LOW TIDAL VOLUME. REPORTEDLY, THE TIDAL VOLUME WAS 120ML, BUT WAS SET FOR 450ML. MANUAL VENTILATION WAS INITIATED CONTINUOUSLY UNTIL IT EXCEEDED THE ALARM UNIT. IT WAS REPORTED THAT RESUSCITATION EFFORTS FAILED AND THE PATIENT DIED. GE HEALTHCARE WILL SUBMIT A FOLLOW-UP REPORT WHEN THE INVESTIGATION HAS BEEN COMPLETED. | GE HEALTHCARE'S (GEHC) INVESTIGATION HAS BEEN COMPLETED AND THE ROOT CAUSE COULD NOT BE DETERMINED. BASED ON THE REVIEW OF THE DEVICE LOGS, PART INSPECTION, LAB TESTING, THE ROOT CAUSE OF THE FAULTY EXPIRATORY FLOW SENSOR WAS DETERMINED TO BE CONTAMINATION DUE TO THE USER NOT USING AN EXPIRATORY FILTER. THIS WOULD ONLY AFFECT THE DISPLAYED VALUES AND NOT DELIVERY. THE ROOT CAUSE OF THE PATIENT DEATH IS UNDETERMINED.
EDWARDS ESHEATH+ INTRODUCER SET | K200258 | AS REPORTED FROM A VOLUNTARY MEDWATCH, TAVR PROCEDURE LEFT COMMON CAROTID ARTERY CUTDOWN PERFORMED AND A 6 FR SHEATH WAS PLACED. THE AORTIC VALVE WAS THEN CROSSED UTILIZING AN AL1 CATHETER AND A STRAIGHT TIP AMPLATZ WIRE. THIS WAS THEN EXCHANGED FOR A PRECURVED SAFARI WIRE UTILIZING A PIGTAIL CATHETER. SUBSEQUENTLY, A 29 MM EDWARDS S3 VALVE WAS DEPLOYED UNDER RAPID VENTRICULAR PACING. TRANSESOPHAGEAL ECHOCARDIOGRAPHY DEMONSTRATED NO PARAVALVULAR LEAK OR AORTIC REGURGITATION. NEXT THE TAVR DELIVERY SYSTEM WAS REMOVED FROM THE LEFT COMMON CAROTID ARTERY. DUE TO A SPLIT SHEATH WITH THE DELIVERY SYSTEM UNABLE TO BE WITHDRAWN COMPLETELY INTO THE SHEATH THIS WAS REMOVED AS ONE PIECE. THIS THEN RESULTED IN FURTHER DISRUPTION OF THE CAROTID ARTERIOTOMY. A 10X20 MM BALLOON WAS INFLATED IN THE PROXIMAL LEFT COMMON CAROTID ARTERY TO OBTAIN PROXIMAL CONTROL. DUE TO THE COMPLEXITY OF THE REPAIR WITH CONCERNS REGARDING FURTHER LOSS OF HEMOSTASIS AND THAT THE PATIENT HAD NO EVIDENCE OF IMPAIRED SIGNAL ON NEURO MONITORING, THE CAROTID ARTERY WAS LIGATED BELOW THE CLAVICLE. THE PATIENT WAS TRANSFERRED TO THE ICU FOR FURTHER CARE. IT WAS REPORTED THAT THE PATIENT WAS MOVING ALL EXTREMITIES IMMEDIATELY POST TAVR. A HEAD CT WAS PERFORMED THE DAY FOLLOWING SURGERY AND DID NOT SHOW OBVIOUS SIGNS OF STROKE, HOWEVER A REPEAT CT WAS PERFORMED 12 HOURS LATER AND SHOWED THE PATIENT HAD SUFFERED A LEFT HEMISPHERIC STROKE. THE PATIENT REMAINED INTUBATED AND SEDATED AND ALSO WENT IN TO RENAL FAILURE REQUIRING CRRT. NEUROSURGERY CONSULTED ON THE PATIENT AND DETERMINED THERE WAS NOTHING THEY COULD OFFER THE PATIENT. THE PATIENT'S FAMILY ESSENTIALLY MADE THE PATIENT A DNR AND HE WAS ALLOWED TO EXPIRE COMFORTABLY. | THE INVESTIGATION IS IN PROGRESS. A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION. THIS IS ONE OF TWO MANUFACTURER REPORTS BEING SUBMITTED FOR THIS CASE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGATION OF ASTHMA (NEW OR WORSENING); LIVER DISEASE/TOXICITY; LUNG DISEASE; DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION STATING THAT THE PATIENT HAS PASSED AWAY. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: EVALUATION METHOD CODE, EVALUATION RESULTS CODE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | K200480 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE PASSED AWAY AND THE PATIENT¿S DAUGHTER IS UNABLE TO RETURN THE AFFECTED DEVICE BECAUSE THEY NO LONGER HAVE IT. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER BECAME DUE TO A USER OF DREAMSTATION CPAP. DAUGHTER CALLED IN STATING PATIENT PASSED AWAY LAST (B)(6) 2023. DID NOT PASS DUE TO THE USE OF CPAP. SHE WANTS TO STOP ALL MAIL. DON'T KNOW WHERE THE CPAP DEVICE WENT THAT WAS RECALLED AND WILL NOT BE RETURNING FOR EVALUATION. REPLACEMENT DEVICE WAS DONATED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2024. PER THE CONTACT, THE CAUSE OF DEATH IS UNKNOWN, BUT ALLEGED THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE EVENT. NO ADDITIONAL INFORMATION WAS PROVIDED. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
RECERT SIMPLYGO MINI | K111885 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE RECERT SIMPLYGO MINI HAD PASSED AWAY AND THE DEVICE WAS ELIGIBLE FOR BUYBACK. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
OPTIFLUX F180NRE HF 12/CS 1.7SA EBEAM | K162488 | ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM 29/DEC/2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER. | ADDITIONAL INFORMATION: A1, A2, A3, A4, B5, D10, G2, H6, H10. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HD THERAPY UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, COMBI-SET BLOODLINES AND THE SERIOUS ADVERSE EVENTS OF HYPOTENSION, LOSS OF CONSCIOUSNESS, AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT¿S EXPIRATION OCCURRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON THE POST-EVENT TESTING, THERE IS EVIDENCE TO SUPPORT THE EXISTENCE OF INTERNAL DEVICE FAILURES, WHICH MAY HAVE CAUSED THE 2008T HEMODIALYSIS SYSTEM TO PERFORM OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR APPROXIMATELY 45% OF ALL ESRD DEATHS, AND APPROXIMATELY 25% OF THOSE DEATHS ARE DUE TO SUDDEN CARDIAC DEATH. BASED ON THE INFORMATION AVAILABLE, THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS ACTIVELY UNDERGOING HD THERAPY WHEN HE WAS FOUND UNRESPONSIVE, THE FACILITY¿S REQUEST FOR POST-EVENT FUNCTIONAL COMPLIANCE TESTING, THE UF PUMP FAILURE REQUIRING RECALIBRATION, THE PRESSURE HOLDING TEST FAILURE, AND A PENDING MANUFACTURER EVALUATION OF THE DIALYZER/BLOODLINES, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
CUSTOM COMBI SET | K962081 | ON 10/DEC/2024, FRESENIUS BECAME AWARE THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, AND COMBI-SET BLOODLINES FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). FOLLOW-UP INFORMATION RECEIVED BY FRESENIUS ON 29/DEC/2024 CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2024 DURING HD THERAPY. THE PATIENT ARRIVED FOR THEIR REGULARLY SCHEDULED HD TREATMENT IN STABLE CONDITION, BLOOD PRESSURE (BP) = 158/77, AND A HEART RATE (HR) = 55 BPM (ALL PRE-TREATMENT MACHINE TESTING PASSED). THE PATIENT¿S TREATMENT BEGAN WITHOUT ISSUE, HOWEVER APPROXIMATELY 12-16 MINUTES INTO A 3.0-HOUR TREATMENT, THE PATIENT WAS FOUND UNRESPONSIVE (BP = 62/46, HR = 78). THE TREATMENT WAS IMMEDIATELY TERMINATED, AND THE PATIENT¿S BLOOD WAS REINFUSED. EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, AND THE PATIENT WAS PROVIDED WITH OXYGEN SUPPORT (VOLUME NOT PROVIDED), A FLUID BOLUS (VOLUME NOT PROVIDED), AND CARDIOPULMONARY RESUSCITATION (CPR) EFFORTS WERE INITIATED. ALTHOUGH THE EXACT TIMELINE AND SPECIFICS OF THE SERIOUS ADVERSE EVENT IS LARGELY UNKNOWN, IT APPEARS CPR WAS UNSUCCESSFUL, AND EMS PRONOUNCED THE PATIENT DEAD AT THE OUTPATIENT DIALYSIS CLINIC. FOLLOWING THE SERIOUS ADVERSE EVENTS THE 2008T HEMODIALYSIS SYSTEM WAS REMOVED FROM SERVICE, AND FUNCTIONAL COMPLIANCE TESTING REVEALED ALARM AND PRESSURE-HOLDING TEST FAILURES (CORRECTED WITH THE REPLACEMENT OF THE INTERNAL 9-VOLT BATTERY). ADDITIONALLY, THE ULTRAFILTRATION (UF) PUMP REQUIRED RECALIBRATION, AS IT FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE UF PUMP ENCOUNTERED AN ERROR IN WHICH MORE FLUID WAS REMOVED THAN WAS PROGRAMED INTO THE 2008T HEMODIALYSIS SYSTEM. THE EXPECTED FLUID REMOVAL WAS 1000 ML/HOUR, HOWEVER POST-EVENT TESTING REVEALED THE ACTUAL FLUID REMOVED 1020 ML/HOUR. THE AIR SEPARATOR WAS ALSO FOUND TO BE LEAKING A SMALL AMOUNT OF FLUID, AND THE UF PUMP STROKES FELL OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS, ACTUAL RESULT 1.029 ML PER STROKE (ALLOWABLE LIMITS = 0.996 ¿ 1.004 ML). THE REQUIRED REPAIRS WERE SUCCESSFULLY COMPLETED, HOWEVER BASED ON THE DOCUMENTATION FROM 29/DEC/2024, THE 2008T HEMODIALYSIS SYSTEM HAS NOT BEEN RETURNED TO SERVICE. ADDITIONALLY, THE OPTIFLUX 180NRE DIALYZER AND COMBI-SET BLOODLINES ARE IN THE PROCESS OF BEING RETURNED TO THE MANUFACTURER. | ADDITIONAL INFORMATION: A1, A2, A3, A4, B5, G2, H6, H10 CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN HD THERAPY UTILIZING A 2008T HEMODIALYSIS SYSTEM, AN OPTIFLUX 180NRE DIALYZER, COMBI-SET BLOODLINES AND THE SERIOUS ADVERSE EVENTS OF HYPOTENSION, LOSS OF CONSCIOUSNESS, AND DEATH, AS THE PATIENT WAS UNDERGOING TREATMENT WHEN THE PATIENT¿S EXPIRATION OCCURRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON THE POST-EVENT TESTING, THERE IS EVIDENCE TO SUPPORT THE EXISTENCE OF INTERNAL DEVICE FAILURES, WHICH MAY HAVE CAUSED THE 2008T HEMODIALYSIS SYSTEM TO PERFORM OUTSIDE OF THE MANUFACTURERS¿ RECOMMENDED SPECIFICATIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR APPROXIMATELY 45% OF ALL ESRD DEATHS, AND APPROXIMATELY 25% OF THOSE DEATHS ARE DUE TO SUDDEN CARDIAC DEATH. BASED ON THE INFORMATION AVAILABLE, THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS ACTIVELY UNDERGOING HD THERAPY WHEN HE WAS FOUND UNRESPONSIVE, THE FACILITY¿S REQUEST FOR POST-EVENT FUNCTIONAL COMPLIANCE TESTING, THE UF PUMP FAILURE REQUIRING RECALIBRATION, THE PRESSURE HOLDING TEST FAILURE, AND A PENDING MANUFACTURER EVALUATION OF THE DIALYZER/BLOODLINES, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE 2008T HEMODIALYSIS SYSTEM, THE OPTIFLUX 180NRE DIALYZER, AND THE COMBI-SET BLOODLINES FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | IN THE PREVIOUS REPORT PATIENT OUTCOME GRID WAS NOT ADDED. IN THIS REPORT PATIENT OUTCOME CODE GRID (BOX H) HAS BEEN UPDATED OR CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY AND HE HAD LUNG CANCER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, ASTHMA (NEW OR WORSENING), INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF KIDNEY DISEASE/TOXICITY, CANCER, DEATH, SORE THROAT AND COUGH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED WRONG INFORMATION IN BOX B. AFTER REVIEW, IT WAS DETERMINED THAT BOX B SHOULD BE CORRECTED. OUTCOMES ATTRIBUTED TO AE HAS BEEN CORRECTED AS AND REPORTED AS BELOW.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS HEADACHE, NAUSEA, VOMITING, KIDNEY DISEASE REDUCED CARDIOPULMONARY RESERVE AND DEATH PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AFTER THE MULTIPLE ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED OR RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE (UNSPECIFIED), DEATH AND OTHER (UNSPECIFIED EVENT). THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
ALARIS SYSTEM | K133532 | IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON 06 DECEMBER 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915. | OMIT : G07001 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. ADDITIONAL INFORMATION: IMDRF ANNEX G CODE.
ALARIS SYSTEM | K133532 | IT WAS INITIALLY REPORTED AN UNSPECIFIED DEVICE ISSUE. PER CUSTOMER'S REPORT, THE PATIENT "DIED". NO FURTHER INFORMATION HAS BEEN PROVIDED AT THE TIME. BD THEN RECEIVED A COPY OF THE CUSTOMER'S MEDWATCH REPORT FROM FDA WHICH STATES, "THE PUMP CAUSED A RATE INCREASE FROM 2.5 MILLILITERS PER HOUR TO 25.6 MILLILITERS PER HOUR WITHOUT EXPLANATION. THIS PUMP WAS NEVER REMOVED ONCE THE INCIDENT WAS NOTICED. THIS PATIENT WAS EXTREMELY UNSTABLE, ON THE SAME PUMP WAS VERSED RUNNING AT 1 MILLILITER PER HOUR, LEVOPHED RUNNING AT 0.5 MILLILITERS PER HOUR (MAX) AND EPINEPHRINE AT 0.3 MILLILITERS PER HOUR (MAX), MAKING EXCHANGING THIS PUMP AT THIS TIME RISKY AND UNSAFE. IT WOULD ALSO REQUIRE US TO WASTE A NEAR-FULL BAGS OF VERSED AND FENTANYL (COSTLY TO THE PATIENT). AND UNSAFE FOR US TO LEAVE THE PATIENT TO WASTE THESE MEDICATIONS PROPERLY. AS WE WERE ARRANGING THE TRANSFER OF PUMPS, I NOTICED THAT THE FENTANYL AND VERSED WERE BOTH ATTACHED TO EXTENSION FILTER TUBING WITH PORTS INSTEAD OF THE PORT-LESS FILTER TUBING. THE PUMP WAS REMOVED FROM THE PATIENT¿S ROOM AND SENT FOR MAINTENANCE." THE PATIENT WAS REPORTEDLY ADMITTED FOR PULMONARY HYPERTENSION. BD RECEIVED ADDITIONAL INFORMATION AFTER MULTIPLE DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT OCCURRED ON 06 DECEMBER 2024. THE MEDICATION THAT REPORTEDLY HAD A RATE INCREASE WAS FENTANYL. PER CUSTOMER, IT IS NOT BELIEVED THAT THE DEVICE CAUSE OR CONTRIBUTED TO THE DEATH. WHEN ASKED FOR THE PRIMARY AND SECONDARY CAUSES OF DEATH, THE CUSTOMER STATED, "PATIENT CONDITION." NO ADDITIONAL DETAILS WERE PROVIDED. THE PATIENT REPORTEDLY PASSED ON (B)(6) 2024 AT 0915. | OMIT: G07001 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE. ADDITIONAL INFORMATION: IMDRF ANNEX G CODE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE CALLER INDICATED THAT THE DEVICE IS NO LONGER IN THEIR POSSESSION. NO ADDITIONAL DETAILS WERE PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION CPAP PRO DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT'S WIFE STATES ON BEHALF OF HER LATE HUSBAND THAT HER HUSBAND HAD SLEEP APNEA, AND HE PASSED AWAY TWO YEARS AGO. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | IN THIS REPORT, SECTION BOX H: DEVICE MANUFACTURERS (EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, CONCLUSION CODE GRID) HAS BEEN UPDATED/CORRECTED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING OF PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
UNIVERSAL MICROINTRODUCER KIT PTFE PEEL APART SHEATH 4.5F | K915167 | PATIENTS NEED TO ENTER THE CHEMOTHERAPY DRUGS, SO THE PLACEMENT NURSE ULTRASOUND UNDER THE PLACEMENT OF GROSHONG, THE USE OF PUNCTURE AND THE USE OF THE GUIDE WIRE INTO THE BLOOD VESSEL ARE VERY SMOOTH, BEFORE THE DELIVERY OF THE GUIDE WIRE INTACT, WITHDRAWN AFTER THE DISCOVERY OF THE WINDING WIRE SCATTERED, THE LENGTH OF THE LEADING WIRE MEASURED 34.7CM. THE PATIENT WENT INTO CRITICAL CARE WITH ACUTE HEART FAILURE THE NIGHT AFTER PLACEMENT, WHETHER IT WAS RELATED TO THE GUIDEWIRE IS CURRENTLY UNKNOWN. ADDITIONAL INFORMATION RECEIVED 12/24/2024: AT PRESENT, IT IS NOT POSSIBLE TO CONFIRM WHETHER THERE ARE FRAGMENTS IN THE PATIENT'S BODY, THE PATIENT HAS PASSED AWAY. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF A BROKEN GUIDEWIRE IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH OF A GUIDEWIRE WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED THE CORE WIRE OF THE GUIDEWIRE WAS BROKEN, CURVED SLIGHTLY LEADING UP TO THE BREAK, AND PROTRUDING FROM THE COILED WIRE. THE ENTIRE LENGTH OF THE COILED WIRE OF THE GUIDEWIRE APPEARED TO BE TIGHTLY COILED, EXCEPT AT THE LOCATION OF THE PROTRUDING CORE WIRE. NO DETAILS OF THIS LOCATION IN THE COILED WIRE AND NO DETAILS OF THE BREAK IN THE CORE WIRE COULD BE DISCERNED FROM THE RETURNED PHOTOGRAPH. THE DISTAL END OF THE COILED WIRE WAS OBSERVED TO BE MORE LUSTROUS COMPARED TO THE REST OF THE WIRE, WHICH SUGGESTS THAT WELD TIP OF THE GUIDEWIRE WAS INTACT; HOWEVER, THIS COULD NOT BE CONFIRMED FROM INSPECTION OF THE RETURNED PHOTOGRAPH. ADDITIONALLY, THE OTHER END OF THE GUIDEWIRE WAS NOT VISIBLE IN THE RETURNED PHOTOGRAPH; THEREFORE, IT COULD NOT BE CONFIRMED IF BOTH WELD TIPS WERE PRESENT FROM INSPECTION OF THE RETURNED PHOTOGRAPH. WHILE THE CORE WIRE OF THE GUIDEWIRE WAS OBSERVED TO BE BROKEN, IT WAS NOT POSSIBLE TO DETERMINE THE CAUSE OF THE BREAK FROM INSPECTION OF THE RETURNED PHOTOGRAPH, AND IT COULD NOT BE DETERMINED IF PIECES OF THE GUIDEWIRE WERE MISSING AS ONLY A PORTION OF THE GUIDEWIRE COULD BE SEEN IN THE RETURNED PHOTOGRAPH. FURTHERMORE, IT COULD NOT BE DETERMINED IF THE OBSERVED BREAK IN THE GUIDEWIRE WAS RELATED TO THE REPORT OF THE PATIENT¿S DEATH, AND NO CLINICAL INFORMATION OR DATA WAS PROVIDED TO SUGGEST THE GUIDEWIRE BREAK CONTRIBUTED TO THE PATIENT¿S DEATH. NO BATCH INFORMATION WAS RECEIVED FROM THE CUSTOMER AND, CONSEQUENTLY, IT WAS NOT POSSIBLE TO PERFORM A HISTORICAL REVIEW OF SIMILAR COMPLAINTS NOR WAS IT POSSIBLE TO PERFORM A REVIEW OF MANUFACTURING RECORDS AND PROCESSES FOR THE BATCH INVOLVED. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT, AND THIS EVALUATION WILL BE UPDATED UPON RECEIPT OF A PHYSICAL SAMPLE FOR EVALUATION AND/OR ANY ADDITIONAL INFORMATION PROVIDED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL, DREAMSTATION ST30 DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH." THIS ISSUE HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A RENTAL, DREAMSTATION ST30, DEVICE IS RETURNING DUE TO A PATIENT PASSING AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS LOCATED IN THE LOGS, AND NO FAULTS WERE FOUND. THE DEVICE'S FILTERS, TUBING, AND MANUAL WERE REPLACED AND THE DEVICE SENT TO BE PUT BACK INTO LOAN STOCK.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE PHILIPS CPAP DEVICE HAD PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | (DEVICE) PROBLEM CODE GRID AND PATIENT OUTCOME CODE GRID HAVE BEEN CORRECTED IN THIS FOLLOW-UP REPORT.
WATCHMAN FXD CURVE ACCESS SYSTEM | K212228 | IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT. PRIOR TO THE WDS BEING INSERTED OR DEPLOYED IN THE PATIENT A PERICARDIAL EFFUSION WAS NOTED. THE PATIENT WAS TAKEN TO OPEN HEART SURGERY TO TREAT THE EFFUSION. THE SAME DAY, THE PATIENT DIED.
DREAMSTATION BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2024, THIS PATIENT UNDERWENT ENDOVASCULAR TREATMENT USING DEBRANCHING AND CHIMNEY TECHNIQUE FOR A DISTAL ARCH AORTIC ANEURYSM WITH GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM, GORE® EXCLUDER® AAA ENDOPROSTHESIS, AND GORE® DRYSEAL FLEX INTRODUCER SHEATH. AFTER DEBRANCHING, THE LEFT SUBCLAVIAN ARTERY WAS PLUG EMBOLISED AND A TGM343415J WAS IMPLANTED DISTALLY. A TGM404020J WAS THEN INSERTED, BUT TGM404020J WAS NOT DEPLOYED AT THIS POINT AS IT WAS TO BE DEPLOYED SIMULTANEOUSLY WITH A CONTRALATERAL LEG ENDOPROSTHESIS (PLC161200J) TO BE IMPLANTED IN THE BRACHIOCEPHALIC ARTERY. AN 8FR NON-GORE SHEATH, A RADIFOCUS¿ GUIDEWIRE, AND AN ANGIOGRAPHIC CATHETER WERE INSERTED FROM THE RIGHT SUBCLAVIAN ARTERY AND SUCCESSFULLY PASSED TO THE ASCENDING AORTA, BUT THE BRACHIOCEPHALIC ARTERY WAS HIGHLY TORTUOUS AND REMAINED TORTUOUS EVEN AFTER THE GUIDEWIRE WAS REPLACED WITH A STIFF TYPE. THE SHEATH WAS REPLACED WITH A DSF1233 AND INSERTED, WHICH PASSED THROUGH WITHOUT RESISTANCE, BUT THE BRACHIOCEPHALIC ARTERY REMAINED TORTUOUS. AN ATTEMPT WAS MADE TO UNSHEATHE THE PLC161200J AFTER DELIVERY TO THE ASCENDING AORTA THROUGH THE DSF1233, BUT THE DSF1233 COULD NOT BE FULLY WITHDRAWN, POSSIBLY BECAUSE THE DSF1233 WAS FOLLOWING THE TORTUOSITY, AND THE PLC161200J WAS RETURNED TO THE DSF1233. THE DSF1233 AND THE PLC161200J WERE PUSHED WITH FORCE AND ADVANCED AGAIN. SUBSEQUENTLY, WHILE ATTEMPTING TO UNSHEATHE AGAIN, THE DSF1233 WAS OBSERVED TO BE OUT OF THE BRACHIOCEPHALIC ARTERY. ANGIOGRAPHY SHOWED A RUPTURE OF THE BRACHIOCEPHALIC ARTERY AND HYPOTENSION WAS OBSERVED, SO A BLOOD TRANSFUSION WAS GIVEN. THE BRACHIOCEPHALIC ARTERY WAS TORN. IT WAS REPORTED THAT THE PLC161200J WAS THOUGHT TO HAVE PASSED THROUGH THE RUPTURE SITE WHILE STILL IN THE DSF1233 AND THAT THE BRACHIOCEPHALIC ARTERY WAS THOUGHT TO HAVE BEEN TORN BY THE MANUPIRATION OF THE DSF1233. BLOOD FLOW IN THE RIGHT SUBCLAVIAN ARTERY WAS UNAFFECTED. ANOTHER ATTEMPT WAS MADE TO INSERT A RADIFOCUS¿ GUIDEWIRE, BUT IT STRAYED INTO A FALSE LUMEN THAT APPEARED TO HAVE FORMED AS A RESULT OF THE VESSEL RUPTURE. AN ATTEMPT TO CREATE A PULL-THROUGH WAS ALSO UNSUCCESSFUL, SO THE PROCEDURE WAS CONVERTED TO OPEN SURGERY, BUT THE VESSEL REPAIR WAS DIFFICULT. THE HEARTBEAT DID NOT RETURN DUE TO HAEMORRHAGE AND THE CHEST WAS INEVITABLY CLOSED WITHOUT FURTHER TREATMENT. THE PATIENT DIED THE SAME DAY AFTER BEING TRANSFERRED TO INTENSIVE CARE. THE CAUSE OF DEATH WAS REPORTED TO BE EXSANGUINATION DUE TO RUPTURE OF THE BRACHIOCEPHALIC ARTERY. IT WAS NOT AVAILABLE WHETHER AN AUTOPSY WAS PERFORMED. REPORTEDLY, TGMR404020J AND PLC161200J WERE REMOVED UNDEPLOYED. THE PHYSICIAN REPORTEDLY STATED AS FOLLOWS: THERE MAY HAVE BEEN AN ERROR OF JUDGEMENT BY THE SURGEON. PERHAPS A STIFFER STIFF GUIDE WIRE SHOULD HAVE BEEN USED. | H3: CODE "OTHER" WAS SELECTED AS THE MEDICAL DEVICE WAS DISCARDED AT FACILITY. RETURN NOT POSSIBLE. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H: PATIENT OUTCOME CODE GRID, EVALUATION METHOD CODE, EVALUATION RESULTS CODE AND CONCLUSION CODE GRID HAS BEEN UPDATED. BOX B: ADVERSE EVENT/PRODUCT PROBLEM HAS BEEN UPDATED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, INFLAMMATORY RESPONSE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, LIVER DISEASE/TOXICITY, AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION GO AUTO | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION GO AUTO DEVICE HAS PASSED AWAY WHILE USING THE DEVICE. THE PATIENT¿S SPOUSE REPORTED THAT THE PATIENT UNEXPECTEDLY PASSED AWAY WHILE USING THE CPAP. ACCORDING TO THE SPOUSE, THE PATIENT APPLIED THE MASK AT APPROXIMATELY 9:30 PM, AND SHE WENT TO BED AROUND 10:15 PM. AT 10:45 PM, SHE NOTICED THE PATIENT SNORING LOUDLY AND NUDGED HIM. HE CEASED SNORING, AND SHE BELIEVED HE HAD ROLLED OVER AND FALLEN BACK ASLEEP, AS WAS HIS USUAL PATTERN. THE SPOUSE AWOKE AT 5:30 AM TO SHOWER, AND UPON EXITING THE SHOWER, NOTICED SOMETHING WAS NOT QUITE RIGHT. SHE TURNED ON THE LIGHT AND OBSERVED THE PATIENT, WITH HIS HAND UNDER HIS HEAD, APPEARING UNRESPONSIVE. THE SPOUSE IMMEDIATELY CALLED 911 AT APPROXIMATELY 5:45-5:50 AM. EMERGENCY RESPONDERS ARRIVED WITHIN 5-10 MINUTES AND PRONOUNCED THE PATIENT DECEASED. THE OFFICIAL CAUSE OF DEATH WAS LISTED AS NATURAL CAUSES. THE PATIENT¿S OFFICIAL DATE OF DEATH IS RECORDED AS (B)(6) 2024; HOWEVER, THE SPOUSE BELIEVES THE PATIENT MAY HAVE PASSED THE NIGHT PRIOR. IT WAS REPORTED THAT THE PATIENT CLEANED THE DEVICE BY PUTTING THE MASK, TUBING AND WATER CHAMBER IN A BLUE BAG AND IT WOULD CLEAN IT SOMEHOW. THIS WAS COMPLETED ONCE A WEEK. DISTILLED WATER WAS USED FOR THE WATER RESERVOIR. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT PATENT HAS EXPIRED DUE TO EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, CANCER AND OTHER PANCREATIC ISSUE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE (B)(4).
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H6: PATIENT OUTCOME CODE GRID WAS CORRECTED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION AND OTHER (UNSPECIFIED INFORMATION). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY DUE TO CHRONIC LUNG DISEASE. THE PATIENT'S WIFE DOES NOT BELIEVE THE DEATH WAS RELATED TO THE USE OF THE CPAP DEVICE AND HAS REQUESTED A RETURN LABEL TO RETURN THE REPLACEMENT DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAD PASSED AWAY DUE TO CHRONIC LUNG DISEASE. THE PATIENT'S WIFE DID NOT BELIEVE THE DEATH WAS RELATED TO THE USE OF THE CPAP DEVICE AND HAD REQUESTED A RETURN LABEL TO RETURN THE REPLACEMENT DEVICE. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE, AND THE EVALUATION NOTES STATED, "DEVICE DOES NOT DUPLICATE CUSTOMER COMPLAINT." AT THE TIME OF INSPECTION, THE DEVICE RECORDED 2,408.2 TOTAL MACHINE HOURS, WITH ZERO BLOWER HOURS. ADDITIONALLY, IT WAS NOTED THAT THE PORT COVER (ACCESSORY DOOR) WAS MISSING. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H CODING IN THIS REPORT.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, HYPERSENSITIVITY, NAUSEA / VOMITING, ASTHMA, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER, DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO BIPAP | K131982 | IT WAS REPORTED THAT THE DREAMSTATION AUTO BIPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
NXSTAGE SYSTEM ONE | K050525 | A REPORT WAS RECEIVED ON (B)(6) 2024 FROM THE HOME THERAPY NURSE (HTN) OF A 62-YEAR-OLD MALE PATIENT APPROVED FOR PERFORMING SOLO HOME HEMODIALYSIS THERAPY, WITH A MEDICAL HISTORY INCLUDING DIABETES, CARDIOVASCULAR ISSUES (NOS) AND END STAGE RENAL DISEASE WHO STATED THE PATIENT EXPIRED AT AN UNSPECIFIED TIME DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 02 JAN 2025 FROM THE HTN WHO STATED THE PATIENT WENT TO USE THE BATHROOM AT AN UNSPECIFIED TIME DURING TREATMENT AND WAS LATER WAS FOUND ON THE FLOOR WITH HIS ARTERIAL LINE DISCONNECTED AND CLAMP TO BOTH THE ARTERIAL LINE AND CATHETER PORT OPEN. PER THE HTN, THE CAUSE OF DEATH WAS AIR EMBOLISM, UNRELATED TO NXSTAGE PRODUCT AND THERAPY. | THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
PILLCAM | K211684 | IT WAS REPORTED THAT THE PATIENT PASSED AWAY AFTER IT WAS REPORTED THAT THE CAPSULE WAS RETAINED. THE CAUSE OF DEATH IS UNKNOWN. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S WIFE, TO SEE WHAT THEY COULD DO FOR THE ELS BECAUSE THEY THREW AWAY THE DEVICE. PATIENT PASSED AWAY AND PATIENT'S WIFE WANTED TO KNOW IF THEY ARE STILL ELIGIBLE. THE CUSTOMER REPRESENTATIVE EXPLAINED SINCE THE DEVICE WAS DISCARDED, IT WOULD BE MARKED AS DISCARDED IN PHILIPS SYSTEM AND THEY ARE NOT ELIGIBLE FOR DEVICE RETURN AWARD. THERE IS NO ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT.
ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | K141952 | IT WAS REPORTED THAT DURING AN UNKNOWN LIVER PROCEDURE THAT ON THE SECOND FIRING NEAR THE PORTAL VEIN THE DEVICE MISFIRED. WHEN THE DEVICE WAS REMOVED THE TISSUE STARTED BLEEDING PROFUSELY AND PATIENT DIED 2HR LATER. DURING THE PROCEDURE THE PATIENT WAS GIVEN 4 ROUNDS OF MTP. | (B)(4). DATE SENT 1/9/2025, D4 BATCH # UNK, B3: ONLY EVENT YEAR KNOWN: 2024. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WERE THE INDICATIONS FOR SURGERY? WHAT IS THE SURGEON¿S EXPERIENCE WITH THE PVS DEVICE? WHAT WAS THE APPROXIMATE WIDTH OF THE COMPRESSED VESSEL? DID THE SURGEON WAIT 15 SECONDS PRIOR TO FIRING? DID THE SURGEON ENSURE THAT THE COMPLETE WIDTH OF THE COMPRESSED VESSEL WAS PROXIMAL TO THE CUT LINE ON THE DEVICE? DID THE SURGEON CONFIRM THAT NO EXTRANEOUS TISSUE WAS DISTAL TO THE CUT LINE WITHIN THE JAWS? WAS THERE ANY TISSUE IN THE JAW DISTAL TO THE CUT LINE? WHAT WAS THE EXACT ANATOMICAL STRUCTURE THAT WAS WITHIN THE JAWS OF THE DEVICE AT THE TIME OF FIRING? WERE THERE ANY DIFFICULTIES WITH THE DEVICE OR STAPLE FORMATION DURING THE INITIAL PROCEDURE? WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WAS THE BLEEDING INTRALUMINAL OR EXTRALUMINAL? WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? WOULD THE SURGEON BE INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING TEAM? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
EVIS EXERA III DUODENOVIDEOSCOPE | K220587 | IT WAS REPORTED THAT FOLLOWING AN ERCP (ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY) PROCEDURE IN (B)(6) 2023, THE PATIENT EXPERIENCED A VRE (VANCOMYCIN RESISTANT ENTEROCOCCI) INFECTION AND LATER PASSED AWAY. NO ADDITIONAL INFORMATION WAS AVAILABLE. | THIS REPORT HAS BEEN SUBMITTED BY THE IMPORTER UNDER THIS MDR REPORT NUMBER (B)(4). THE INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED OR IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
DAVINCI SI | K081137 | IT WAS REPORTED THAT A PATIENT EXPIRED THE DAY AFTER A DA VINCI-ASSISTED RADICAL PROSTATECTOMY WITH LYMPHADENECTOMY PROCEDURE. DURING THE PROCEDURE, AN UNSPECIFIED FORCEPS INSTRUMENT WAS USED TO GRASP A BLOOD VESSEL, AND ¿DUE TO ITS STRONG GRIPPING FORCE¿ DAMAGE WAS CAUSED TO THE VESSEL RESULTING IN BLEEDING THAT WAS CONTROLLED, AND THE PROCEDURE WAS COMPLETED. POST-OPERATIVELY, UNSPECIFIED SIGNS OF SUSPECTED BLEEDING WERE OBSERVED ¿SUCH AS A DROP IN BLOOD PRESSURE¿. FOLLOW-UP WAS INITIATED AND NO ADDITIONAL INFORMATION WAS RECEIVED TO DATE. | A REVIEW OF THE DA VINCI SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS, MONOPOLAR CURVED SCISSORS, TWO LARGE NEEDLE DRIVER INSTRUMENTS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A PROSTATECTOMY PROCEDURE IN WHICH AN UNSPECIFIED FORCEPS INSTRUMENT WAS USED TO GRASP A VESSEL WHICH BECAME DAMAGED. THE EXACT VESSEL, HOW THE INSTRUMENT MAY HAVE DAMAGED THE VESSEL, AND THE DETAILS AROUND HOW THIS INJURY HAPPENED WERE NOT PROVIDED. THE PATIENT LATER DIED. HOW THEY DIED AND THE EVENTS LEADING TO THEIR DEATH WERE NOT PROVIDED. HOW THE VESSEL INJURY MAY OR MAY HAVE NOT CONTRIBUTED TO THIS DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DELUXE 230V 50HZ INT. UK PLUG | K170853 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DELUXE 230V 50HZ INT. UK PLUG THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES NO VISIBLE FOAM DEGRADATION. EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, AND CONCLUSION CODE GRID.
INNOSPIRE DELUXE RENTAL UNIT | K170853 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF INNOSPIRE DELUXE RENTAL UNIT THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP, DOM | K131982 | IT WAS REPORTED THAT THE REP DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WAS NO REPORTED MEDICAL INTERVENTION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DELUXE 230V 50HZ IN | K170853 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION OF DELUXE 230V THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE.
NXSTAGE SYSTEM ONE | K170469 | A REPORT WAS RECEIVED ON 19 DEC 2024 FROM A HEALTH CARE PROFESSIONAL (HCP) OF A 46 YEAR OLD MALE PATIENT APPROVED FOR PERFORMING SOLO HOME HEMODIALYSIS TREATMENT WITH A MEDICAL HISTORY INCLUDING DIABETES AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2024. ADDITIONAL INFORMATION WAS RECEIVED ON 27 DEC 2024 FROM THE HOME THERAPY MANAGER (HTM) WHO STATED THE PATIENT WAS TREATING SOLO AT THE TIME OF THE EVENT AND WAS FOUND DECEASED ON (B)(6) 2024 CONNECTED TO THE DEVICE. PER THE HTM THE CAUSE OF DEATH IS UNKNOWN. | THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
REMSTAR AUTOA-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
MYOSURE REACH TISSUE REMOVAL DEVICE | K152723 | IT WAS REPORTED THAT A PATIENT RECEIVED A NOVASURE/MYOSURE PROCEDURE ON (B)(6) 2024. INITIALLY THE PHYSICIAN RESECTED A FIBROID WITH THE FLUENT PRO. THE PHYSICIAN WAS ABLE TO RESECT THE FIBROID AND ALMOST HIT THE 2500 DEFICIT LIMIT. MYOSURE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PHYSICIAN PERFORMED A NOVASURE PROCEDURE WITH A SUCCESSFUL ABLATION. THE PATIENT WAS DISCHARGED AT THE TIME AND APPEARED FINE. 2 DAYS LATER THE PHYSICIAN WAS EXAMINING THE PATHOLOGY SAMPLES AND NOTED THAT THE PATIENT HAD DECEASED DUE TO A CARDIAC ARREST ON (B)(6) 2024. NO OTHER INFORMATION IS AVAILABLE, MULTIPLE ATTEMPTS AT OBTAINING INFORMATION WERE NOT RESPONDED. | HOLOGIC IS SUBMITTING THIS REPORT IN ACCORDANCE WITH 21. CFR PART 803. THE SUBMISSION IS BASED UPON THE CURRENTLY AVAILABLE INFORMATION. THE SUBMISSION OF THIS REPORT DOES NOT REPRESENT A CONFIRMATION OF DEVICE MALFUNCTION INVOLVING THE HOLOGIC PRODUCTS IN THE EVENT DESCRIBED. LOT AND SERIAL NUMBER OF THE DEVICE NOT PROVIDED BY THE COMPLAINANT; THEREFORE, THE UDI, EXPIRATION AND MANUFACTURING DATES ARE NOT KNOWN. DEVICE HISTORY RECORD (DHR) REVIEW WAS UNABLE TO BE CONDUCTED FOR THE DISPOSABLE DEVICE AS THE IDENTIFICATION NUMBERS WERE NOT PROVIDED BY THE COMPLAINANT. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
FLUENT PRO FLUID MANAGEMENT SYSTEM | K240886 | IT WAS REPORTED THAT A PATIENT RECEIVED A NOVASURE/MYOSURE PROCEDURE ON DECEMBER (B)(6) 2024. INITIALLY THE PHYSICIAN RESECTED A FIBROID WITH THE FLUENT PRO. THE PHYSICIAN WAS ABLE TO RESECT THE FIBROID AND ALMOST HIT THE 2500 DEFICIT LIMIT. MYOSURE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PHYSICIAN PERFORMED A NOVASURE PROCEDURE WITH A SUCCESSFUL ABLATION. THE PATIENT WAS DISCHARGED AT THE TIME AND APPEARED FINE. 2 DAYS LATER THE PHYSICIAN WAS EXAMINING THE PATHOLOGY SAMPLES AND NOTED THAT THE PATIENT HAD DECEASED DUE TO A CARDIAC ARREST ON (B)(6) 2024. NO OTHER INFORMATION IS AVAILABLE, MULTIPLE ATTEMPTS AT OBTAINING INFORMATION WERE NOT RESPONDED. | THE CONSOLE WAS RETURNED FOR INVESTIGATION AN INFLOW SPEED FAULT ON THE FLUENT PRO CAN ONLY BE DISPLAYED DURING STARTUP OR DURING PRIMING. TO TEST FOR THIS FAILURE, THE CONSOLE WAS POWERED ON, SET UP WITH THE DISPOSABLES, AND WAS PRIMED. HERE, IT WAS OBSERVED THAT THE CONSOLE WORKED AS INTENDED AND DID NOT DISPLAY AN INFLOW SPEED FAULT. HENCE, THE FAILURE MODE REPORTED BY THE CUSTOMER WAS NOT REPLICATED DURING INVESTIGATION. A VISUAL INSPECTION WAS CONDUCTED AND THE WHEELS ON THE SYSTEM WERE DAMAGED. IT IS DETERMINED THAT THE DAMAGE TO THE WHEELS DOES NOT RELATE TO THE REPORTED ISSUE. TO FURTHER INVESTIGATE THE REPORTED FAILURE, THE BACK COVER WAS REMOVED TO VISUALLY INSPECT THE CABLES INVOLVED WITH THE INFLOW MOTOR. HERE, IT WAS OBSERVED THAT THERE WERE NO SIGNS OF DAMAGES ON THE INFLOW MOTOR AND ITS CABLES. ADDITIONALLY, THE SYSTEM LOG WAS REVIEWED FOR THE DATE OF THE PROCEDURE AND IT WAS OBSERVED THAT AN INFLOW MOTOR SPEED FAULT WAS PRESENT, AS SHOWN IN THE FILE ATTACHED. AN INSPECTION ONTO THE INFLOW MOTOR SHAFT WAS PERFORMED TO DETERMINE IF THERE IS EVIDENCE OF EXCESSIVE WEAR OR LOOSE SHAFT BEARINGS IN THE FLUID DECK. THERE WERE NO SIGNS OF DAMAGES OR ISSUES ON THE FLUID DECK, SPECIFICALLY THE INFLOW DECK. FOR THE REASON THAT THE REPORTED ISSUE CANNOT BE REPRODUCED, THE CAUSE OF THE REPORTED FAILURE COULD NOT BE FOUND. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE REPORTED SERIAL NUMBER. THE DEVICE WAS RELEASED MEETING ALL QA SPECIFICATIONS. WE ARE CURRENTLY UNABLE TO ESTABLISH A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED.
CADD DISPOSABLE | K111275 | IT WAS REPORTED THAT THE PATIENT HAD EXPIRED WHILE THE PUMP WAS IN USE. PER REPORTER, THE PATIENT WAS FOUND AT HOME AND THE CAUSE OF DEATH IS UNKNOWN. THE PUMP HAD BEEN INFUSING AT A CONTINUOUS RATE OF 5 ML/HOUR WITHOUT ANY REPORTED ISSUES. | E1: FACILITY NAME (B)(6). H3: NEITHER SAMPLES NOR PICTURES WERE RECEIVED FOR THE ANALYSIS OF THIS COMPLAINT. NO DEVICE HISTORY RECORD WAS REVIEWED SINCE LOT NUMBER WAS NOT PROVIDED. WITHOUT THE RETURN OF THE USED SAMPLES A COMPREHENSIVE FAILURE INVESTIGATION CANNOT BE PERFORMED, AND A PROBABLE CAUSE CANNOT BE DETERMINED.
BNX | K103668 | IT WAS REPORTED THAT A PATIENT DIED AFTER A FINE NEEDLE ASPIRATION PROCEDURE. PATIENT WAS ADMITTED THE DAY AFTER BEACON FNA 22G 1 PASS PROCEDURE CYST EMPTYING DUE TO FULMINANT ACUTE PANCREATITIS. PATIENT DIED THE FOLLOWING DAY AFTER ADMITTED, WITH MULTI-ORGAN FAILURE. PATIENT FINAL DIAGNOSIS ATYPIC SIMPLE CYST (CYTOHISTOLOGIC AND LAB NEGATIVE FOR CANCER) OF PANCREAS/MESENTERY. IMMEDIATE MANAGEMENT FOR FULMINANT ACUTE PANCREATITIS AND COMORBIDITY OF HYPERTENSION (IN TARGET WITH TREATMENT), MEDITERREAN ANEMIA TRAIT (NORMAL HGB LEVELS), NO CONCOMITANT TREATMENTS. | ADDITIONAL INFORMATION: A2(AGE AT EVENT), A3A, B2, B3, B5, G3 D10 CONCOMITANT PRODUCT: UNK-BEACON, UNKNOWN BEACON (LOT#UNKNOWN). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
TM80 | K220058 | NO ALLEGED MALFUNCTION WAS REPORTED BY THE CUSTOMER. IT WAS REPORTED THAT ON (B)(6) 2024, AT 16:54 (4:54 PM), A PATIENT WAS PLACED ON THE TM80 TELEMETRY MONITOR CONNECTED TO A BENEVISION DMS, HAD VENTRICULAR TACHYCARDIA (V-TACH)/VENTRICULAR FIBRILLATION (V-FIB) TO ASYSTOLE EPISODES, CONTINUED UNTIL 17:56 (5:56 PM). THE ALARM WAS ALLEGEDLY SILENCED AT THE WORKSTATION. CODE WAS CALLED AND CHEST COMPRESSIONS STARTED. PATIENT EXPIRED. | MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND PATIENT DATABASE FOR INVESTIGATION. MINDRAY EVALUATED THE DATA AND CONFIRMED THAT THE TM80 USED WITH THE BENEVISION DMS PERFORMED PER SPECIFICATIONS. MINDRAY'S TM80 TELEMETRY MONITOR AND THE BENEVISION DMS SERVER GENERATED THE CORRESPONDING ALARMS NORMALLY. HOWEVER, BECAUSE THE USER HAD THE WORKSTATION IN LOCAL STANDBY MODE FOR THE TM80 THAT WAS ACTIVELY MONITORING THE PATIENT, THE CORRESPONDING ALARMS WERE NOT DISPLAYED AT THE WORKSTATION.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. IT IS UNKNOWN IF THE PATIENT WAS USING THE DEVICE AT THE TIME OF DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | THIS PATIENT'S SURGICAL PROCEDURE WAS A SCHEDULED NON-ROBOTIC SURGERY; THE REPORTED PATIENT EVENT WAS ALLEGED TO BE INDIRECTLY RELATED BY THE SURGEON WHO REPORTED THAT THE PATIENT SUFFERED A PROBABLE THROMBOEMBOLIC ACCIDENT EVENT RELATED TO INTERRUPTION OF THE THE PATIENT¿S ANTICOAGULANT TREATMENT, THEN RESUMED FOLLOWING THE SURGEON'S EARLIER DA VINCI IESU ERBE EVENT, LEADING TO POSTPONEMENT OF THEIR SURGERY. BASED ON THE INVESTIGATION, THERE IS NO INDICATION THAT THE DEVICE DIRECTLY CAUSED THE REPORTED ADVERSE EVENT; THE CASCADE LED TO THE SURGEON'S SUBSEQUENT PROCEDURE(S) BEING POSTPONED BY THE SURGEON. PLEASE SEE SECTION H10 RELATED MEDWATCH REPORT FOR THE IESU ERBE DEVICE USED DURING A DIFFERENT PATIENT PROCEDURE. AN INTUITIVE SURGICAL, INC. (ISI) FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE CUSTOMER SITE TO INVESTIGATE THE SURGEON'S UNRELATED EARLIER ROBOTIC CASE WHERE THE ERBE GENERATOR WAS REPLACED AS A RESULT OF MONOPOLAR POWER ACTIVATION PROBLEMS AND C-35 ERRORS NOTED DURING TESTING. ISI DID RECEIVE THE IESU ERBE GENERATOR FOR FAILURE ANALYSIS. DURING THE POWER-ON TEST, IESU CAUTERY ACTIVATION, AND VISUAL INSPECTION, THE ERROR WAS FOUND, CONFIRMING THE FAULT OCCURRED IN THE FIELD. UPON VISUAL INSPECTION, NO ISSUES WERE FOUND THAT WOULD BE RELATED TO THE REPORTED EVENT. THE COMPLAINT WAS CONFIRMED BASED ON THE FIELD EVALUATION AND FAILURE ANALYSIS, WHICH INDICATES THAT THE DEVICE MAY HAVE CONTRIBUTED TO THE CUSTOMER REPORTED ISSUE. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THE FOLLOWING POSSIBLY RELATED SYSTEM ERRORS: C-34 - ENERGY ACTIVATION HALTED BY AN INSTRUMENT OR INSTRUMENT CABLE CONNECTED TO THE ERBE WHILE USING THE COAG FUNCTION. THE PROBABLE CAUSE OF THE ERRORS IS ATTRIBUTED TO A FAULTY ELECTRICAL COMPONENT INSIDE THE IESU. ONE ERROR INDICATES A LOWER VOLTAGE THAN EXPECTED DURING ENERGY ACTIVATION. THE OTHER ERROR INDICATES A HIGHER CURRENT MEASUREMENT VALUE THAN EXPECTED WHILE THE UNIT IS IN AN OFF STATE. ADDITIONAL B, C, D IMDRF ANNEX CODES ADDED REFLECT THE ISSUE THAT OCCURRED WITH THE IESU ERBE DURING A DIFFERENT SURGERY, RESULTING IN THIS PATIENT'S NON-ROBOTIC SURGERY BEING POSTPONED. C20 - NO FINDINGS AVAILABLE AND D15 - CAUSE NOT ESTABLISHED REFER TO THE IMPACT FOR THE PATIENT IN THIS REPORT.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
NOVOPEN 6 | K123766 | CASE DESCRIPTION: THIS SERIOUS SPONTANEOUS CASE FROM FRANCE WAS REPORTED BY A CONSUMER AS "KETOACIDOSIS (PATIENT HAS DIABETES)(DIABETIC KETOACIDOSIS)" BEGINNING ON (B)(6)2024, "BROKEN NOVOPEN 6(DEVICE BREAKAGE)" BEGINNING ON (B)(6)2024, AND CONCERNED A 42 YEARS OLD MALE PATIENT WHO WAS TREATED WITH NOVOPEN 6 (INSULIN DELIVERY DEVICE) FROM UNKNOWN START DATE FOR "DEVICE THERAPY", NOVORAPID PENFILL (INSULIN ASPART 100 U/ML) (DOSE, FREQUENCY & ROUTE USED- UNK) FROM UNKNOWN START DATE FOR "PRODUCT USED FOR UNKNOWN INDICATION". BATCH NUMBERS: NOVOPEN 6: UNKNOWN. NOVORAPID PENFILL: NR7KV74. ACTION TAKEN TO NOVORAPID PENFILL WAS NOT REPORTED. ON (B)(6)2024 THE OUTCOME FOR THE EVENT "KETOACIDOSIS (PATIENT HAS DIABETES) (DIABETIC KETOACIDOSIS)" WAS FATAL. ON (B)(6)2024 THE OUTCOME FOR THE EVENT "BROKEN NOVOPEN 6(DEVICE BREAKAGE)" WAS FATAL. EVENT ONSET DATE IS NOT REPORTED IN THE CASE. HOWEVER, THE INCIDENT DATES ARE CAPTURED TO ENSURE THE MIR FORM IS GENERATED. SINCE LAST SUBMISSION THE CASE HAS BEEN UPDATED WITH THE FOLLOWING: NEW SUSPECT PRODUCT ADDED(NOVORAPID). EXPECTED DATE OF FOLLOW UP REPORT EXTENDED . NARRATIVE UPDATED ACCORDINGLY. PRELIMINARY MANUFACTURER'S COMMENT: (B)(6)2025: THE SUSPECTED DEVICE NOVOPEN 6 HAS BEEN RETURNED TO NOVO NORDISK FOR EVALUATION. INVESTIGATIONS ARE ONGOING. RELEVANT INFORMATION ON MEDICATION USED WITH NOVOPEN 6, PRECIPITATING FACTORS FOR DIABETIC KETOACIDOSIS, TREATMENT GIVEN IN HOSPITAL AND INVESTIGATION RESULTS OF DEFECTIVE DEVICE ARE UNAVAILABLE FOR COMPLETE CAUSALITY ASSESSMENT. PATIENT'S UNDERLYING MEDICAL CONDITION OF DIABETES MELLITUS IS A SIGNIFICANT CONFOUNDING FACTOR FOR THE DEVELOPMENT OF ACUTE COMPLICATION OF DIABETIC KETOACIDOSIS LEADING TO FATAL OUTCOME. "THIS REPORT IS FOR A FOREIGN DEVICE THAT IS ASSESSED AS "SIMILAR" TO US MARKETED NOVOPEN ECHO".
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND LIVER CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, KIDNEY DISEASE/TOXICITY AND LUNG DISEASE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LEUKEMIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, VOMITING, LIVER DISEASE/TOXICITY AND CANCER, DEATH AND NAUSEA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), REDUCED CARDIOPULMONARY RESERVE, DEATH, HEART CONDITION, AND PNEUMONIA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT WAS DEATH. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | IN PREVIOUS REPORT, THE MANUFACTURER REPORTED INCORRECT INFORMATION IN BOX B: COMMON DEVICE NAME. AFTER REVIEWED, IT WAS CAPTURED CORRECTLY. IN BOX B: COMMON DEVICE NAME HAS BEEN UPDATED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, THROAT, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, DEATH AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH AND ESOPHAGEAL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, COUGH, SINUS INFECTION AND UPPER AIRWAY IRRITATION. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
CADD-SOLIS VIP AMBULATORY INFUSION PUMP | K111275 | A REPORT OF A PATIENT¿S DEATH WAS RECEIVED BY ICU MEDICAL ON (B)(6)2024 WITH UPDATE ON (B)(6)2024, THIS PATIENT IS REPORTED TO HAVE A INFUSION ONGOING AT THE TIME OF THE REPORTED INCIDENT. LACK OF DETAILS REGARDING THIS CLINICAL EVENT PRECLUDES A COMPREHENSIVE MEDICAL ASSESSMENT. THERE IS A REPORT FROM COLORADO BLOOD CANCER INSTITUTE OF A CADD INFUSION PUMP SYSTEM WAS DELIVERING A MEDICATION AT 5 ML/HR AT HOME WITH NO REPORTED ISSUE. THERE IS NO HISTORY OF WHAT MEDICATION WAS BEING DELIVERED TO THE PATIENT, NO PAST MEDICAL HISTORY AND NO HISTORY OF OTHER MEDICATIONS THAT THE PATIENT WAS CONCURRENTLY TAKING. IT IS REPORTED THAT THIS WAS A FEMALE PATIENT, AND THE PATIENT WAS FOUND BY HER SISTER AT HOME AND THE DATE OF THE PATIENT¿S DEATH WAS REPORTED AS (B)(6)2024. NO FURTHER DETAILS OF THE FUNCTIONALITY OF THE PUMP OR THE CONDITION OF THE PATIENT ARE PROVIDED. THERE IS A GENERAL LACK OF INFORMATION PROVIDED IN THIS CASE TO MAKE A DETERMINATION OF CAUSALITY. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THIS MEDICAL ASSESSMENT WILL BE FURTHER REVISED AS NEEDED. | E1 INITIAL REPORTER FULL PHONE NUMBER: (B)(6). H3, H6: INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT AND PATIENT'S WIFE WANTED TO RETURN THE REPLACEMENT DEVICE BACK AS IT WAS NO LONGER NEEDED. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION PATIENT PASSED AWAY. THERE WAS REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURE FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), LUNG DISEASE AND DEATH WHILE USING THE DEVICE. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/ OR HEADACHE, HYPERSENSITIVITY, LIVER DISEASE/ TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, LUNG CANCER AND LIVER CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CHRONIC BRONCHITIS, RECURRING PNEUMONIA, COPD. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAM STATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAM STATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, ASPIRATION PNEUMONITIS AND CHRONIC OBSTRUCTIVE PULMONARY DISEASE (COPD). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DIZZINESS, HEADACHE, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ESOPHAGEAL ADAMOU CARCINOMA, EYE IRRITATION, SKIN IRRITATION CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, DEATH AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, REDUCED CARDIOPULMONARY RESERVE, DEATH, NOSE BLEEDS, LUNG INFLAMMATION AND OTHER (UNSPECIFIED EVENT). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, DEATH AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, KIDNEY DISEASE/TOXICITY AND RESPIRATORY FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING HEART ATTACK AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, LIVER DISEASE/TOXICITY, LUNG DISEASE, CANCER, OTHER (UNSPECIFIED EVENT), BREAST, MULTIPLE MYELOMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE, LUNG DISEASE, CANCER AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, SKIN IRRITATION, CANCER, DEATH, CANCER, BLADDER AND PROSTATE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND SHORTNESS OF BREATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED/RECERTIFIED DREAMSTATION CPAP WITH AN ALLEGATION OF DEATH, LUNG DISEASE AND CANCER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY REPORTED INCORRECT INFORMATION IN PREVIOUS REPORT. THE FOLLOWING CORRECTION HAS BEEN UPDATED IN THIS REPORT. BOX H: TYPE OF REPORTED COMPLAINT WAS CAPTURED SERIOUS INJURY IN PREVIOUS REPORT. TYPE OF REPORTED COMPLAINT WAS CORRECTED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, LYMPHOMA AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, DIZZINESS, HEADACHE, INFLAMMATORY RESPONSE AND KIDNEY DISEASE/TOXICITY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND ARDS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021 AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, DIZZINESS AND HEADACHE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, DIZZINESS, HEADACHE, HYPERSENSITIVITY, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, REDUCED PULMONARY RESERVE, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND HEART FAILURE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION HUMIDIFIER, GB-RENT | K131982 | THIS REPORT IS BEING SUBMITTED TO CORRECT THE DESCRIPTION OF THE EVENT OR PROBLEM PREVIOUSLY REPORTED. IT WAS REPORTED THAT THE DREAMSTATION HUMIDIFIER DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE IS NO EVIDENCE THAT THE DREAMSTATION HUMIDIFIER DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR. THIS COMPLAINT IS NOT REPORTABLE TO ANY REGULATORY AGENCIES. | THIS REPORT IS BEING SUBMITTED TO CORRECT THE DESCRIPTION OF THE EVENT OR PROBLEM PREVIOUSLY REPORTED. IT WAS REPORTED THAT THE DREAMSTATION HUMIDIFIER DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE IS NO EVIDENCE THAT THE DREAMSTATION HUMIDIFIER DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR. THIS COMPLAINT IS NOT REPORTABLE TO ANY REGULATORY AGENCIES.
DREAM STATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, LUNG DISEASE, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING STOMACH ISSUES, CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY AND DEATH, THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DURAN ANCORE | K032810 | IT WAS REPORTED THAT A PATIENT IMPLANTED WITH A 27MM MITRAL DURAN ANCORE ANNULOPLASTY BAND DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE ANNULOPLASTY BAND COULD NOT BE EXCLUDED. IT WAS STATED THAT PRIOR TO DEATH, THE PATIENT WAS IN CARDIAC ARREST, AND ACTIVE CARDIOPULMONARY RESUSCITATION WAS BEING PERFORMED. UPON ARRIVAL TO THE FACILITY, THERE WAS NO HEART RATE DETECTED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, APPENDIX, COLON, LIVER PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
COMPOSIX MESH E/X | K002684 | ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) 2004, COMPOSIX E/X ON (B)(6) 2008 AND COMPOSIX L/P ON (B)(6) 2011. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST COMPOSIX KUGEL HERNIA PATCH AND COMPOSIX E/X. ATTORNEY ALLEGES THAT THE PATIENT HAD SUBSEQUENT SURGICAL INTERVENTION DUE TO THE HERNIA MESH DEVICE AND WRONGFUL DEATH OF THE PATIENT. IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. | NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING SUBSEQUENT SURGICAL INTERVENTION AND DEATH. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. THIS EMDR REPRESENTS BARD/DAVOL COMPOSIX MESH E/X (DEVICE #2). AN ADDITIONAL EMDR WAS SUBMITTED TO REPRESENT BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: NOT RETURNED.
MESH ¿ COMPOSIX KUGEL | K061314 | ATTORNEY ALLEGES THAT THE PATIENT UNDERWENT SURGERY FOR IMPLANT OF UNSPECIFIED BARD/DAVOL COMPOSIX KUGEL HERNIA PATCH ON (B)(6) 2004, COMPOSIX E/X ON (B)(6) 2008 AND COMPOSIX L/P ON (B)(6) 2011. AS REPORTED, THE PATIENT IS MAKING A CLAIM FOR AN ADVERSE PATIENT OUTCOME AGAINST COMPOSIX KUGEL HERNIA PATCH AND COMPOSIX E/X. ATTORNEY ALLEGES THAT THE PATIENT HAD SUBSEQUENT SURGICAL INTERVENTION DUE TO THE HERNIA MESH DEVICE AND WRONGFUL DEATH OF THE PATIENT. IT IS ALSO ALLEGED THAT THE PATIENT EXPERIENCED EMOTIONAL DISTRESS AND THE DEVICE WAS DEFECTIVE. | NO CONCLUSIONS CAN BE MADE. THE PATIENT'S ATTORNEY ALLEGES ADVERSE PATIENT OUTCOME ASSOCIATED WITH THE HERNIA MESH USED TO TREAT THE PATIENT INCLUDING SUBSEQUENT SURGICAL INTERVENTION AND DEATH. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. SHOULD ADDITIONAL INFORMATION BE PROVIDED, A SUPPLEMENTAL EMDR WILL BE SUBMITTED. THIS EMDR REPRESENTS BARD/DAVOL MESH ¿ COMPOSIX KUGEL (DEVICE #1). AN ADDITIONAL EMDR WAS SUBMITTED TO REPRESENT BARD/DAVOL COMPOSIX MESH E/X (DEVICE #2). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. H3 OTHER TEXT: NOT RETURNED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, METASTATIC LUNG CANCER AND BRAIN AND BONE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
VERSACROSS ACCESS SOLUTION | K150709 | IT WAS REPORTED A PATIENT DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A WATCHMAN FLX PRO PROCEDURE, A VERSACROSS ACCESS SOLUTION KIT WAS SELECTED FOR USE. A TRACE EFFUSION WAS NOTED PRIOR TO PROCEDURE. THE PHYSICIAN OBTAINED RIGHT FEMORAL VEIN ACCESS AND AN 18GUAGE SHEATH WAS PLACED. THE VERSACROSS SHEATH WAS INSERTED. THE PHYSICIAN DROPPED BELOW THE SEPTUM ON FIRST ATTEMPT, SO IT WAS RE-ATTEMPTED TO RE-WIRE THE SUPERIOR VENA CAVA (SVC) TO COME BACK DOWN TO THE SEPTUM. THE TRANSSEPTAL PUNCTURE WAS PERFORMED AND THE VERSACROSS SHEATH WAS EXCHANGED FOR THE WATCHMAN TRUSTEER SHEATH OVER THE WIRE. A PIGTAIL CATHETER WENT IN OVER THE WIRE. AT THIS POINT, AN EFFUSION WAS NOTED. AUTO TRANSFUSION BEGAN WITH THAT PIGTAIL CATHETER WHILE A PERICARDIOCENTESIS WAS PERFORMED. AFTER THAT THERE WERE TWO PIGTAIL CATHETERS PLACED IN THE PERICARDIAL SPACE. 12 UNITS OF PACKED RED BLOOD CELLS (PRBCS), 8 UNITS OF FRESH FROZEN PLASMA (FFP) AND 1 UNIT OF PLATELETS WERE GIVEN AND THE PATIENT WAS TAKEN TO THE OPERATING ROOM (OR). THE PATIENT ARRESTED BEFORE GETTING TRANSFERRED THE OR TABLE, PATIENT CHEST WAS OPENED BUT WERE UNABLE TO RESUSCITATE. IT WAS DETERMINED THE CAUSE OF DEATH WAS DUE TO PERFORATION AT THE BACK WALL OF THE HEART LEADING TO PERICARDIAL EFFUSION, CARDIAC TAMPONADE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. IT WAS FURTHER CONFIRMED VIA GFE DATED 10JAN2025, THERE WAS NO IMAGING/VISUALIZATION ISSUES NOTED. THE TSP WAS WAS DIFFICULT BECAUSE OF THE REACHABILITY. THE PHYSICIAN BROUGHT THE SHEATH OUT OVER THE WIRE AND PUT MORE OF A BEND ON THE VERSACROSS AND WENT BACK IN TO THE SVC AND BACK DOWN. THE TSP WAS VERY INFERIOR AND POSTERIOR. THE PERFORATION COULD HAVE PARTIALLY CAUSED BY THE WIRE HOWEVER THE PHYSICIAN PUT THE TRUSTEER OVER THE WIRE AND INTO THE PERICARDIAL SPACE. THERE WAS NO MALFUNCTION WITH THE VERSACROSS SYSTEM. ACTIVATED CLOTTING TIME (ACT) GREATER THAN 400 WITH AN INR OF 2.5 ON CASE DAY. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING THYROID AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, NAUSEA, VOMITING, ASTHMA (NEW OR WORSENING), CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
HEARTSTART XL+ DEFIBRILLATOR/MONITOR | K110825 | PHILIPS RECEIVED A COMPLAINT ON THE XL+ DEVICE INDICATING THAT THE DEVICE BURNED THE PATIENT'S SKIN DURING USE. THE PATIENT HAD NO HEARTBEAT AND WAS UNCONSCIOUS IN THE PROCESS OF THE BURNING OF THE PATIENTS SKIN. THE PATIENT DIED. | THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS REMOTE SERVICE ENGINEER (RSE) AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. MULTIPLE ATTEMPTS WERE MADE TO REQUEST INFORMATION REGARDING RESOLUTION OF THE REPORTED PROBLEM. NO RESPONSE WAS RECEIVED, SO NO INFORMATION IS AVAILABLE. THIS WILL BE DOCUMENTED AS A MALFUNCTION, THE CAUSE OF WHICH WAS NOT DETERMINED. THE DEVICE REMAINS AT THE CUSTOMER SITE AND NO FURTHER EVALUATION IS WARRANTED AT THIS TIME. BASED ON THE INFORMATION AVAILABLE THERE IS INSUFFICIENT INFORMATION TO CONFIRM THE REPORTED PROBLEM. BASED ON THE INFORMATION AVAILABLE AND RESULTS OF ADDITIONAL ANALYSIS, NO FURTHER ACTION IS NECESSARY AT THIS TIME. WE ARE UNABLE TO CONFIRM THE FINAL DISPOSITION OF THE DEVICE BECAUSE THE CUSTOMER DID NOT RESPOND TO REQUESTS FOR ADDITIONAL INFORMATION. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE PASSED AWAY. THERE WERE NO FURTHER DETAILS PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ACIST | K010390 | DURING A PERCUTANEOUS CORONARY INTERVENTION USING THE ACIST CVI INJECTION SYSTEM, AIR WAS INJECTED INTO A PATIENT. THE AIR INJECTION OCCURRED AFTER INSERTION OF A CORONARY GUIDING CATHETER AND ENGAGEMENT INTO THE LEFT MAIN CORONARY ARTERY AND UPON INJECTION OF CONTRAST MEDIA. THE PATIENT EXPERIENCED CARDIOGENIC CRASH AND SHOCK AND EXPIRED THE SAME DAY. | THE ACIST ANGIOGRAPHIC INJECTION SYSTEM, MODEL CVI, SYSTEM SERIAL NUMBER (B)(6), WAS RECEIVED FOR EVALUATION ON FEBRUARY 4, 2025. THE INJECTION SYSTEM WAS FUNCTIONALLY TESTED AND MET THE PRE-ESTABLISHED SPECIFICATIONS. THERE WAS NO EVIDENCE OF DEVICE MALFUNCTION RELATED TO THE REPORTED EVENT. THE INSTRUCTIONS FOR USE HAVE BEEN REVIEWED AND NO INADEQUACIES WERE IDENTIFIED REGARDING WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR USE OF THE DEVICE. PER THE ACIST CVI USER'S MANUAL, THE AIR COLUMN DETECT SENSOR IS DESIGNED TO AID THE USER IN THE DETECTION OF AIR COLUMNS IN THE INJECTION LINE, BUT IT IS NOT DESIGNED TO REPLACE THE VIGILANCE AND CARE REQUIRED OF THE OPERATOR IN VISUALLY INSPECTING FOR AIR AND CLEARING AIR FROM THE ENTIRE PATIENT KIT AND ANGIOGRAPHIC CATHETER. THE AIR COLUMN DETECT MECHANISM IS TO BE USED IN CONJUNCTION WITH AND TO COMPLEMENT THE USER'S OTHER PROCEDURES FOR PREVENTING AIR INJECTIONS. THIS REPORT IS CLOSED.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY, AND THE DEVICE WAS RETURNED. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY,DEATH,KIDNEY DAMAGE, HEART ATTACK. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA /VOMITING, ASTHMA (NEW OR WORSENING), REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER, DEATH PROBLEMS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THE CENTRAL STATION DID NOT GENERATE AN ALARM WHEN THE PATIENT WENT INTO CARDIAC ARREST AND PASSED AWAY. THE CUSTOMER REQUESTED ASSISTANCE IN DETERMINING WHETHER ALARMS WERE TRIGGERED AT THE TIME OF THE EVENT. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | PATIENT INVOLVEMENT WAS REPORTED, THE PATIENT EXPERIENCED A CARDIAC ARREST AND PASSED AWAY. THE DEVICE, INCLUDING THE ASSOCIATED BEDSIDE DEVICE(S), WERE TESTED AND DETERMINED TO BE FUNCTIONING AS DESIGNED. IT WAS ALSO INDICATED DURING TESTING THAT THE PIC IX SPEAKER WAS PUSHED TO THE BACK OF THE SHELF. THE PROVIDED LOGS WERE REVIEWED BY A PHILIPS PRODUCT SUPPORT ENGINEER AND CLINICAL SPECIALIST. IT WAS DETERMINED THAT THE DEVICE DID ALARM FOR ASYSTOLE DURING THE INDICATED TIMEFRAMES. THE DEVICE WAS TESTED AND DETERMINED TO BE FUNCTIONING AS DESIGNED. REVIEW OF THE LOGS ALSO INDICATED THE DEVICE WAS ALARMING AS EXPECTED. IT WAS CONFIRMED THAT THE DEVICE WAS TESTED AND IS STILL IN CLINICAL USE. REPORTING INSTITUTION PHONE#: (B)(6).
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, AND/OR HEADACHE, NAUSEA / VOMITING, INFLAMMATORY RESPONSE, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | BOX H WAS UPDATED FROM SERIOUS INJURY TO DEATH.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
TUBING SET | K112360 | COMPLAINT ID# (B)(4). | THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT A HLS SET WAS REPLACED AND AFTER THE REPLACEMENT (A FEW HOURS LATER) THE CUSTOMER NOTICED A DECREASE IN OXYGENATION VALUES. THE PATIENT OXYGENATION WAS IN THE LOW 80`S. THE CUSTOMER CHECKED THE GAS SOURCE TO ENSURE PROPER CONNECTION AND THEY INCREASED THE SWEEP TEMPORARILY TO REMOVE ANY POTENTIAL CONDENSATION. THE POST OXYGENATION (PAO2) DECREASED FURTHER AND THE CUSTOMER EXCHANGED THE HLS SET AGAIN. NEW INFORMATION RECEIVED ON 2025-01-27 THAT THE PATIENT IS 33 YEARS OLD, WEIGHT 149.9KG AND A MALE. NO VISIBLE CLOTS WERE DETECTED. AFTER THE SECOND EXCHANGE TO THE THIRD SET THE PATIENT IMPROVED SLIGHTLY BUT THAT DID NOT LAST LONG. THE PATIENT HAD HIGH LEVELS OF TRIGLYCERIDES AND CONTINUED TO HAVE LOW OXYGENATION AND PASSED AWAY LATER IN THE AFTERNOON. ON 2025-02-03 THE INFORMATION WAS RECEIVED THAT THE FIRST AND THIRD USED SET WERE DISCARDED AND THE LOT NUMBERS ARE UNKNOWN. THE SECOND USED SET WAS AVAILABLE FOR INVESTIGATION. AS THE PATIENT SUFFERED A LOW OXYGENATION AND PASSED AWAY, A REPORT IS REQUIRED. THE AFFECTED PRODUCT (SECOND SET) WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-04-28 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. ALL PERFORMANCE TESTS RELATED TO THE REPORTED FAILURE WERE FUNCTIONING AS PER SPECIFICATION. FURTHER, DURING VISUAL INSPECTION THIRD PARTY ARTICLES, CONNECTED TO THE MODULE, WERE FOUND. FURTHER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-05-22 WITH FOLLOWING CONCLUSION: "ACCORDING TO THE COMPLAINT NARRATIVE, THE INCIDENT IN SCOPE INVOLVES A 33-YEAR-OLD MALE PATIENT ON VV-ECLS USING A CARDIOHELP SYSTEM WITH AN HLS SET EXPERIENCED A DECLINE IN OXYGENATION FOLLOWING A CIRCUIT EXCHANGE ON DAY SIX OF SUPPORT. WITHIN TEN HOURS OF INITIATING THE NEW CIRCUIT, POST-OXYGENATOR PAO2 DROPPED FROM 227 TO 210 MMHG (FIO2 1.0, SWEEP 6 L/MIN, BLOOD FLOW 5.8 - 6.28 L/MIN), AND THE PATIENT'S SPO2 FELL INTO THE LOW 80S. A SUBSEQUENT CIRCUIT REPLACEMENT DID NOT RESTORE ADEQUATE OXYGENATION, AND THE PATIENT SUBSEQUENTLY EXPIRED. INVESTIGATION OF THE RETURNED HLS SET - THE SECOND OF THREE CIRCUITS - REVEALED DAMAGE OF A GAS FILTER PART AND A CRACKED THIRD-PARTY STOPCOCK. THESE WERE ATTRIBUTED TO POST-USE HANDLING, WITH NO CLINICAL LEAKAGE REPORTED. INTERNAL INSPECTION SHOWED NO CLOTS OR MECHANICAL OBSTRUCTIONS AS DESCRIBED IN THE INVESTIGATION REPORT, AND THE DEVICE PASSED ALL LEAK AND FUNCTION TESTS AFTER REMOVING THE DAMAGED EXTERNAL STOPCOCK PART. ALTHOUGH, THE OBSERVED CRACKING OF THE THIRD-PARTY STOPCOCK MAY BE ASSOCIATED WITH ENVIRONMENTAL STRESS CRACKING OF POLYCARBONATE, AS DESCRIBED IN THE HANTEL TECHNOLOGIES WHITE PAPER ON MATERIAL SELECTION. THE REPORT HIGHLIGHTS THAT PC COMPONENTS USED IN MEDICAL DEVICES CAN FAIL WHEN EXPOSED TO LIPOPHILIC AGENTS, WHICH ARE COMMONLY USED IN CRITICAL CARE AND INTERVENTIONAL PROCEDURES. SUCH EXPOSURE - ESPECIALLY UNDER HOOP STRESS FROM HAND-TIGHTENING CONNECTIONS - CAN RESULT IN CRAZING AND CRACKING, EVEN IN THE ABSENCE OF VISIBLE DEFECTS UNDER NORMAL CONDITIONS. SPECIFICALLY, THE PAPER DEMONSTRATED THAT PC LUERS AND STOPCOCKS SUBJECTED TO LIPID-BASED AGENTS SUCH AS LIPIODOL OR SIMULATED CONTRAST MEDIA (1PERCENTAGE POPPY SEED OIL IN ETHYL STEARATE) SHOWED MICROCRACKS AND, EVENTUALLY, FULL STRUCTURAL FAILURE AFTER MECHANICAL STRESS. THE COMBINATION OF CHEMICAL EXPOSURE AND MECHANICAL FORCE AMPLIFIED THE HOOP STRESS BEYOND THE MATERIAL'S RESISTANCE THRESHOLD, LEADING TO ESC, A POTENTIAL MECHANISM CONSISTENT WITH THE CRACKED STOPCOCK OBSERVED IN THIS CASE. GAS PERMEABILITY WAS CONFIRMED DURING THE PRODUCT INVESTIGATION, INDICATING INTACT CO2 REMOVAL AND O2 DELIVERY FUNCTIONS. CLINICALLY, NO THROMBI WERE NOTED, ROUTINE FLUSHING WAS PERFORMED, AND TRANSMEMBRANE PRESSURE REMAINED WITHIN NORMAL LIMITS. THE DEVICE HAD BEEN PRIMED NINE DAYS PRIOR TO USE. HOWEVER, ACCORDING TO THE INSTRUCTIONS FOR USE, THE SET SHOULD ONLY BE PRIMED ¿[¿] SHORTLY BEFORE USE [¿]¿. NONETHELESS, THIS DEVIATION IS NOT EXPECTED TO HAVE IMPACTED THE INCIDENT DESCRIBED. ALTHOUGH THE TECHNICAL INVESTIGATION CONCLUDED THAT THERE WERE NO INDICATIONS OF POTENTIAL IMPAIRMENT TO OXYGENATION PERFORMANCE, SEVERAL CLINICAL AND CONTEXTUAL FACTORS MERIT CONSIDERATION AS POSSIBLE CONTRIBUTORS TO THE OBSERVED OXYGENATION DECLINE: 1) THE INVESTIGATION IDENTIFIED A CRACK IN THE THIRD-PARTY THREE-WAY STOPCOCK CONNECTED TO THE SAMPLE LINE AT THE QUICK-VENT LUER. WHILE DEEMED LIKELY TO HAVE OCCURRED POST-USE, POLYCARBONATE COMPONENTS LIKE THIS ARE KNOWN TO BE SUSCEPTIBLE TO ENVIRONMENTAL STRESS CRACKING (ESC), PARTICULARLY WHEN EXPOSED TO LIPOPHILIC AGENTS AND MECHANICAL STRESS. THE USE OF NON-ORIGINAL PARTS INTRODUCES MATERIAL-RELATED RISKS THAT CANNOT BE FULLY EXCLUDED AS CONTRIBUTING FACTORS 2) ANOTHER POTENTIAL CONTRIBUTOR, THOUGH CONSIDERED UNLIKELY, RELATES TO THE GAS SUPPLY SYSTEM. A LACK OF GAS PRESSURE FROM THE WALL SOURCE OR IMPROPER CALIBRATION OF THE SECHRIST GAS BLENDER COULD, HYPOTHETICALLY, IMPAIR OXYGEN TRANSFER. HOWEVER, SUCH ISSUES WOULD TYPICALLY AFFECT MULTIPLE PATIENTS USING THE SAME SYSTEM INFRASTRUCTURE. IN THIS CASE, NO ABNORMAL GAS SUPPLY CONDITIONS WERE NOTED, AND NO OTHER COMPLAINTS WERE REPORTED FROM THE FACILITY, MAKING THIS SCENARIO LESS PROBABLE. 3) THE CUSTOMER REPORTED OBSERVING "WHITE MATERIAL" DURING BLOOD SAMPLING, WHICH MAY BE RELATED TO THE PATIENT'S EXCEPTIONALLY HIGH TRIGLYCERIDE LEVEL (987 MG/DL) IN COMBINATION WITH THE USE OF PROPOFOL FOR SEDATION. BOTH HYPERTRIGLYCERIDEMIA AND PROLONGED PROPOFOL INFUSION ARE RECOGNIZED RISK FACTORS FOR MEMBRANE LUNG DYSFUNCTION AND IMPAIRED GAS EXCHANGE IN ECMO THERAPY. THIS ASSOCIATION IS SUPPORTED BY SEVERAL PUBLISHED REPORTS AND CASE STUDIES HIGHLIGHTING MECHANICAL FAILURE OF ECMO CIRCUITS DUE TO HYPERTRIGLYCERIDEMIA, PARTICULARLY IN THE SETTING OF CONTINUOUS PROPOFOL ADMINISTRATION WHILE THE DEVICE TESTING CONFIRMED THAT THE HLS SET MET PERFORMANCE SPECIFICATIONS AND NO TECHNICAL FAILURE WAS DETECTED, THE CLINICAL CONTEXT - ALONG WITH THE PUBLICATION EVIDENCE PRESENTED HERE - SUGGESTS THAT PATIENT-SPECIFIC METABOLIC FACTORS, INCLUDING HYPERTRIGLYCERIDEMIA AND THE USE OF PROPOFOL FOR SEDATION, MAY HAVE PLAYED A ROLE IN THE OBSERVED REDUCTION IN OXYGENATION. NOTABLY, THE NEED FOR MULTIPLE HLS SET EXCHANGES DUE TO ELEVATED TRIGLYCERIDE LEVELS FURTHER SUPPORTS THE LIKELIHOOD OF UNDERLYING METABOLIC STRESS. THESE CONDITIONS APPEAR TO BE THE MORE PLAUSIBLE CONTRIBUTORS TO THE PATIENT¿S CLINICAL DETERIORATION RATHER THAN A DEFICIENCY IN THE DEVICE ITSELF. IN CONCLUSION, FROM A TECHNICAL STANDPOINT, THE RETURNED HLS SET DID NOT DEMONSTRATE ANY INHERENT DEVICE MALFUNCTION THAT COULD EXPLAIN THE REPORTED OXYGENATION IMPAIRMENT. ALL RELEVANT FUNCTIONAL TESTS, AFTER REMOVAL OF THIRD-PARTY COMPONENTS, WERE PASSED SUCCESSFULLY. THE BROKEN GAS FILTER FITTING AND CRACKED STOPCOCK, WHILE NON-CONFORMING, WERE NOT LINKED TO GAS EXCHANGE PERFORMANCE FAILURE. CLINICALLY, HOWEVER, THE PRESENCE OF HIGH TRIGLYCERIDES AND LIPID-BASED SEDATIVE (PROPOFOL) USE MAY HAVE CONTRIBUTED TO COMPROMISED MEMBRANE PERFORMANCE AND REDUCTION OF AGGREGATE MEMBRANE AREA VIA LIPID FOULING - A PHENOMENON NOT DIRECTLY DETECTABLE THROUGH STANDARD POST-USE TECHNICAL TESTING. THE PATIENT¿S SEVERE ILLNESS, INSTABILITY, AND COMPLEX CLINICAL PICTURE (INCLUDING FLU A AND MULTIPLE CIRCUIT EXCHANGES WITHIN A SHORT TIMEFRAME) LIKELY CONTRIBUTED TO THE EXPIRATION OF THE PATIENT. LAST, NO DIMINUTION IN PRODUCT PERFORMANCE WAS DETECTED IN INTERNAL TESTING OR SUSPECTED OUTSIDE OF THE REVELATION OF SEVERE HYPERLIPIDEMIA (987 MG/DL)." ACCORDING TO THE INSTRUCTIONS FOR USE OF THE HLS SET, IN CHAPTER "SAFETY INSTRUCTIONS FOR THE HLS SET ADVANCED", IT IS STATED THAT " AN INSUFFICIENT SUPPLY OF POWER AND OXYGEN CAN LEAD TO INADEQUATE PATIENT SUPPORT. ENSURE THERE IS A SUFFICIENT SUPPLY OF MEDICAL OXYGEN. ENSURE THAT THE BATTERIES OF THE DRIVE ARE FULLY CHARGED", AS WELL AS "USE A MEDICAL GAS SUPPLY WITH DRY AIR AND OXYGEN. DO NOT USE AN AIR HUMIDIFIER IN THE GAS SUPPLY". IN CHAPTER "SAFETY INSTRUCTIONS FOR THE OXYGENATOR", IT IS INDICATED THAT "EXCEEDING THE PERMISSIBLE MAXIMUM VALUES DAMAGES THE OXYGENATOR. THIS CAN LEAD TO EMBOLISMS AND INADEQUATE PATIENT SUPPORT. OBSERVE THE PERMISSIBLE MAXIMUM VALUES FOR BLOOD FLOW AND GAS FLOW AS WELL AS PRESSURE ON THE BLOOD SIDE AND ON THE GAS SIDE", AS WELL "FLUSH THE OXYGENATOR REGULARLY TO MAINTAIN OR INCREASE THE GAS EXCHANGE CAPACITY. IF USED FOR LONGER THAN 6 HOURS: FLUSH THE OXYGENATOR AT LEAST ONCE A DAY. CARRY OUT FLUSHING ONLY WHEN THE PATIENT'S BLOOD GAS LEVELS PERMIT. MONITOR THE PATIENT'S BLOOD GAS PARAMETERS CAREFULLY DURING THE FLUSHING PROCESS. DISCONTINUE FLUSHING THE GAS FIBERS IF SIGNS OF HYPOCAPNIA OCCUR". LASTLY, IN THE SAME CHAPTER IT IS SAID THAT "THE OXYGENATOR PERFORMANCE IS RESTRICTED IF THE OXYGENATOR OPERATING POSITION IS INCORRECT OR IF THE GAS FLOW IS OBSTRUCTED. THIS CAN LEAD TO INADEQUATE PATIENT SUPPORT OR AIR EMBOLISMS IN THE PATIENT. ONLY OPERATE THE OXYGENATOR WITH THE GAS INLET AT THE TOP. DO NOT OCCLUDE OR BLOCK THE GAS OUTLET". THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT (SECOND SET) WERE REVIEWED ON 2025-05-26. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "LOW OXYGENATION" COULD NOT BE LINKED TO A DEVICE DEFECT. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING INFLAMMATORY RESPONSE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER AND COPD (CHRONIC OBSTRUCTIVE PULMONARY DISEASE). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAD BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE AND CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
TUBING SET | K112360 | THE EVENT OCCURRED IN USA DURING TREATMENT. IT WAS REPORTED THAT AN HLS SET WAS REPLACED AND AFTER THE REPLACEMENT (A FEW HOURS LATER) THE CUSTOMER NOTICED A DECREASE IN OXYGENATION VALUES. THE PATIENT OXYGENATION WAS IN THE LOW 80S. THE CUSTOMER CHECKED THE GAS SOURCE TO ENSURE PROPER CONNECTION AND THEY INCREASED THE SWEEP TEMPORARILY TO REMOVE ANY POTENTIAL CONDENSATION. THE POST OXYGENATION (PAO2) DECREASED FURTHER AND THE CUSTOMER EXCHANGED THE HLS SET AGAIN. AS THE PATIENT SUFFERED A LOW OXYGENATION, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, SKIN IRRITATION, REDUCED CARDIOPULMONARY RESERVE AND 2 SURGERIES. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED TO FRESENIUS THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP DOCUMENTATION RECEIVED FROM THE PATIENT¿S HOME PROGRAM FACILITY ADMINISTRATOR (FA) CONFIRMED THE PATIENT EXPIRED ON (B)(6) 2025 WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. PER THE FA, THE PATIENT¿S CAUSE OF DEATH IS UNCONFIRMED DUE TO THE PENDING AUTOPSY RESULTS, HOWEVER IT WAS LIKELY A CARDIAC ARREST. ALTHOUGH THE SPECIFICS ARE LARGELY UNKNOWN, IT APPEARS THE PATIENT COMPLETED THEIR LAST CCPD TREATMENT ON (B)(6) 2025 WITHOUT REPORTED ISSUE. PER THE FA¿S RESPONSE, THERE WAS NO ALLEGATION THE SERIOUS ADVERSE EVENTS WERE RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN PD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF A PROBABLE CARDIAC ARREST AND DEATH, AS THE PATIENT WAS CONNECTED TO THE DEVICE WHEN SHE EXPIRED. THE DEFINITIVE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN (AUTOPSY RESULTS AND DEATH CERTIFICATE UNAVAILABLE); THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. HOWEVER, BASED ON FOLLOW-UP DOCUMENTATION FROM THE HOME PROGRAMS¿ FA, THERE WAS NO ALLEGATION OR INDICATION THAT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) WAS DEFICIENT OR MALFUNCTIONED IN ANY WAY. CARDIAC DISEASE IS THE LEADING CAUSE OF DEATH IN DIALYSIS PATIENTS, AND MORTALITY RATES ARE UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION PRESENTLY AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN PATIENT¿S CARDIAC ARREST AND DEATH. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE PATIENT WAS CONNECTED TO THE DEVICE WHEN SHE EXPIRED, THE PENDING AUTOPSY REPORT, AND NO CONFIRMED CAUSE OF DEATH; THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE LIBERTY SELECT CYCLER FROM HAVING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS EXPERIENCED BY THE PATIENT. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
SAFARI2¿ | K151244 | EVENT DESCRIPTION: WHEN THE SAFARI WAS PRESSED AGAINST A PATIENT WITH A SMALL LEFT VENTRICLE, WIRE PERFORATION OCCURRED. PHYSICIAN COMMENTS: PATIENT OUTCOME: NO PATIENT INJURY. INTRODUCER SHEATH USED: NOT AVAILABLE AT THIS MOMENT. GUIDEWIRE USED: YES - UNKNOWN - SAFARI 2 GUIDE. CATHETER USED: NOT AVAILABLE AT THIS MOMENT. BALLOON CATHETER USED: NOT AVAILABLE AT THIS MOMENT. STENT USED: NONE. INFLATION DEVICE USED: NOT AVAILABLE AT THIS MOMENT. IMAGING CATHETER: UNKNOWN. ADDITIONAL INFORMATION 17JAN2025: QUESTION: PLEASE DESCRIBE THE PATIENT'S CURRENT CONDITION. ANSWER: DECEASED ABOUT 1 MONTH AFTER SURGERY (DUE TO ANOTHER COMPLICATION WHILE RECUPERATING THAT WAS NOT DIRECTLY RELATED TO THIS EVENT). QUESTION: PLEASE PROVIDE DETAILS OF THE CIRCUMSTANCES LEADING UP TO THE PERFORATION. ANSWER: AFTER TAVI VALVE IMPLANTATION, WHEN THE TAVI DEVICE IS REMOVED, THE SAFARI WIRE ADVANCES TOWARD THE APEX OF THE HEART, EXPOSING THE WINDING OF THE TIP OF THE SAFARI WIRE TO THE OUTSIDE OF THE HEART. QUESTION: WHAT IS THE LOCATION OF THE PERFORATION? ANSWER: SEVERAL CENTIMETERS HEADWARD FROM THE APEX OF THE HEART. QUESTION: PLEASE DESCRIBE IN DETAIL WHAT KIND OF PROCEDURE WAS PERFORMED OR PLANNED. (DRAINAGE, HEMOSTASIS OF THE PERFORATION BY THORACOTOMY, ETC.) DRAINAGE WAS PERFORMED, HEMOSTASIS OF PERFORATED AREA BY OPEN CHEST, ETC.). ANSWER: HEMOSTASIS BY OPEN CHEST. QUESTION: WHAT IS THE OPINION OF THE DOCTOR IN CHARGE REGARDING THE CAUSE OF THE PERFORATION? ANSWER: THE SAFARI WIRE WAS HELD IN PLACE TO PREVENT IT FROM BEING PULLED OUT WHEN THE TAVI DEVICE WAS REMOVED, BUT THE SAFARI WIRE WENT TOO DEEP. THE PATIENT WAS A VERY ELDERLY WOMAN WITH A SMALL BUILD AND A VERY SMALL LEFT VENTRICLE, SO THE RISK WAS PROBABLY HIGH. QUESTION: WAS THERE ANY PROBLEM WITH THE APPEARANCE OR FUNCTIONALITY OF THE PRODUCT IN QUESTION? ANSWER: NOTHING IN PARTICULAR. QUESTION: SHAPING? ANSWER: NONE. PROCEDURE NAME: TAVI BY LEFT CAROTID APPROACH. TREATMENT SITE: A FEW CENTIMETERS HEADWARD FROM THE APEX. THE DEVICE IN QUESTION WILL NOT BE RETURNED. | PRODUCT WAS NOT RECEIVED FOR EVALUATION; THEREFORE, NO VISUAL OR PHYSICAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE (DFU) WARNINGS SECTION: MONITOR WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. NEVER ADVANCE THE GUIDEWIRE AGAINST RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED IN THE VENTRICLE. THE CURVE OF THE SAFARI2 GUIDEWIRE SHOULD BE CONSTRAINED WITH IN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 1. ENSURE A CATHETER IS APPROPRIATELY PLACED IN THE VENTRICLE OR OTHER INTENDED TREATMENT AREA. 2. CAREFULLY REMOVE THE SAFARI2 GUIDEWIRE, FROM THE HOOP BY GRASPING THE PROXIMAL END OF THE J-STRAIGHTENER SEPARATING THE STRAIGHTENER FROM THE HOOP. 3. AFTER INSPECTION OF THE SAFARI2 GUIDEWIRE, ADVANCE THE J-STRAIGHTENER OVER THE DISTAL SECTION OF THE SAFARI2 WIRE TO STRAIGHTEN THE CURVE. 4. INSERT THE SAFARI2 GUIDEWIRE INTO THE HUB OF THE CATHETER WITH THE AID OF THE J- STRAIGHTENER. 5. CAREFULLY ADVANCE THE DISTAL TIP OF THE SAFARI2 GUIDEWIRE INTO THE LUMEN OF THE CATHETER. 6. REMOVE THE SAFARI2 GUIDEWIRE J- STRAIGHTENER BY WITHDRAWING IT OVER THE LENGTH OF THE SAFARI2 GUIDEWIRE. 7. ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8. CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9. MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. 10. TO REMOVE THE GUIDEWIRE FROM THE TREATMENT AREA: A. A CATHETER MUST BE ADVANCED INTO THE TREATMENT AREA OVER THE SAFARI2 GUIDEWIRE PRIOR TO WITHDRAWING THE WIRE. B. THE SAFARI2 GUIDEWIRE IS PULLED BACK COMPLETELY INTO THE CATHETER. C. THE SAFARI2 GUIDEWIRE MAY THEN BE WITHDRAWN FROM THE CATHETER. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT CLINICAL AND/OR PROCEDURAL FACTORS HAVE CONTRIBUTED TO THE EVENT. CARDIAC PERFORATION, ADDITIONAL SURGICAL PROCEDURE, AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS AND ARE LISTED IN THE SAFARI2 INSTRUCTIONS FOR USE (IFU). DESPITE ATTEMPTS TO GATHER FURTHER DETAILS, NO FURTHER INFORMATION HAS BEEN RECEIVED TO DATE. PER THE DISTRIBUTOR, NO ADDITIONAL INFORMATION IS AVAILABLE REGARDING THIS EVENT. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED. THE SECTIONS LEFT BLANK OR UNKNOWN ON THIS FORM COULD NOT BE COMPLETED DUE TO MISSING INFORMATION. DATE OF DEATH WAS ESTIMATED BASED ON THE REPORT OF THE PATIENT EXPIRED ABOUT ONE MONTH AFTER SURGERY. ALTHOUGH ANNEX G INSTRUCTIONS STATE THAT WITH STAND-ALONE DEVICES ANNEX G TERM SHOULD NOT BE USED, CODE 4755 WAS SELECTED DUE TO COMPONENT CODE (G) SHOWING UP ON THE MISSING DATA REPORT.
GORE-TEX® STRETCH VASCULAR GRAFT | K880167 | REPORTEDLY: UNEXPECTED DEATH, POSSIBLY RELATED TO AV (ARTERIOVENOUS) GRAFT HEMORRHAGE AT ACCESS SITE. PATIENT WAS FOUND UNRESPONSIVE IN ROOM, HEMORRHAGE EVIDENT. DNR (DO-NOT-RESUSCITATE ORDER). 2200-A:-OTHER USED TO CAPTURE HEMORRHAGE. | W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CHRONIC BRONCHITIS AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING COLON AND ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAM STATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, LUNG DISEASE, LUNG CANCER, BLADDER CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAM STATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS AND/OR HEADACHE, NAUSEA / VOMITING, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S SISTER HAS REQUESTED A RETURN LABEL FOR THE DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION, BUT THE EVALUATION IS NOT YET COMPLETE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE DEVICE IS NO LONGER AVAILABLE FOR RETURN, AS IT WAS RECYCLED FOLLOWING THE USER'S PASSING. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT YET BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE USER'S WIFE HAS STATED THAT SHE IS UNSURE OF THE DEVICE'S CURRENT LOCATION AND IS UNABLE TO RETURN IT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
INFINITY DELTA | K152407 | IT WAS REPORTED THAT THE DEVICE DID NOT SOUND THE ASYSTOLE ALARM FAST ENOUGH. IT WAS FURTHER REPORTED THE PATIENT DIED BUT THE CONNECTION BETWEEN THE REPORTED EVENT AND THE PATIENT DEATH COULD NOT BE CLARIFIED YET. | THE LOCAL DRÄGER S&S ORGANIZATION HAS PERFORMED AN ON-SITE EVALUATION OF THE PATIENT MONITOR AND THE CENTRAL MONITORING STATION. THE PATIENT MONITOR HAD A MECHANICAL DAMAGE OF THE HOUSING AND WAS EQUIPPED WITH A NON-DRÄGER INTERNAL BATTERY WHICH WAS DEFECTIVE OR DID OTHERWISE NOT ALLOW OPERATION WITHOUT MAINS SUPPLY. BESIDES THESE NON-CONFORMITIES, MONITORING WAS FUNCTIONAL. THE LOGS FOR THE PARTICULAR MONITORING EPISODE HAD BEEN OVERWRITTEN ALREADY SINCE THE DEVICE REMAINED IN FURTHER USE. THE LOGS OF THE CENTRAL MONITORING STATION HAD BEEN ERASED AS WELL. IN GENERAL, THE REPORTED ASPECT OF DELAYED ALARMING CAN NEITHER BE CONFIRMED NOR DENIED DUE TO THE FACT THAT THE INFORMATION RELEVANT FOR AN ASSESSMENT WAS NOT AVAILABLE. THERE IS HOWEVER NO INDICATION FOR THE POTENTIAL PRESENCE OF A DEVICE MALFUNCTION. THE NON-CONFORMITIES OBSERVED WITH THE DEVICE SHOULD HAVE PREVENTED FROM PUTTING IT INTO USE, BUT THESE HAVE NO RELATION TO THE REPORTED EVENT; MONITORING WAS FUNCTIONAL WHEN THE DEVICE WAS TESTED IN FOLLOW-UP OF THE EVENT. ANOTHER ASPECT THAT SUBSTANTIATES THE POSTULATION THAT NO DEVICE-RELATED ERRORS HAVE LIKELY OCCURRED DURING THE COURSE OF EVENT IS THAT THE DEVICE WAS FURTHER USED AFTERWARDS - AN OPERATOR WOULD PROBABLY NOT DO THIS WHEN SIGNIFICANT CONCERNS AGAINST THE SAFETY AND PERFORMANCE OF THE DEVICE EXIST. A RELIABLE CONCLUSION IN REGARD TO THE EXACT ROOT CAUSE FOR THE REPORTED EVENT CANNOT BE MADE. A NUMBER OF CONTRIBUTING FACTORS MUST BE TAKEN INTO CONSIDERATION WHILE - ON THE OTHER HAND - A DIFFERENTIATION IS NOT POSSIBLE DUE TO LACK OF INFORMATION. FOR EXAMPLE, IF A SENSOR FOR A PARTICULAR MONITORING CHANNEL BECOMES DETACHED FROM THE PATIENT (ECG ELECTRODES, SPO2 SENSOR), THE DEVICE WILL POST A TECHNICAL ALARM. IT IS IN THE NATURE OF THINGS THAT LIFE-THREATENING CONDITIONS WHICH OCCUR IN SENSOR-OFF CONDITION CANNOT BE DETECTED ANYMORE VIA THIS CHANNEL; THE DEVICE WILL NOT POST PHYSIOLOGICAL ALARMS. IT WOULD ALSO BE PLAUSIBLE THAT A LOSS OF NETWORK CONNECTION HAS INTERRUPTED THE FORWARDING OF ALARMS TO THE CENTRAL STATION WHILE THE ALARM POSTED BY THE BEDSIDE MONITOR WAS RECOGNIZED WITH DELAY ONLY. CORRECTION - THE INITIAL REPORT WAS SENT UNDER MDR NUMBER 2510954-2024-00025 IN ERROR. THE CORRECT MDR FOR THIS SUBMISSION IS: 1220063-2025-00008. AFTER CONFIRMATION FROM THE BIOMED ABOUT THE EVENT THAT OCCURRED ON FEBRUARY 2ND. IT TURNS OUT THAT THE DOCTORS HAVE RULED THAT THESE WERE NOT UNDESIRABLE EVENTS, SO THERE IS NO REASON TO QUESTION THE MONITORS. THE CUSTOMER IS NO LONGER ALLEGING THAT A MALFUNCTION OCCURRED. NO FURTHER INFORMATION WAS PROVIDED.
SOLITAIRE FR | K113455 | ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS NOT THOUGHT TO BE RELATED TO MEDTRONIC DEVICE OR THERAPY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REACT CATHETER | K182097 | ADDITIONAL INFORMATION RECEIVED REPORTED THE PATIENT DIED AT HOME. THE CAUSE OF DEATH WAS NOT THOUGHT TO BE RELATED TO MEDTRONIC DEVICE OR THERAPY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
WATCHMAN TRUSTEER ACCESS SYSTEM | K240018 | IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN GAINED ACCESS TO THE PATIENT AND THEN ATTEMPTED TO PERFORM THE TRANSSEPTAL PUNCTURE. THE FIRST ATTEMPT WAS UNSUCCESSFUL, BUT A SECOND ATTEMPT WAS MADE, AND THE PHYSICIAN SUCCESSFULLY COMPLETED THE TRANSSEPTAL PUNCTURE. HOWEVER, IT WAS BELIEVED THE BACK WALL OF THE HEART WAS PUNCTURED AT THIS TIME, BUT NOT SEEN. THE PHYSICIAN THEN EXCHANGED THE TRANSEPTAL SHEATH FOR THE WAS. THE WAS WENT ACROSS THE SEPTUM AND THE UNKNOWN PIGTAIL CATHETER WENT IN. AT THIS TIME PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE WAS NOTED. THE PHYSICIAN PERFORMED A PERICARDIOCENTESIS AND AUTO TRANSFUSED THE BLOOD BACK INTO THE PATIENT. IN ADDITION, THE PATIENT RECEIVED TWELVE (12) UNITS OF BLOOD, EIGHT (8) UNITS OF PLASMA AND ONE (1) UNIT OF PLATELETS. THE PATIENT WAS TAKEN TO THE OPERATING ROOM TO BE TREATED AND ON THE WAY THERE THE PATIENT WENT INTO CARDIAC ARREST. TWO HOURS POST PROCEDURE THE PATIENT WAS UNABLE TO BE RESUSCITATED, AND THE PATIENT DIED.
WATCHMAN FXD CURVE ACCESS SYSTEM | K212228 | IT WAS REPORTED THAT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS DONE IMMEDIATELY PRIOR TO PLACING THE PATIENT UNDER TWILIGHT CONSCIOUS SEDATION. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS PERFORMED WHILE A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) WAS INSERTED WITH VERSACROSS CONNECT (VCC) TO PERFORM TRANSSEPTAL PUNCTURE (TSP). TSP WENT SMOOTHLY AND NO SUSPICION OF EVENT FROM TSP. THE INTRAARTERIAL SEPTUM (IAS) WAS FLOSSED WITH THE WAS AND VCC AND THE WAS AND ICE CATHETER WAS POSITIONED ACROSS THE IAS INTO THE LEFT ATRIUM (LA). DURING A FIVE (5) MINUTE PERIOD, THE WAS AND ICE REMAINED IN THE LA UNATTENDED AS THE PHYSICIAN NEEDED TO LEAVE THE ROOM BRIEFLY, AND WHEN THE PHYSICIAN RETURNED THE WAS NOT ABLE TO BE VISUALIZED ON IMAGING. THE QRS INTERVAL BEGAN WIDENING, BRADYCARDIA WAS NOTED AND THEN ASYSTOLE. AS STAFF WAS PREPARING TO PERFORM CARDIOPULMONARY RESUSCITATION (CPR), VENTRICULAR FIBRILLATION WAS OBSERVED. THE STAFF AND PHYSICIAN CODED THE PATIENT FOR TWO (2) HOURS USING A TEMPORARY PACEMAKER, INTRA-AORTIC BALLOON PUMP (IABP), DIAGNOSTIC CARDIAC CATHETERIZATION (CC), AND DEFIBRILLATION. TEE WAS REINSERTED AND THERE WAS NO EVIDENCE OF PERICARDIAL EFFUSION. THE TEMPORARY PACEMAKER WAS FIRST UNABLE TO CAPTURE BETWEEN COMPRESSION CYCLES AND ULTIMATELY ONE LEAD WAS PLACED AND CAPTURED SUCCESSFULLY. THE DIAGNOSTIC CC REVEALED NO SIGN OF CORONARY DISEASE. ALTHOUGH IABP WAS SUCCESSFULLY PLACED, THE PATIENT MAINTAINED ONLY A WEAK BLOOD PRESSURE AND POOR OXYGEN SATURATION MAINTENANCE AT 70 OR BELOW. DUE TO THE LONG PERIOD OF TIME OF LOW OXYGEN SATURATION, THE PHYSICIAN DECIDED TO STOP RESUSCITATING THE PATIENT AND THE PATIENT DIED. IT WAS NOTED THAT THERE WAS NOT ANY HEAVY SNORING, NO GASPS OF AIR TAKEN BY THE PATIENT, AND NO CONTRAST INJECTIONS PERFORMED WITHIN THE LA. THE CAUSE OF DEATH IS UNKNOWN.
DREAMSTATION BIPAP AUTOSV | K090539 | A DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAD DECEASED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION DUE TO A USER OF DREAMSTATION AUTO CPAP. WIFE CALLED IN STATING PATIENT PASSED AWAY LAST 2022. SHE SAID THAT THE DEVICE HAS BEEN RETURNED AND NO LONGER HAS THE UNIT OR TRACKING DETAILS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
GORE® TRI-LOBE BALLOON CATHETER | K081799 | THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT OF A THORACIC AORTIC ARCH ANEURYSM USING GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM AND A NON-GORE DEVICE (ZENITH ALPHA EXTENSION). WHEN A 22FR GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS ADVANCED, REPORTEDLY THERE WAS RESISTANCE. TGMR313120J WAS IMPLANTED FIRST, THEN TGM343420J WAS IMPLANTED IN ZONE 2. HOWEVER, TGM343420J MOVED DISTALLY AND IMPLANTED IN ZONE 3. THEN, THE NON-GORE DEVICE WAS IMPLANTED PROXIMALLY. WHEN A TOUCH-UP USING A GORE® TRI-LOBE BALLOON CATHETER WAS PERFORMED TO THE NON-GORE DEVICE, THE BLOOD PRESSURE DROPPED, THE PATIENT LOST CONSCIOUSNESS AND THE PATIENT WAS INTO CARDIAC ARREST. A DRAINAGE AND A CARDIAC MASSAGE WERE ATTEMPTED, BUT THE PATIENT DIED. THE PHYSICIAN STATED THAT A RETROGRADE THORACIC AORTIC DISSECTION OCCURRED DUE TO THE TOUCH-UP TO THE NON-GORE DEVICE USING THE GORE® TRI-LOBE BALLOON CATHETER AND IT CAUSED A CARDIAC TAMPONADE. REPORTEDLY, THE GORE® TRI-LOBE BALLOON CATHETER WAS INFLATED WITH THE BALLOON WAS LOCATED INSIDE THE STENT GRAFT. | H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE VERIFIED THAT THE LOT MET ALL PRE-RELEASE SPECIFICATIONS. H6: CODE D12: ACCORDING TO THE GORE® TRI-LOBE BALLOON CATHETER INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO: TRAUMA TO THE VESSEL WALL, INCLUDING SPASM, DISSECTION, PERFORATION OR RUPTURE, DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
PHILIPS CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTIONS WERE REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | ADDITIONAL INFORMATION WAS RECEIVED ON 05/26/2025 AND H11 IS UPDATED AS: THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE PREVIOUSLY CALLED TO ASK FOR ALL COMMUNICATION TO STOP, BUT SHE IS STILL RECEIVING EMAILS AND LETTERS. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR THE PATIENT TO BE REMOVED FROM THE MAILING LIST. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY THREE YEARS AGO. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS BEEN MISPLACED AND WILL NOT BE RETURNED FOR INVESTIGATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
INTELLIVUE MX450 PATIENT MONITOR | K130849 | IT WAS REPORTED THE STAFF REPORTED NOT HEARING ALARMS FOR DESATURATION FROM THE STANDALONE MONITOR, AND THE NICU PATIENT PASSED AWAY. THE CUSTOMER REQUESTED ASSISTANCE WITH THE AUDIT LOGS. THE LOGS WERE INITIALLY REVIEWED BY WITH THE CUSTOMER, REVEALING THE CAUSE OF NO ALARM WAS THE ALARMS HAD BEEN PAUSED BY THE USER. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. THIS INFORMATION INDICATED THE PATIENT WAS AN EMERGENCY 32-WEEK GESTATION C-SECTION DELIVERY WITH ALVEOLAR HEMORRHAGING AND CARDIAC ARREST. THE PATIENT ALMOST PASSED IN THE OR BUT STAFF WAS ABLE TO GET THE PATIENT TO THE NICU. THE PATIENT WAS MONITORED FOR ECG AND SPO2 AND STAFF INDICATED THEY DID NOT HEAR ALARMS FROM THIS STANDALONE MONITOR AND THE PATIENT SUBSEQUENTLY PASSING AWAY. AN INTERNAL INVESTIGATION WAS PERFORMED WITH THE MONITOR WITH ITS EXISTING LEADS AND CABLES USED DURING THE TIME OF THE EVENT. ALARMS WERE GENERATED NORMALLY WITH APPROPRIATE VOLUME SETTINGS; HOWEVER, IT WAS NOTED DURING LOG REVIEW THAT THE ALARMS HAD BEEN PAUSED BY USER, WITH THE PAUSE BEING THE DEFAULT OF 2 MINUTES. THE PAUSE TIME COULD NOT BE ADJUSTED AS THE CONFIGURED DEFAULT WAS 2 MINUTES. BASED ON THIS INFORMATION, THE DEVICE FUNCTIONED AS CONFIGURED; HOWEVER, USE ERROR AND ALARM MANAGEMENT MAY HAVE BEEN FACTORS. IF ADDITIONAL INFORMATION IS OBTAINED, PLEASE REQUEST RE-ASSESSMENT OF THE RECORD BY THE PMS CLINICAL EXPERT. | PHILIPS RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT WAS AN EMERGENCY 32 WEEK DELIVERY THAT HAD ALVEOLI HEMORRHAGING AND NEONATE CARDIAC ARREST. THE PATIENT HAD TO BE DELIVERED PREMATURELY VIA C-SECTION DUE TO COMPLICATIONS. PATIENT ALMOST PASSED IN THE OR ROOM, BUT THEY WERE ABLE TO STABILIZE AND GET THE PATIENT TO THE NICU. THE FIELD SERVICE ENGINEER (FSE) MET WITH THE HOSPITAL RISK MANAGER TO EVALUATE THE DEVICE FROM THE EVENT. THE MONITOR SPEAKER WAS FUNCTIONING NORMALLY, AND THE SPEAKER SETTING WAS AT 7 AND RED ALARM WAS +1. THEY TESTED THE EXISTING CABLES AND LEADS THAT WAS USED DURING THE TIME OF THE EVENT. THE MONITORS SPO2 AND ECG ALARMED NORMALLY. RISK MANAGER ALSO WANTED TO KNOW WHAT THE "PAUSE ALARM" TIME LAPSE WAS AND/OR IF IT COULD BE CHANGED AFTER PRESSING THE "PAUSE ALARM" KEY. THE "PAUSE ALARM" WAS SET AT THE DEFAULT OF 2 MINUTES, AND THIS MONITOR CONFIG DIDN''T PROVIDE THE ABILITY TO CHOSE THE "PAUSE ALARM" TIME. IT WAS ALSO STATED THE DEVICE WAS NOT CONNECTED TO A PIC IX NETWORK. THIS DEVICE WAS CONNECTED TO A CAPSULE DEVICE. IT WAS FOUND THAT THE STAFF NEVER ADMITTED THE PATIENT INTO THE MONITOR AND NEVER DID AN "END CASE" AFTER THIS PATIENT OR THE PATIENTS PRIOR TO OR AFTER THIS EVENT. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL PRODUCT SPECIALIST REVIEWED THE LOG DATA. THE ALARMS SHOWED THAT THE STAFF ACKNOWLEDGED THE ALARMS. THE ALARM REVIEW HAS THE HISTORY OF THE ALARM NOTIFICATIONS WITH A MAXIMUM OF 300 ENTRIES. FIRST IN FIRST OUT STORAGE OF ALARMS. SO TIME IS OF ESSENCE, SINCE THE OLDEST ALARMS WILL BE OVERWRITTEN ONCE THE MAXIMUM IS REACHED. A REPORT CAN BE PRINTED IN BOTH PORTRAIT AND LANDSCAPE ORIENTATION. CAUSE: ALARMS WERE PAUSED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED WE WERE UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE PAUSE TIME COULD NOT BE ADJUSTED AS THE CONFIGURED DEFAULT WAS 2 MINUTES. BASED ON THIS INFORMATION, THE DEVICE FUNCTIONED AS CONFIGURED; HOWEVER, USE ERROR AND ALARM MANAGEMENT MAY HAVE BEEN FACTORS. THE CUSTOMER WAS PROVIDED WITH INFORMATION FROM REVIEW. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DS2ADV AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION, AND NO ISSUES WERE IDENTIFIED DURING THE ASSESSMENT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS BEING USED IN A NURSING HOME, AND A FAMILY MEMBER INDICATED THAT, HAD THEY STILL HAD POSSESSION OF IT, THEY WOULD HAVE RETURNED IT FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
MITRACLIP | K190167 | IT WAS REPORTED THAT ON (B)(6) 2025, A SCHEDULED MITRACLIP PROCEDURE WAS PERFORMED TO TREAT FUNCTIONAL MITRAL REGURGITATION (MR) GRADE 4. PT WAS ALLERGIC TO HEPARIN. ANGIOMAX WAS USED TO ANTICOAGULATE PATIENT AND ANTICOAGULANT CITRATE DEXTROSE SOLUTION USP (ACD) WAS USED TO PREP THE STEERABLE GUIDE CATHETER (SGC). THE SGC WAS INSERTED INTO THE PATIENT. ONCE THE SGC CROSSED THE SEPTUM AND THE DILATOR WAS REMOVED, AIR WAS NOTED IN THE SGC. THE SGC WAS TAKEN OFF THE STABILIZER AND HELD BELOW THE LEVEL OF THE PATIENT'S HEART AND ASPIRATION WAS PERFORMED UNTIL BLOOD FILLED THE HEMOSTASIS VALVE CHAMBER. PATIENT DEVELOPED ST ELEVATION IN ALL LEADS WITH BRADYCARDIA AND HYPOTENSION 70/40. ANESTHETIC MEDICATION WAS ADMINISTERED TO TREAT THE LOW BLOOD PRESSURE, BUT DURING THIS THE PATIENT'S RHYTHM SLOWED INTO ASYSTOLE. CARDIO-PULMONARY RESUSCITATION (CPR) WAS PERFORMED, AFTER WHICH THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA. CARDIOVERSION WAS THEN PERFORMED WHICH RECOVERED THE RHYTHM BUT IT WAS TACHYCARDIC GREATER THAN 130 BEATS PER MINUTE (BPM). BLOOD PRESSURE STILL REMAINED LOW IN 70/40 RANGE. ANESTHESIA CONTINUED TO WORK ON THE HYPOTENSION WHILE A CORONARY ANGIOGRAPHY WAS PERFORMED WHICH REVEALED CORONARIES WERE PATENT. ECHOCARDIOGRAM WAS ALSO PERFORMED WHICH SHOWED NO EFFUSION AND DECREASED VENTRICULAR FUNCTION FROM APPROXIMATELY 40% AT BASELINE TO 10-15%. BASAL SECTION OF THE HEART WAS CONTRACTING WELL BUT BOTH THE RIGHT AND LEFT VENTRICULAR FREE WALLS AND APEX CONTRACTILITY WAS DIMINISHED. NO AIR WAS OBSERVED IN THE ANATOMY. ANESTHESIA WAS ABLE TO RECOVER BLOOD PRESSURE SO THE PATIENT WAS SENT TO THE INTENSIVE CARE UNIT (ICU) FOR RECOVERY. LATER ON THE SAME NIGHT, THE PATIENT DIED. THE FAMILY DID NOT WANT EXTRAORDINARY MEASURES TO BE TAKEN. THE CAUSE OF THE PATIENT DETERIORATION DURING THE PROCEDURE WAS REPORTED AS UNKNOWN. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY DID NOT INDICATE A LOT-SPECIFIC QUALITY ISSUE. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED LEAK, DEATH, ST ELEVATION IN ALL LEADS, BRADYCARDIA, AND SUBSEQUENT HYPOTENSION. THE PATIENT EFFECTS OF MYOCARDIAL INFARCTION, CARDIA ARREST, TACHYCARDIA, AND HEART FAILURE APPEAR TO BE CASCADING EFFECTS. DEATH, BRADYCARDIA, HYPOTENSION, MYOCARDIAL INFARCTION, CARDIA ARREST, TACHYCARDIA, HEART FAILURE, AND NON-SPECIFIC EKG/ECG CHANGES ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED MEDICATION REQUIRED AND UNEXPECTED MEDICAL INTERVENTIONS WERE RESULTS OF CASE SPECIFIC CIRCUMSTANCES AS ANESTHETIC MEDICATION WAS ADMINISTERED TO TREAT THE LOW BLOOD PRESSURE, ADDITIONAL ASPIRATION WAS PROVIDED, CARDIO-PULMONARY RESUSCITATION (CPR) AND CORONARY ANGIOGRAPHY WERE PERFORMED. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE USER PASSED AWAY YEARS AGO. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE WAS GIVEN TO A FRIEND AND IS NOT AVAILABLE TO BE RETURNED FOR EVALUATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP AS THE PATIENT PASSED AWAY LAST YEAR. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO BIPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | IT WAS REPORTED THAT A PATIENT WITH BREAST CANCER UNDERWENT A DA VINCI-ASSISTED BILATERAL NIPPLE-SPARING MASTECTOMY WITH IMMEDIATE RECONSTRUCTION IN (B)(6) 2018. THE PATIENT PASSED AWAY 2 YEARS POST-PROCEDURE. THE REPORTER DID NOT PROVIDE A CAUSE OF DEATH. THERE WAS NO REPORT THAT ANY DA VINCI DEVICE MALFUNCTION OR INTRAOPERATIVE ISSUES OCCURRED DURING THE PROCEDURE. REQUESTS FOR ADDITIONAL INFORMATION FROM THE OPERATING SURGEON HAVE BEEN SENT; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE DEVICE LOG FOUND THAT THE FOLLOWING INSTRUMENTS WERE USED IN THE PROCEDURE: ENDOSCOPE, TWO MONOPOLAR CURVED SCISSORS, TWO FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. ALL INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES WITH THE EXCEPTION OF TWO OF THE INSTRUMENTS; ONE SINGLE-USE INSTRUMENT, AND ONE INSTRUMENT THAT HAD REACHED ITS FINAL LIFE USE LIMIT WITH THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. A PATIENT DEATH WAS REPORTED 2 YEARS AFTER A BILATERAL NIPPLE SPARING MASTECTOMY. NO OTHER INTRAOPERATIVE OR POST OPERATIVE DETAILS ARE PROVIDED. HOW THEY DIED OR THE EVENTS LEADING TO THEIR DEATH ARE NOT PROVIDED. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THE PATIENT EVENT. A REVIEW OF THE DEVICE LOG FOUND THAT THE FOLLOWING INSTRUMENTS WERE USED IN THE PROCEDURE: ENDOSCOPE, TWO MONOPOLAR CURVED SCISSORS, TWO FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. ALL INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES EXCEPT FOR TWO OF THE INSTRUMENTS THAT WERE EXPIRED AFTER THIS PROCEDURE AND SINGLE-USE INSTRUMENT. SECTION H10 OF THIS REPORT (RELATED REPORT NUMBER) FOR THE RECEIVED VOLUNTARY MEDWATCH REPORT NUMBER (MW5164582). SECTION E: THE USER FACILITY WHERE THE SURGERY WAS PERFORMED IS DOCUMENTED. THE REPORTER'S NAME AND CORRESPONDING INFORMATION IS NOT AFFILIATED WITH THE USER FACILITY.
DEVILBISS HEALTHCARE | K071397 | DEVILBISS HEALTHCARE WAS NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY SERVICE OF A LEGAL COMPLAINT, WHICH ALLEGES THAT THE CONCENTRATOR WAS EMITTING NOISE AND ALARMING, BUT THAT THE DME PROVIDER DID NOT COME TO SERVICE, REPAIR OR REPLACE THE UNIT UNTIL SEVEN DAYS AFTER BEING NOTIFIED OF THE PROBLEM, WHICH ALLEGEDLY RESULTED IN THE PATIENT HAVING LOW OXYGEN SATURATION, DESPITE THE PATIENT ALSO DEMONSTRATING LOW SATURATION ON A DIFFERENT PORTABLE DEVICE. THE ALARMING DEVICE WAS REPLACED, AND IT WAS REPORTED THAT THE PATIENT PASSED AWAY THREE DAYS LATER. DEVILBISS HEALTHCARE WILL COOPERATE AND PARTICIPATE IN THE LITIGATION PROCESS TO GATHER FURTHER INFORMATION.
CLOTTRIEVER XL CATHETER | K242557 | THE PATIENT WAS A 27-YEAR-OLD MALE WITH DEEP VEIN THROMBOSIS. THE MAJORITY OF THE PATIENT'S THROMBUS HAD BEEN REMOVED USING A TRIEVER20. IMAGING WAS PERFORMED WHICH REVEALED RESIDUAL CLOTS THE PHYSICIAN CHOSE TO ADDRESS WITH THE CLOTTRIEVER XL CATHETER (CTXL). NO INFERIOR VENA CAVA (IVC) FILTER OR PROTECTION WAS SEEN. THE CTXL WAS USED TO PERFORM A PASS IN THE COMMON FEMORAL VEIN. AFTER REMOVAL AND CLEANING THE BASKET, THE PATIENT'S O2 DROPPED TO THE LOW 90S. AFTER A COUPLE OF DEEP BREATHS, THE PATIENT'S O2 CAME BACK UP, AND THE PHYSICIAN PERFORMED ANOTHER PASS. AFTER THIS OCCURRED, THE PATIENT'S O2 DROPPED AGAIN BUT MORE SIGNIFICANTLY. A CODE WAS CALLED, AND THE PATIENT WAS FLIPPED ONTO HIS STRETCHER. MULTIPLE ROUNDS OF RESUSCITATION MEDICATIONS WERE ADMINISTERED TO THE PATIENT. RESUSCITATION EFFORTS (CPR) WERE CONTINUED FOR 45 MINUTES BEFORE THE CASE WAS CALLED AND THE PATIENT EXPIRED. | THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE PROVIDED, AND THE LOT HISTORY RECORDS WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, OR DESIGN. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS THE RESULT OF INADVERTENT CLOT EMBOLISM. DISTAL EMBOLIZATION OF BLOOD CLOTS IS IDENTIFIED IN THE DEVICE LABELING AS POSSIBLE ADVERSE EVENTS/COMPLICATIONS. MANUFACTURER REFERENCE #: (B)(4).
TRILOGY EVO | K181166 | THE MANUFACTURER WAS NOTIFIED OF A PATIENT DEATH WHILE THEY WERE ON VENTILATOR THERAPY USING A TRILOGY EVO VENTILATOR. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, IT WAS ALLEGED THAT THE PATIENT WITH PRIMARY DIAGNOSES OF CHRONIC RESPIRATORY FAILURE, COPD EXPIRED ON (B)(6) 2024 DURING CLINICAL AND THERAPEUTIC USE OF THE DEVICE. THE PATIENT WAS FOUND DECEASED WITH THE NON-INVASIVE INTERFACE ATTACHED AND THE MACHINE POWERED OFF. THERE HAS BEEN NO ALLEGATION OF DEVICE MALFUNCTION AT THIS TIME. IT IS UNKNOWN AT THIS TIME IF THE DEVICE ALARMED AT THE TIME OF THE EVENT. THE DEVICE SETTINGS AT THE TIME OF THE ALLEGED INCIDENT WERE AVAPS MODE VT500 8-10 RATE AUTO PS MIN 8 PS MAX 25. THE SECURE DATA CARD WILL NOT BE RETURNING WITH THE DEVICE BECAUSE THIS ITEM IS NOT APPLICABLE. THE CIRCUIT WILL NOT BE RETURNED WITH THE DEVICE. THE DEVICE HAS NOT BEEN RETRIEVED FOR FURTHER INVESTIGATION THUS FAR. | THE DEVICE IN QUESTION HAS BEEN RETRIEVED AND EVALUATED BY THE MANUFACTURER. DEVICE DRPT (DIAGNOSTIC REPORT) LOGS WERE INVESTIGATED FINDING NO MALFUNCTION OR FAILURE OF THE DEVICE TO PERFORM TO MANUFACTURER DECLARED SPECIFICATIONS. FURTHER ANALYSIS OF THE DEVICE INTERACTION TIME SIGNATURES FOUND THAT DURING THE DATE OF THE EVENT IN QUESTION (26-DEC-2024) THE DEVICE HAD BEEN IN STANDBY MODE AND WAS NOT PROVIDING THERAPY FROM 0730 (B)(6) 2024 UNTIL THERAPY WAS INITIATED AT 0905 (B)(6) 2024. THE DEVICE UNDERWENT FURTHER EVALUATION AND TESTING WITH THERAPY BEING PERFORMED FOR APPROXIMATELY 15.5 HOURS WITH NO FAILURE OR MALFUNCTIONS NOTED. BASED UPON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, NO CAUSAL RELATIONSHIP OF THE DEVICE TO THE PATIENT OUTCOME OF DEATH HAS BEEN NOTED. HOWEVER, CONTRIBUTION OF THE DEVICE TO THE PATIENT OUTCOME OF DEATH HAS BEEN CONFIRMED: FORESEEABLE USE ERROR - FAILURE TO TRANSITION THE DEVICE FROM STANDBY TO THERAPY STATUS-POST PATIENT CONNECTION TO THE DEVICE. DEVICE INHERENT SAFETY DESIGN PERMITS THE PATIENT TO BREATH SPONTANEOUSLY IN THE EVENT SUCH A USE ERROR OCCURS.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT WAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFICIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER AND KIDNEY SHUTDOWN AND DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | ADDITIONAL INFORMATION RECEIVED FROM THE HOSPITAL RISK MANAGEMENT DEPARTMENT STATED, "[THE INFORMATION SUMMARY YOU PROVIDED] "IS NOT ACCURATE. THERE WAS NO INJURY TO A PATIENT INVOLVING DA VINCI EQUIPMENT." NO FURTHER INFORMATION WAS PROVIDED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT'S WIFE NO LONGER HAS THE DEVICE AS HER HUSBAND PASSED AWAY TWO YEARS AGO. NO ADDITIONAL DETAILS HAVE BEEN PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER HAS BEEN NOTIFIED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT'S WIFE CONTACTED US TO REPORT THAT SHE HAD RECEIVED A LETTER REGARDING A REPLACEMENT DEVICE FOR HER HUSBAND, WHICH HE NEVER RECEIVED. SHE CLARIFIED THAT HER HUSBAND HAS PASSED AWAY AND NO LONGER REQUIRES A REPLACEMENT. NO FURTHER DETAILS HAVE BEEN PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
SERVO-S BASE UNIT | K123149 | IT WAS REPORTED THAT THE VENTILATOR WAS NOT CONNECTED TO O2 SUPPLY WHILE ON PATIENT. THE PATIENT DIED THE DAY AFTER THE EVENT. MANUFACTURER'S REF.#: (B)(4). | THE UDI INFORMATION IS NOT AVAILABLE SINCE THE DEVICE WAS MANUFACTURED BEFORE 2015.
2008T HD SYS. CDX W/BIBAG BLUE STAR | K173972 | IT WAS REPORTED TO FRESENIUS THAT THIS UNKNOWN PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON HEMODIALYSIS (HD) UTILIZING A 2008T HEMODIALYSIS SYSTEM FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING OUTPATIENT HD THERAPY (DATE NOT PROVIDED DURING THE INTAKE PROCESS). ON (B)(6) 2025, THIS UNKNOWN PATIENT REQUESTED HD THERAPY BE PAUSED TO UTILIZE THE RESTROOM. WHILE RETURNING THE PATIENT¿S BLOOD, IT WAS REPORTED THE PATIENT BECAME UNRESPONSIVE. THE HD REGISTERED NURSE (HDRN) REPORTEDLY INITIATED CARDIOPULMONARY RESUSCITATION (CPR) MEASURES, HOWEVER THE PATIENT EXPIRED. THE 2008T HEMODIALYSIS SYSTEM WAS SEQUESTERED POST-EVENT, HOWEVER NO DOCUMENTATION REGARDING THE ANY POST-EVENT DEVICE TESTING WAS OBTAINED. IT SHOULD BE NOTED THAT PRIOR TO TREATMENT, THE 2008T HEMODIALYSIS SYSTEMS¿ CONDUCTIVITY WAS WITHIN ACCEPTABLE LIMITS, NO PRETREATMENT ALARMS WERE NOTED, AND THERE WERE NO DEFECTS ON EITHER THE BLOODLINES OR DIALYZER NOTED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, DEATH CERTIFICATE, TREATMENT RECORDS, MACHINE RECORDS) WERE UNSUCCESSFUL. | PLANT INVESTIGATION: NO PARTS WERE RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. ADDITIONALLY, NO ON-SITE EVALUATION WAS PERFORMED BY A FRESENIUS FIELD SERVICE TECHNICIAN (FST). A RECORDS REVIEW WAS PERFORMED ON THE REPORTED SERIAL NUMBER. AN INVESTIGATION OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO NON-CONFORMANCES, OR ANY ASSOCIATED REWORK IDENTIFIED DURING THE MANUFACTURING PROCESS WHICH COULD BE RELATED TO THE REPORTED EVENT. IN ADDITION, THE DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. THE INVESTIGATION INTO THE CAUSE OF THE REPORTED PROBLEM WAS NOT ABLE TO BE CONFIRMED. A DEFINITIVE CONCLUSION REGARDING THE COMPLAINT INCIDENT CANNOT BE REACHED WITHOUT A PHYSICAL EXAMINATION OF THE COMPLAINT DEVICE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | ON (B)(6) 2025, A HOME CARE FACILITY ADMINISTRATOR AND PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PDRN, IT WAS REPORTED THIS PATIENT EXPIRED ON (B)(6) 2025 AT HOME DUE TO A CARDIAC ARREST, ATTRIBUTED TO MULTIMORBIDITY THAT INCLUDED A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE. IT WAS EXPLAINED THAT THE PATIENT HAD BEEN IN AND OUT OF THE HOSPITAL IN RECENT MONTHS FOR A DECLINE IN OVERALL HEALTH AND FOR REASONS UNRELATED TO PD THERAPY OR THE USE OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). THE PATIENT WAS FOUND UNRESPONSIVE BY FAMILY ON THE MORNING OF (B)(6) 2025 WHILE STILL CONNECTED TO HER CYCLER. EMERGENCY SERVICES WERE NOT CONTACTED, AND THE PATIENT WAS PRONOUNCED DECEASED IN HER HOME WITHOUT TRANSPORT TO THE HOSPITAL. IT WAS CONFIRMED THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO HER DEATH. THE CAUSE OF THIS PATIENT¿S CARDIAC ARREST AND SUBSEQUENT DEATH CAN BE ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE WITH MULTIMORBIDITY. THERE WAS NO INDICATION OF AN ASSOCIATED DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS CONFIRMED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOVASCULAR DISEASE. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE CUSTOMER ALLEGES THE PATIENT HAS PASSED AWAY, DID NOT RECEIVE REPLACEMENT DEVICE. NO MEDICAL INTERVENTION WAS REQUIRED BY THE PATIENT. THE MANUFACTURER¿S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION IS COMPLETE.
EDWARDS ESHEATH+ INTRODUCER | K200258 | AS REPORTED THROUGH THE JAPANESE TAVI REGISTRY, DURING A TRANSFEMORAL TAVR PROCEDURE WITH A 23 MM SAPIEN 3 ULTRA RESILIA VALVE, THERE WAS DAMAGE FROM THE RIGHT COMMON ILIAC ARTERY (CIA) TO THE ABDOMINAL AORTA OBSERVED, AND A STANFORD TYPE A AORTIC DISSECTION. THERE WAS ALSO RETROPERITONEAL BLEEDING. AN EMERGENCY LAPAROTOMY ARTIFICIAL BLOOD VESSEL REPLACEMENT WAS PERFORMED. AFTER THE PROCEDURE, THE PATIENT'S HYPOTENSION PERSISTED, AND A NON-OCCLUSIVE MESENTERIC INFARCTION OCCURRED. THE PATIENT DIED 3 DAYS AFTER THE TAVR PROCEDURE. | HEALTH EFFECT - CLINICAL CODE: AORTIC DISSECTION. PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. ACCORDING TO THE DEVICE TRAINING MANUALS, RISK FACTORS FOR AORTIC DISSECTION, CARDIAC/AORTIC HEMATOMA, OR ANNULAR RUPTURE DURING THE TVR PROCEDURE INCLUDE SIGNIFICANT VALVE OVER SIZING, SEVERELY OBLITERATED SINUSES OF VALSALVA, PORCELAIN AORTA AND/OR PRESENCE OF BULKY CALCIFICATION, AND NARROW CALCIFIED SINOTUBULAR JUNCTION. IN ADDITION, ADVANCED AGE, FEMALE GENDER, SMALL BODY WEIGHT, AND STEROID DEPENDENCY CAN ALSO BE CONTRIBUTING FACTORS. THE SAPIEN TRANSCATHETER HEART VALVE (ALL MODELS) RELIES CALCIUM TO SECURELY ANCHOR IN THE LANDING ZONE. DESPITE THIS BENEFICIAL ASPECT OF CALCIUM, BULKY CALCIUM CAN INCREASE THE RISK OF CALCIFIC NODULE DISPLACEMENT INTO THE VASCULATURE, WHICH CAN LEAD TO VASCULAR INJURY. AT TIMES, THE EXTENT AND DISTRIBUTION OF CALCIUM CAN IMPAIR EASE OF DELIVERY OF THE VALVE, CORRECT POSITIONING OF THE VALVE, DEPLOYMENT OF THE VALVE, AND PROCEDURAL SUCCESS. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THIS ADVERSE EVENT. ALTHOUGH THE EXACT CAUSE AND SOURCE OF THE AORTIC DISSECTION/CARDIOVASCULAR INJURIES CANNOT BE DETERMINED, IT IS POSSIBLE THE MECHANISMS OF THE TAVR PROCEDURE DESCRIBED ABOVE OR PATIENT FACTORS NOT PROVIDED MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
SERVO-S BASE UNIT | K123149 | IT WAS REPORTED THAT THE VENTILATOR WAS NOT CONNECTED TO O2 SUPPLY WHILE ON PATIENT. THE PATIENT DIED THE DAY AFTER THE EVENT. MANUFACTURER'S REF.#: (B)(4). | THE UDI INFORMATION IS NOT AVAILABLE SINCE THE DEVICE WAS MANUFACTURED BEFORE 2015.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON SHIPPING BACK A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION, AN ERROR 53 WAS DISCOVERED IN THE DEVICE LOGS. THE PCA WAS REPLACED. DEVICE WAS TESTED AND ALL TESTING PASSED. THE DEVICE WAS RETURNED TO STOCK.
INTELLIVUE MP70 | K151681 | IT WAS REPORTED THE MONITOR DID NOT ALARM FOR ST CHANGES WHEN THE PATIENT WAS HAVING A STEMI (ST-ELEVATION MYOCARDIAL INFARCTION). WHEN NURSING NOTICED CHANGES IN ECG, A 12-LEAD ECG WAS PERFORMED TO CONFIRM THE STEMI, BUT THE PATIENT SUBSEQUENTLY PASSED AWAY. IT WAS INDICATED THE USERS NOTICED ST ELEVATION ON THE MONITOR AND, AFTER REVIEWING PATIENT DATA, DETERMINED THE PATIENT HAD BEEN HAVING ST ELEVATION FOR APPROXIMATELY 11 HOURS WITH NO ALARMS, THOUGH ST WAS TURNED ON. DURING THIS TIME, OTHER ALARMS WERE BEING GENERATED FOR LOW SPO2, LOW BLOOD PRESSURE, AND HEART RATE. | THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED. THE FSE USED A SIMULATOR TO TEST THE FUNCTIONALITY OF ST AND THE DEVICE ALARMED AS EXPECTED. THE SPEAKER TESTING PASSED, AS WELL. THE CLINICAL CONFIGURATION AND CLINICAL AUDIT LOGS WERE OBTAINED BY THE FSE, HOWEVER THE PHYSIO DATA WAS NO LONGER AVAILABLE SINCE OVER SEVEN DAYS ELAPSED. THE LOG DATA WAS REVIEWED BY THE FSE AND THE CSDC AND IT WAS VERIFIED THAT THERE WAS NO ST ALARMING FOUND IN THE LOGS. THE CSDC EXPLAINED TO THE CUSTOMER THAT ST ALARMS ARE ALWAYS YELLOW ALARMS AND REQUIRE 60 SECONDS OF ST ELEVATION BEFORE ALARMING. THE CSDC FURTHER EXPLAINED TO THE CUSTOMER THAT LEAD PLACEMENT OF THE V LEAD MAY HAVE BEEN INACCURATE AND, BECAUSE OF THIS, THE ST MAY NOT HAVE BEEN AS HIGH HAS +2.0. DUE TO THE LIMITED WAVE DATA, LESS THAN 60 FULL SECONDS, FURTHER ANALYSIS IS NOT POSSIBLE AND THE ISSUE CANNOT BE CONFIRMED. THE CUSTOMER DECLINED TO QUARANTINE THE DEVICE AND FOUND THE TESTING AND EXPLANATION SATISFACTORY. THE DEVICE REMAINS IN USE.
SPECTRUM IQ | K222048 | IT WAS REPORTED FROM THE HOSPITAL¿S BIOMED DEPARTMENT, A SPECTRUM IQ INFUSION PUMP DID NOT HAVE A BATTERY CONNECTED DURING AN INFUSION AND THE PATIENT SUBSEQUENTLY DIED. DURING AN UNSPECIFIED INFUSION, THE SPECTRUM IQ PUMP WAS PLUGGED INTO THE WALL SOCKET. AT AN UNKNOWN POINT DURING THE INFUSION, THE PATIENT HAD TO BE ¿MOVED¿, AND THE PUMP WAS UNPLUGGED FROM THE WALL SOCKET. AT THIS TIME, IT WAS NOTED THERE WAS NO BATTERY PRESENT INSIDE THE SPECTRUM PUMP. THE CUSTOMER COULD NOT CLARIFY IF THE PUMP ALARMED THE ¿INSTALL BATTERY" ALARM. THIS ALARM WILL TRIGGER WHEN THE PUMP IS TURNED ON AND WILL CONTINUE TO ALARM UNTIL THE USER ACKNOWLEDGES THE ALERT. IF THE USER ACKNOWLEDGES THIS ALERT AND THERE IS STILL NO BATTERY INSTALLED; THE ONSCREEN ICON WILL FLASH CONTINUOUSLY UNTIL A BATTERY IS INSTALLED. AT SOME POINT IN TIME, THE NURSE NOTED THE PUMP WAS NOT WORKING AND ¿FOUND¿ A BATTERY. AFTER AN UNKNOWN AMOUNT OF TIME THE PATIENT DIED. THE REPORTER WAS NOT INFORMED HOW LONG AFTER THE PUMP WAS UNPLUGGED BEFORE THE PATIENT DIED. IT WAS NOT REPORTED WHAT SOLUTION WAS BEING INFUSED, IF ANY LIFE SUSTAINING TREATMENT WAS PROVIDED, THE CAUSE OF DEATH, OR IF AN AUTOPSY WAS PERFORMED. NO FURTHER INFORMATION WAS AVAILABLE AT THE TIME OF THIS REPORT. | THE DEVICE HAS BEEN RECEIVED, AND THE EVALUATION IS IN PROGRESS. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION FROM A DEALER REGARDING INFORMATION THEY HAD RECEIVED FROM A FATHER OF THE PATIENT WHO HAD PASSED AWAY. THE DEALER CONTACTED THE FATHER OF THE PATIENT TO ASK THE VENTILATOR'S DATA, BUT "THE FATHER SAID THAT HE WANTED TO CONFIRM IF THERE WERE SOME MALFUNCTIONS SINCE THE PATIENT'S CONDITION WAS TEMPORARILY STABLE, AND IT SEEMED THAT HIS TONE WAS CALM." THE DEALER CONTACTED THE CLINIC FOR INFORMATION, BUT THERE WAS NO INFORMATION ON THE CAUSE OF DEATH. A SALES REPRESENTATIVE RECEIVED ADDITIONAL INFORMATION ON THIS ISSUE. ACCORDING TO THE SALE REPRESENTATIVE, THE DATE OF THE PATIENT PASSING AWAY WAS CONFIRMED. THE SALES REPRESENTATIVE STATED THAT "A VISITING NURSE ARRIVED BEFORE THE DOCTOR, AND THE MOTHER WHO WAS AT THE SCENE REPORTED NO COMPLAINTS OF DEVICE MALFUNCTION (THE PATIENT'S PARENTS DIVORCED THREE DAYS BEFORE HIS DEATH. THE COMPLAINT WAS RECEIVED FROM THE FATHER, WHO WAS NOT PRESENT AT THE TIME AND SUSPECTS THE FAMILY MIGHT HAVE TURNED OFF THE DEVICE.)". THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION FROM A DEALER REGARDING INFORMATION THEY HAD RECEIVED FROM A FATHER OF THE PATIENT WHO HAD PASSED AWAY. THE DEALER CONTACTED THE FATHER OF THE PATIENT TO ASK THE VENTILATOR'S DATA, BUT "THE FATHER SAID THAT HE WANTED TO CONFIRM IF THERE WERE SOME MALFUNCTIONS SINCE THE PATIENT'S CONDITION WAS TEMPORARILY STABLE, AND IT SEEMED THAT HIS TONE WAS CALM." THE DEALER CONTACTED THE CLINIC FOR INFORMATION, BUT THERE WAS NO INFORMATION ON THE CAUSE OF DEATH. A SALES REPRESENTATIVE RECEIVED ADDITIONAL INFORMATION ON THIS ISSUE. ACCORDING TO THE SALE REPRESENTATIVE, THE DATE OF THE PATIENT PASSING AWAY WAS CONFIRMED. THE SALES REPRESENTATIVE STATED THAT "A VISITING NURSE ARRIVED BEFORE THE DOCTOR, AND THE MOTHER WHO WAS AT THE SCENE REPORTED NO COMPLAINTS OF DEVICE MALFUNCTION (THE PATIENT'S PARENTS DIVORCED THREE DAYS BEFORE HIS DEATH. THE COMPLAINT WAS RECEIVED FROM THE FATHER, WHO WAS NOT PRESENT AT THE TIME AND SUSPECTS THE FAMILY MIGHT HAVE TURNED OFF THE DEVICE.)" THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER¿S COMPLAINT WAS NOT CONFIRMED DURING AN OPERATIONAL CHECK OF THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN EVALUATED AND OBSERVED NO ERRORS IN THE SYSTEM LOG. THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED A FUNCTIONAL TEST ON THIS DEVICE, BUT WAS NOT ABLE CONFIRM THE ISSUE ON THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN DETERMINED THE DEVICE WAS OPERATING AS NORMAL. THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED THE FINAL AND RUN-IN TESTS, ALONG WITH THE BATTERY AND VALVE CHECKS FOR THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN DETERMINED THE DEVICE WAS OPERATIONAL AND IT WAS CLEANED.
DREAMSTATION | K102465 | IT WAS REPORTED THAT THE DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. AT THIS TIME, THE DEVICE WAS RETURNED AND A FOLLOW-UP WILL BE FILED WHEN THE DEVICE EVALUATION HAS BEEN COMPLETED.
DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | K213919 | IT WAS REPORTED BY AN ADMINISTRATOR THAT THE PATIENT PASSED AWAY WHILE USING THE DEXCOM G7 CONTINUOUS GLUCOSE MONITOR SYSTEM. ACCORDING TO THE REPORTER, PRIOR TO THE PATIENT'S DEATH, THE PATIENT HAD EXPRESSED CONCERNS TO HER FRIENDS AND FAMILY THAT THE CGM HAD BEEN PROVIDING INCORRECT READINGS. WHETHER AND HOW THESE INACCURATE READINGS WERE RELATED TO THE PATIENT'S DEATH, HOWEVER, WERE NOT SPECIFIED, NOR WERE ANY OTHER DETAILS PROVIDED. DEXCOM REPRESENTATIVES HAVE ATTEMPTED TO FOLLOW UP WITH BOTH THE REPORTER AND THE REPORTING ORGANIZATION MULTIPLE TIMES FOR FURTHER DETAILS, BUT WERE TOLD THAT IT WAS AN ONGOING INVESTIGATION, AND NO NEW INFORMATION REGARDING THE EVENT, NOR PATIENT INFORMATION WERE RECEIVED. NO DATA WAS RETURNED FOR THE ALLEGED DATE OF INCIDENT ON (B)(6) 2024. WITHOUT PERFORMANCE DATA TO INVESTIGATE, THE COMPLAINT OF DISCREPANCIES REMAINS UNCONFIRMED. THE PROBABLE CAUSE OF THE ALLEGED EVENT COULD NOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. | (B)(4).
SECHRIST INDUSTRIES INC. | K140559 | ON JANUARY 31, 2025, SECHRIST BECAME AWARE OF A DEATH, TO A CHILD, THAT WAS BEING TREATED IN A HYPERBARIC CHAMBER. OFFICIALS IN (B)(6), MICHIGAN CONFIRMED A CHILD DIED AFTER THE INCIDENT. | THIS REPORT IS BASED SOLELY ON THE AVAILABLE INFORMATION AT THIS TIME. THIS INCIDENT IS CURRENTLY UNDER INVESTIGATION BY THE (B)(6), MICHIGAN OFFICIALS AND SECHRIST HAS NOT YET BEEN OFFICIALLY NOTIFIED. AT THIS TIME, THERE IS NOT ENOUGH INFORMATION TO DETERMINE THE ROOT CAUSE OR IF ANY DEVICE MALFUNCTION OCCURED. MANUFACTURER REFERENCE FILE NO: (B)(4).
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT ON (B)(6) 2024, AN UMBILICAL ARTERIAL CATHETER (UAC) AND UMBILICAL VENOUS CATHETER (UVC) WERE PLACED ON A 27-WEEK-OLD INFANT SHORTLY AFTER ARRIVAL TO THE NEONATAL INTENSIVE CARE UNIT (NICU). AN INFUSION PUMP WAS USED WITH TUBING REF# (B)(4) FOR BOTH UAC AND UVC LINES. THE UAC LINE HAD AN ADDED TRANSDUCER ATTACHED TO THE END WITH ONE, 3-WAY STOPCOCK PRIOR TO THE PATIENT LINE CONNECTION. ALL IV LINES WERE REPORTEDLY PRIMED AND FLUSHED SEVERAL TIMES, DURING SET UP AND THEN AGAIN BEFORE CONNECTING TO THE PATIENT¿S CATHETERS. THE CONNECTION TO THE INFANT OCCURRED APPROXIMATELY 1959-2000 HRS. IT WAS REPORTED THAT THE INFANT'S CONDITION "ABRUPTLY DETERIORATED WITHIN MINUTES"; "HEART RATE DROPPED"; "CYANOSIS WAS NOTED". ADVANCED CARDIOVASCULAR LIFE SUPPORT (ACLS) MEASURES WERE IMPLEMENTED INCLUDING INTUBATION AND MEDICATION ADMINISTRATION. ONCE STABILIZED, A REPEAT CHEST X-RAY WAS TAKEN WHICH REPORTEDLY SHOWED "AIR IN THE LEFT VENTRICLE OF THE HEART." THIS WAS NOT SEEN ON A PREVIOUS CHEST X-RAY TAKEN AFTER LINE PLACEMENT. DURING THE ABOVE ACTIONS, PROVIDERS NOTED AIR IN THE UAC LINE WHICH THEY ASPIRATED AND REMOVED APPROXIMATELY 2 TO 3ML. THE PATIENT REPORTEDLY DIED 8 DAYS LATER. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "NO AUDIBLE ALARMS SOUNDED, PER STAFF IN THE ROOM. THE LOCATION OF THE AIR WAS DISTAL TO THE PUMP. PUMP FAILURE WAS NOT A CONCERN. PUMPS WERE EVALUATED BY AGILITI BIOMED AND NO DEFICIENCIES WERE FOUND..." THE CUSTOMER (REGISTERED NURSE QUALITY IMPROVEMENT: RL ADMIN LITE & MEDSUN REPORTER) STATED THAT THE PRIMARY AND SECONDARY CAUSES OF DEATH WERE "COMPLICATIONS OF PREMATURE BIRTH ¿ 27 WEEKS AND 4 DAYS. ANOXIC BRAIN INJURY, MULTIPLE ORGAN DYSFUNCTION." | A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
CARDIOHELP | K133598 | COMPLAINT ID# (B)(4). | IT WAS REPORTED THAT A BACKFLOW PREVENTION MESSAGE OCCURRED ON THE SAME TIME AS PATIENT WENT INTO VENTRICULAR TACHYCARDIA (VT), WHICH IS A FAST, ABNORMAL HEART RHYTHM (ARRHYTHMIA). FURTHER A ZERO FLOW MESSAGE APPEARED. DUE TO THE LOW FLOW THE PATIENT REQUIRED CPR (REANIMATION). THE LOW FLOW STATE COULD NOT BE RESOLVED. THE CARDIOHELP WAS EXCHANGED. USING THE EMERGENCY DRIVE RESTORED THE FLOW AND THE FLOW REMAINED ACCEPTABLE ON THE REPLACEMENT CARDIOHELP. ONE OF THE NURSING STAFF HAS MENTIONED THAT THE CONSOLE HAD ¿FROZEN¿. THE PATIENT PASSED AWAY. NEW INFORMATION WAS RECEIVED ON (B)(6) 2025 FROM CUSTOMER THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION, WHICH WAS FOUND OUT AFTER THE PATIENT WAS PUT ONTO ECMO. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON (B)(6) 2025 AND (B)(6) 2025. NO FAULT WAS FOUND AND NO PART WAS REPLACED. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOGFILES WERE ANALYZED BY GETINGE LIFE-CYCLE-ENGINEERING ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: THE MESSAGE 'BACKFLOW PREVENTION' WAS TRIGGERED DURING THERAPY. THE USER WAS UNABLE TO START THE PUMP. THE REASON FOR THIS WAS THE ERROR ¿33006 PUMP DISPOSABLE STOP¿ REPORTED SEVEN SECONDS EARLIER. THIS CAUSED THE PUMP TO SWITCH OFF DUE TO A MAGNETIC GAP BETWEEN THE MOTOR AND THE PUMP THAT WAS DETERMINED TO BE TOO LARGE. THE ARTERIAL PRESSURE IN THE BODY CAUSED A NEGATIVE FLOW WHEN THE PUMP STOPPED. THIS TRIGGERED THE 'BACKFLOW PREVENTION' WITH THE 'ZERO FLOW' MODE AFTER 6 SECONDS. BECAUSE THE 33006 ERROR WAS NOT RESET, THE 'ZERO FLOW' MODE WAS UNABLE TO COUNTERACT THE NEGATIVE FLOW BY BLOCKING THE PUMP. THE USER SHOULD HAVE MADE SURE THAT THE HLS IS SECURELY LOCKED IN THE DRIVE. ADDITIONALLY IT WAS REQUIRED TO ADJUST THE 'USER TARGET SPEED' DOWN TO 0 RPM TO RESET THE 33006 ERROR MODE (THE MESSAGE 'PUMP DISPOSABLE ERROR STOP¿ DISAPPEARS FROM THE DISPLAY). THEN EITHER THE 'BACKFLOW PREVENTION' WOULD HAVE BEEN ACTIVATED OR THE PUMP COULD HAVE BEEN CONTROLLED BY THE USER. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON (B)(6) 2025 WITH FOLLOWING CONCLUSION: "THE LCE INVESTIGATION REPORT EXPLAINS BACKFLOW PREVENTION APPEARED SECONDARY TO DISPLACEMENT/DISLODGEMENT OF THE HLS MODULE FROM THE PUMP WELL, WHICH ELICITED THE ERROR ¿PUMP DISPOSABLE STOP¿ 6 SECONDS BEFORE THE REPORTED EVENT. FURTHER, THE DESCRIPTION OF A ¿FROZEN¿ SCREEN/SYSTEM MAY BE EXPLAINED BY AN UNRESOLVED THE ALARM/ERROR ¿PUMP DISPOSABLE STOP¿. THE IFU OF THE CARDIOHELP DESCRIBES THE FOLLOWING REGARDING THIS PUMP DISPOSABLE STOP¿ MESSAGE. TO RESOLVE THE ERROR, PROPER SEATING OF THE HLS MODULE IN THE CARDIOHELP PUMP WELL IS REQUIRED WHICH IS ENSURED BY SECURELY LOCKING THE MODULE INTO PUMP DRIVE. ONCE THE HLS MODULE IS PROPERLY SEATED, THE RPMS SHOULD BE RESET TO ZERO (0) TO CLEAR THE ERROR MODE. PRESSING THE ZERO FLOW MODE BUTTON IS REQUIRED TO RESET BACKFLOW PREVENTION AND REESTABLISH FLOW. RPM SHOULD BE INCREASED (AFTER THE ZERO SETTING) TO CREATE REQUISITE TARGET BLOOD FLOW. THE IFU ADDRESSES THIS PROCEDURE IN FOLLOWING MANNER: ¿THE BACKFLOW PREVENTION CAN DETECT AND RESPOND TO A BACKFLOW OF BLOOD. FOR THIS, THE CARDIOHELP-I MONITORS THE BLOOD FLOW, DISPLAYS ANY ALARMS AND ACTIVATES THE ZERO FLOW MODE AUTOMATICALLY TO PREVENT ANY BACKFLOW¿. IT REMAINS UNCLEAR WHAT INITIALLY CAUSED THE ERROR MESSAGE ¿PUMP DISPOSABLE STOP¿. THE INVESTIGATION REPORT EXPLAINS THAT THIS ALARM IS TRIGGERED IF A MAGNETIC GAP BETWEEN THE MOTOR AND THE PUMP IS DETERMINED TO BE TOO LARGE. A POSSIBLE ROOT CAUSE FOR THE REPORTED EVENT THAT MAY BE PROPOSED IS DISLODGMENT OF THE HLS MODULE DURING TRANSPORT AND/OR REPOSITIONING OF THE PATIENT WHICH MAY HAVE EXERTED EXCESS STRESS ON THE TUBES/LINES, LEADING TO DISLODGEMENT OF THE DISPOSABLE. ALTERNATIVELY, THE DISPOSABLE MAY NOT HAVE BEEN FULLY ENGAGED IN THE PUMP RECEIVER WHICH EASILY DISLODGED UPON MINIMAL STRESS APPLIED TO THE TUBES/LINES AND/OR THE HLS MODULE. IN EITHER CASE, THE ERROR MESSAGE ¿PUMP DISPOSABLE STOP¿ WOULD HAVE BEEN ELICITED UPON HLS MODULE DISLODGEMENT. LAST, STRUCTURAL DAMAGE OF THE HLS MODULE IS ANOTHER POSSIBLE ROOT CAUSE (AS A PROPOSED ROOT CAUSE UNDERLYING THE EVENT) WHICH MAY HAVE MANIFESTED AS BENT OR DAMAGED ANCHORING FLANGE(S). DAMAGE TO THE ANCHORING FLANGES MAY HAVE NOT ALLOWED THE HLS SET TO PROPERLY ENGAGE WITH THE PUMP DRIVE OF THE CARDIOHELP HARDWARE THEREBY DISALLOWING THE HLS MODULE TO SECURELY ATTACH TO THE CARDIOHELP PUMP DRIVE. THAT SAID, THIS PROPOSAL COULD BE RULED OUT AS THE HLS SET DID FUNCTION IN A REPLACEMENT CARDIOHELP AND BLOOD FLOW WAS REESTABLISHED. THE CARDIOHELP INVOLVED IN THE COMPLAINT WAS INVESTIGATED BY GETINGE FIELD SERVICE TECHNICIAN (FST). THE INVESTIGATION WAS UNABLE TO IDENTIFY ANY FAULTS OR ISSUES RELATING TO PERFORMANCE WITH THE UNIT. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS PASSED AS DESIGNED. THE CUSTOMER STATED THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM THE PROCEDURE. IT IS ASSUMED THAT THE STATEMENT OF NONRECOVERY WAS PREDICATED ON THE DISPOSITION IO THE PATIENT BEFORE SUPPORT USING CARDIOHELP. UNFORTUNATELY, NO DETAILS OF THE PATIENT CONSTITUTION WERE DISCLOSED BY THE CUSTOMER OR WITHIN THE COMPLAINT NARRATIVE. THAT SAID, THE ABSENCE OR REDUCTION OF BLOOD DURING THE REPORTED EVENT MAY HAVE CONTRIBUTED TO THE REPORTED DETERIORATION OF THE PATIENT. HOWEVER, A DIMINUTION OF PRODUCT PERFORMANCE OR MALPERFORMANCE OF THE PRODUCT COULD NOT BE REPRODUCED DURING INSPECTION AND EVALUATION OF THE CARDIOHELP HARDWARE. FURTHER, EXAMINATION OF THE ASSOCIATED DISPOSABLE WAS UNABLE TO BE PERFORMED DUE TO COMPONENT DISPOSAL; THEREFORE, COMMENTARY PERTAINING TO EITHER PERFORMANCE OR MALPERFORMANCE CANNOT BE ADVANCED IN THIS REVIEW." ACCORDING TO THE INSTRUCTION FOR USE (IFU) CHAPTER "CONNECTING THE COMBINED FLOW/BUBBLE SENSOR" THE BUBBLE MONITORING FUNCTION TEST AND FLOW OFF-SET CALIBRATION HAS TO BE PERFORMED BEFORE EVERY USE. THUS A DEFECTIVE FLOW/BUBBLE SENSOR SHOULD BE DETECTED PRIOR TO USE, DURING PRIMING. IN ADDITION AS THE CARDIOHELP INCLUDES PRESSURE SENSORS AND A VENOUS PROBE IT IS ABLE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. IN THE IFU CHAPTER "USING THE EMERGENCY DRIVE WITH THE DISPOSABLE HLS RETAINER" IS STATED THAT THE EMERGENCY DRIVE CAN BE USED TO MANUALLY CONTROL THE BLOOD FLOW IN CASE OF A FAILED CARDIOHELP. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED ON(B)(6) 2025 FOR THE PERIOD OF (B)(6) 2019 TO (B)(6) 2025. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED PRODUCT AND FAILURE. THERE IS NO INDICATION THAT MANUFACTURING ISSUES OCCURRED DURING THIS TIME, THUS PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE DEVICE WAS MANUFACTURED ON 2019-10-09. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. BASED ON THE RESULTS THE REPORTED FAILURE "BACKFLOW PREVENTION" COULD BE CONFIRMED, BUT WAS NOT RELATED TO A DEVICE MALFUNCTION. FURTHER IT WAS STATED BY THE CUSTOMER THAT THE PATIENT WAS NOT EXPECTED TO RECOVER FROM HIS CONDITION. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY'S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE AND DECEASED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
CARDIOHELP | K133598 | THE EVENT OCCURRED IN AUSTRALIA DURING TREATMENT. IT WAS REPORTED THAT A BACKFLOW PREVENTION MESSAGE OCCURRED ON THE SAME TIME AS THE PATIENT WENT INTO VENTRICULAR TACHYCARDIA. FURTHER A ZERO FLOW MESSAGE APPEARED. DUE TO THE LOW FLOW THE PATIENT REQUIRED CPR. THE LOW FLOW STATE COULD NOT BE RESOLVED. THE CARDIOHELP WAS REPLACED. USING THE EMERGENCY DRIVE RESTORED THE FLOW AND THE FLOW REMAINED ACCEPTABLE ON THE REPLACEMENT CARDIOHELP. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. | A GETINGE SERVICE TECHNICIAN WILL INVESTIGATE THE AFFECTED CARDIOHELP. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE AUSTRALIAN MARKET ON A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP", WHICH IS SOLD IN THE US. FOR D4 UNIQUE IDENTIFIER (UDI)#, PRIMARY DI NUMBER HAS BEEN USED FOR BASE UNIT, CARDIOHELP WITH CATALOG NUMBER 701048012, WHICH IS SOLD IN THE US. PROVIDED D4 SERIAL NUMBER AND H4 DEVICE MANUFACTURER DATE CORRESPOND TO THE DEVICE INVOLVED IN THE EVENT, NOT AVAILABLE IN US.
EFFICIA DFM100 | K110825 | PHILIPS RECEIVED A COMPLAINT ON THE EFFICIA DFM100 DEFIBRILLATOR MONITOR INDICATING THE DEVICE SUGGESTED SHOCK ON AED WHEN NOT NEEDED. AFTER THE CUSTOMER CHECKED THE ECG STRIP, THE ECG SHOWED VENTRICULAR FIBRILLATION WITH A PACED RHYTHM, THE DEVICE WORKED PROPERLY, AND SHOCK WAS ADVISED TWO TIMES BUT THE STAFF WAITED UNTIL TIMEOUT (AUTO DISCHARGE) OVERALL. SIX OF EIGHT SHOCKS WERE PROVIDED DURING CPR. THE PATIENT DIED. THE EFFICIA DFM100 DEFIBRILLATOR, MODEL# 866199, IS SUBSTANTIALLY SIMILAR TO THE HEARTSTART XL+ DEFIBRILLATOR (MODEL# 861290) AND WILL BE REPORTED IN THE UNITED STATES UNDER DEVICE MODEL# 861290. | A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
MICROTARGETING¿ ELECTRODE | K033173 | A CASE WAS COMPLETED ON (B)(6) 2025 AT (B)(6) MEDICAL CENTER, (B)(6) WITH DR. (B)(6), WHERE A 57 Y/O PATIENT HAD BOSTON SCIENTIFIC LEADS IMPLANTED BILATERALLY IN STN. MER WAS PERFORMED ON BOTH HEMISPHERES SIMULTANEOUSLY USING FHC 44975Z MICROELECTRODES (A SINGLE TRACK PER SIDE) AND A GUIDELINE 4000 LP+ SYSTEM. OTHER FHC PRODUCTS USED DURING THE CASE WERE STARFIX PLATFORM, 66-IT-03P INSERTION TUBES, 66-ST-12P STYLET. PATIENT RESPONDED WELL TO STIMULATION AND O-ARM IMAGING SHOWED ACCURATE TARGETING. POST-OP CT SHOWED NO ABNORMALITIES AND PATIENT WAS DISCHARGED THE SAME DAY OF SURGERY. THE PATIENT WAS DOING WELL AND HAD GONE BACK TO WORK. FACILITY DID NOT RETAIN LOT# OF PRODUCT USED AND DISPOSED OF PRODUCTS. ON 01/16/2025 FHC, INC. REPRESENTATIVE WAS ATTENDING ANOTHER CASE AT (B)(6) MEDICAL CENTER WITH DR. (B)(6) WHERE DURING THE CASE, THE DR. INFORMED THE FHC REP, THAT THE PATIENT THEY OPERATED ON (B)(6) 2025 WAS FOUND DEAD THIS WEEK AND THAT THEY WERE GOING TO CONDUCT AN AUTOPSY TO DETERMINE THE CAUSE OF DEATH. THE AUTOPSY WAS CONDUCTED ON (B)(6) 2025 WHERE AN INTRACEREBRAL HEMORRHAGE (ICH) AT DBS-LEAD ENTRY SITE AND THALAMUS WAS DETERMINED TO BE CAUSE OF DEATH. SINCE POST-OP CT WAS CLEAR, IT WAS CONSIDERED A DELAYED ICH. | FHC IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY FHC, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. FHC HAS MADE REASONABLE EFFORTS TO OBTAIN MORE COMPLETE INFORMATION IN THE TIME ALLOTTED AND HAS PROVIDED AS MUCH INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, FHC, OR ITS EMPLOYEES THAT THE DEVICE, FHC OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE AN ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY "DEFECTS" OR HAS "MALFUNCTIONED". THESE WORDS ARE INCLUDED IN THE FDA 3500A FORM AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. FHC OBJECTS TO THE USE OF THESE WORDS AND OTHERS LIKE IT BECAUSE OF THE LACK OF DEFINITION AND THE CONNOTATIONS IMPLIED BY THESE TERMS. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
OPTIMA IGS 320 | K122457 | THE CUSTOMER REPORTED THAT DURING AN EMERGENT VASCULAR PROCEDURE, THE GE HEALTHCARE SYSTEM DISPLAYED A COOLING FAILURE OF THE CHILLER AND X-RAY BECAME UNAVAILABLE A FEW MINUTES LATER. THE PROCEDURE COULD NOT BE COMPLETED AND THE PATIENT EXPIRED LATER THE SAME DAY. GE HEALTHCARE INVESTIGATION HAS BEEN OPENED AND IS IN PROGRESS. | LEGAL MANUFACTURER: HCS BEIJING - WEST AREA OF BUILDING NO.3, NO.1 YONGCHANG NORTH ROAD, BEIJING ECONOMIC AND TECHNOLOGICAL DEVELOPMENT AREA CHINA BEIJING BEIJING, 100176. D4: UDI NOT REQUIRED AS DEVICE IS MANUFACTURED PRIOR TO UDI COMPLIANCE DATE. DEVICE INSTALLATION DATE IS 29-FEB-2016. GE HEALTHCARE'S INVESTIGATION INTO THE REPORTED OCCURRENCE IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION HAS BEEN COMPLETED.
VERSACROSS ACCESS SOLUTION | K183655 | IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A RE-DO ATRIAL FIBRILLATION ABLATION PROCEDURE, A VERSACROSS ACCESS KIT WAS SELECTED FOR USE. THE PATIENT WAS INTUBATED AND IN ATRIAL FLUTTER. THE PHYSICIAN MAPPED THE RIGHT ATRIUM (RA) WITH A NON-BOSTON SCIENTIFIC MAPPING CATHETER. THEN, THE PHYSICIAN CROSSED TO LEFT SIDE USING VERSACROSS SHEATH AND VERSACROSS RF WIRE AND MAPPED LEFT ATRIUM (LA). AFTER MAPPING, THE PHYSICIAN DECIDED TO USE FARAPULSE. THE VERSACROSS RF WIRE WAS PLACED INTO THE LEFT SUPERIOR PULMONARY VEIN (LSPV) AND THEN EXCHANGED FOR THE FARADRIVE SHEATH. ONCE THE SHEATH WAS ACROSS, THE PHYSICIAN PASSED IT BACK AND FORTH A COUPLE OF TIMES BEFORE PULLING IT INTO THE RA AND ADVANCING THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER INTO THE LA. SOMETIME DURING THIS PERIOD, THE PATIENT BLOOD PRESSURE DROPPED, AND ICE IMAGINING BECAME VERY FOGGY. AN EFFUSION CHECK WAS PERFORMED AND CONFIRMED THE PATIENT HAD A LARGE EFFUSION ALREADY. THE PHYSICIAN IMMEDIATELY WORKED TO PERFORM A PERICARDIOCENTESIS. A CARDIOTHORACIC (CT) SURGERY CAME IN AND RESUSCITATION WAS STARTED. WHEN CT SURGEON OPENED THE CHEST, IT WAS FOUND THE VERSACROSS THROUGH THE APPENDAGE. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH IS BLOOD LOSS FROM PERFORATED LEFT ATRIAL APPENDAGE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). IN THE PHYSICIAN OPINION, FLOSSING THE SHEATH TO DILATE THE SEPTUM MAY HAVE PULLED THE WIRE AND ADVANCED THE SHEATH AND WIRE THROUGH THE APPENDAGE. THERE WERE NO VERSACROSS MALFUNCTION OR CONTRIBUTION TO PATIENT COMPLICATION REPORTED. THE VERSACROSS DILATOR WAS ALSO PRESENT INSIDE THE SHEATH. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS RF WIRE. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
VERSACROSS ACCESS SOLUTION | K150709 | IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A RE-DO ATRIAL FIBRILLATION ABLATION PROCEDURE, A VERSACROSS ACCESS KIT WAS SELECTED FOR USE. THE PATIENT WAS INTUBATED AND IN ATRIAL FLUTTER. THE PHYSICIAN MAPPED THE RIGHT ATRIUM (RA) WITH A NON-BOSTON SCIENTIFIC MAPPING CATHETER. THEN, THE PHYSICIAN CROSSED TO LEFT SIDE USING VERSACROSS SHEATH AND VERSACROSS RF WIRE AND MAPPED LEFT ATRIUM (LA). AFTER MAPPING, THE PHYSICIAN DECIDED TO USE FARAPULSE. THE VERSACROSS RF WIRE WAS PLACED INTO THE LEFT SUPERIOR PULMONARY VEIN (LSPV) AND THEN EXCHANGED FOR THE FARADRIVE SHEATH. ONCE THE SHEATH WAS ACROSS, THE PHYSICIAN PASSED IT BACK AND FORTH A COUPLE OF TIMES BEFORE PULLING IT INTO THE RA AND ADVANCING THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) CATHETER INTO THE LA. SOMETIME DURING THIS PERIOD, THE PATIENT BLOOD PRESSURE DROPPED, AND ICE IMAGINING BECAME VERY FOGGY. AN EFFUSION CHECK WAS PERFORMED AND CONFIRMED THE PATIENT HAD A LARGE EFFUSION ALREADY. THE PHYSICIAN IMMEDIATELY WORKED TO PERFORM A PERICARDIOCENTESIS. A CARDIOTHORACIC (CT) SURGERY CAME IN AND RESUSCITATION WAS STARTED. WHEN CT SURGEON OPENED THE CHEST, IT WAS FOUND THE VERSACROSS THROUGH THE APPENDAGE. THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH IS BLOOD LOSS FROM PERFORATED LEFT ATRIAL APPENDAGE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). IN THE PHYSICIAN OPINION, FLOSSING THE SHEATH TO DILATE THE SEPTUM MAY HAVE PULLED THE WIRE AND ADVANCED THE SHEATH AND WIRE THROUGH THE APPENDAGE. THERE WERE NO VERSACROSS MALFUNCTION OR CONTRIBUTION TO PATIENT COMPLICATION REPORTED. THE VERSACROSS DILATOR WAS ALSO PRESENT INSIDE THE SHEATH. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. A SEPARATE PROBLEM REPORT WILL BE FILED FOR THE VERSACROSS DILATOR. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
STELLANT DUAL CT INJECTOR | K082905 | THE CUSTOMER REPORTED THAT AN ALLEGED AIR INJECTION AND SUBSEQUENT PATIENT DEATH OCCURRED WHILE A MEDRAD® STELLANT CT INJECTION SYSTEM (SN (B)(6)) WAS IN USE. A REQUEST FOR ADDITIONAL INFORMATION, INCLUDING AN AUTOPSY REPORT AND/OR CAUSE OF DEATH REPORT, HAS BEEN MADE. HOWEVER, THE HOSPITAL INDICATED THAT A POLICE INVESTIGATION WAS UNDERWAY AND THEREFORE, LIMITED INFORMATION IS AVAILABLE AT THIS TIME. | A BAYER SERVICE INJECTOR SYSTEM CHECKOUT, AS WELL AS A REQUEST FOR THE RETURN OF THE DISPOSABLES AND/OR LOT NUMBERS USED IN THE REPORTED INCIDENT, HAS BEEN MADE AND IS AWAITING THE CUSTOMER RESPONSE. IN THE EVENT THAT ADDITIONAL INFORMATION IS OBTAINED, A FOLLOW-UP REPORT WILL BE SUBMITTED. THIS INFORMATION DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE, THE COMPANY, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO A REPORTABLE EVENT.
DREAMSTATION | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
NXSTAGE SYSTEM ONE | K170469 | A REPORT WAS RECEIVED ON 14 JAN 2025 FROM THE CAREGIVER OF A 68 YEAR OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING HISTORY OF HEART ATTACK, MULTI VESSEL CORONARY ARTERY DISEASE, HYPERTENSION AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT LOST BLOOD (AMOUNT AND SOURCE OF BLOOD LOSS NOT PROVIDED) AND EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON 20 JAN 2025 FROM THE HOME THERAPY NURSE (HTN) WHO STATED EMERGENCY MEDICAL SERVICES (EMS) WERE CONTACTED, THEY ARRIVED ON SCENE AND MEDICAL INTERVENTION WAS ATTEMPTED (NOS). NO FURTHER DETAILS OF THE EVENT WERE PROVIDED. | THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
CARDIOSAVE HYBRID W/ E/F PLUG | K112372 | IT WAS REPORTED THAT ON (B)(6) 2025 CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) WAS USED ON A PATIENT IMMEDIATELY AFTER THE BATTERY MAINTENANCE PROGRAM WAS COMPLETED. DEVICE HAD AN 'OVERTEMPERATURE IN THE SYSTEM' ALARM AND ERROR CODE #3, HOWEVER; ON (B)(6) 2025 DEVICE SHOWED AN ALARM MESSAGE OF "OVERTEMPERATURE IN THE SYSTEM". THE USER ENDED THE IABP THERAPY BECAUSE THE PATIENT'S OVERALL THERAPY HAD BEEN STOPPED. THE PATIENT HAS SINCE DIED, BUT ACCORDING TO THE USER, THERE IS NO DIRECT CONNECTION TO THE IABP ALARM SITUATIONS. | DUE TO CHARACTER LIMITATION, BELOW FIELD ARE ADDED FROM E. EVENT SITE TELEPHONE- (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K063377 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED INFECTION AND LATER PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6)1900 FOR THE MDR SUBMISSION REQUIREMENT. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
INTELLIVUE MP5 | K150310 | IT WAS REPORTED ST ANALYSIS NEEDED TO BE ADJUSTED FOR EACH NEW EMERGENCY ROOM PATIENT; HOWEVER, CONSULTATION WITH BIOMED INDICATED ST ANALYSIS DEFAULT WAS SET TO OFF. A PATIENT DEATH OCCURRED AFTER ST ELEVATION WAS NOTICED BY STAFF, WHICH MAY HAVE BEEN RELATED TO NOT BEING ABLE TO HEAR ALARMS. IT REMAINS UNKNOWN WHETHER ST ANALYSIS WAS TURNED ON AT THE TIME OF THE EVENT. THE DEVICE WAS IN USE ON A PATIENT. THERE WAS A REPORT OF PATIENT OR USER HARM. | A FOLLOW-UP REPORT WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION.
HEMOSPRAY ENDOSCOPIC HEMOSTAT | K200972 | COOK ENDOSCOPY WAS NOTIFIED OF THIS EVENT VIA A CLINICAL LITERATURE ARTICLE. THIS ARTICLE WAS PUBLISHED IN 2024. PLEASE SEE BELOW FOR RELEVANT EXCERPTS OF THIS ARTICLE. PERFORATION DUE TO CO2 INSUFFLATION GASTROINTESTINAL PERFORATION HAS BEEN REPORTED AS A RARE ADVERSE EVENT OF TC-325 [HEMOSPRAY] USE. A 90-YEAR-OLD WOMAN WITH A HISTORY OF STROKE (ON ASPIRIN) WAS SENT TO THE EMERGENCY DEPARTMENT FROM SUB-ACUTE REHABILITATION FOR HYPOTENSION, LETHARGY, AND DARK STOOL. UPON ARRIVAL, HER VITAL SIGNS WERE TEMPERATURE 36.4C, BLOOD PRESSURE 76/45 MMHG, HEART RATE 143 BPM, OXYGEN SATURATION 94% ON A NON-REBREATHER. HER LABS SHOWED: HEMOGLOBIN 4.4 G/DL, PLATELETS 210 X 103 PLATELETS/UL, BLOOD UREA NITROGEN 51 MG/DL, CREATININE 0.64 MG/DL, PROTHROMBIN TIME 14.5 SECONDS, ACTIVATED PARTIAL THROMBOPLASTIN TIME 27.3 SECONDS. SHE RECEIVED INTRAVENOUS FLUID RESUSCITATION AND 3 UNITS OF PACKED RED BLOOD CELLS WITH STABILIZATION OF HER VITAL SIGNS. THE FOLLOWING DAY, ESOPHAGOGASTRODUODENOSCOPY SHOWED ACTIVE BLEEDING IN THE PRESENCE OF SEVERE EROSIVE ESOPHAGITIS IN THE HIATAL HERNIA SAC AND 4 NONOBSTRUCTING OOZING LINEAR GASTRIC ULCERS IN THE GASTRIC FUNDUS AND LESSER CURVATURE OF THE GASTRIC BODY. A FORREST IB LINEAR FUNDIC ULCER, 10MM IN LARGEST DIMENSION, WAS NOTED AND SUCCESSFULLY APPROXIMATED WITH 2 HEMOSTATIC CLIPS. GIVEN THE PERSISTENCE OF ACTIVE BLEEDING FROM AN UNCLEAR SOURCE THE PATIENT WAS INTUBATED FOR AIRWAY PROTECTION AND TC-325 [HEMOSPRAY} WAS APPLIED BROADLY IN THE PROXIMAL STOMACH AND HERNIA SAC. DURING THIS PHASE OF THE PROCEDURE THE PATIENT ACUTELY DEVELOPED A DISTENDED, TYMPANITIC ABDOMEN. HER VENTILATION WORSENED REQUIRING HIGHER POSITIVE END-EXPIRATORY PRESSURE, AND SHE EXPERIENCED INTRAPROCEDURAL CARDIAC ARREST. ADVANCED CARDIAC LIFE SUPPORT MEASURES WERE INITIATED; NEEDLE DECOMPRESSION WAS PERFORMED FOR TENSION CAPNOPERITONEUM WHICH IMMEDIATELY IMPROVED HER VENTILATION. RETURN OF SPONTANEOUS CIRCULATION WAS ACHIEVED AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. THE PATIENT LATER EXPIRED AND AUTOPSY WAS DECLINED. GASTRIC PERFORATION AND TENSION CAPNOPERITONEUM FOLLOWING HEMOSTATIC POWDER SPRAY APPLICATION BRIAN MCSTEEN, MD THE AMERICAN JOURNAL OF GASTROENTEROLOGY VOLUME 119 - SUPPLEMENT - OCTOBER 2024 WWW.AMJASTRO.COM | INVESTIGATION EVALUATION: A PRODUCT EVALUATION WAS NOT PERFORMED IN RESPONSE TO THIS REPORT BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT PROVIDED TO COOK FOR EVALUATION. THE REPORT COULD NOT BE CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE CONDUCTED BECAUSE THE LOT NUMBER WAS NOT PROVIDED. INVESTIGATION CONCLUSION: WE COULD NOT CONDUCT A COMPLETE INVESTIGATION BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT RETURNED FOR EVALUATION. THERE IS NO EVIDENCE TO SUGGEST THE REPORT IS DUE TO A DEVICE FAILURE. THE IFU PRECAUTIONS "ENSURE GASTROINTESTINAL LUMEN IS NOT DISTENDED BECAUSE HEMOSPRAY ADDS VOLUME IN EXCESS OF INSUFFLATION VOLUMES DURING PROCEDURE. CLOSELY MONITOR BOWEL DISTENSION AND BALANCE INSUFFLATION AND HEMOSPRAY VOLUMES AS NECESSARY." ADDITIONALLY, THE IFU PRECAUTIONS "LIKE OTHER MODALITIES, HEMOSPRAY MAY NOT BE EFFECTIVE FOR ALL TYPES OF BLEEDS. GASTROINTESTINAL BLEEDING MAY EXACERBATE EXISTING COMORBIDITIES, INCREASING THE POTENTIAL FOR ADVERSE EVENTS INCLUDING PATIENT MORTALITY." PRIOR TO DISTRIBUTION, ALL HEMOSPRAY ENDOSCOPIC HEMOSTAT ARE SUBJECTED TO A VISUAL INSPECTION TO ENSURE DEVICE INTEGRITY. CORRECTIVE ACTION: BASED ON THE QUALITY ENGINEERING RISK ASSESSMENT NO CORRECTIVE ACTION IS WARRANTED AT THIS TIME. A REVIEW OF THE COMPLAINT HISTORY WAS CONDUCTED. THIS REPORT REPRESENTS AN UNUSUAL OCCURRENCE. BASED ON THIS REVIEW, THE LIKELIHOOD OF THIS TYPE OF REPORT IS RARE. QUALITY ASSURANCE WILL CONTINUE TO MONITOR FOR COMPLAINT TRENDS.
PATIENT INFORMATION CENTER IX | K183387 | PHILIPS RECEIVED A COMPLAINT ON A PATIENT INFORMATION CENTER IX. IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN OBTAINING THE CLINICAL AUDIT LOGS FOLLOWING A PATIENT INCIDENT. IT WAS INDICATED THE PATIENT DEVELOPED AN SPO2 ISSUE AND PASSED AWAY; HOWEVER, IT WAS ALSO INDICATED THERE WAS NO IMPACT TO THE PATIENT. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. | A PHILIPS TECHNICAL CONSULTANT (TC) SPOKE TO THE CUSTOMER BIOMEDICAL ENGINEER (BIOMED) MANAGER. THE BIOMED CONTACTED THE NURSE/CLINICAL QUALITY IMPROVEMENT SPECIALIST WHO STATED THEY WERE REQUESTING THE LOGS FILES TO CONFIRM THAT THERE WERE NO DELAYS IN NOTIFICATION OR TREATMENT OF THE PATIENT AND TO HOPEFULLY VERIFY THAT THE PATIENT WAS ON CONTINUOUS PULSE OXIMETRY IF POSSIBLE. THE BIOMED LATER CONFIRMED WITH THE TC THAT THERE WAS NO DELAY; THEY JUST WANTED TO SEE WHAT O2 SENSOR WAS READING AT THE MOMENT. BASED ON THE INFORMATION AVAILABLE, THERE WAS NO ALLEGED PRODUCT FAILURE; THE CUSTOMER JUST NEEDED THE LOG FILES FOR REVIEW. THIS COMPLAINT IS DEEMED NOT REPORTABLE WITH PHILIPS DUE THERE IS NO ALLEGATION OF PRODUCT DEFICIENCY OR OTHER INDICATION THAT A PRODUCT PERFORMANCE ISSUE OR MALFUNCTION HAS OCCURRED. THE CUSTOMER HAS REQUESTED ASSISTANCE FROM PHILIPS ONLY.
PATIENT INFORMATION CENTER IX | K153702 | IT WAS REPORTED THE PATIENT PASSED AWAY WHILE CONNECTED TO THE MONITOR AND THE CUSTOMER REQUESTED ASSISTANCE WITH THE ALARM LOGS AND CONFIGURATION SETTINGS. THE HOSPITAL DID NOT SHARE THE CAUSE OF DEATH AS THEY FELT IT WAS PROTECTED HEALTH INFORMATION (PHI). BIOMED INDICATED THE BEDSIDE MONITOR ALARMED AS EXPECTED BUT THE CLINICAL MANAGER WANTED TO REVIEW THE AUDIT LOG, SPECIFICALLY RELATED TO SPO2 SENSOR OFF INOP TECHNICAL ALARMS, SPO2 SENSOR OFF CONFIGURATION, AND ECG AND HEART RATE EXTREME ALARMS. GOOD FAITH EFFORTS (GFE) CONFIRMED THE DEVICE DID NOT CONTRIBUTE TO THE DEATH SINCE THE ISSUE IN QUESTION WAS SPO2 SENSOR OFF. THE SYSTEM GENERATED EKG ALARMS AND OTHER ALARMED APPROPRIATELY. | A PHILIPS REMOTE APPLICATION SPECIALIST (RAS) INTERVIEWED THE CUSTOMER, AND THEY REPORTED ON (B)(6) 2024 (TIME 2200 TO (B)(6) 2024 1:30AM) THEIR PATIENT (ICU BED239) PASSED AWAY. THE CUSTOMER STATED THAT THE PATIENT'S BEDSIDE MONITOR DID ALARM AS EXPECTED; HOWEVER, THE CLINICAL MANAGER (CM) HAD INQUIRIES REGARDING THE ALARM LOG AND THE APPLICATION CONFIGURATION SETTINGS. THE CM HAD THE FOLLOWING QUESTIONS: 1) THE CLINICAL AUDIT LOG DOES NOT SHOW THE "SPO2 SENSOR OFF" INOP ALARM EVENT WHY? 2) IS IT POSSIBLE TO CHANGE THE "SPO2 SENSOR OFF" ALARM SEVERITY CONFIG SETTING? 3) HOW IT THE ECG/HR EXTREME ALARM BEHAVIORS CONFIGURED? THE PHILIPS RAS REVIEWED FOLLOWING QUESTIONS AND NOTED A CENTRAL STATION WITH SOFTWARE B.02.X CANNOT LOG A "SENSOR OFF INOP" ALARM AS AN EVENT IN THE CLINICAL AUDIT LOG. THE CLINICAL AUDIT CAN LOG THE FOLLOWING INOPS EVENTS: ALL SERVER YELLOW OR RED INOP MESSAGES. ALL BATTERY INOP MESSAGES. ECG LEADS OFF. TRANSMITTER OFF. NO SPO2T. BATT LOW. THE CENTRAL STATION SERVER WITH SOFTWARE C AND GREATER WILL TRACK ALL THE TECHNICAL INOP ALARMS EVENT ON THE AUDIT LOG. THE RAS EXPLAINED THE SPO2 INOP TECHNICAL ALARMS ARE NOT CONFIGURABLE. THESE ALARMS CANNOT BE CONFIGURED TO YELLOW OR RED ALARM SEVERITY. THE CENTRAL STATION LOCAL SERVER INOPERATIVE SOUND VOLUME CAN BE SET TO BE LOUDER THAN THE CURRENT VOLUME. THE RAS EMAILED REFERENCES TO THE CUSTOMER, AND THEY CONFIRMED THE MONITOR GENERATED THE PATIENT'S HEART RATE LOW ALARMS AND THE ALARMS WERE SILENCED. CENTRAL STATION LOGS WERE REQUESTED BUT COULD NOT BE PROVIDED. THE CUSTOMER NOTED THE SYSTEM WORKED AS DESIGNED SINCE SPO2 SENOR BEING OFF WAS A BLUE ALARM AND DOES NOT RE-ALARM LIKE YELLOW AND RED ALARMS. THE QUESTIONS FROM THE CLINICAL MANGER WERE ANSWERED TO THEIR SATISFACTION AS THEY NOW UNDERSTAND THAT SPO2 SENSOR ALARM IS A BLUE ALARM. NO FURTHER INVESTIGATION OR ACTION IS WARRANTED AT THIS TIME.
EMERGE? | K113220 | IT WAS REPORTED THAT SHAFT BREAK AND THROMBOSIS OCCURRED, RESULTING IN PATIENT DEATH. THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT. CORONARY ANGIOGRAPHY REVEALED THAT THE LEFT ANTERIOR DESCENDING (LAD) ARTERY WAS COMPLETELY OCCLUDED FROM THE PROXIMAL SEGMENT, WITH CALCIFICATION VISIBLE DISTAL TO THE OCCLUSION. A 2.50MM X 15MM EMERGE BALLOON CATHETER WAS ADVANCED TO TREAT THE LESION DILATED TO 12 ATMOSPHERES FOR 30 SECONDS. WHEN ANOTHER PRE-DILATION NON-BOSTON SCIENTIFIC BALLOON WAS INSERTED, THE PHYSICIAN BELIEVED THAT THE EMERGE BALLOON TIP HAD FRACTURED WITHIN THE GUIDING CATHETER AND THE NON-BSC BALLOON PUSHED THE DEVICE FRAGMENT INTO THE DISTAL LAD. ACUTE THROMBOSIS WAS IDENTIFIED AND NON-BOSTON SCIENTIFIC STENTS WERE IMPLANTED IN THE LAD. THE PATIENT EXPERIENCED CHEST PAIN AND DISCOMFORT. MEDICATION WAS ADMINISTERED AND ANGIOGRAPHY SHOWED GOOD BLOOD FLOW IN THE LAD. THE PATIENT THEN EXPERIENCED A DROP IN BLOOD PRESSURE AND OXYGEN SATURATION. TRACHEAL INTUBATION WAS PERFORMED, AND AN INTRA-AORTIC BALLOON PUMP WAS INSERTED. ADDITIONAL NON-BOSTON SCIENTIFIC STENTS WERE IMPLANTED, AND POST DILATION WAS PERFORMED WITH A QUANTUM BALLOON. REPEAT ANGIOGRAPHY REVEALED GOOD BLOOD FLOW IN THE LAD, HOWEVER, THE PATIENT'S BLOOD PRESSURE CONTINUED TO DROP AND HEART RATE DECREASED SIGNIFICANTLY. EXTERNAL CHEST COMPRESSIONS WERE PERFORMED, AND ADDITIONAL MEDICATION WAS ADMINISTERED. THE PROCEDURE WAS COMPLETED SUCCESSFULLY, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU), WHERE ONE HOUR LATER WAS CONFIRMED DECEASED. IT WAS FURTHER CONFIRMED THAT RE-EXAMINATION ANGIOGRAPHY REVEALED ACUTE THROMBOSIS AFTER BALLOON PRE-TREATMENT WITH THE NON-BSC BALLOON AND THE EMERGE BALLOON IN THE PROXIMAL TO MID-SEGMENT OF THE LAD, WITH A 90% STENOSIS IN THE DISTAL SEGMENT. THEREFORE, AN EMERGE BALLOON WAS INSERTED INTO THE DISTAL SEGMENT LAD FOR PRE-DILATATION PRIOR TO THE NON-BSC STENT IMPLANTATIONS. | G4 - PREMARKET: K163174. DEVICE EVALUATED BY MFR: THE DEVICE FG EMERGE MR, US 2.50MM X 15MM WAS RETURNED FOR ANALYSIS. VISUAL AND TACTILE INSPECTION IDENTIFIED MULTIPLE KINKS ALONG THE HYPOTUBE SHAFT. VISUAL, TACTILE AND MICROSCOPIC ANALYSIS IDENTIFIED A BREAK IN THE INNER LUMEN AND BALLOON, 22.8CM FROM THE PORT EXCHANGE WITH THE DISTAL SECTION NOT RETURNED. MICROSCOPIC ANALYSIS OF THE BALLOON PROFILE REVEALED A COMPLETE CIRCUMFERENTIAL TEAR 22.8CM FROM THE PORT EXCHANGE WITH THE DISTAL SECTION NOT RETURNED. THE BUMPER TIP SHOWED NO SIGNS OF DISTAL TIP DAMAGE. THE DISTAL MARKERBAND WAS NOT RETURNED, HOWEVER, NO ISSUES WERE NOTED WITH THE PROXIMAL MARKERBAND. NO OTHER DEVICE ISSUES WERE IDENTIFIED DURING RETURNED PRODUCT ANALYSIS.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS DUE TO DIABETIC KETOACIDOSIS. ACCORDING TO THE PHYSICIAN, A REVIEW OF THE PUMP HISTORY REVEALED THAT THE CUSTOMER CHANGED THE CARTRIDGE TWO DAYS PRIOR TO THE EVENT, RESULTING IN BLOOD GLUCOSE (BG) READINGS DISPLAYED AS "HIGH." REPORTEDLY, THE PUMP FUNCTIONED AS INTENDED AND DELIVERED AUTOMATIC BOLUSES IN RESPONSE TO THE HIGH BG LEVELS. THE CUSTOMER ALSO DELIVERED MANUAL OVERRIDE BOLUSES VIA THE PUMP IN AN ATTEMPT TO ADDRESS THE HIGH BG LEVELS; HOWEVER, THE BG VALUES REMAINED HIGH. AS A RESULT OF THE CURRENT FINDINGS, THE PHYSICIAN STATED THAT "THE ALGORITHM IN THE PUMP HAS NOT BEEN THE PROBLEM, BUT PROBABLY THE DELIVERY" AND THAT THE EVENT WAS DUE TO "A VERY UNFORTUNATE USER ERROR." NO ADDITIONAL INFORMATION COULD BE CONFIRMED AT THE TIME OF THIS REPORT, AS PUMP DATA IS NOT AVAILABLE FOR REVIEW. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
BD ALARIS PUMP MODULE SMARTSITE INFUSION SET | K221327 | NO ADDITIONAL INFORMATION. | THE CUSTOMER REPORTED AIR-IN-LINE, AND RETURNED ONE SAMPLE OF MATERIAL 2426-0007, LOT 24109078. UPON INITIAL VISUAL EXAMINATION, THE SET HAD NO DEFECTS OR ABNORMALITIES. THE SET WAS RUN AT 120ML/HR FOR 30 MINUTES IN ORDER TO REPLICATE AN AGGRESSIVE INFUSION. THERE WAS NO AIR IN LINE, LEAKS, OR OTHER ISSUES OBSERVED DURING THIS TIME. THE COMPLAINT OF AIR IN LINE COULD NOT BE VERIFIED. THE ROOT CAUSE FOR THE ISSUE COULD NOT BE FOUND WITHOUT A REPLICATED FAILURE. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED. THERE WERE NO QUALITY NOTIFICATIONS ISSUED FOR THE FAILURE MODE REPORTED BY THE CUSTOMER DURING THE PRODUCTION BUILD OF THIS SET.
BD ALARIS PUMP MODULE SMARTSITE INFUSION SET | K221327 | NO ADDITIONAL INFORMATION. | THE CUSTOMER REPORTED AIR-IN-LINE, AND RETURNED ONE SAMPLE OF MATERIAL 2426-0007, LOT 24109078. UPON INITIAL VISUAL EXAMINATION, THE SET HAD NO DEFECTS OR ABNORMALITIES. THE SET WAS RUN AT 120ML/HR FOR 30 MINUTES IN ORDER TO REPLICATE AN AGGRESSIVE INFUSION. THERE WAS NO AIR IN LINE, LEAKS, OR OTHER ISSUES OBSERVED DURING THIS TIME. THE COMPLAINT OF AIR IN LINE COULD NOT BE VERIFIED. THE ROOT CAUSE FOR THE ISSUE COULD NOT BE FOUND WITHOUT A REPLICATED FAILURE. A DEVICE HISTORY RECORD REVIEW WAS PERFORMED. THERE WERE NO QUALITY NOTIFICATIONS ISSUED FOR THE FAILURE MODE REPORTED BY THE CUSTOMER DURING THE PRODUCTION BUILD OF THIS SET.
HEART LUNG MACHINE | K991864 | NEW INFORMATION HAS BEEN PROVIDED BY THE SSU DATED ON 2025-03-20 THAT THE PATIENT PASSED AWAY. COMPLAINT ID: (B)(4). | THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
EVITA 4 | K980642 | IT WAS REPORTED THAT THE EVITA 4 WENT OUT DURING PREPARATION FOR TRANSFER TO HELICOPTER TRANSPORT AND WOULD NOT HAVE CONTINUED TO VENTILATE. IT WAS REPORTED THAT DUE TO THE SITUATION (INCLUDING A CHANGE IN VITAL SIGN MONITORING), THE DETERIORATION IN VITAL SIGNS WAS NOT NOTICED IMMEDIATELY AND THE PATIENT PASSED AWAY. THE DOCTORS RESPONSIBLE COULD NEITHER CONFIRM NOR COMPLETELY RULE OUT WHETHER THE FAILURE OF THE VENTILATOR WAS THE CAUSE OF DEATH. ACCORDING TO THE DOCTORS, THE VENTILATOR DISPLAY WAS BLACK. THE DEVICE COULD BE RESTARTED IMMEDIATELY AND FUNCTIONED WITHOUT ANY PROBLEMS. IT WAS NOT POSSIBLE TO SAY FOR CERTAIN WHETHER THE DEVICE HAD ALARMED THE MALFUNCTION. THE DEVICE HAS NO UDI AS AN UDI WAS NOT REQUIRED AT THE TIME THE DEVICE WAS MANUFACTURED. | THE LOG FILE OF THE AFFECTED DEVICE WAS ANALYSED REGARDING THE REPORTED EVENT. BASED ON THE LOG FILE ANALYSIS THE REPORTED EVENT COULD BE TRACED. THE DEVICE WAS OPERATED IN BIPAP ASB VENTILATION MODE DURING THE EVENT. DURING VENTILATION, THE DEVICE ISSUED THE ALARMS ¿MV LOW!!!¿, ¿MV HIGH!!!¿, ¿TIDAL VOLUME HIGH!!!¿, ¿FIO2 HIGH!!!¿ AND ¿ETCO2 HIGH!!!¿ OUTPUT. AT 12:13 P.M. THE DEVICE WAS SWITCHED TO STANDBY MODE, WHICH WAS ALARMED AS SPECIFIED WITH ¿STANDBY ACTIVATED!!!¿ AND THEN CONFIRMED WITH THE "ALARM RESET" BUTTON. AT 12:19 P.M., THE DEVICE WAS SWITCHED BACK INTO OPERATION AND AGAIN EMITTED THE VENTILATION-RELEVANT ALARM MESSAGES. THE DEVICE WAS SWITCHED BACK TO STANDBY MODE AND SWITCHED OFF BY THE USER AT 12:45 P.M.. THE ANALYSIS OF THE LOGFILE COULD NOT FIND ANY INDICATION OF A DEVICE MALFUNCTION. IF THE "STANDBY" HARDKEY IS PRESSED CONTINUOUSLY FOR 3 SECONDS BY THE USER DURING OPERATION, EVITA 4 SWITCHES TO "STANDBY" MODE AND ISSUES A VISUAL AND AUDIBLE ALARM WITH THE WARNING "STANDBY ACTIVATED !!!". THIS ALARM REMAINS ACTIVE UNTIL THE USER CONFIRMS THE ALARM BY PRESSING THE "ALARM RESET" HARDKEY. IN STANDBY MODE, THE PATIENT CAN BREATHE SPONTANEOUSLY VIA THE EMERGENCY BREATHING DEVICE. THE DEVICE HAS BEHAVED AS SPECIFIED AND ALERTED THE USER ACOUSTICALLY AND VISUALLY TO THE CHANGE TO STANDBY MODE. THE NUMBER OF SIMILAR CASES, RELATED TO THE SAME ROOT CAUSE, IS WITHIN THE EXPECTED RANGE OF THE RESPECTIVE RISK ASSESSMENT AND THUS ACCEPTED.
CANNULA & CATHETER | K102532 | COMPLAINT # (B)(4). | IT WAS REPORTED THAT VENOUS CANNULA WAS DISCONNECTED DURING PROCEDURE. THE PATIENT HAS BEEN IN PALLIATIVE CARE UNDER BIVENTRICULAR SUPPORT SINCE 23/11. ON 24/11, THE CANNULA INITIALLY PLACED IN THE RIGHT FEMORAL WAS REPLACED BY A CANNULA PLACED IN THE LEFT FEMORAL DUE TO SUSPECTED NERVE COMPRESSION BY THE FIRST CANNULA. ON 29/11, ACCIDENTAL VENOUS DECANNULATION OCCURRED WITH HEMORRHAGIC SHOCK AND AIR ENTRAINMENT IN THE CIRCUIT. THE VENOUS CANNULA WAS FIXED WITH ONLY ONE SUTURE POINT. THE REMOVABLE FIXATION DEVICE PROVIDED WITH THE CANNULA IS RARELY USED AS IT IS DEEMED INSUFFICIENT. PATIENT HAS EXPIRED DUE TO ADVERSE EVENT. SAMPLE INVESTIGATION COULD NOT BE PERFORMED SINCE THE PRODUCT WAS DISCARDED BY CUSTOMER. MORE INFORMATION HAS BEEN REQUESTED FROM CUSTOMER BY MEDICAL AFFAIRS IN ORDER TO INVESTIGATE THE ISSUE, HOWEVER IT WAS REPORTED BY GETINGE SALES & SERVICE REPRESENTATIVE THAT THE CUSTOMER IS NOT ABLE TO PROVIDE FURTHER INFORMATION. MEDICAL REVIEW HAS BEEN PERFORMED ON 2025-02-26. THE DESCRIBED DECANNULATION OF THE HLS CANNULA (EXPLAINED AS ACCIDENTAL) RESULTED IN HAEMORRHAGIC SHOCK WHICH LIKELY CONTRIBUTED TO THE EXPIRATION OF THE PATIENT. IT WAS REPORTED IN THE COMPLAINT THAT THE VENOUS CANNULA WAS FIXED WITH A SINGLE STITCH AND THAT THE FIXATION DEVICE PROVIDED WITH THE CANNULA WAS NOT USED AS IT IS DEEMED INSUFFICIENT. THE INSTRUCTION FOR USE OF THE HLS CANNULA THAT WAS PLACED ON THE PATIENT EXPLAINS THE FOLLOWING REGARDING WARNINGS AND PRECAUTIONS. IFU, HLS CANNULAE, G-139, V05 5.3 SAFETY INSTRUCTIONS FOR CANNULAE -IF THE PATIENT IS REPOSITIONED OR TRANSPORTED, THERE IS A RISK OF DECANNULATION CAUSED BY STRAIN ON THE TUBING AND MECHANICAL DAMAGE. THE GREATEST CARE SHOULD THEREFORE BE EXERCISED WHEN CARRYING OUT THESE MEASURES. -ENSURE THAT THERE IS NO STRAIN ON THE CANNULA. -AVOID SUBJECTING THE CANNULA TO MECHANICAL LOADS. IFU, HLS CANNULAE, G-139, V05 7 APPLICATION -PERFORM A CAREFUL VISUAL INSPECTION OF THE DEVICE BEFORE USE. IN PARTICULAR, ENSURE THERE IS NO DAMAGE TO THE MATERIAL, CRACKS, BURRS OR FISSURES. CAUTION! ATTACH THE CANNULA SECURELY TO THE VESSEL OR SURROUNDING TISSUE TO PREVENT ACCIDENTAL DISLOCATION DURING THE EXTRACORPOREAL CIRCULATION. CAUTION! IF A SUTURE IS DIRECTLY LOOPED AROUND THE WIRE-REINFORCED SECTION, THE CANNULA CAN BE CUT INTO, KINKED OR DAMAGED. SECURE THE CANNULA AROUND THE RIGID BARBED CONNECTORS IN SUCH A WAY AS TO ENSURE THAT IT IS NOT KINKED AT THE INSERTION SITE. -ATTACH THE CANNULA SECURELY TO THE VESSEL. FOR THIS PURPOSE, USE THE REMOVABLE SKIN ATTACHMENT ACCESSORY SUPPLIED OR ELSE A SLIT AT THE END OF THE CANNULA FOR THE SUTURE MATERIAL. THE REMOVABLE SKIN ATTACHMENT ACCESSORY HAS HOLES TO TAKE THE SUTURE MATERIAL. CAUTION! CHECK THE CONNECTION BETWEEN CANNULA AND TUBING SET REGULARLY. IN ABSENCE OF MORE DETAILS, A PROBABLE ROOT CAUSE OF THE REPORTED DECANNULATION CANNOT BE DETERMINED. FURTHER, WITHOUT FURTHER INFORMATION REGARDING THE EVENT, POSSIBLE ROOT CAUSES CANNOT BE PROPOSED. BASED ON THE INVESTIGATION RESULTS, COMPLAINT COULD BE CONFIRMED HOWEVER PRODUCT RELATED INFLUENCES ARE UNLIKELY. BESIDES, THE LOT NUMBER OF THE AFFECTED HLS CANNULAE WAS NOT PROVIDED BY CUSTOMER. HOWEVER, THE FOLLOWING BATCH NUMBERS WERE FOUND VIA SAP THAT ARE SOLD TO CHU BORDEAUX - HOPITAL PELLEGRIN WITHIN THE SHELF LIFE: 3000434622, 3000432796, 3000427010, 3000423314, 3000414511, 3000404875, 3000401849, 3000367020, 3000353387, 3000332654, 3000304537. THE PRODUCTION HISTORY RECORDS (DHR) OF THE AFFECTED BE-PVL 2155#BE-HLS CANNULA 21F VL WITH THE ABOVE LOT #S WERE REVIEWED ON 2025-01-29. ACCORDING TO THE DHR RESULT, THE PRODUCT BE-PVL 2155#BE-HLS CANNULA 21F VL PASSED THE DEFINED MANUFACTURING AND FINAL RELEASE SPECIFICATIONS. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED FAILURE AND PRODUCT AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. THE EXECUTOR OF THE ESTATE CONTACTED US, MENTIONING THAT SHE HAD RECEIVED A LETTER FROM PHILIPS REGARDING A REPLACEMENT DEVICE. HOWEVER, SINCE THE PATIENT HAD ALREADY PASSED AWAY, SHE WAS UNCERTAIN WHETHER THE DEVICE WAS IN THEIR POSSESSION. SHE STATED THAT SHE WOULD FOLLOW UP IF THE DEVICE IS LOCATED SO THAT IT CAN BE RETURNED. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS INVOLVED IN THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
CANNULA & CATHETER | K102532 | COMPLAINT- # (B)(4). | IT WAS REPORTED THAT VENOUS CANNULA WAS DISCONNECTED DURING PROCEDURE. THE PATIENT HAS BEEN IN PALLIATIVE CARE UNDER BIVENTRICULAR SUPPORT SINCE (B)(6). ON (B)(6), THE CANNULA INITIALLY PLACED IN THE RIGHT FEMORAL WAS REPLACED BY A CANNULA PLACED IN THE LEFT FEMORAL DUE TO SUSPECTED NERVE COMPRESSION BY THE FIRST CANNULA. ON (B)(6), ACCIDENTAL VENOUS DECANNULATION OCCURRED WITH HEMORRHAGIC SHOCK AND AIR ENTRAINMENT IN THE CIRCUIT. THE VENOUS CANNULA WAS FIXED WITH ONLY ONE SUTURE POINT. THE REMOVABLE FIXATION DEVICE PROVIDED WITH THE CANNULA IS RARELY USED AS IT IS DEEMED INSUFFICIENT. PATIENT HAS EXPIRED DUE TO ADVERSE EVENT. SAMPLE INVESTIGATION COULD NOT BE PERFORMED SINCE THE PRODUCT WAS DISCARDED BY CUSTOMER. MORE INFORMATION HAS BEEN REQUESTED FROM CUSTOMER BY MEDICAL AFFAIRS IN ORDER TO INVESTIGATE THE ISSUE, HOWEVER IT WAS REPORTED BY GETINGE SALES & SERVICE REPRESENTATIVE THAT THE CUSTOMER IS NOT ABLE TO PROVIDE FURTHER INFORMATION. MEDICAL REVIEW HAS BEEN PERFORMED ON (B)(6)2025. THE DESCRIBED DECANNULATION OF THE HLS CANNULA (EXPLAINED AS ACCIDENTAL) RESULTED IN HAEMORRHAGIC SHOCK WHICH LIKELY CONTRIBUTED TO THE EXPIRATION OF THE PATIENT. IT WAS REPORTED IN THE COMPLAINT THAT THE VENOUS CANNULA WAS FIXED WITH A SINGLE STITCH AND THAT THE FIXATION DEVICE PROVIDED WITH THE CANNULA WAS NOT USED AS IT IS DEEMED INSUFFICIENT. THE INSTRUCTION FOR USE OF THE HLS CANNULA THAT WAS PLACED ON THE PATIENT EXPLAINS THE FOLLOWING REGARDING WARNINGS AND PRECAUTIONS. IFU, HLS CANNULAE, G-139, V05 5.3 SAFETY INSTRUCTIONS FOR CANNULAE -IF THE PATIENT IS REPOSITIONED OR TRANSPORTED, THERE IS A RISK OF DECANNULATION CAUSED BY STRAIN ON THE TUBING AND MECHANICAL DAMAGE. THE GREATEST CARE SHOULD THEREFORE BE EXERCISED WHEN CARRYING OUT THESE MEASURES. -ENSURE THAT THERE IS NO STRAIN ON THE CANNULA. -AVOID SUBJECTING THE CANNULA TO MECHANICAL LOADS. IFU, HLS CANNULAE, G-139, V05 7 APPLICATION -PERFORM A CAREFUL VISUAL INSPECTION OF THE DEVICE BEFORE USE. IN PARTICULAR, ENSURE THERE IS NO DAMAGE TO THE MATERIAL, CRACKS, BURRS OR FISSURES. CAUTION! ATTACH THE CANNULA SECURELY TO THE VESSEL OR SURROUNDING TISSUE TO PREVENT ACCIDENTAL DISLOCATION DURING THE EXTRACORPOREAL CIRCULATION. CAUTION! IF A SUTURE IS DIRECTLY LOOPED AROUND THE WIRE-REINFORCED SECTION, THE CANNULA CAN BE CUT INTO, KINKED OR DAMAGED. SECURE THE CANNULA AROUND THE RIGID BARBED CONNECTORS IN SUCH A WAY AS TO ENSURE THAT IT IS NOT KINKED AT THE INSERTION SITE. -ATTACH THE CANNULA SECURELY TO THE VESSEL. FOR THIS PURPOSE, USE THE REMOVABLE SKIN ATTACHMENT ACCESSORY SUPPLIED OR ELSE A SLIT AT THE END OF THE CANNULA FOR THE SUTURE MATERIAL. THE REMOVABLE SKIN ATTACHMENT ACCESSORY HAS HOLES TO TAKE THE SUTURE MATERIAL. CAUTION! CHECK THE CONNECTION BETWEEN CANNULA AND TUBING SET REGULARLY. IN ABSENCE OF MORE DETAILS, A PROBABLE ROOT CAUSE OF THE REPORTED DECANNULATION CANNOT BE DETERMINED. FURTHER, WITHOUT FURTHER INFORMATION REGARDING THE EVENT, POSSIBLE ROOT CAUSES CANNOT BE PROPOSED. BASED ON THE INVESTIGATION RESULTS, COMPLAINT COULD BE CONFIRMED HOWEVER PRODUCT RELATED INFLUENCES ARE UNLIKELY. BESIDES, THE LOT NUMBER OF THE AFFECTED HLS CANNULAE WAS NOT PROVIDED BY CUSTOMER. HOWEVER, THE FOLLOWING BATCH NUMBERS WERE FOUND VIA SAP THAT ARE SOLD TO (B)(6) WITHIN THE SHELF LIFE: 3000434622, 3000432796, 3000427010, 3000423314, 3000414511, 3000404875, 3000401849, 3000367020, 3000353387, 3000332654, 3000304537. THE PRODUCTION HISTORY RECORDS (DHR) OF THE AFFECTED BE-PVL 2155#BE-HLS CANNULA 21F VL WITH THE ABOVE LOT #S WERE REVIEWED ON (B)(6)2025. ACCORDING TO THE DHR RESULT, THE PRODUCT BE-PVL 2155#BE-HLS CANNULA 21F VL PASSED THE DEFINED MANUFACTURING AND FINAL RELEASE SPECIFICATIONS. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED FAILURE AND PRODUCT AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S WIFE REPORTED RECEIVING A LETTER REGARDING THE RETURN OF HER HUSBAND'S DEVICE, BUT SINCE HE PASSED AWAY TWO YEARS AGO, SHE IS UNSURE OF THE DEVICE'S WHEREABOUTS OR WHAT HAPPENED TO IT. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S HUSBAND CONTACTED US, EXPRESSING HIS INTENTION TO RETURN THE DEVICE BUT MENTIONED HE WAS UNFAMILIAR WITH HOW TO USE FEDEX. HE WAS REFERRED TO THE DME FOR FURTHER ASSISTANCE. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
CAREEVENT | K161164 | THE CUSTOMER REPORTED THAT THE ALARMS DID NOT RING ON THE SMARTPHONE AND REQUESTED A LOG REVIEW. THE LOGS INDICATED THAT THE ALARMS OCCURRED ON THE MONITOR AND CENTRAL STATION AND WERE TRANSMITTED TO THE SMARTPHONE. ADDITIONAL INFORMATION IS REQUESTED FOR FURTHER CLARIFICATION. THE DEVICE WAS IN CLINICAL USE AT TIME OF EVENT AND A PATIENT DIED. | THE FOLLOWING FUNCTIONAL TESTS WERE PERFORMED: PHILIPS WAS NOTIFIED ON (B)(6) 2025, BUT ALL THE LOGS FROM THE EVENT DATE (31 OCT 2024) HAD BEEN OVERWRITTEN AND THE ONLY LOGS THAT COULD BE RETRIEVED WAS THE ONE COLLECTED BY THE CUSTOMER. THE FIELD SERVICE ENGINEER (FSE) PROVIDED AN ANALYSIS OF THOSE LOGS. FOR THE EVENT TABLE FOR (B)(6) 2024 BETWEEN 08:00 AND 09:00. 223 EVENTS, 0 EVENTS NOT TRANSMITTED, 22 ESCALATIONS. THE PHILIPS FIELD SERVICE ENGINEER, SUPPORT (B)(6), PHILIPS TECHNICAL MANAGER IN CHARGE AND THE CUSTOMER SENIOR BIOMEDICAL TECHNICIAN CONDUCTED PERFORMANCE VERIFICATION TESTING ONSITE. THE CUSTOMER'S ALARM CONFIGURATION OF CAREEVENT WAS TESTED AND VERIFIED. IN ADDITION, THE REVIEW STATED THAT THE "AUTO REMOVE ENDED EVENTS" FEATURE WAS UNCHECKED. AS CONFIGURED, THIS SETTING AUTOMATED THE DELETION OF COMPLETED EVENTS IN THE PHONE HISTORY, RESULTING IN THE DISAPPEARANCE OF SOME EVENTS FROM THE HISTORY. THIS COULD EXPLAIN THE CUSTOMER'S OBSERVATIONS. AFTER RESTARTING THE VIRTUAL SERVER AND CHECKING THE POINTS STATED IN THE SERVICE MANUAL, IT SHOWED THAT THE CAREEVENT SYSTEM WAS OPERATING PER SPECIFICATIONS. ALARMS THAT ARE SET TO BE SENT TO PHONES ARE BEING SENT. ESCALATIONS ARE TRIGGERED AFTER THE CONFIGURED TIME OF ONE MINUTE ON THE DEFINED PHONES. USERS DESIGNATED FOR DELEGATION ARE NOTIFIED AT THE SAME TIME AS FIRST-LEVEL USERS. THE SYSTEM STOPS SENDING NOTIFICATIONS IF USERS ARE LOGGED OUT OR IF THE PHONE IS RUNNING OUT OF BATTERY. LOGOUT/RECONNECT NOTIFICATIONS ARE GENERATED ON THE PHONE AND ARE FOUND IN THE HISTORY. NO MALFUNCTION WAS OBSERVED DURING THE TESTS. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED WE WERE UNABLE TO REPLICATE THE REPORTED PROBLEM. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE CUSTOMER WAS PROVIDED WITH INFORMATION FROM REVIEW. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
DS2ADV AUTO CPAP W/HUMID CELL/BT | K200480 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S HUSBAND CONTACTED US WANTING TO RETURN THE DEVICE AS HIS WIFE PASSED AWAY 2 YEARS AGO. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | K141952 | IT WAS REPORTED THAT DURING A LOBECTOMY PROCEDURE WHERE THE STAPLER WAS STUCK ON THE PULMONARY ARTERY AND COULD NOT BE OPENED. REP WALKED THE SURGEON THROUGH THE STEPS OVER THE PHONE BUT THEY WERE STILL HAVING TROUBLE. SALES REP THEN DROVE TO THE FACILITY WHERE HE WAS TOLD THE PATIENT PASSED PRIOR TO HIS ARRIVAL. SURGEON STATED TO THE REP THAT HE JUST COULDN'T OPEN IT AND THE PATIENT BLED OUT. SURGEON DESCRIBED THE SITUATION TO THE REP THAT HE FIRED THE DEVICE AND IT SOUNDED LIKE IT FIRED HALFWAY. SURGEON TRIED TO OPEN BUT COULDN'T OPEN THE DEVICE. REP ASKED THE SURGEON IF HE USED THE REVERSE BUTTON, SURGEON REPLIED THAT HE DID. SURGEON SAID HE ENGAGED THE MANUAL OVERRIDE HOWEVER HE WAS OPERATING ALONE. REP INFORMED THE SURGEON THAT WHEN ENGAGING THE MANUAL OVERRIDE SOMEONE NEEDS TO HOLD ON TO THE SHAFT TO ENGAGE BECAUSE ITS A PRETTY HARD RATCHET TO RATCHET THE KNIFE BLADE BACK. SURGEON SAID HE TRIED THE MANUAL OVERRIDE BUT WAS UNABLE TO GET IT BACK. THERE WAS A LOT OF BLEEDING. A SECOND STAPLER WAS USED TO STAPLE UNDERNEATH WHEN THEY COULDN'T OPEN THE FIRST STAPLER. IT WAS DESCRIBED BY THE SURGICAL TECH THAT AFTER THE SECOND STAPLING THERE WAS BLOOD EVERYWHERE AND THE PASSED VERY QUICKLY AFTER THAT. THE SURGEON WALKED THE REP TO THE OR WHERE HE WAS ABLE TO TAKE A LOOK AT THE STAPLER, THE MANUAL OVERRIDE HAD BEEN ENGAGED BECAUSE THE CAP WAS OFF AND THE WRENCH WAS UP. | (B)(4). DATE SENT: 7/21/2025. D4: BATCH # 874C51. INVESTIGATION SUMMARY: THE PRODUCT WAS RETURNED TO ETHICON ENDO-SURGERY FOR EVALUATION. VISUAL INSPECTION WAS CONDUCTED ON THE RETURNED DEVICE. VISUAL ANALYSIS OF THE RETURNED SAMPLE DETERMINED THAT ONE PVE35A DEVICE WAS RECEIVED WITH NO APPARENT DAMAGE, WITH THE CLOSING TRIGGER AND ANVIL IN THE CLOSED POSITION, AND WITH A RELOAD LOADED ON THE DEVICE. THE RELOAD WAS RECEIVED PARTIALLY FIRED 1/10TH IN THE LOCK-OUT POSITION WITH THE REMAINING STAPLES INSIDE THE RELOAD POCKETS. IT CANNOT BE CONFIRMED WHY THE RELOAD LOCKED OUT; HOWEVER, IT IS POSSIBLE THAT THE ONE-PIECE SLED COMPONENT WITHIN THE RELOAD MAY HAVE BEEN MOVED OUT OF POSITION DURING THE LOADING PROCESS LEADING TO THE INADVERTENT LOCKOUT. PLEASE REFER TO THE INSTRUCTIONS FOR USE FOR MORE INFORMATION. THE MANUAL OVERRIDE DOOR WAS NOT PRESENT ON THE DEVICE AND WAS NOT RETURNED FOR ANALYSIS; THE OVERRIDE LEVER WAS UP WHICH DENOTES THAT THE KNIFE WAS MANUALLY RETURNED TO HOME POSITION. IT SHOULD BE NOTED AFTER THE MANUAL OVERRIDE SYSTEM IS USED THE INSTRUMENT IS DISABLED AND CANNOT BE USED FOR ANY SUBSEQUENCE FIRINGS. THE DEVICE WAS CT SCANNED, AND THE ARTICULATION LOCK TOOTH WAS NOTED TO BE BROKEN. IT IS POSSIBLE THAT AN OUTSIDE THE NORMAL USE FORCE WAS APPLIED ON THE DISTAL JAW CREATING A TORQUE ON THE ARTICULATION COMPONENTS AND BREAKING THE ARTICULATION BAR. THIS CONDITION PREVENTS THE DEVICE FROM HOLDING THE ARTICULATION SETTING AND DOES NOT AFFECT THE OPENING MECHANISM. AN ATTEMPT WAS MADE TO OPEN THE INSTRUMENT, AND IT OPENED WITHOUT DIFFICULTY WHEN THE ANVIL RELEASE BUTTON WAS ACTUATED WHILE SQUEEZING THE CLOSING TRIGGER. PLEASE NOTE THAT TO OPEN THE JAWS, SQUEEZE THE CLOSING TRIGGER, THEN SIMULTANEOUSLY PRESS THE ANVIL RELEASE BUTTON ON EITHER SIDE OF THE INSTRUMENT. WHILE PRESSURE IS STILL ON THE ANVIL RELEASE BUTTON, SLOWLY RELEASE THE CLOSING TRIGGER. FOR FURTHER INSTRUCTIONS PLEASE REFER TO THE INSTRUCTIONS FOR USE. TO VERIFY THE CONDITION OF THE INTERNAL COMPONENTS, THE DEVICE WAS DISASSEMBLED, AND THE SLIDE LEVER WAS NOTED TO BE PAST THE FULLY FIRED POSITION AND THE TAB THAT RIDES IN THE DRIVE BAR SLOT WAS BROKEN. THE DAMAGE TO THE TAB WAS MOST LIKELY CAUSED BY APPLYING AN EXCESS OF FORCE TO THE REVERSE BUTTON SWITCH WHEN ATTEMPTING TO OPEN THE INSTRUMENT AFTER THE DRIVE BAR HAD ALREADY RETURNED TO THE HOME POSITION. THIS DAMAGE FOUND WOULD NOT HAVE PREVENTED THE INSTRUMENT FROM OPENING. NO FUNCTIONAL TEST WAS PERFORMED DUE TO THE CONDITION OF THE DEVICE. THE EVENT DESCRIBED OF DIFFICULT TO OPEN COULD NOT BE CONFIRMED AS THE DEVICE OPENED AND CLOSED WITHOUT ANY DIFFICULTIES NOTED. THE INSTRUMENT OPENED NORMALLY WHEN THE UNCLAMP LOCKOUT WAS PRESSED WHILE SQUEEZING THE CLAMPING TRIGGER. DEVICE HISTORY REVIEW: A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE BATCH NUMBER 874C51, AND NO NON-CONFORMANCES WERE IDENTIFIED.
PLUM 360 | K161469 | A COMPLAINT WAS RECEIVED WITH REGARDS TO A PLUM 360 THAT THERE WAS A DEATH WITHOUT KNOWN PRODUCT PROBLEM. IT WAS REPORTED THAT THE PATIENT HAD EXPIRED. ADDITIONALLY, IT WAS STATED THAT THEY DO NOT BELIEVE THE PUMP IS THE PROBLEM, BUT THE PATIENT'S FAMILY TOOK A PICTURE OF THE PUMP AND THEY QUESTIONED WHAT THEY SAW ON THE PUMP TO IMPLY THE CLINICIAN MAY NOT HAVE DONE SOMETHING RIGHT/CORRECT. IT WAS NOTED THAT THE EVENT OCCURRED DURING INFUSION. | THIS IS A GENERAL COMPLAINT, AND NO DEVICE WAS RETURNED TO THE SERVICE HUB FOR ANALYSIS AND EVALUATION. THEREFORE, NO VISUAL INSPECTION AND FUNCTIONAL TESTING COULD BE PERFORMED TO DETERMINE IF THE DEVICE PLAYED A ROLE IN THE ADVERSE EVENT. A 12-MONTH REVIEW WAS NOT ABLE TO BE CONDUCTED AS NO SERIAL NUMBER WAS PROVIDED TO ICU MEDICAL. PUMP EVENT LOGS WERE NOT PROVIDED. CONCLUSION: IN CONCLUSION, SINCE THE DEVICE WAS NOT RETURNED TO THE SERVICE HUB FOR ANALYSIS AND EVALUATION, NO VISUAL INSPECTION AND FUNCTIONAL TESTING COULD BE PERFORMED TO DETERMINE IF THE DEVICE HAD A ROLE IN THE ADVERSE EVENT.
TD60 TELEMETRY TRANSMITTER | K220058 | NO PRODUCT MALFUNCTION WAS REPORTED. IT WAS REPORTED THAT ON (B)(6) 2025, AT 20:15, THE PATIENT IN ROOM (B)(6), CONNECTED TO THE TD60 TELEMETRY MONITOR (ID TELE 22), WAS EXPERIENCING RHYTHM CHANGES AND MULTIPLE PVCS. THE PATIENT WAS ALSO REPORTED TO BE OFF TELE 45 MINUTES TO AN HOUR BEFORE ASYSTOLE. THE PATIENT EXPIRED. | MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED SYSTEM LOGS AND ECG STRIPS FOR INVESTIGATION. THE TD60 WAS MISPLACED BY THE USER AND WILL NOT AVAILABLE FOR PHYSICAL EVALUATION. THE INVESTIGATION IS ONGOING.
EDWARDS ESHEATH PLUS | K200258 | AS REPORTED BY A FIELD CLINICAL SPECIALIST, A PATIENT UNDERWENT A SUBCLAVIAN TAVR PROCEDURE WITH A 23MM SAPIEN 3 ULTRA RESILIA VALVE IN AORTIC POSITION. PRIOR TO THE PROCEDURE, A PROCTOR REVIEW RECOMMENDED A CAROTID APPROACH DUE TO A TORTUOUS RIGHT SUBCLAVIAN BUT THE PHYSICIAN PROCEEDED WITH SUBCLAVIAN. DURING THE PROCEDURE, THE ESHEATH WAS POINTING LATERAL WHEN THE DILATOR WAS REMOVED. THE TEAM THEN ADVANCED THE PIGTAIL CATHETER FOR ANGIOGRAM AND PERFORMED ANGIO IN THE DISSECTION PLANE. THE PATIENT HAD AN ASCENDING AORTIC DISSECTION UP TO THE INNONIMATE. VASCULAR SURGERY WAS CALLED TO CONSULT; HOWEVER, THE PATIENT WAS NOT A CANDIDATE. THE SURGEON CALLED ANOTHER FACILITY WITH VASCULAR SURGERY TO TAKE OVER CARE. THE PATIENT WAS TRANSFERRED TO PACU AWAITING TRANSFER IN STABLE CONDITION; HOWEVER, THE PATIENT EXPIRED IN THE ICU OVERNIGHT. AS PER FOLLOW-UP WITH THE FCS, THE OPERATOR ALSO PULLED BOTH THE DILATOR AND THE WIRE AFTER THE EDWARDS ESHEATH WAS INTRODUCED. THE SHEATH POPPED UP WITHOUT THE SUPPORT OF THE WIRE WHICH CAUSED THE DISSECTION. THERE WAS NO ALLEGATION THE ESHEATH MALFUNCTIONED. THE PERCEIVED ROOT CAUSE OF THE DEATH IS DUE TO THE AORTIC DISSECTION. THE DEVICES WERE DISCARDED. | A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO EVALUATION FINDINGS. SECTIONS B4, G3, G6, H2 AND H11 HAS BEEN UPDATED. THE COMPLAINT FOR DIFFICULTY INTRODUCING SHEATH WAS UNABLE TO BE CONFIRMED AS NO DEVICE AND/OR RELEVANT IMAGERY WERE PROVIDED. THE REPORTED EVENT DOES NOT ALLEGE A MALFUNCTION THAT COULD BE RELATED TO AN EDWARDS MANUFACTURING DEFICIENCY. A REVIEW OF IFU/TRAINING MATERIALS REVEALED NO DEFICIENCIES. AS REPORTED, "DURING THE PROCEDURE, THE ESHEATH WAS POINTING LATERAL WHEN THE DILATOR WAS REMOVED. THE TEAM THEN ADVANCED THE PIGTAIL CATHETER FOR ANGIOGRAM AND PERFORMED ANGIO IN THE DISSECTION PLANE. THE PATIENT HAD AN ASCENDING AORTIC DISSECTION UP TO THE INNONIMATE." ADDITIONAL FOLLOW-UP CLARIFIED, "THE OPERATOR ALSO PULLED BOTH THE DILATOR AND THE WIRE AFTER THE EDWARDS ESHEATH WAS INTRODUCED. THE SHEATH POPPED UP WITHOUT THE SUPPORT OF THE WIRE AND CAUSED THE DISSECTION." DURING TAVR, A GUIDEWIRE IS USED TO FACILITATE SAFE SHEATH PLACEMENT AS WELL AS TO KEEP THE SHEATH ORIENTATION SECURE AND PARALLEL WITH RESPECT TO PATIENT-S VASCULATURE. RETRIEVAL OF THE GUIDEWIRE ALONG WITH THE INTRODUCER, WITHOUT HOLDING THE SHEATH IN PLACE, COULD CAUSE THE SHEATH TO DISLOCATE AND DISSECT THE VESSEL, AS REPORTED. AS SUCH, AVAILABLE INFORMATION SUGGESTS THAT USE ERROR (DEVICE MANIPULATION TECHNIQUE) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT A TECH CALLED IN TO REPORT A PATIENT INCIDENT WHERE IT WAS NOTED THAT THE PATIENT PASSED AWAY. AT THE TIME, THE TECH DIDN'T HAVE ALL THE DEVICES INVOLVED AND SAID THEY WOULD CALL BACK WITH THE OTHER DEVICE INFORMATION AND WHETHER THEY WANTED THE INCIDENT INVESTIGATED. HOWEVER, IT WAS NOTED THAT THE TECH NEVER CALLED BACK. AFTER NUMEROUS ATTEMPTS OF REACHING OUT TO THE TECH, THEY NEVER RESPONDED. | THE ACTUAL DATE OF EVENT IS UNKNOWN. THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS SHOULD THE DEVICE BE REPAIRED OR THE DEVICE/LOGS BE RECEIVED FOR EVALUATION. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
MICRA | K132030 | IT WAS REPORTED DURING IMPLANT THAT THE PATIENT BECAME HYPOTENSIVE AFTER THE INTRODUCER WAS INITIALLY PLACED AND INTRA-ABDOMINAL BLEEDING WAS NOTED. MEDICATION WAS REQUIRED TO IMPROVE THE BLOOD PRESSURE. CATECHOLAMINES WERE PRODUCED AND THE PATIENT APPEARED IN BETTER CONDITION FOR SEVERAL MINUTES. THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) WAS PLACED AND THE PATIENT BECAME HYPOTENSIVE AGAIN. AN ECHO WAS PERFORMED AND BLOOD WAS OBSERVED IN THE RIGHT HIP. THE PATIENT PASSED AWAY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN DETERMINING WHETHER AN ALARM WAS ACTIVELY ACKNOWLEDGED RELATED TO A PATIENT DEATH. IT WAS INDICATED THE ALARMS FOR 'ECG LEADS OFF' WERE SILENCED, WHICH LED TO A DELAY IN CARE. THE PATIENT WAS FOUND AND SUCCESSFULLY RESUSCITATED BUT LATER PASSED AWAY; HOWEVER, THE CAUSE OF DEATH REMAINS UNKNOWN. NO OTHER CLINICAL INFORMATION WAS PROVIDED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL SPECIALIST REVIEWED THE LOGS PROVIDED. THE CLINICAL AUDIT LOG PROVIDES A CHRONOLOGICAL RECORD OF THE ALARMS AND ACTIONS. THE CLINICAL AUDIT LOG SHOWS FREQUENT SPO2 INOPS ALARM DURING THE TIME IN QUESTION, SUCH AS SPO2 SENSOR OFF, SPO2 NO PULSE, SPO2 PULSE? AND SPO2 NOISY SIGNAL. THESE ALARMS WERE ACKNOWLEDGED AT THE PIC IX. INOPS ARE TECHNICAL ALARMS, THEY INDICATE THAT THE MONITOR CANNOT MEASURE OR DETECT ALARM CONDITIONS RELIABLY. IF AN INOP INTERRUPTS MONITORING AND ALARM DETECTION (FOR EXAMPLE, LEADS OFF), THE MONITOR PLACES A QUESTION MARK IN PLACE OF THE MEASUREMENT NUMERIC AND AN AUDIBLE INDICATOR TONE WILL BE SOUNDED. THESE INOPS APPEAR IN THE CLINICAL AUDIT LOG. CANNOT ANALYZE ECG INOPS ARE ALSO SHOWN IN THE LOG STARTING AT 07:05:45 AND TRANSITED TO !!ECG LEADS OFF AT 7:33:52. CANNOT ANALYZE ECG INOP IS GENERATED IF 20 OF THE LAST 30 SECONDS ARE CLASSIFIED AS EITHER NOISY OR QUESTIONABLE. DURING THIS INOP/TECHNICAL CONDITION, THE ARRHYTHMIA ANALYSIS CONTINUES AND AN ALARM WILL BE ANNOUNCED IF AN ALARM CONDITION IS MET. HOWEVER, THE CANNOT ANALYZE ECG INOP/TECHNICAL ALARM INDICATES THAT THE EFFECTIVENESS OF THE ARRHYTHMIA MONITORING FOR THE PATIENT IS COMPROMISED AND A QUICK RESPONSE TO THIS ALARM IS IMPORTANT.!!! ECG LEADS OFF INDICATES THAT THE REQUIRED ECG LEADS ARE NOT AVAILABLE. AT 7:37:02, ALL ALARMS WERE PAUSED AT THE MONITOR. IN ADDITION, TESTS CARRIED OUT ON SITE CONFIRMED THAT THE DEVICE WAS FUNCTIONING WITHIN MANUFACTURER'S SPECIFICATIONS AND ARE READY FOR CLINICAL USE. BASED ON A REVIEW OF THE INFORMATION PROVIDED AND THE TESTING CONDUCTED ON SITE, THE PIC IX DEVICE WAS FUNCTIONING AS DESIGNED.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THAT A (B)(6) Y/O MALE PATIENT WAS BEING MONITORED WHEN INSTANCES OF V-FIB DID NOT ALARM ON THE MONITOR AND THE PATIENT PASSED AWAY. THE PATIENT HAD AN IMPLANTED ICD/PACEMAKER, WHICH ATTEMPTED TO PACE UNSUCCESSFULLY. THE EVENT OCCURRED IN THE MED SURG UNIT AND THE SPEAKERS WERE TESTED FOR FUNCTIONALITY. THE REMAINING QUESTIONS WERE RESPONDED TO WITH ASKU (ASKED BUT UNKNOWN), NOT ALL CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. AT THIS TIME IT CANNOT BE DETERMINED IF THIS DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. | DIAGNOSTIC/FUNCTIONAL TESTING WAS PERFORMED PER A PHILIPS FIELD SERVICE ENGINEER (FSE) ONSITE. THE FSE COLLECTED LOG FILES AND DATA FOR THE EVENT AND VERIFIED THE INTELLIVUE MX40 AND THE PATIENT INFORMATION CENTER IX (PIC IX) ENTERPRISE SYSTEM FOR PROPER OPERATION. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION TO THE PHILIPS SENIOR PRODUCT SUPPORT ENGINEER AND PHILIPS CLINICAL PRODUCT SPECIALIST TO VERIFY IF THE (B)(6) MX40 SN: (B)(6) PATIENT INFORMATION CENTER IX FUNCTIONED AS DESIGNED AND DID ALARM. THE INTELLIVUE MX40 AND THE PATIENT INFORMATION CENTER IX (PIC IX) DEVICE WERE CONFIRMED TO HAVE ALARMED FOR THE PATIENT¿S ECG CHANGES (BOTH YELLOW AND RED LEVEL ALARMS). THE EQUIPMENT WAS WORKING AS DESIGNED, CONFIGURED, AND OPERATED. THE UNACKNOWLEDGED ASYSTOLE ALARM SUPPRESSED VF/VT DETECTION. NO PRODUCT MALFUNCTION WAS DETECTED. THE RESULTS OF THE CLINICAL AUDIT LOG REVIEW SHOW: THE AUDIT LOGS SHOW THAT FOR THE 1-HOUR TIMEFRAME ON (B)(6)2025, FROM 13:00:00 UNTIL 14:00:00: A RED LEVEL CARDIAC ALARM WAS ACTIVE FOR 0:41:36 (41 MINUTES AND 36 SECONDS: ASYSTOLE FOR 40 MINUTES AND 13 SECONDS, VENTRICULAR FIBRILLATION/TACHYCARDIA FOR 1 MINUTE AND 23 SECONDS). A YELLOW LEVEL CARDIAC ALARM (IRREGULAR HR OR PAUSE) WAS ACTIVE FOR 15 MINUTES. THE AUDIT LOGS ALSO REVEAL: ISSUES IMPACTING ECG ANALYSIS: CANNOT ANALYZE ECG INOPERATIVE CONDITION FOR 27 MINUTES AND 29 SECONDS. LL LEAD OFF INOPERATIVE CONDITION FOR 20 MINUTES AND 26 SECONDS. ECG LEADS OFF (NO MONITORING POSSIBLE) FOR 3 MINUTES AND 26 SECONDS. BASED ON THE INFORMATION AVAILABLE AND THE RESULT OF ADDITIONAL ANALYSIS CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS DUE TO USER KNOWLEDGE. THE REPORTED PROBLEM WAS NOT CONFIRMED. ANALYSIS WAS PERFORMED AND INVESTIGATION HAS BEEN COMPLETED. INFORMATION WAS PROVIDED TO THE CUSTOMER OF THE ANALYSIS FINDINGS AND CONFIRMED THE DEVICE WAS WORKING TO ITS SPECIFICATION. THERE IS NO DEVICE MALFUNCTION AND THE REPORTED ISSUE WAS DUE TO USER KNOWLEDGE. THE DEVICE REMAINS AT THE CUSTOMER SITE. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED. - PLEASE SEE SPLIT CASE FOR INTELLIVUE MX40 (PRODUCT PROBLEM) REPORTED UNDER MFR REPORT NUMBER 1218950-2025-000075.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | IT WAS REPORTED THAT THE PATIENT OF THE REP DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY AND THE PATIENT'S WIFE NEEDED A RETURN LABEL TO SEND BACK THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | IT WAS REPORTED THAT THE DREAMSTATION AUTO CPAP DEVICE WAS NO LONGER AVAILABLE TO RETURN AND HAD BEEN DONATED. ACCORDING TO THE PATIENT'S DAUGHTER, HER FATHER PASSED AWAY LAST YEAR AND STOPPED USING THE DEVICE PRIOR TO PASSING. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. THE PATIENT'S DAUGHTER REQUESTED THAT NO FURTHER CORRESPONDENCE BE SENT. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | IT WAS REPORTED THAT THE PATIENT OF THE REMSTAR AUTO A-FLEX DEVICE HAD PASSED AWAY AND THE PATIENT'S WIFE NEEDED A RETURN LABEL TO SEND BACK THE DEVICE. CURRENTLY, THERE IS NO INFORMATION THAT THE DEATH IS RELATED TO THE DEVICE. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ESSENCE RENTAL WITH SSD 230V50HZ UK | K042655 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE ESSENCE RENTAL HAD STATES PATIENT HAS PASSED AWAY AND LOST THE DEVICE. NO MEDICAL INTERVENTION WAS SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
MONARCH ENDOSCOPY PLATFORM | K211493 | IT WAS REPORTED THAT THE PATIENT CODED DURING PROCEDURE, THE PHYSICIAN HAD CONDUCTED 4-BIOPSIES AND ON THE 5TH BIOPSY AN ABNORMAL AMOUNT OF BLOOD WAS COMING THROUGH THE ET TUBE. THE SCOPE AND THE PATIENT INTRODUCER WERE REMOVED AND THEY WERE ABLE TO STABILIZE THE PATIENT IN THE ROOM. THE PATIENT WAS MOVED TO ICU WHERE THEY LATER CODED AND PASSED AWAY. THE PHYSICIAN STATED THAT THE MONARCH DID NOT CAUSE OR CONTRIBUTE TO THE DEATH.
OXYLOG 3000PLUS | K103625 | IT WAS REPORTED THAT THE DEVICE HAD BEEN LEFT IN DEVICE CHECK SCREEN WHILE ON CHARGE, AT SOME POINT THE MAINS LEAD HAD SEPARATED FROM THE POWER SUPPLY TO NO LONGER MAKE CONTACT. ONCE PUT INTO USE THE BATTERIES WERE FLAT. FURTHERMORE, IT WAS REPORTED BY THE USER THAT THE DEVICE' "BATTERY WENT FLAT DURING RESUSCITATION IN EU, RESULTING IN A DELAY IN PROVIDING THERAPY AS MANUAL VENTILATION VIA A BAG AND MASK WAS REQUIRED. IT IS POSSIBLE THE INCIDENT HASTENED THE PATIENT'S DEMISE BUT ONLY BY MINUTES OR AT BEST HOURS IN THE OPINION OF THE MEDICS." | THE LOG FILE OF THE AFFECTED DEVICE WAS PROVIDED AND ANALYZED. IN ADDITION, THE POWER SUPPLY INCLUDING A POWER CORD WERE RETURNED FOR FURTHER INVESTIGATION. BASED ON THE LOG FILE ANALYSIS, A DEVICE MALFUNCTION COULD BE EXCLUDED AS ROOT CAUSE OF THE REPORTED EVENT (NEITHER DEVICE NOR INTERNAL BATTERY FAULT). ON THE (B)(6) 2025 THE LOG FILE SHOWS THAT THE DEVICE WAS TURNED ON AT 00:43 A.M. AND A DEVICE CHECK WAS SUCCESSFULLY PERFORMED. AT 07:46 A.M. A SECOND DEVICE CHECK WAS CARRIED OUT SUCCESSFULLY. AT 08:40:48 A.M., THE CONFIGURATION MODE WAS EXITED INDICATING THE START OF THE VENTILATION MODE. IMMEDIATELY AFTER SETTING THE VENTILATION MODE, AT 08:40:49 A.M., THE DEVICE ALARMED FOR ¿CHARGE INT. BATTERY¿ WITH MID PRIORITY (10 MINS OF BATTERY RUNTIME LEFT ACC. TO THE IFU) AND AGAIN AT 08:50:49 A.M.. AT 08:54:16 A.M., THE ALARM ¿INT. BATTERY DISCHARGED¿ WAS GENERATED BY THE DEVICE AND THE VENTILATION STOPPED DUE TO LOSS OF POWER. AT 08:56:12 A.M. THE DEVICE WAS MANUALLY SHUTDOWN BY USER. THE DEVICE IS EQUIPPED WITH DIFFERENT INDICATORS IN TERMS OF INTERNAL BATTERY CHARGE STATE AND EXTERNAL POWER SUPPLY CONNECTION. THE DISPLAY SYMBOLS OF THE POWER SUPPLY ARE LOCATED AT THE FRONT PANEL OF THE DEVICE AND INDICATE THE CHARGE STATUS OF THE INTERNAL BATTERY (A THREE COLORED INDICATOR) AND THE GREEN INDICATOR LIGHTS UP WHEN AN EXTERNAL POWER SOURCE IS CONNECTED. FURTHERMORE, THE DEVICE DISPLAYS THE REMAINING CAPACITY OF THE INTERNAL BATTERY IN INCREMENTS OF 25 % IN THE LOWER RIGHT SECTION OF THE INFORMATION WINDOW (¿BATTERY METER¿) WHILE IN OPERATION. ALARM MESSAGES OF CORRESPONDING PRIORITIES ARE GENERATED IF THE DEVICE DETECTS AN ISSUE WITH THE FUNCTION OF THE INTERNAL BATTERY OR THE NEED TO CHARGE IT. CHARGING THE BATTERY MIGHT BE REQUIRED TO ENSURE THAT THE DEVICE CAN CONTINUOUSLY BE OPERATED WITHOUT CONNECTING IT TO AN EXTERNAL POWER SOURCE. IN ADDITION TO THE LOG FILE ANALYSIS, THE POWER SUPPLY INCLUDING THE POWER CORD WAS SUBJECTED TO AN ENGINEERING TEST (TEST OF PULLING FORCES). THE POWER SUPPLY AS RETURNED FOR INVESTIGATION WAS WITH A POWER CORD OTHER THAN THE ORIGINAL POWER CORD (1844369), WHICH IS INSIDE THE PACKAGING OF THE OXYLOG 3000PLUS AS PART OF THE DEVICE CONFIGURATION. THE POWER CORD AS RETURNED HAS A STRAIGHT CONNECTION LEAD, WHICH HAS A HIGHER POSSIBILITY OF BEING PULLED OUT OF THE SOCKET. THE PULLING FORCES OF A POWER CORD DOES NOT DIRECTLY EXERT FORCE ON THE CONNECTION LEAD WHEN IT IS POSITIONED AT A 90-DEGREE ANGLE; THE ANGLE ALTERS THE DIRECTION OF THE FORCE VECTOR AND REDUCES THE PULLING FORCE ACTING ALONG THE CONNECTION LEAD PREVENTING UNINTENTIONAL UNPLUGGING. THIS ASSESSMENT COULD BE CONFIRMED BY THE PULLING FORCES AS MEASURED DURING THE ENGINEERING TEST. THE PULLING FORCE WHICH HAD TO BE APPLIED FOR PULLING THE ANGLED CONNECTION LEAD OUT OF THE POWER SUPPLY WAS SIGNIFICANTLY HIGHER (APPROX. 3 TIMES) COMPARED TO THE STRAIGHT CONNECTION LEAD. DRÄGER CONSIDERS THE CURRENT PRODUCT DESIGN OF THE POWER SUPPLY INCLUDING THE POWER CORD (1844369) TO BE PROVEN AND SAFE. AN ADDITIONAL RETENTION CLIP IS THEREFORE NOT NECESSARY. IN THE CURRENT EVENT, THE DEVICE REACTED AS SPECIFIED. BASED ON THE INVESTIGATION, THERE WAS NO DEVICE FAILURE FOUND. THE NUMBER OF SIMILAR CASES IS WITHIN THE EXPECTED RANGE OF THE RESPECTIVE RISK ASSESSMENT AND THUS ACCEPTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO HAD BEEN USING A DREAMSTATION AUTO CPAP DEVICE. THE CALLER STATED THE DEVICE WAS DISCARDED AND CANNOT BE RETURNED. NO FURTHER DETAILS WERE PROVIDED, AND THERE IS NO INDICATION THAT THE DEVICE WAS RELATED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
ALPHAVAC | K211081 | AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26 YEAR OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT . (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT (B)(4) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION (B)(4). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. (B)(6) PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE (B)(4). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING . THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING . AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 1 OF 3. | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION THAT ON (B)(6) 2025 AT APPROXIMATELY 10:19 A NURSE AT AN INSTITUTIONAL FACILITY DISCOVERED THE PATIENT WAS NOT BREATHING WHILE ON A TRILOGY EVO DEVICE. THE PATIENT WAS NOT BREATHING AT THE TIME, SO THE NURSE TOUCHED THE RESPIRATOR TO CHECK IF IT WAS WORKING, BUT WAS SO UPSET THAT THEY DID NOT REMEMBER HOW THEY OPERATED IT. THE CAUSE OF DEATH IS REPORTED TO BE RESPIRATORY FAILURE DUE TO THE PROGRESSION OF AMYOTROPHIC LATERAL SCLEROSIS (ALS). NO DIRECT ALLEGATION OF DEVICE CAUSE AND/OR CONTRIBUTION HAS BEEN LODGED WITHIN THE COMPLAINT RECORD. FURTHER ANALYSIS OF THE DEVICE HAS BEEN REQUESTED TO DETERMINE IF THERE WAS A PRESENCE OF A MALFUNCTION, HUMAN RELATED USE-ERROR, AND/OR DEVICE FULL FUNCTIONALITY. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT ON 27-JAN-2025 AT APPROXIMATELY 1019 A NURSE AT AN INSTITUTIONAL FACILITY DISCOVERED THE PATIENT WAS NOT BREATHING WHILE ON A TRILOGY EVO DEVICE. THE PATIENT WAS NOT BREATHING AT THE TIME, SO THE NURSE TOUCHED THE RESPIRATOR TO CHECK IF IT WAS WORKING, BUT WAS SO UPSET THAT THEY DID NOT REMEMBER HOW THEY OPERATED IT. THE CAUSE OF DEATH IS REPORTED TO BE RESPIRATORY FAILURE DUE TO THE PROGRESSION OF AMYOTROPHIC LATERAL SCLEROSIS (ALS). NO DIRECT ALLEGATION OF DEVICE CAUSE AND/OR CONTRIBUTION HAS BEEN LODGED WITHIN THE COMPLAINT RECORD. FURTHER ANALYSIS OF THE DEVICE HAS BEEN REQUESTED TO DETERMINE IF THERE WAS A PRESENCE OF A MALFUNCTION, HUMAN RELATED USE-ERROR, AND/OR DEVICE FULL FUNCTIONALITY. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER'S SERVICE CENTER, REVIEW OF THE ERROR LOG, CONFIRMED THAT THE TREATMENT OPERATION WAS STOPPED BY PRESSING A BUTTON ON THE DATE AND TIME SPECIFIED. NO ABNORMALITY IN THE DEVICE WAS CONFIRMED BY ANY OTHER INVESTIGATIONS. ADDITIONALLY, A PERIODIC MAINTENANCE 1 WAS PERFORMED. THE DEVICE'S AIR INLET FOAM FILTER AND PARTICLE FILTER WERE REPLACED TO PREVENT FAILURE. PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY. CONFIRMED THE SOFTWARE VERSION IS 1.06.10.00.
ALPHAVAC | K211081 | AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26 YEAR OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT. (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT (B)(4) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION (B)(4). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. (B)(6) PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE (B)(4). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING. THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING. AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 2 OF 3. | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
ALPHAVAC | K211081 | AN ANGIODYNAMICS SENIOR TERRITORY MANAGER REPORTED THAT A PATIENT EXPIRED DURING AN OFF-LABEL ALPHAVAC PROCEDURE. IT WAS REPORTED THAT THE MASS THAT WAS BEING REMOVED HAD EMBOLIZED TO THE LUNG, RESULTING IN THE PATIENT'S EXPIRATION. THE FOLLOWING WAS REPORTED FROM THE TECH WHO WAS SCRUBBED IN FOR THE PROCEDURE: "THE PATIENT WAS A 26 YEAR OLD MALE WITH ADVANCED METASTATIC CANCER AND HE HAD A TUMOR THAT WAS BLOCKING HIS IVC/RA. HE WAS TERMINAL BUT THIS TUMOR WAS CAUSING RENAL AND LIVER FAILURE AND THEY WANTED TO IMPROVE HIS QUALITY OF LIFE BEFORE HE DIED. DR. (B)(6) STARTED OUT BY TRYING TO SNARE THE TUMOR PIECE BY PIECE WITH AN ENSNARE AND THE ONO DEVICE. EARLY IN THE PROCEDURE WE SAW ON THE ECHO HOW HE HAD PULLED OFF A BIG PIECE, BUT IT FELL OUT OF THE SNARE AND WE DIDN'T KNOW WHERE I WENT. (B)(6) WENT ON AND WAS TAKING PIECES OF THE TUMOR OUT. HE SWITCHED FROM THE SNARE TO THE ALPHAVAC AT SOME POINT (22FR) AND IT WORKED OKAY. WE HAD AN ISSUE WITH IT AT ONE POINT (B)(4) BUT I CAN'T REMEMBER WHAT HAPPENED WITH IT WHICH IS WHY WE OPENED ANOTHER ONE. THE PATIENT STARTED CODING AT ONE POINT AND WE OPENED THE 18FR VERSION (B)(4). THE DILATOR WOULDN'T PASS THROUGH THE WHOLE DEVICE AS MUCH AS DR. (B)(6) PUSHED. THEY GOT IT TO WORK AT ONE POINT AND THE PATIENT RECOVERED (ON ROCKET FUEL THOUGH). THEN LATER ON ANESTHESIA SAID HE WAS HAVING A PE SO WE TOOK A PIC OF THE PA (HE ONLY HAD AN LPA AS HIS RIGHT LUNG WAS REMOVED) AND 2/3 OF IT WAS BLOCKED. PATIENT THEN STARTED CODING AGAIN. WHILE PATIENT WAS RECEIVING CPR, WE WANTED TO PUT THE 18 ALPHAVAC UP THERE BUT COULDN'T GET THE DILATORS TO WORK. WE OPENED A SECOND ONE (B)(4). THEY BASICALLY HAD THE SAME DILATOR PROBLEMS BUT HE JUST SHOVED IT IN I THINK AND IT DIDN'T GET ANYTHING OUT. PRETTY SURE WE PUT THE 22FR ONE UP THERE TOO BUT I CAN'T REMEMBER. ALL OF THEM GOT NOTHING . THEN WE OPENED THE PENUMBRA INDIGO AND IT ALSO GOT NOTHING. AFTER A LOT OF CPR AND THE FACT THAT HE WASN'T A CANDIDATE FOR ADVANCED MSC THEY CALLED IT. IT WAS PROBABLY A BIG PIECE OF THE TUMOR THAT EMBOLIZED AND GOT LODGED UP THERE AND WOULDN'T COME OUT. IT WAS A VERY HIGH STRESS CASE. BASICALLY THE PHYSICIAN HAD ISSUES GETTING THE DILATORS TO GO IN. ALSO THE WAY THE 22FR ONE IS FLARED AT THE TIP RESULTING IN THE PHYSICIAN HAVING A HAD A HARD TIME MAKING IT GO IN." DESPITE GOOD FAITH EFFORT ATTEMPTS, NO FURTHER DETAILS HAVE BEEN PROVIDED. DEVICE 3 OF 3. | THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. (B)(4) REFERENCE REPORT 1317056-2025-00019- DEVICE 1 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00020- DEVICE 2 OF 3, (B)(4) REFERENCE REPORT 1317056-2025-00021- DEVICE 3 OF 3.
VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE | K150709 | IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. IT WAS REPORTED THAT VERSACROSS CONNECT ACCESS SOLUTION FOR FARADRIVE WAS SELECTED FOR A PULSE FIELD ABLATION (PFA) PROCEDURE. THE PHYSICIAN COMPLETED AN OFF-LABEL CTI ABLATION AND POSTERIOR WALL ABLATION. NITRO AND NEO WAS GIVEN PRIOR TO ABLATING. THERE WERE NO ISSUES WITH TRANSSEPTAL, AND THE PATIENT WAS STABLE. AFTER ABLATING, THE PHYSICIAN CHECKED A HIS WHICH WAS 105. THE CASE WAS CONCLUDED, AND ALL CATHETERS WERE WITHDRAWN OUT OF THE BODY. THE PATIENT WAS CLOSED, AND ANESTHESIA WAS TOLD TO WAKE THE PATIENT UP. THE PATIENT WOULD NEED TO COME BACK LATER FOR A PACEMAKER. IT WAS THEN REPORTED THAT THE PATIENT WAS HAVING BRADYCARDIA EPISODES, A TRANSVENOUS PACING WIRES WERE PLACED AND AN ARTERIAL LINE. AN ULTRASOUND WAS PERFORMED AND RULED OUT ANY PERICARDIAL EFFUSION. THE PATIENT HAD A LOW EJECTION FRACTION (EF) PRIOR TO ABLATION. A SWAN CATHETER WAS PLACED, AND THE PATIENT WAS SENT TO INTENSIVE CARE UNIT (ICU) WHERE WENT INTO VENTRICULAR TACHYCARDIA AND DID NOT MAKE IT. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. THERE WAS NO DEVICE MALFUNCTIONS REPORTED, NOR COMPLICATIONS DURING THE WHOLE PROCEDURE. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
GELSOFT | K955230 | EVENT DATE: (B)(6) 2022. IMPLANT DATE: (B)(6) 2020. EVENT REPORTED FROM PANTHER POST MARKET STUDY (B)(4). GELSOFT PLUS STRAIGHT IMPLANTED IN A MALE PATIENT (FOR ABOVE KNEE FEMOROPOPLITEAL) WITH PRE-CONDITION REPORTED AS CORONARY ARTERY DISEASE, CHRONIC OBSTRUCTIVE PULMONARY DISEASE, HYPERTENSION (CONTROLLED BY MEDICATION), HYPERLIPIDEMIA (CONTROLLED BY MEDICATION), TYPE I DIABETES (CONTROLLED BY MEDICATION), RENAL INSUFFICIENCY, PERIPHERAL ARTERIAL OCCLUSIVE DISEASE ("PAD" - INTERPRETED AS PERIPHERAL ARTERIAL DISEASE), ASA (PHYSICAL STATUS CLASSIFICATION SYSTEM), CLASS III, CURRENT SMOKER. INDICATION: OCCLUSIVE DISEASE, TYPE: DISABLING INTERMITTENT CLAUDICATION. THE STUDY SITE REPORTED THE EVENT OF DEATH. CAUSE IS UNKNOWN. THERE WAS NO AUTOPSY. DATE OF DEATH WAS REPORTED AS (B)(6) 2022. NO DEVICE DEFICIENCY REPORTED BY THE SITE IN THE STUDY DATABASE. THIS REPORT IS BEING SUBMITTED AS FOLLOW UP#1 FOR MANUFACTURING REPORT NUMBER 9612515-2025-00013 TO PROVIDE EVENT CLOSURE INFORMATION FOR TAG0152. | MANUFACTURER NARRATIVE: CLINICAL CODE: 4580 - INSUFFICIENT INFORMATION - EVENT OF PATIENT DEATH REPORTED WITH NO CAUSE OF DEATH AND NO DEVICE DEFICIENCY. IMPACT CODE: 1802 - DEATH - REPORTED AS DEATH EVENT NO AUTOPSY WAS PERFORMED. DEVICE PROBLEM CODE : 2993 -ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM- NO CAUSE OF DEATH AND NO DEVICE FAILURE /DEFICIENCY HAS BEEN REPORTED NO AUTOPSY WAS PERFORMED. COMPONENT CODE: 4755 - PART/COMPONENT/SUB-ASSEMBLY TERM NOT APPLICABLE TYPE OF INVESTIGATION: 4110 - TREND ANALYSIS: A 4 YEAR SIMILAR EVENT REVIEW COULD NOT BE PERFORMED AS THIS IS THE ONLY EVENT OF GELSOFT /MEDICAL EVENT / DEATH (UNKNOWN CAUSE NO DEVICE DEFICIENCY REPORTED (B)(6) 2021 - (B)(6)2025) 4111 - COMMUNICATION INTERVIEW: ADDITIONAL INFORMATION WAS RECEIVED STATING THE PATIENT SUFFERED POST-OPERATIVE COMPLICATIONS OF ACUTE PULMONARY EDEMA (OAP) DUE TO OVERLOAD. GOOD PROGRESS AND WEANING OFF OXYGEN AFTER DAILY DIALYSIS FOR 3 DAYS + PROBABILISTIC ANTIBIOTICS DUE TO PROBABLE SUPERINFECTION. ON (B)(6)2021 PATIENT HAD A FOLLOW-UP AWAITING KIDNEY TRANSPLANT. GOOD PERMEABILITY ON THE RIGHT BUT STENOSES ON THE LEFT. HYPERIMMUNIZATION IN (B)(6) 2021, CLASS 1, TGI 73%. - CONTRAINDICATION MAINTAINED FOR PERIOPERATIVE BYPASS. THE SITE HAVE CONFIRMED THEY HAVE NO MORE INFORMATION. THEY WERE AWARE OF THE DEATH VIA A DATABASE IN FRANCE, WHERE DATE OF DEATH IS WRITTEN. AS IT DID NOT HAPPEN AT THE HOSPITAL, THEY ARE NOT ABLE TO RETRIEVE FURTHER INFORMATION, WHICH IS WHY THEY SELECTED "UNKNOWN". NO AUTOPSY WAS PERFORMED. ADDITIONAL INFORMATION ON THE POSSIBLE INFECTION AND IS SWABS WERE TAKEN WAS REQUESTED ON (B)(6)2025 SITE RESPONDED ON (B)(6) - THE SITE HAVE RESPONDED TO THE QUESTION ABOUT THE INFECTION: "ACUTE RESPIRATORY DISTRESS DUE TO FLUID OVERLOAD IN THE CONTEXT OF DELAYED DIALYSIS, WITH A POSSIBLE INFECTIOUS COMPONENT UNDER INVESTIGATION. HE ALSO RECEIVED PROBABILISTIC ANTIBIOTIC THERAPY FOR A PROBABLE PULMONARY SUPERINFECTION WITH TAZOCILLINE AND ROVAMYCINE. NO CLEAR INFECTION IDENTIFIED; THE TREATMENT IS PROBABILISTIC. A BACTERIOLOGICAL SAMPLE WAS TAKEN, BUT NO RESULTS ARE REFERENCED IN THE MEDICAL RECORD. THE COVID PCR TEST IS NEGATIVE. SCANS WERE RECEIVED AND REVIEWED FOR THIS EVENT , HOWEVER THE IMAGING PROVIDED ONLY COVERS THE CHEST AND THE GRAFT WAS USED IN THE LEG. ADDITIONAL SCANS WERE REQUESTED AND UPON RECEIPT THESE WERE FOUND TO BE ULTRASOUND SCANS SO NO REVIEW WAS POSSIBLE. 3331 - ANALYSIS OF PRODUCTION RECORDS: FULL BATCH REVIEW WAS PERFORMED FROM BASE FABRIC TO FINISHED PRODUCT WHICH CONFIRMED THE GRAFT WAS MANUFACTURED TO SPECIFICATION. 4117 - DEVICE NOT RETURNED: DEVICE REMAINS IMPLANTED. INVESTIGATION FINDINGS: 213-NO DEVICE PROBLEM FOUND - NO DEVICE DEFICIENCY OR CAUSE OF DEATH WAS REPORTED AND FULL BATCH REVIEW SHOWED GRAFT WAS MANUFACTURED TO SPECIFICATION AS SCANS WERE ULTRASOUND NO REVIEW OF THESE COULD BE PERFORMED. INVESTIGATION CONCLUSION: 67 - NO PROBLEM DETECTED, IT WAS NOT CONFIRMED THERE WAS A PROBLEM WITH THE DEVICE -NO CAUSAL LINK BETWEEN THE EVENT AND THE DEVICE COULD BE ESTABLISHED.
EDWARDS ESHEATH PLUS INTRODUCER SHEATH | K200258 | AS REPORTED BY THE FIELD CLINICAL SPECIALIST (FCS), DURING IMPLANT OF A 26 MM SAPIEN 3 ULTRA VALVE IN THE AORTIC POSITION, THE VALVE WAS UNABLE TO PASS THROUGH THE ESHEATH+ AND A STRUT BENT WHILE TRYING TO PUSH THE VALVE THROUGH IT. A SECOND VALVE WAS ALSO USED IN THE SAME MANNER AND ALSO DAMAGED. THE DELIVERY SYSTEM WAS UNABLE TO ADVANCE THE FIRST TIME OUT OF THE PARTIAL EXPANDING PORTION AND THE SECOND TIME IT WAS STUCK JUST DISTAL TO THE PARTIAL EXPANDING PORTION. IT'S NOTED THERE WAS A TEAR AT THE PARTIALLY EXPANDING TO FULLY EXPANDING TRANSITION POINT OF THE FIRST ESHEATH+. THE SECOND ESHEATH+ WAS DAMAGED AFTER THE CASE IN THE PARTIALLY EXPANDING PORTION DUE TO TRYING TO GET THE VALVE OUT TO EXAMINE. THE PATIENT HAD RAPID DECOMPENSATION WITH NO EFFUSION NOTED IN THE CHEST. THE PHYSICIAN BELIEVES CAUSE WOULD POSSIBLY BE FROM RETROPERITONEAL BLEED FROM THE ILIAC ARTERY, HOWEVER NO ANGIO WAS TAKEN DURING CPR AND LITTLE TO NO FLOW MADE IT INTO THE ILLIACS TO EXAMINE IF/ WHERE THE PATIENT WAS BLEEDING FROM. ADDITIONAL INFORMATION RECEIVED NOTED IT'S STILL AN UNKNOWN CAUSE OF DEATH AND AN AUTOPSY IS BEING DONE, THOUGH THE RESULTS HAVE NOT YET BEEN MADE AVAILABLE. AFTER REVIEWING THE CASE WITH THE CT SURGEON, IT WAS NOTED IN THE ANESTHESIA CHART THAT THE PATIENT'S PRESSURE WAS MAPPING IN THE 30-40 FOR A TOTAL OF 8 MINUTES BEFORE ANY ACLS OR ACTION WAS TAKEN. IT'S UNKNOWN IF THE PATIENT HAD A BLEED SOMEWHERE OR IF THEY JUST DECOMPENSATED AND COULD NOT RECOVER. THE PERCEIVED ROOT CAUSE OF THE BENT STRUTS WAS THAT THE STRUTS CAUSE ON THE ILIAC STENTS. THE PERCEIVED ROOT CAUSE OF THE RESISTANCE WAS EITHER THE STENTS IN THE ILIAC ARTERY OR CALCIUM BURDEN. | A SUPPLEMENTAL MDR IS BEING SUBMITTED FOR A COMPLETION TO THE ENGINEERING EVALUATION. THE FOLLOWING SECTIONS OF THIS REPORT HAS BEEN UPDATED OR CORRECTED: B4, G3, G6, H2, H3, H6, H11. THE PRODUCT WAS RETURNED; THEREFORE, A PRODUCT EVALUATION INVESTIGATION WAS COMPLETED. FUNCTIONAL TESTING WAS PERFORMED ALONG WITH A VISUAL EXAMINATION. THE ASSOCIATED COMMANDER DELIVERY SYSTEM WAS ADVANCED THROUGH THE SHEATH AND THE REMAINING SHEATH EXPANDED AND THE TIP OPENED, AS DESIGNED. AS THE ASSOCIATED DELIVERY SYSTEM WAS ABLE TO BE FULLY ADVANCED THROUGH THE SHEATH WITH NO ISSUES, DIMENSIONAL TESTING INCLUDING MEASURING THE DELIVERY SYSTEM FLEX TIP OUTER DIAMETER (OD) WAS NOT NECESSARY TO BE PERFORMED. THE RETURNED DEVICES WERE VISUALLY EXAMINED, AND THE FOLLOWING WAS OBSERVED: SHEATH SHAFT CURVATURE. EXPOSED THV THROUGH PUNCTURED STRAIN RELIEF APPROXIMATELY 1.5" FROM COMNUT. ADDITIONAL SHEATH PUNCTURE APPROXIMATELY 0.7" FROM DISTAL STRAIN RELIEF. LINER PARTIALLY EXPANDED APPROXIMATELY 4.5" FROM STRAIN RELIEF. REMAINING LINER UNEXPANDED AND DISTAL TIP UNOPENED. SHEATH SHAFT KINK APPROXIMATELY 1" FROM DISTAL STRAIN RELIEF. PROVIDED IMAGERY WAS REVIEWED, CONFIRMING THE COMPLAINT EVENT. THE 3MENSIO IMAGERY WAS PROVIDED AND THE FOLLOWING WAS OBSERVED: PATIENT'S ACCESS VESSELS HAD PRESENCE OF TORTUOSITY, CALCIFICATION/STENTS, AND UNDERSIZED VESSEL DIAMETERS. THE 14F ESHEATH+ IS INDICATED FOR USE WITH VESSEL DIAMETERS '5.5MM. AS THE RETURNED DEVICE MEETS SPECIFICATION WITH NO EVIDENCE TO SUPPORT A MANUFACTURING/DESIGN DEFECT HAS CONTRIBUTED TO THE COMPLAINT, A MANUFACTURING MITIGATION REVIEW IS NOT REQUIRED. WHILE THE IFU/TRAINING MANUALS SPECIFIC TO THIS COMPLAINT EVENT ARE NOT LISTED IN THE TECHNICAL SUMMARY WRITTEN BY EDWARDS LIFESCIENCES, REVIEW OF THE INSTRUCTIONS REVEALS SIMILAR CONTENT AS DOCUMENTED IN THE TECHNICAL SUMMARY. THEREFORE, THE REVIEW OF THE INSTRUCTIONS/MANUALS WITHIN THE TECHNICAL SUMMARY REMAIN EQUIVALENT TO THE INSTRUCTIONS/MANUALS FOR THIS COMPLAINT EVENT. THE CONTENT ADEQUATELY PROVIDES WARNINGS AND INSTRUCTIONS FOR USE REGARDING THE REPORTED COMPLAINT EVENT. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. CURRENT RISK MITIGATIONS INCLUDE DESIGN AND MANUFACTURING CONTROLS, IFU WARNINGS AND CAUTIONS, AND PHYSICIAN TRAINING. NO EVIDENCE OF PRODUCT NON-CONFORMANCES OR LABELING/IFU INADEQUACIES WERE IDENTIFIED IN THE EVALUATION.' THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR INABILITY TO ADVANCE THROUGH SHEATH WAS CONFIRMED BASED ON THE EVALUATION OF THE RETURNED DEVICE AND PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (UNDERSIZED VESSEL, CALCIFICATION/STENTS, TORTUOSITY) LIKELY CONTRIBUTED TO THE EVENT AS THESE PATIENT FACTORS WERE CONFIRMED VIA 3MENSIO IMAGERY. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. UNDERSIZED VESSELS (E.G. DUE TO ANATOMICAL VARIATION, PRE-EXISTING STENT OR GRAFT, SCAR TISSUE, OR FIBROSIS) CAN CREATE A RESTRICTED PATHWAY OR CONSTRAINED CONDITION RESULTING IN DIFFICULTY DURING SHEATH EXPANSION AND INCREASED RESISTANCE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM. THE COMPLAINT FOR SHEATH PUNCTURED WAS CONFIRMED BASED ON THE EVALUATION OF THE RETURNED DEVICE AND PROVIDED IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (UNDERSIZED VESSEL, CALCIFICATION, TORTUOSITY) AND/OR PROCEDURAL FACTORS (DEVICE MANIPULATION) LIKELY CONTRIBUTED TO THE EVENT AS INDICATED BY THE CONDITION OF THE RETURNED DEVICE. IT IS POSSIBLE THAT ADDITIONAL DEVICE MANIPULATION TO OVERCOME THE RESISTANCE MAY BE APPLIED DURING THE PROCEDURE RESULTING IN DAMAGE TO THE SHEATH. DURING DELIVERY SYSTEM ADVANCEMENT THROUGH A CHALLENGING PATHWAY, IT IS POSSIBLE THAT THE DEVICES MAY NOT BE COAXIALLY ALIGNED. THIS CAN LEAD TO THE CRIMPED VALVE TO CATCH ONTO THE SHEATH HDPE AND/OR STRAIN RELIEF AND LEAD TO DAMAGE. EXCESSIVE DEVICE MANIPULATION APPLIED TO OVERCOME THE RESISTANCE CAN FURTHER DAMAGE THE SHEATH THROUGH CONTINUED INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. VESSEL CALCIFICATION AND/OR TORTUOSITY IF PRESENT CAN EXASPERATE THE INTERACTION BETWEEN THE CRIMPED VALVE AND SHEATH. SHARP CALCIFIED NODULES CAN DIRECTLY WEAKEN THE SHEATH SHAFT MAKING IT MORE SUSCEPTIBLE TO DAMAGE AS THE DELIVERY SYSTEM WITH CRIMPED VALVE IS ADVANCED THROUGH. TORTUOSITY CAN SUBJECT THE SHEATH TO SUBOPTIMAL ANGLES THAT CAN LEAD TO SHAFT KINKS. EXCESSIVE MANIPULATION CAN LEAD TO SHEATH SHAFT DAMAGE (I.E. KINKS, SCRATCHES, TIP DAMAGE) IF COMPOUNDED WITH VESSEL CALCIFICATION AND/OR TORTUOSITY. THE PRODUCT RISK ASSESSMENT (PRA) IS CAPTURED IN THE TECHNICAL SUMMARY, WHICH APPLIES TO THIS COMPLAINT EVENT. THE RISKS OUTLINED IN THE PRA REMAIN THE SAME. THE PRA DOCUMENTS THE DIFFICULTY ADVANCING THE DELIVERY SYSTEM THROUGH THE SHEATH, RESULTING IN DIFFICULTY OR THE INABILITY TO INTRODUCE DELIVERY SYSTEM/LOADER AND ADVANCE THROUGH SHEATH, PROLONGING PROCEDURE, WHICH DID OCCUR DURING THIS EVENT. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. AS THE COMPLAINT DID NOT EXCEED THE PRA RE-ESCALATION THRESHOLD, NO FURTHER ACTION IS REQUIRED. A CAPA/SCAR IS NOT REQUIRED BECAUSE AN EDWARDS/SUPPLIER DEFECT WHICH COULD HAVE RESULTED IN THE COMPLAINT WAS NOT CONFIRMED. PRIOR CORRECTIVE ACTION, DETAILED IN THE TECHNICAL SUMMARY, CAPTURES PRIOR HIGH PUSH FORCE INVESTIGATION AND CORRECTIVE/PREVENTATIVE ACTION. TRENDING ANALYSIS INDICATES COMPLAINT RATES ARE WITHIN THE CONTROL LIMIT. NO FURTHER ESCALATION IS NEEDED AT THIS TIME.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K150709 | DURING THE PROCEDURE, A PATIENT DEATH OCCURRED. IT HAS BEEN REPORTED THAT A VERSACROSS CONNECT LAAC ACCESS SOLUTION KIT WAS SELECTED FOR USE FOR A WATCHMAN PROCEDURE. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE) WAS DONE IMMEDIATELY PRIOR TO PLACING THE PATIENT UNDER TWILIGHT CONSCIOUS SEDATION. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING WAS PERFORMED WHILE A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) WAS INSERTED WITH VERSACROSS CONNECT TO PERFORM TRANSSEPTAL PUNCTURE (TSP). TSP WENT SMOOTHLY AND NO SUSPICION OF EVENT FROM TSP. THE INTRAARTERIAL SEPTUM (IAS) WAS FLOSSED WITH THE WAS AND VCC AND THE WAS AND ICE CATHETER WAS POSITIONED ACROSS THE IAS INTO THE LEFT ATRIUM (LA). DURING A FIVE (5) MINUTE PERIOD, THE WAS AND ICE REMAINED IN THE LA UNATTENDED AS THE PHYSICIAN NEEDED TO LEAVE THE ROOM BRIEFLY, AND WHEN THE PHYSICIAN RETURNED THE WAS WAS NOT ABLE TO BE VISUALIZED ON IMAGING. THE QRS INTERVAL BEGAN WIDENING, BRADYCARDIA WAS NOTED AND THEN ASYSTOLE. AS STAFF WAS PREPARING TO PERFORM CARDIOPULMONARY RESUSCITATION (CPR), VENTRICULAR FIBRILLATION WAS OBSERVED. THE STAFF AND PHYSICIAN CODED THE PATIENT FOR TWO (2) HOURS USING A TEMPORARY PACEMAKER, INTRA-AORTIC BALLOON PUMP (IABP), DIAGNOSTIC CARDIAC CATHETERIZATION (CC), AND DEFIBRILLATION. TEE WAS REINSERTED AND THERE WAS NO EVIDENCE OF PERICARDIAL EFFUSION. THE TEMPORARY PACEMAKER WAS FIRST UNABLE TO CAPTURE BETWEEN COMPRESSION CYCLES AND ULTIMATELY ONE LEAD WAS PLACED AND CAPTURED SUCCESSFULLY. THE DIAGNOSTIC CC REVEALED NO SIGN OF CORONARY DISEASE. ALTHOUGH IABP WAS SUCCESSFULLY PLACED, THE PATIENT MAINTAINED ONLY A WEAK BLOOD PRESSURE AND POOR OXYGEN SATURATION MAINTENANCE AT 70 OR BELOW. DUE TO THE LONG PERIOD OF TIME OF LOW OXYGEN SATURATION, THE PHYSICIAN DECIDED TO STOP RESUSCITATING THE PATIENT AND THE PATIENT DIED. IT WAS NOTED THAT THERE WAS NOT ANY HEAVY SNORING, NO GASPS OF AIR TAKEN BY THE PATIENT, AND NO CONTRAST INJECTIONS PERFORMED WITHIN THE LA. PRODUCT IS NOT EXPECTED TO RETURN FOR ANALYSIS. IN THE PHYSICIAN'S OPINION, THE DEVICE/PROCEDURE DID NOT CONTRIBUTE TO THE PATIENT COMPLICATION. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION IS IN PROGRESS, BUT IT WAS NOT AVAILABLE AS OF YET. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
HT70 VENTILATOR | K111146 | IT WAS REPORTED THAT WHILE IN USE ON AN ELDERLY, TERMINALLY ILL PATIENT, THE HT70 VENTILATOR GENERATED AN ALARM AND VENTILATION STOPPED. WHILE BEING TRANSFERRED TO AN ALTERNATE VENTILATOR, THE PATIENT EXPERIENCED CARDIAC ARREST AND CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, BUT THE PATIENT DID NOT SURVIVE. | H3: IMAGES OF THE SINGLE BOARD COMPUTOR (SBC) BOARD AND THE MAIN CONTROL BOARD WERE RECEIVED FOR REVIEW. THERE WAS THERMAL DAMAGE OBSERVED WHICH INCLUDED THE C92 CAPACITOR ON THE MAIN CONTROL BOARD. THERE IS AN EXISTING PREVIOUS INTERNAL INVESTIGATION REGARDING THIS ISSUE. THE ACTUAL SBC AND MAIN CONTROL BOARDS WERE RECEIVED BY MEDTRONIC AND ARE UNDERGOING ANALYSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | K213919 | IT WAS REPORTED BY A HEALTHCARE PROVIDER VIA MAUDE (MW5164857) AND BY THE FOOD AND DRUG ADMINISTRATION THAT A PATIENT PASSED AWAY ON (B)(6) 2025. THE MAUDE REPORT STATES THAT THE PATIENT WAS SUPPOSED TO COME TO THE HEALTHCARE PROVIDER (HCP) OFFICE BUT CANCELLED AND DID NOT RESCHEDULE. THE DATE OF DEATH WAS NOT PROVIDED. THERE IS NO REPORTER INFORMATION TO GATHER ADDITIONAL INFORMATION AS THE HCP DID NOT WISH TO FOLLOW UP AND THERE IS NO PATIENT INFORMATION AVAILABLE. NO DATA OR PRODUCT WAS PROVIDED FOR INVESTIGATION. THE ALLEGATION WAS UNDETERMINED. THE PROBABLE CAUSE COULD NOT BE DETERMINED. | (B)(4).
ION | K182188 | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT EXPERIENCED A CARDIAC ARREST AND THE PROCEDURE WAS ABORTED. ADVANCED CARDIAC LIFE SUPPORT (ACLS) WAS ADMINISTERED, RESULTING IN THE RETURN OF SPONTANEOUS CIRCULATION (ROSC). HOWEVER, THE PATIENT CODED FIVE MORE TIMES IN THE PROCEDURE ROOM. EACH TIME, ACLS WAS PERFORMED UNTIL THE PATIENT WAS STABLE ENOUGH FOR TRANSFER TO THE INTENSIVE CARE UNIT. THE PATIENT CODED AGAIN IN THE ICU AND AT AN UNSPECIFIED TIME LATER PASSED AWAY. THE INTUITIVE ENDOLUMINAL TERRITORY ASSOCIATE (ETA) WAS PRESENT DURING THE ION PROCEDURE. THE PHYSICIAN INFORMED THE ETA THAT THE CARDIAC ARREST AND THE PATIENT OUTCOME WERE NOT RELATED TO THE ION SYSTEM AND WOULD LIKELY HAVE OCCURRED WITH ANOTHER MODALITY. THE PHYSICIAN ATTRIBUTED THE EVENT TO THE PATIENT'S EXTENSIVE MULTIPLE COMORBID CONDITIONS; THE PATIENT'S MEDICAL HISTORY WAS NOT PROVIDED. THERE WAS NO DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. | THE PHYSICIAN STATED THAT THE PATIENT OUTCOME WAS RELATED TO EXTENSIVE, MULTIPLE CO-MORBIDITIES. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND THERE WERE NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH EXTENSIVE MEDICAL COMORBIDITIES NOT OTHERWISE SPECIFIED UNDERWENT AN ION LUNG BIOPSY. DURING THE PROCEDURE THE PATIENT SUFFERED MULTIPLE EPISODES OF CARDIAC ARREST WITH EVENTUAL ROSC THEN TRANSFERRED TO THE ICU WHERE THE PATIENT ARRESTED AGAIN AND ULTIMATELY DIED. THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT OR ACCESSORIES. THE AVAILABLE DATA SUGGESTS THE ADVERSE EVENT WAS RELATED TO THE PROCEDURE IN A PATIENT WITH SIGNIFICANT UNDERLYING CO-MORBIDITIES. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 4 TOTAL DEATHS (0.02%). A MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO DEATHS ASSOCIATED WITH 581 CASES. A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). ANOTHER SURVEY-BASED STUDY OF 103,978 BRONCHOSCOPIES DESCRIBED 71 CASES (0.068%) OF ASSOCIATED CARDIOVASCULAR EVENTS. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. ---FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). ---OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. ---FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. ---KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. ---BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. FOR SECTION B2: DATE OF DEATH IS ESTIMATED AS SAME DAY AS PROCEDURE(B)(6) 2025 OR FOLLOWING DAY (B)(6) 2025. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LEFT ADRENALECTOMY PROCEDURE, BLEEDING OCCURRED AND THE PROCEDURE WAS CONVERTED TO OPEN. IT WAS REPORTED THAT THE PATIENT EXPIRED. IT IS UNKNOWN IF THE PATIENT EXPIRED INTRA-OPERATIVELY OR POST-OPERATIVELY. THE CUSTOMER REPORTED THAT THIS EVENT HAD NOTHING TO DO WITH THE DA VINCI PRODUCTS USED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION HAVE BEEN UNSUCCESSFUL. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, PERMANENT CAUTERY HOOK, SMALL GRASPING RETRACTOR, CADIERE FORCEPS, FENESTRATED BIPOLAR FORCEPS, LONG BIPOLAR GRASPER, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN ADRENALECTOMY. NEITHER THE DIAGNOSIS, NOR COMORBIDITIES, DETAILS OF THE CASE, POST OPERATIVE COURSE, OR CAUSE OF DEATH ARE PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN INFORMED OF THE PASSING OF A PATIENT WHO WAS USING A DREAMSTATION AUTO CPAP DEVICE. A REQUEST WAS MADE FOR ALL COMMUNICATION TO STOP. NO FURTHER DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE PLAYED A ROLE IN THE PATIENT'S DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
NXSTAGE SYSTEM ONE | K170469 | A REPORT WAS RECEIVED ON 21 JAN 2025 FROM THE HOME THERAPY NURSE (HTN) OF A 58-YEAR-OLD FEMALE PATIENT APPROVED FOR PERFORMING SOLO HEMODIALYSIS, WITH A MEDICAL HISTORY INCLUDING DIABETES, CARDIAC COMORBIDITIES AND END STAGE RENAL DISEASE WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED ON 28 JAN 2025 FROM THE HOME THERAPY NURSE (HTN) WHO STATED THAT THE PATIENT WAS FOUND EXPIRED WITH RINSEBACK COMPLETED AND STILL CONNECTED TO THE DEVICE. PER THE HTN, THE CAUSE OF DEATH IS LISTED AS 'ESRD IN DIABETES' AND WAS UNRELATED TO NXSTAGE PRODUCT OR THERAPY. | THERE WAS NO DEVICE MALFUNCTION. THE CYCLER WAS RECEIVED FOR EVALUATION AND SUCCESSFULLY PASSED TESTING. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. AVAILABLE LOG FILES WERE RETRIEVED AND ANALYZED WHICH SHOWED DEVICE PERFORMANCE WAS WITHOUT DEFICIENCY AND WAS UNREMARKABLE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. THE NXSTAGE SYSTEM ONE USER GUIDE SUPPLEMENT FOR SOLO HOME HEMODIALYSIS OUTLINES RISKS, WARNINGS, AND REQUIREMENTS FOR THE SOLO HOME HEMODIALYSIS PATIENT. UDI: (B)(4).
DEVILBISS HEALTHCARE | K071397 | DEVILBISS HEALTHCARE WAS RECENTLY NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY HEALTH CANADA, WHICH STATED "OXYGEN COMPRESSOR (USED AT PATIENT'S HOME) HAD LOW FLOW DUE TO BACTERIAL OVERGROWTH. THE USER APPARENTLY CONTACTED ONTARIO MEDICAL SUPPLY AS THE WARNING LIGHT WAS FLASHING "YELLOW". AT THE TIME OF DEATH ((B)(6) 2022) THERE HAD BEEN NO RESPONSE, APPARENTLY, FROM ONTARIO MEDICAL SUPPLY. ALSO, NO MAINTENANCE RECORDS WERE AVAILABLE FROM ONTARIO MEDICAL SUPPLY DURING THE INVESTIGATION (OMS REPRESENTATIVE STATED THEY HAD BEEN DESTROYED IN A HACKING INCIDENT)." THE DEVICE IS EQUIPPED WITH A YELLOW INDICATOR LIGHT WITH INTERMITTENT AUDIBLE BEEPING THAT ACTIVATES WHEN THE OXYGEN CONCENTRATION BEGINS TO FALL BELOW SPECIFICATION. THIS LIGHT IS AN INDICATOR THAT THE EQUIPMENT MAY BE IN NEED OF ROUTINE MAINTENANCE. IT IS UNCLEAR HOW BACTERIAL OVERGROWTH WAS DETERMINED OR HOW SUCH A CONDITION WOULD CONTRIBUTE TO THE STATED LOW FLOW. THE DEVICE CAN BE USED WITH A HUMIDIFIER ACCESSORY TO PROVIDE MOISTURE TO THE OXYGEN. IT IS POSSIBLE THAT THE "BACTERIAL OVERGROWTH" WAS IN THE HUMIDIFIER ACCESSORY AS THAT IS VISIBLE TO USER. THE DEVICE IS NOT A LIFE-SUPPORTING OR LIFE-SUSTAINING DEVICE. THE DURABLE MEDICAL EQUIPMENT (DME) SUPPLIER WAS NOTIFIED OF THE INCIDENT WITH NO RESPONSE. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. DEVILBISS HEALTHCARE WILL FILE AN UPDATE IF ADDITIONAL INFORMATION BECOMES AVAILABLE.
STEALTHSTATION¿ S8 SYSTEM | K162309 | ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE HEALTH CARE PROFESSIONAL NOTICED A LACK OF PRECISION IN THEIR NAVIGATION AND THOUGHT IT WOULD BE RELATED TO THE BLINKING LEDS DURING THE SCAN ACQUISITION. BUT THEY STATED ALSO THAT THE REFERENCE FRAME COULD HAVE MOVED BY MISTAKE. | H2: ADDITIONAL INFORMATION WAS SUBMITTED IN 3006544299-2024-01068 THAT WAS ADDED TO B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED INFORMATION FROM (B)(6) OF THE (B)(6) IN (B)(6) THAT A 66-YEAR-OLD, MALE PATIENT ON VENTILATION THERAPY USING A TRILOGY EVO VENTILATOR WAS FOUND DECEASED ON (B)(6) 2025, IN (B)(6). THE PATIENT WAS REPORTEDLY FOUND DECEASED WHILE STILL WEARING THE [FULL FACE] VENTILATOR MASK. HOWEVER, THE VENTILATOR WAS UNPLUGGED, AND THE INTERNAL BATTERY AND THE DETACHABLE BATTERY WERE BOTH FULLY DISCHARGED (DEAD). IT IS UNKNOWN AT THIS TIME IF THE DEVICE ALARMED AND WHAT THE DEVICE SETTINGS WERE. THE CAUSE OF DEATH IS ALSO UNKNOWN AT THIS TIME. THERE HAS BEEN NO ALLEGATION OF A DEVICE MALFUNCTION AT THIS TIME. DEVICE EVENT AND ALARM LOGS ARE BEING REQUESTED. RETURN OF THE DEVICE AND CIRCUIT HAVE BEEN REQUESTED FOR EVALUATION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION; THEREFORE, THE FUNCTIONALITY OF THE DEVICE IS NOT ABLE TO BE CONFIRMED AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION HAS BEEN COMPLETED.
ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | K141952 | IT WAS REPORTED THAT DURING AN UNKNOWN PROCEDURE, THE SURGEON WAS UNABLE TO DISENGAGE THE STAPLER. THERE WAS NO PATIENT CONSEQUENCE NOR SURGICAL DELAY REPORTED. ADDITIONAL INFORMATION REQUESTED. | (B)(4). DATE SENT: 2/28/2025. UPON REVIEW OF THE ADDITIONAL INFORMATION PROVIDED, IT WAS CONCLUDED THAT THIS EVENT IS A DUPLICATE OF THE EVENT REPORTED UNDER MDR#3005075853-2025-01040.
DXM 1040 MICROSCAN WALKAWAY | K911400 | THE CUSTOMER REPORTED GETTING HIGH PROBABILITY MISIDENTIFICATION (HPID) OF CITROBACTER FARMERI ON PATIENT SAMPLES USING NEG URINE COMBO 90 (NUC90, P/N C48640, LOT # 2025-11-11) PANELS PROCESSED IN THEIR DXM 1040 MICROSCAN WALKAWAY INSTRUMENT S/N (B)(6). THE CUSTOMER STATED THE MISIDENTIFICATION IMPACTED ONE PATIENT SAMPLE, MULTIPLE CULTURES. THE CUSTOMER NOTIFIED THE PHYSICIAN OF THE ERRONEOUS RESULTS OBTAINED FROM URINE SAMPLE. THE CUSTOMER STATED THE PATIENT WAS PUT ON HOSPICE AND PASSED AWAY ON (B)(6) 2025. IT WAS FURTHER NOTED THAT THE PATIENT DID NOT HAVE CAPACITY AND WAS REFUSING TREATMENT. THE PATIENT¿S FAMILY OPTED FOR COMFORT FOCUSED MEASURES. THE CUSTOMER INDICATED THEY REACHED OUT TO THEIR INTERNAL TEAM, AND THEY STATED THE MISIDENTIFICATION CAUSED A DELAY IN APPROPRIATE TREATMENT FOR THE PATIENT. | THE CAUSE OF DEATH IS UNKNOWN. THE INFORMATION AVAILABLE SUGGESTS THE MICROSCAN SYSTEM FAILED TO IDENTIFY THE ORGANISM CORRECTLY BASED ON THE ATYPICAL BEHAVIOR OF THE ISOLATE ITSELF AND NEGATIVE LYS REACTION, WHICH IS A KEY BIOCHEMICAL REACTION USED BY THE SYSTEM TO IDENTIFY AN E. COLI. AS INSTRUMENT AND PANEL QUALITY CONTROL TESTING WERE ALL IN RANGE, A PRODUCT PERFORMANCE ISSUE OR MALFUNCTION IS NOT INDICATED. PANEL DATA DID NOT INDICATE AN OBVIOUS INSTRUMENT OR PANEL ISSUE AND CONFIRMED THE ATYPICAL BEHAVIOR OF THE SAMPLE. THEREFORE, THE PROBABLE CAUSE OF THE REPORTED ORGANISM MISIDENTIFICATION WAS AN ATYPICAL ISOLATE. SECTION A4 AND A5: INFORMATION NOT PROVIDED BY THE CUSTOMER. THE BECKMAN COULTER INTERNAL IDENTIFIER IS (B)(4).
EQUISTREAM ALPHACURVE DL CATHETER 14.5FR, 28CM | K090101 | IT WAS REPORTED THAT WHILE THE PATIENT'S CAREGIVER WAS CHANGING THE DRESSING AROUND THE DIALYSIS CATHETER, THE TUBING WAS INADVERTENTLY CUT BENEATH ONE OF THE PORTS. IT WAS FURTHER REPORTED THAT THE PATIENT BEGAN BLEEDING OUT THROUGH THE CUT CATHETER END AND COLLAPSED. REPORTEDLY, THE CAREGIVER CALLED 911 AND THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM. THE BLEEDING COULD NOT BE STOPPED UNTIL THE TUBE WAS CLAMPED IN THE EMERGENCY ROOM. DESPITE ALL EFFORTS INCLUDING TRANSFUSION, THE PATIENT COULD NOT BE RESUSCITATED. THE PATIENT EXPIRED DUE TO EXSANGUINATION. | H11: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED CUT, FLUID LEAK, AND PATIENT DEATH, AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. FROM THE REPORT, THE EVENT IS STATED TO BE DUE TO INADVERTENT CUTTING OF THE CATHETER. THEREFORE, IT APPEARS THE MOST LIKELY ROOT CAUSE FOR THE REPORTED MATERIAL CUT, LEAKAGE, AND PATIENT DEATH IS SHARP INSTRUMENT DAMAGE DUE TO USER ERROR. IT IS UNKNOWN WHETHER PATIENT OR PRODUCT ISSUES CONTRIBUTED TO THE REPORTED EVENT. AT THIS TIME, A RELATIONSHIP BETWEEN THE CATHETER AND THE REPORTED MATERIAL CUT AND DEATH CANNOT BE CONFIRMED. LABELING REVIEW: A REVIEW OF PRODUCT LABELING DOCUMENTATION (E.G., PROCEDURAL INSTRUCTIONS, INDICATIONS, WARNINGS, PRECAUTIONS, CAUTIONS, POSSIBLE COMPLICATIONS, CONTRAINDICATIONS, NURSING GUIDE, AND UNIT LABEL) DID NOT FIND ANY PRODUCT LABELING INADEQUACY. D4 (UNIQUE IDENTIFIER (UDI) #), D5, E3, G3, H4, H6 (DEVICE, COMPONENT, METHOD). SECTION A THROUGH F: THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
GLIDEPATH | K051748 | IT WAS REPORTED THAT DURING A SCHEDULED HEMODIALYSIS TREATMENT THROUGH THE HEMODIALYSIS CENTRAL VENOUS CATHETER, THE NURSE IDENTIFIED A POOL OF BLOOD OF APPROXIMATELY 100-300ML UNDER THE PATIENT'S CHAIR DURING THE FIRST SEVERAL MINUTES OF TREATMENT INITIATION. IT WAS FURTHER REPORTED THAT THE BLOOD PUMP WAS STOPPED, AND IT WAS DISCOVERED THAT THE BLOOD LOSS WAS ORIGINATING FROM THE VENOUS BLOODLINE/VENOUS LUMEN CONNECTION SITE. REPORTEDLY, THE PATIENT'S CAROTID PULSE WAS PALPABLE, HEMODIALYSIS TREATMENT WAS TERMINATED, AND BLOOD WAS RETURNED USING ARTERIAL LUMEN OF THE CATHETER. THE PATIENT BECAME UNRESPONSIVE, THE EMERGENCY RESPONSE SERVICES WAS CALLED, AND CPR WAS INITIATED. UPON ARRIVAL, THE EMERGENCY RESPONSE SERVICES CONTINUED RESUSCITATIVE EFFORTS WHILE TRANSPORTING THE PATIENT TO THE HOSPITAL. THE PATIENT REPORTEDLY EXPIRED SHORTLY AFTER. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
KUGEL PATCH | K963141 | PER INFORMATION PROVIDED BY PATIENT'S ATTORNEY: (B)(6) 2003 - PATIENT WITH A COMPLICATED HISTORY OF MULTIPLE HERNIA REPAIRS WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF KUGEL PATCH (DEVICE #1). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS DISSECTED, AND ADHESIONS WERE TAKEN DOWN. A KUGEL PATCH (DEVICE #1) WAS PLACED AND SUTURED.¿ (B)(6) 2003 - PATIENT WAS DIAGNOSED WITH SEROMA. (B)(6) 2004 - PATIENT HAD CELLULITIS AROUND THE WOUND THAT HAD BEEN TREATED WITH ANTIBIOTICS WAS COMPLETELY GONE AWAY. (B)(6) 2004 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT REPAIR WITH IMPLANT OF COMPOSIX KUGEL (DEVICE #2). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS OPENED AND MULTIPLE FASCIAL DEFECTS IDENTIFIED. MULTIPLE BOWEL ADHESIONS FROM THE ANTERIOR ABDOMINAL WALL TO OLD GORE-TEX MESH AND OLD PACKS WAS FOUND. ONCE ALL THE ADHESIONS WERE CLEARED, A COMPOSIX KUGEL (DEVICE #2) WAS PLACED AND SUTURED.¿ (B)(6) 2005 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF SEPRAMESH IP (DEVICE #3). PER OPERATIVE NOTES, ¿THE BOWEL ADHESIONS TO THE OLD KUGEL MESHES (DEVICE #1 & #2) PLACED SUPERIORLY WERE TAKEN DOWN. THE HERNIA SAC WAS EXCISED AND A SEPRAMESH IP (DEVICE #3) WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH, VENTRAL HERNIA THEREBY UNDERWENT REMOVAL OF ALL ABDOMINAL WALL MESHES (DEVICE #1, #2, & #3) AND IMPLANT OF BIOLOGICAL MESH. PER OPERATIVE NOTES, ¿THERE WAS AN ABSCESS AND FISTULA NOTED. BOWEL INTERLOOP ADHESIONS FROM PREVIOUS ABDOMINAL WALL MESH WERE TAKEN DOWN. TH MESH ALONG WITH THE PROLENE SUTURES, OLD METAL TACKS WERE ISOLATED, AND HERNIAS REDUCED. ALL THE ABDOMINAL WALL MESHES (DEVICE #1, #2 & #3) WERE RESECTED AND A BIOLOGICAL MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH WITH COLONIC FISTULA THEREBY UNDERWENT REMOVAL OF BIOLOGICAL MESH AND IMPLANT OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿THE BIOLOGICAL MESH AND COLONIC FISTULA WERE REMOVED. SMALL BOWEL RESECTION AND ANASTOMOSIS WAS DONE. A SYNTHETIC MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH PNEUMONIA. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT REMOVAL OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿A LARGE AMOUNT OF SUCCUS COMING FROM THE INTRAABDOMINAL REGION. THERE WERE INTRA-ABDOMINAL LOOP ABSCESSES. THE END ILEOSTOMY BECAME A MUCOUS FISTULA WAS SUTURED. THE SYNTHETIC MESH WAS REMOVED AND DEBRIDED ABDOMINAL WALL, THE MUSCLE, SKIN, AND FASCIA THAT WAS NECROTIC. A TEMPORARY VAC DRESSING PLACED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN THEREBY UNDERWENT ABDOMEN WASHOUT, VAC PLACEMENT AND ABDOMINAL WALL DEBRIDEMENT. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN AND ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT OPEN REPAIR OF ABDOMINAL WOUND, ABDOMEN WASHOUT AND ENTEROCUTANEOUS FISTULA. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, PLUS RETRACTION OF ILEOSTOMY THEREBY UNDERWENT ABDOMINAL WASHOUT, ILEOSTOMY REVISION, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT ABDOMINAL WASHOUT, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTERO ATMOSPHERIC FISTULA THEREBY UNDERWENT WOUND VAC CHANGE. (B)(6) 2018 - PATIENT PASSED AWAY. PER DEATH SUMMARY, ¿76 Y/O FEMALE WITH INFECTED ABDOMINAL MESH WAS TAKEN TO THE OPERATING ROOM FOR EXPLORATORY LAPAROTOMY MESH REMOVAL COMPLEX ABDOMINAL WALL REPAIR, BOWEL RESECTIONS ON (B)(6) 2017 WAS TAKEN BACK TO THE OPERATING ROOM ON (B)(6) 2017 FOR INTRA-ABDOMINAL SEPSIS ON ENTEROCUTANEOUS FISTULA. PATIENT DEVELOPED ANOTHER FISTULA SUBSEQUENT TO THIS AND HAD MULTIPLE TRIPS TO THE OPERATING ROOM FOR ABDOMINAL WALL WASHOUTS AND ATTEMPTS AT ENTEROCUTANEOUS FISTULA REPAIR. PATIENT DID DEVELOP A PNEUMONIA AND WAS ON MULTIPLE ANTIBIOTICS FOR SEVERAL WEEKS WHILE IN THE HOSPITAL. PATIENT WAS UNABLE TO BE WEANED FROM THE VENTILATOR AND THE FAMILY ULTIMATELY MADE A DECISION TO TERMINALLY EXTUBATE THE PATIENT AND MAKE HER COMFORT MEASURES ONLY AFTER SEVERAL WEEKS OF BEING ON THE VENTILATOR. PATIENT WENT ALMOST 48 HOURS AFTER EXTUBATION AND DIED AT 12:36 PM ON (B)(6) 2018.¿ ATTORNEY ALLEGED THAT THE PATIENT HAD ADHESIONS, BOWEL OBSTRUCTION, DEATH, FISTULAE, INFECTION, PAIN AND HERNIA RECURRENCE. IT WAS ALSO ALLEGED THAT THE PATIENT HAD BULGE, WOUND VAC, TRANSVERSE COLOSTOMY, ABDOMINAL WALL DEBRIDEMENT, COLON REPAIR, SEPARATION OF ABDOMINAL WALL COMPONENTS, END-ILEOSTOMY, ENDOSOTOMY. | BASED ON THE INFORMATION RECEIVED, NO CONCLUSIONS CAN BE MADE. PER MEDICAL RECORDS REVIEW, ABOUT 2 YEARS POST IMPLANT OF SINGLE USE STERILE KUGEL PATCH, PATIENT WAS DIAGNOSED WITH FISTULA, ABSCESS, INFECTION, SEROMA, HERNIA RECURRENCE, CELLULITIS AND ADHESIONS THEREBY UNDERWENT REPAIR WITH REMOVAL OF MESH. PATIENT WAS PASSED AWAY ON (B)(6) 2018. NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. THE INSTRUCTIONS-FOR-USE SUPPLIED WITH THE DEVICE LISTS FISTULA, SEROMA, HERNIA RECURRENCE AND ADHESIONS AS POSSIBLE COMPLICATIONS. IN REGARDS TO INFECTION THE WARNING SECTION STATES, "IF AN INFECTION DEVELOPS, TREAT THE INFECTION AGGRESSIVELY. CONSIDERATION SHOULD BE GIVEN REGARDING THE NEED TO REMOVE THE MESH. AN UNRESOLVED INFECTION MAY REQUIRE REMOVAL OF THE DEVICE". NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. THIS EMDR REPRESENTS THE KUGEL PATCH (DEVICE #1). AN ADDITIONAL SUPPLEMENTAL EMDR WAS SUBMITTED TO REPRESENTS THE MESH ¿ COMPOSIX KUGEL (DEVICE #2) AND AN EMDR WAS SUBMITTED TO REPRESENT THE SEPRAMESH IP (DEVICE #3). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
SEPRAMESH IP | K063739 | PER INFORMATION PROVIDED BY PATIENT'S ATTORNEY: (B)(6) 2003 - PATIENT WITH A COMPLICATED HISTORY OF MULTIPLE HERNIA REPAIRS WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF KUGEL PATCH (DEVICE #1). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS DISSECTED, AND ADHESIONS WERE TAKEN DOWN. A KUGEL PATCH (DEVICE #1) WAS PLACED AND SUTURED.¿ (B)(6) 2003 - PATIENT WAS DIAGNOSED WITH SEROMA. (B)(6) 2004 - PATIENT HAD CELLULITIS AROUND THE WOUND THAT HAD BEEN TREATED WITH ANTIBIOTICS WAS COMPLETELY GONE AWAY. (B)(6) 2004 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT REPAIR WITH IMPLANT OF COMPOSIX KUGEL (DEVICE #2). PER OPERATIVE NOTES, ¿THE HERNIA SAC WAS OPENED AND MULTIPLE FASCIAL DEFECTS IDENTIFIED. MULTIPLE BOWEL ADHESIONS FROM THE ANTERIOR ABDOMINAL WALL TO OLD GORE-TEX MESH AND OLD PACKS WAS FOUND. ONCE ALL THE ADHESIONS WERE CLEARED, A COMPOSIX KUGEL (DEVICE #2) WAS PLACED AND SUTURED.¿ (B)(6) 2005 - PATIENT WAS DIAGNOSED WITH RECURRENT INCISIONAL HERNIA THEREBY UNDERWENT OPEN REPAIR WITH THE IMPLANT OF SEPRAMESH IP (DEVICE #3). PER OPERATIVE NOTES, ¿THE BOWEL ADHESIONS TO THE OLD KUGEL MESHES (DEVICE #1 & #2) PLACED SUPERIORLY WERE TAKEN DOWN. THE HERNIA SAC WAS EXCISED AND A SEPRAMESH IP (DEVICE #3) WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH, VENTRAL HERNIA THEREBY UNDERWENT REMOVAL OF ALL ABDOMINAL WALL MESHES (DEVICE #1, #2, & #3) AND IMPLANT OF BIOLOGICAL MESH. PER OPERATIVE NOTES, ¿THERE WAS AN ABSCESS AND FISTULA NOTED. BOWEL INTERLOOP ADHESIONS FROM PREVIOUS ABDOMINAL WALL MESH WERE TAKEN DOWN. TH MESH ALONG WITH THE PROLENE SUTURES, OLD METAL TACKS WERE ISOLATED, AND HERNIAS REDUCED. ALL THE ABDOMINAL WALL MESHES (DEVICE #1, #2 & #3) WERE RESECTED AND A BIOLOGICAL MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH INFECTED ABDOMINAL WALL MESH WITH COLONIC FISTULA THEREBY UNDERWENT REMOVAL OF BIOLOGICAL MESH AND IMPLANT OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿THE BIOLOGICAL MESH AND COLONIC FISTULA WERE REMOVED. SMALL BOWEL RESECTION AND ANASTOMOSIS WAS DONE. A SYNTHETIC MESH WAS PLACED AND SUTURED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH PNEUMONIA. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT REMOVAL OF SYNTHETIC MESH. PER OPERATIVE NOTES, ¿A LARGE AMOUNT OF SUCCUS COMING FROM THE INTRAABDOMINAL REGION. THERE WERE INTRA-ABDOMINAL LOOP ABSCESSES. THE END ILEOSTOMY BECAME A MUCOUS FISTULA WAS SUTURED. THE SYNTHETIC MESH WAS REMOVED AND DEBRIDED ABDOMINAL WALL, THE MUSCLE, SKIN, AND FASCIA THAT WAS NECROTIC. A TEMPORARY VAC DRESSING PLACED.¿ (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN THEREBY UNDERWENT ABDOMEN WASHOUT, VAC PLACEMENT AND ABDOMINAL WALL DEBRIDEMENT. (B)(6) 2017 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN AND ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT OPEN REPAIR OF ABDOMINAL WOUND, ABDOMEN WASHOUT AND ENTEROCUTANEOUS FISTULA. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, PLUS RETRACTION OF ILEOSTOMY THEREBY UNDERWENT ABDOMINAL WASHOUT, ILEOSTOMY REVISION, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTEROCUTANEOUS FISTULA THEREBY UNDERWENT ABDOMINAL WASHOUT, WOUND VAC CHANGE. (B)(6) 2018 - PATIENT WAS DIAGNOSED WITH OPEN ABDOMEN, ENTERO ATMOSPHERIC FISTULA THEREBY UNDERWENT WOUND VAC CHANGE. (B)(6) 2018 - PATIENT PASSED AWAY. PER DEATH SUMMARY, ¿76 Y/O FEMALE WITH INFECTED ABDOMINAL MESH WAS TAKEN TO THE OPERATING ROOM FOR EXPLORATORY LAPAROTOMY MESH REMOVAL COMPLEX ABDOMINAL WALL REPAIR, BOWEL RESECTIONS ON (B)(6) 2017 WAS TAKEN BACK TO THE OPERATING ROOM ON (B)(6) 2017 FOR INTRA-ABDOMINAL SEPSIS ON ENTEROCUTANEOUS FISTULA. PATIENT DEVELOPED ANOTHER FISTULA SUBSEQUENT TO THIS AND HAD MULTIPLE TRIPS TO THE OPERATING ROOM FOR ABDOMINAL WALL WASHOUTS AND ATTEMPTS AT ENTEROCUTANEOUS FISTULA REPAIR. PATIENT DID DEVELOP A PNEUMONIA AND WAS ON MULTIPLE ANTIBIOTICS FOR SEVERAL WEEKS WHILE IN THE HOSPITAL. PATIENT WAS UNABLE TO BE WEANED FROM THE VENTILATOR AND THE FAMILY ULTIMATELY MADE A DECISION TO TERMINALLY EXTUBATE THE PATIENT AND MAKE HER COMFORT MEASURES ONLY AFTER SEVERAL WEEKS OF BEING ON THE VENTILATOR. PATIENT WENT ALMOST 48 HOURS AFTER EXTUBATION AND DIED AT 12:36 PM ON (B)(6) 2018.¿ ATTORNEY ALLEGED THAT THE PATIENT HAD ADHESIONS, BOWEL OBSTRUCTION, DEATH, FISTULAE, INFECTION, PAIN AND HERNIA RECURRENCE. IT WAS ALSO ALLEGED THAT THE PATIENT HAD BULGE, WOUND VAC, TRANSVERSE COLOSTOMY, ABDOMINAL WALL DEBRIDEMENT, COLON REPAIR, SEPARATION OF ABDOMINAL WALL COMPONENTS, END-ILEOSTOMY, END OSOTOMY. | BASED ON THE INFORMATION RECEIVED, NO CONCLUSIONS CAN BE MADE. PER MEDICAL RECORDS REVIEW, ABOUT 12 YEARS POST IMPLANT OF SINGLE USE STERILE SEPRAMESH IP, PATIENT WAS DIAGNOSED WITH FISTULA, ABSCESS, INFECTION AND ADHESIONS THEREBY UNDERWENT REPAIR WITH REMOVAL OF MESH. PATIENT WAS PASSED AWAY ON (B)(6) 2018. NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. THE INSTRUCTIONS-FOR-USE SUPPLIED WITH THE DEVICE LISTS FISTULA AND ADHESIONS AS POSSIBLE COMPLICATIONS. IN REGARDS TO INFECTION THE WARNING SECTION STATES, "IF AN INFECTION DEVELOPS, TREAT THE INFECTION AGGRESSIVELY. CONSIDERATION SHOULD BE GIVEN REGARDING THE NEED TO REMOVE THE PROSTHESIS. AN UNRESOLVED INFECTION MAY REQUIRE REMOVAL OF THE PROSTHESIS". NO LOT NUMBER HAS BEEN PROVIDED; THEREFORE, A REVIEW OF THE MANUFACTURING RECORDS IS NOT POSSIBLE. THIS EMDR REPRESENTS THE SEPRAMESH IP (DEVICE #3). AN ADDITIONAL SUPPLEMENTAL EMDR WAS SUBMITTED TO REPRESENTS THE MESH ¿ COMPOSIX KUGEL (DEVICE #2) AND AN EMDR WAS SUBMITTED TO REPRESENT THE KUGEL PATCH (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
SWAN-GANZ CCOMBO V | K160084 | EDWARDS LIFESCIENCES WAS NOTIFIED VIA A DIRECT VOLUNTARY INCIDENT REPORT SUBMISSION TO THE FDA THAT A CUSTOMER EXPERIENCED RESISTANCE WITH REMOVAL OF A SWAN GANZ CATHETER. PATIENT WAS REPOSITIONED WHICH ALLOWED THE CATHETER TO BE REMOVED SMOOTHLY. AFTER REMOVAL, PATIENT'S BLOOD PRESSURE SUDDENLY DROPPED AND CHEST TUBE NOTED TO HAVE LARGELY INCREASED SANGUINEOUS OUTPUT. PATIENT WAS AWAKE THROUGHOUT THE EVENT. PLANS WERE MADE TO GO TO THE OPERATING ROOM. THERE WAS APPROXIMATELY A 3 X 2 CM DEFECT IN THE ANTERIOR WALL OF THE SVC, JUST SUPERIOR TO THE CAVOATRIAL JUNCTION. THE INJURY ENCOMPASSED ABOUT 40% OF THE CIRCUMFERENCE OF THE SUPERIOR VENA CAVA (SVC). THE SVC HAD BEEN CLAMPED FOR A FEW MINUTES AND THERE WAS EVIDENCE OF VENOUS CONGESTION IN THE INNOMINATE VEIN. PATCH REPAIR OF LARGE SUPERIOR VENA CAVA LACERATION WAS DONE. PATIENT PROVED TO HAVE TOLERATED THE PROCEDURE SATISFACTORILY. PATIENT WAS THEN TRANSFERRED TO INTENSIVE CARE UNIT. IT WAS SUSPECTED THAT THE SWAN HAD BEEN LOOPED BY THE SUTURE AT TIME OF OPERATION EITHER IN THE SVC OR RIGHT ATRIUM AS THIS WOULD LEAVE A SMALL DEFECT IN THE WALL. PATIENT EXPERIENCED HEMORRHAGE/BLEEDING, LOW BLOOD PRESSURE/ HYPOTENSION, LACERATIONS, AND DEATH. PATIENT WAS UNDERGOING REMOVAL OF MITRAL VALVE NODULE, A REDO AORTIC VALVE REPLACEMENT, AND TRANSESOPHAGEAL ECHOCARDIOGRAM CENTRAL LINE INSERTION WHEN ISSUE OCCURRED. THE VOLUNTARY INCIDENT REPORT DID NOT INCLUDE CUSTOMER CONTACT INFORMATION, PATIENT INFORMATION OR PRODUCT INFORMATION RELATED TO THE REPORTED EVENT. | EDWARDS LIFESCIENCES WAS UNABLE TO PERFORM DUE DILIGENCE EFFORTS TO RETRIEVE THE SUSPECT DEVICE FOR ANALYSIS AS THERE WAS NO CUSTOMER CONTACT INFORMATION PROVIDED IN THE VOLUNTARY INCIDENT REPORT. THE LOT NUMBER WAS NOT PROVIDED THUS A DEVICE HISTORY RECORD WAS NOT REVIEWED. NO PRODUCT WAS RETURNED FOR EVALUATION; THEREFORE, A DEFINITIVE ROOT CAUSE COULD NOT BE DETERMINED, AND CORRECTIVE OR PREVENTIVE ACTION IS NOT INDICATED AT THIS TIME. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS A PART OF THE MONTHLY REVIEW.
SPACELABS ULTRAVIEW SL COMMAND MODULE | K103142 | THE CUSTOMER REPORTED THAT A CRITICALLY ILL PATIENT THAT OVERDOSED ON CALCIUM CHANNEL BLOCKERS AND A TRANSCUTANEOUS PACER WAS ADMITTED TO A BEDSIDE MONITOR. THE PACER DETECTION OF THE MODULE WAS TURNED ON, HOWEVER, THERE WAS STILL SIGNIFICANT ARTIFACT SEEN ON THE MODULE. A TRANSVENOUS WIRE WAS PLACED AND THE CUSTOMER STATED THAT THE INTERFERENCE IMPACTED THE CLINICAL STAFF'S ABILITY TO CONFIRM PACER FUNCTIONALITY. | THE CUSTOMER BIOMED TESTED THE DEVICE AFTER THE EVENT AND COULD NOT FIND ANY ISSUES WITH THE PACER DETECTION OF THE COMMAND MODULE. A DATABASE INCLUDING THE PATIENT WAVEFORM WAS REQUESTED TO EVALUATE THE ARTIFACT PRESENT AT THE TIME OF THE REPORTED EVENT. AT THIS TIME THE DATABASE HAS NOT BEEN RECEIVED. A FOLLOW UP REPORT WILL BE SUBMITTED IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | K141952 | PER USER FACILITY MEDWATCH FORM, #MW5165696, A PATIENT HAD SCHEDULED PROCEDURE FOR LEFT VIDEO ASSISTED THORACIC SURGERY WITH WEDGE RESECTION, POSSIBLE THORACOTOMY WITH LOBECTOMY, LEFT THORACIC NODE DISSECTION. SURGEON NOTED MALFUNCTION WITH STAPLER DEVICE, EXCESSIVE BLEEDING OCCURRED, AND PATIENT EXSANGUINATED IN THE OPERATING ROOM. DEVICE IS AVAILABLE FOR RETURN. | (B)(4). DATE SENT: 2/20/2025. D4 BATCH#: UNK. #MW5165696. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. NO LOT OR BATCH NUMBER WAS PROVIDED THEREFORE A DEVICE HISTORY COULD NOT BE DONE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT WERE THE INDICATIONS FOR SURGERY? WHAT WAS THE SURGICAL PROCEDURE? WHAT ANATOMICAL STRUCTURE/VESSEL WAS THE DEVICE BEING USED ON WHEN THE BLEEDING OCCURRED? WHAT WAS THE APPROXIMATE COMPRESSED WIDTH OF THE VESSEL THAT THE DEVICE WAS USED ON? CAN MORE DETAIL BE PROVIDED ON THE MALFUNCTION OF DEVICE? WHAT WAS THE APPROXIMATE AMOUNT OF BLOOD LOSS? WHAT BLOOD PRODUCTS WERE ADMINISTERED? HOW WAS THE BLEEDING ADDRESSED? WHAT IS THE SURGEON¿S EXPERIENCE WITH THE PVS DEVICE? WHAT WAS THE APPROXIMATE WIDTH OF THE COMPRESSED VESSEL? DID THE SURGEON WAIT 15 SECONDS PRIOR TO FIRING? DID THE SURGEON ENSURE THAT THE COMPLETE WIDTH OF THE COMPRESSED VESSEL WAS PROXIMAL TO THE CUT LINE ON THE DEVICE? DID THE SURGEON CONFIRM THAT NO EXTRANEOUS TISSUE WAS DISTAL TO THE CUT LINE WITHIN THE JAWS? WERE ANY SURGICAL CLIPS USED IN THE AREA OF THE DEVICE FIRING? WERE THERE ANY DIFFICULTIES WITH THE DEVICE OR STAPLE FORMATION DURING THE INITIAL PROCEDURE? HOW WAS THE POST-OP BLEEDING IDENTIFIED? HOW MANY DAYS POSTOPERATIVE WAS THE BLEEDING IDENTIFIED? WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? WAS A TRANSFUSION REQUIRED? WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? IS THERE A RECORDING AVAILABLE OF THE PROCEDURE? WHAT IS THE PATIENT¿S CURRENT STATUS? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
POWERED 60 ECHELON +, 340MM SHAFT | K110385 | IT WAS REPORTED THAT DURING A LEFT ATRIAL APPENDAGE CLOSURE, THE PATIENT DIED WHEN ECHELON WAS USED. THERE IS A POSSIBILITY THAT OTHER TISSUES WERE INVOLVED AND PROBABLY THERE WAS A PROBLEM WITH THE PROCEDURE NOT BECAUSE OF ECHELON. THE ISSUE HAS OCCURRED BUT THE DETAILS ARE UNKNOWN AT THIS TIME, SO WILL BE ADDED AS ADDITIONAL INFORMATION BECOMES AVAILABLE. NO FURTHER INFORMATION IS AVAILABLE. | (B)(4). B3: UNKNOWN: CAPTURED AS AWARENESS DATE. DATE SENT 2/20/2025. D4 BATCH#: UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. INVESTIGATION SUMMARY: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT OR PRODUCT ID NUMBERS ARE RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: WHAT COLOR CARTRIDGE WAS USED? WERE THERE ANY DIFFICULTIES WITH DEVICE DURING SURGERY CAN A TIMELINE OF EVENTS BE PROVIDED? WHAT WAS THE CAUSE OF THE PATIENT DEATH? WAS AN AUTOPSY PERFORMED, AND CAN THE REPORT BE PROVIDED? DOES THE SURGEON BELIEVE THE PATIENT¿S DEATH WAS RELATED IN ANY WAY TO AN ALLEGED DEFICIENCY WITH THE ETHICON DEVICE? IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING PERSONNEL? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY AT A NURSING HOME FACILITY WHILE CONNECTED TO THE PD CLINIC¿S LIBERTY SELECT CYCLER. THE DEATH WAS REPORTEDLY NOT RELATED TO PD THERAPY. MULTIPLE ATTEMPTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION, AND THUS FAR NO FURTHER DETAILS HAVE BEEN PROVIDED. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS (PD) UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT DEATH. HOWEVER, THERE WAS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE CYCLER. ADDITIONALLY, THERE WAS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THE EVENT. THE INFORMATION INITIALLY REPORTED STATED THAT ALTHOUGH THE PATIENT WAS CONNECTED TO THE CYCLER AT THE TIME OF DEATH, THE DEATH WAS NOT PD RELATED. ALTHOUGH MAINTENANCE DIALYSIS PREVENTS DEATH FROM UREMIA, MORTALITY AMONG PATIENTS WITH END-STAGE KIDNEY DISEASE (ESKD) REMAINS HIGH. AMONG PATIENTS WITH ESKD, CARDIOVASCULAR DISEASE ACCOUNTS FOR APPROXIMATELY 50 PERCENT OF DEATHS. BASED ON THE LIMITED INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
TA | K231491 | THE PRODUCT WAS USED FOR A SMALL BOWEL RESECTION WITH ANASTOMOSIS. DURING SUBSEQUENT SURGERY IT WAS DISCOVERED THAT THE FREE END OF THE SMALL BOWEL STAPLE LINE HAD LEAKAGE ALONG THE STAPLE LINE INDICATIVE OF A FAILED STAPLE LINE AND THE ANASTOMOSIS DID NOT SEAL. IT WAS REPORTED THAT AFTER USE, THE PATIENT EXPERIENCED SURGICAL STAPLER FAILURE, FAILURE TO THE STAPLER DEVICE, BOWEL LEAKAGE, DEFECTIVE DEVICE, PAIN, SHOCK, MENTAL ANGUISH, UNABLE TO PERFORM NORMAL ACTIVITIES, LEUKOCYTOSIS, ASCITES, SEROSITIS, AND DEATH. POST-OPERATIVE PATIENT TREATMENT INCLUDED EXPLORATORY LAPAROTOMY WITH LYSIS OF ADHESIONS, SMALL BOWEL RESECTION WITH PRIMARY ANASTOMOSIS, DIVERTING LOOP ILEOSTOMY, & CT SCAN. INFORMATION RECEIVED INDICATES THE PATIENT IS DECEASED. NO INFORMATION WAS PROVIDED REGARDING THE CIRCUMSTANCES OF EXPIRATION. | H6 IME E2402: LEUKOCYTOSIS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
GIA | K221013 | ACCORDING TO THE REPORTER, DURING AN OPEN SMALL BOWEL ANASTOMOSIS, THE SURGEONS WERE REPAIRING THE FAILED ANASTOMOSIS AND THE FIRST FIRING WAS SATISFACTORY TO THE SURGEONS BUT THE SECOND FIRING HAD A HOLE AND DID NOT SEAL. A TEMPORARY STAPLE LINE WAS CREATED USING AN ULTRA UNIVERSAL STAPLER FOR THE INTRALUMENAL ANASTOMOSIS AND THEN CLOSED THE ANASTOMOSIS WITH A LINEAR STAPLER AND THUS THIS STAPLE LINE WAS CUT OFF AS A NORMAL PART OF CLOSING THE COMMON OPENING WITH THE LINEAR STAPLER. IT WAS NOTED THAT THERE WAS TISSUE DAMAGE. | ADDITIONAL INFO: B5, G1, H6 (REMOVED AND ADDED PATIENT CODES, DEVICE CODES), CHANGED ADDITIONAL CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
PARKER | K984528 | PATIENT WAS SAT UP IN BED FOR A CHEST X-RAY. ET TUBE HAD BEEN IN PLACE SINCE (B)(6) 2025. SUDDEN CHANGE IN OXYGENATION AND VENTILATION NOTED. OBVIOUS CUFF LEAK. RN AND RT INJECTED 3 CC OF AIR WITH IMPROVEMENT AND THEN RAPID SOFTENING OF THE PILOT BALLOON. RECURRENCE OF OBVIOUS CUFF LEAK. UNABLE TO MAINTAIN CUFF PRESSURE. TUBE EMERGENTLY REPLACED. AFTER REINTUBATION, CARDIAC DYSFUNCTION WITH LOSS OF PULSE. CODE CALLED. PILOT LINE WAS CUT DURING EXTUBATION AND REINTUBATION. POST PROCEDURE EXAMINATION OF THE BALLOON SHOWED BALLOON PARTIALLY INFLATED. UTILIZING NEEDLE AND SYRINGE STAFF UNABLE TO INFLATE CUFF. PRESSURE TO CUFF ON THE OUTSIDE OF THE CUFF DOES NOT CREATE A LEAK. UNABLE TO TEST PILOT LINE OR INJECTION PORT. | H6: 4756 (APPROPRIATE IMPACT TERM/CODE NOT AVAILABLE): ENDOTRACEAL TUBE.THE PRODUCT INVOLVED IN THE REPORT HAS NOT BEEN RETURNED. A REVIEW OF THE DEVICE HISTORY RECORD IS IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 20 FEB 2025 HAS BEEN INCLUDED IN THE HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AIRLIFE. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AIRLIFE. AIRLIFE HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AIRLIFE COMPLAINT DATABASE AND IS IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT A AIRLIFE PRODUCT IS DEFECTIVE OR CAUSES SERIOUS INJURY. COMPLAINT WAS CONSIDERED CONFIRMED BASED ON CUSTOMER NARRATIVE OF EVENTS. A 24 MONTH REVIEW WAS CONDUCTED AND 49 ADDITIONAL COMPLAINTS WERE IDENTIFIED RELATED TO THIS ISSUE HOWEVER NO TREND WAS OBSERVED. THE PART AND LOT NUMBER WERE NOT PROVIDED, HOWEVER THE VENDOR WAS NOTIFIED OF THE INCIDENT. THIS INCIDENT WAS NOT IDENTIFIED AS AN INCREASING TREND. NO FURTHER ACTION TO BE TAKEN.
ULTRASONIC PROBE | K982323 | IT WAS REPORTED THAT DURING AN ENDOBRONCHIAL ULTRASOUND-GUIDED TRANSBRONCHIAL NEEDLE ASPIRATION PROCEDURE FOR SUSPECTED PERIPHERAL LUNG CANCER, THE ULTRASONIC PROBE WAS USED IN COMBINATION WITH THE SINGLE-USE GUIDE SHEATH. UPON REMOVAL FROM THE PATIENT, IT WAS OBSERVED THAT APPROXIMATELY 5CM OF THE PROBE'S COATING HAD PEELED OFF. CT AND FLUOROSCOPY WERE IMMEDIATELY PERFORMED TO CHECK FOR RETAINED FRAGMENTS, AND NONE WERE DETECTED. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS DISCHARGED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT'S CONDITION SUDDENLY DETERIORATED, REQUIRING EMERGENCY INTERVENTION. THE PATIENT WAS TRANSFERRED TO THE HOSPITAL WHERE DEATH WAS CONFIRMED THE SAME DAY. THE CAUSE OF DEATH WAS DOCUMENTED AS UNKNOWN, AND NO ADDITIONAL INFORMATION REGARDING AUTOPSY FINDINGS OR CONTRIBUTING FACTORS WAS AVAILABLE. THERE WERE NO REPORTS OF PROCEDURAL COMPLICATIONS, IMMEDIATE POST-PROCEDURAL ADVERSE EVENTS, OR CONFIRMED DEVICE-RELATED INJURY. | THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THE DEVICE WAS RETURNED TO OLYMPUS FOR INSPECTION AND THE REPORTED FAILURE WAS CONFIRMED. IN ADDITION, THE FOLLOWING REPORTABLE MALFUNCTIONS WERE IDENTIFIED DURING THE DEVICE EVALUATION: THERE IS LEAKAGE OF ULTRASONIC MEDIA BASED ON THE RESULTS OF THE INVESTIGATION, IT IS LIKELY THE FOLLOWING LED TO THE MALFUNCTION: THE ELONGATION OF THE MATERIAL AT THE TIP AND THE FACT THAT THE FRACTURE CROSS-SECTION LOOKS LIKE IT HAS BEEN TORN OFF, RATHER THAN CUT WITH A BLADE, INDICATE THAT STRESS WAS APPLIED TO THE TIP SHEATH IN A PULLING MANNER. IN ADDITION, THERE IS NO TWISTING OR CRUSHING OF THE INSERTION SHEATH, INDICATING THAT THE STRESS WAS NOT APPLIED TO THE TIP SHEATH WHILE THE INTERNAL BLADE (FLEXIBLE SHAFT) WAS BEING DRIVEN. THEREFORE, IT IS ASSUMED THAT THE TIP SHEATH WAS DAMAGED AND FELL OFF DUE TO THE STRESS APPLIED TO THE TIP SHEATH WHILE THE INTERNAL BLADE (FLEXIBLE SHAFT) WAS NOT BEING DRIVEN. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
SINGLE USE GUIDE SHEATH KIT | K060243 | IT WAS REPORTED THAT DURING AN ENDOBRONCHIAL ULTRASOUND-GUIDED TRANSBRONCHIAL NEEDLE ASPIRATION PROCEDURE FOR SUSPECTED PERIPHERAL LUNG CANCER, THE ULTRASONIC PROBE WAS USED IN COMBINATION WITH THE SINGLE-USE GUIDE SHEATH. UPON REMOVAL FROM THE PATIENT, IT WAS OBSERVED THAT APPROXIMATELY 5CM OF THE PROBE'S COATING HAD PEELED OFF. CT AND FLUOROSCOPY WERE IMMEDIATELY PERFORMED TO CHECK FOR RETAINED FRAGMENTS, AND NONE WERE DETECTED. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS DISCHARGED ON (B)(6) 2025. ON (B)(6) 2025, THE PATIENT'S CONDITION SUDDENLY DETERIORATED, REQUIRING EMERGENCY INTERVENTION. THE PATIENT WAS TRANSFERRED TO THE HOSPITAL WHERE DEATH WAS CONFIRMED THE SAME DAY. THE CAUSE OF DEATH WAS DOCUMENTED AS UNKNOWN, AND NO ADDITIONAL INFORMATION REGARDING AUTOPSY FINDINGS OR CONTRIBUTING FACTORS WAS AVAILABLE. THERE WERE NO REPORTS OF PROCEDURAL COMPLICATIONS, IMMEDIATE POST-PROCEDURAL ADVERSE EVENTS, OR CONFIRMED DEVICE-RELATED INJURY. THERE WAS NO REPORTED MALFUNCTION OR ISSUE WITH THE SINGLE-USE GUIDE SHEATH. | THIS REPORT IS RELATED TO THE FOLLOWING LINKED PATIENT IDENTIFIERS: (B)(6). THERE WAS NO ALLEGED DEVICE FAILURE. BASED ON THE RESULTS OF THE INVESTIGATION, AND SINCE THE DEVICE WAS NOT RETURNED FOR EVALUATION, THE DEFINITIVE ROOT CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
PROVOX2 | K971244 | THIS IS A CASE REPORT DESCRIBED IN THE LITERATURE PUBLICATION: YOHAN SONG, KRISTA KIYOSAKI, EDWARD J. DAMROSE, EPIDURAL ABSCESS AND PARAPLEGIA: DELAYED SEQUELA OF TRACHEOESOPHAGEAL PUNCTURE, OTOLARYNGOLOGY CASE REPORTS, VOLUME 5, 2017, PAGES 21-23, ISSN 2468-5488, HTTPS://DOI.ORG/10.1016/J.XOCR.2017.10.002. (HTTPS://WWW.SCIENCEDIRECT.COM/SCIENCE/ARTICLE/PII/S2468548817301388). THIS CASE REPORT WAS IDENTIFIED DURING A LITERATURE SEARCH FOR VOICE PROSTHESES CONDUCTED IN ACCORDANCE WITH THE PERIODIC UPDATE OF THE CLINICAL EVALUATION REPORT. THE CASE DESCRIBES A 78-YEAR OLD FEMALE WHO HAD TOTAL LARYNGECTOMY AND PRIMARY PLACEMENT OF A 10 MM PROVOX2 IN OCTOBER 2016. FOUR MONTHS LATER SHE WAS BROUGHT TO THE EMERGENCY DEPARTMENT WITH FOR A TWO-DAY HISTORY OF PROGRESSIVE BILATERAL LOWER EXTREMITY WEAKNESS AND SUBSEQUENT FALL. MR IMAGING OF THE CERVICAL AND THORACIC SPINE SHOWED AN EPIDURAL ABSCESS. THE IMAGING WAS HIGHLY SUSPICIOUS FOR EROSION OF THE VOICE PROSTHESIS THROUGH THE POSTERIOR ESOPHAGUS AT THE LEVEL OF THE ABSCESS. SHE HAD AN OPERATION TO REMOVE THE ABSCESS. POST-OPERATIVELY, THERE WAS NO OBVIOUS PROJECTION OF THE DEVICE THROUGH THE POSTERIOR ESOPHAGUS, AS WAS EXPECTED BY THE IMAGING. THE VOICE PROSTHESIS WAS REMOVED AND EXCHANGED FOR A 4.5MM DEVICE. APPROXIMATELY SEVEN DAYS POSTOPERATIVELY, THE PATIENT DEVELOPED WORSENING DYSPHAGIA. MRI OF THE CERVICAL AND THORACIC SPINE AGAIN SHOWED AN ABSCESS AS WELL AS OSTEOMYELITIS OF THE VERTEBRAL BODIES. GIVEN THE PERSISTENT INFECTION, THE PROGRESSION TO OSTEOMYELITIS, HER IRREVERSIBLE NEUROLOGICAL DEFICITS, THE DIFFICULTY AND UNLIKELY SUCCESS OF FURTHER SURGICAL INTERVENTIONS, THE DECISION WAS MADE TO TRANSITION THE PATIENT TO HOSPICE. THE PATIENT PASSED AWAY FROM SEPSIS SEVEN WEEKS LATER. IN THE DISCUSSION OF THE ARTICLE, THE AUTHORS HYPOTHESIZE THAT HER 10 MM TEP PROSTHESIS MAY HAVE BEEN TOO LONG. FOR THIS REASON, IT WAS EXCHANGED TO A SHORTER 4.5 MM ONE DURING HER HOSPITALIZATION. THE PATIENT ALSO WORE A FENESTRATED LARYNGECTOMY TUBE WITH A HEAT AND MOISTURE EXCHANGER WITHIN THE STOMA (PROVOX LARYTUBE, ATOS MEDICAL, NEW BERLIN, WI, USA). DIGITAL PRESSURE TRANSMITTED FROM THE LARYNGECTOMY TUBE TO THE PROSTHESIS DURING PHONATION MAY HAVE HELPED TO DRIVE THE VOICE PROSTHESIS THROUGH THE POSTERIOR ESOPHAGEAL WALL. PLEASE NOTE THAT THIS IS JUST A SUMMARY OF THE CASE. FOR A FULL CASE DESCRIPTION, PLEASE SEE REFERENCED ARTICLE. | CONCLUSION PRODUCT DEFECT COULD NOT BE CONFIRMED. IT IS ASSESSED THAT THE COMPLICATIONS WERE DUE TO USAGE. THE PATIENT WAS UNDER SUPERVISION. INVESTIGATION THIS IS A THEORETICAL INVESTIGATION; SAMPLES WERE NOT AVAILABLE AND WAS DERIVED FROM A CASE REPORT IN LITERATURE STUDY. THE PATIENT USED A PROVOX2 10 MM VOICE PROSTHESIS FOR FOUR MONTHS. THE PATIENT WAS WEAKENED AND FELL, LEADING TO DISCOVERY OF ABSCESS IN SPINAL CORE. THE VOICE PROSTHESIS WAS REPLACED WITH A 4,5MM TO TRY TO OPTIMIZE FIT. THE PATIENT PASSED AWAY SOME WEEKS LATER, NOT BEING ABLE TO HEAL FROM INFECTION. THE PATIENT ALSO USED A LARYTUBE. THE CASE REPORT SUGGESTS A HYPOTHESIS THAT THE 10MM PROSTHESIS WAS TOO LONG, EXCESSIVE DIGITAL PRESSURE APPLIED TO THE HEAT-MOISTURE EXCHANGER (HME) AND LARYTUBE CAUSED THE PROSTHESIS TO MOVE WITHIN THE PUNCTURE (BACKWARDS AND FORWARDS) AND PERFORATE AND ERODE THE POSTERIOR ESOPHAGEAL WALL. THE DAMAGE TO THE ESOPHAGEAL WALL COULD DERIVE FROM TRACHEO-ESOPHAGEAL PROTHESIS INSERTION ACCORDING TO THE CASE REPORT. THERE WAS NO SIGN OF DAMAGE TO THE ESOPHAGEAL WALL AFTER FIRST SURGERY. FOLLOW UP QUESTIONS ARE SENT TO THE AUTHORS/HOSPITAL: - DOES THE HOSPITAL ROUTINELY REVIEW THE SIZE OF THE VOICE PROSTHESIS AT EACH REPLACEMENT OR REPLACE LIKE FOR LIKE ? ANSWER: YES, WE ROUTINELY REVIEW THE SIZE OF THE PROSTHESIS AT TIME OF REPLACEMENT. THE PHYSICIAN COULD SEE THE DEVICE IN THE TRACHEA AND IT APPEARED TO BE FLUSHED WITH THE TRACHEAL WALL. PHYSICIAN DID NOT PERCEIVE THAT THE DEVICE WAS BEING PUSHED THROUGH THE POSTERIOR ESOPHAGEAL WALL. - WAS THERE ANYONE FROM ATOS MEDICAL PRESENT WITH TROUBLE SHOOTING? EDUCATIONAL TEAM? ANSWER: NO ONE FROM ATOS WAS INVOLVED, WE DID NOT CONSULT. WE DO NOT ROUTINELY DO THIS. - WAS THE RESPONSIBLE CLINICIAN TRAINED WITH THE ATOS MEDICAL PRODUCTS? ANSWER: YES, THE RESPONSIBLE PHYSICIAN (ME) IS TRAINED WITH ATOS PRODUCTS. - WAS THERE FOLLOW UP MEETINGS WITH THE PATIENT AFTER FIRST SURGERY FOR TRAINING/ADVISING THE PATIENT? ANSWER: YES, THERE WERE MANY MEETINGS WITH THE PATIENT. ADDITIONAL FROM AUTHOR: ¿THE REASON WE THINK FOR THE COMPLICATION WAS THAT THE PATIENT WAS WEARING A FENESTRATED LARYNGECTOMY TUBE WHICH WAS APPLYING PRESSURE ONTO THE ANTERIOR SURFACE OF THE VOICE PROSTHESIS. AS SPEECH BECAME MORE DIFFICULT SHE WOULD PUSH HARDER, DRIVING THE DEVICE FURTHER BACK INTO THE ESOPHAGEAL WALL.¿ ROOT CAUSE USE ERROR, IN THE SENSE THAT IT WAS NOT THE PRODUCT BUT CIRCUMSTANCES WITH PUNCTURE AND USAGE WITH LARYTUBE THAT CAUSED THE INFECTION. DISCUSSION IN THE IFU 10877 THE FOLLOWING IS STATED: ¿DISLODGEMENT OR EXTRUSION OF THE PROVOX2 VOICE PROSTHESIS¿ AND SUBSEQUENT INGESTION, ASPIRATION OR TISSUE DAMAGE MAY OCCUR. TO REDUCE THE RISK¿ : - SELECT PROPER PROSTHESIS SIZE - IF LARYNGECTOMY TUBES¿ ARE USED, CHOOSE SIZE THAT DO NOT EXERT PRESSURE ON THE PROSTHESIS¿ - INSTRUCT THE PATIENT TO CONSULT A PHYSICIAN IMMEDIATELY IF THERE ARE ANY SIGNS OF TISSUE EDEMA AND/OR INFLAMMATION/INFECTION TRACHEAL AND/OR ESOPHAGEAL TISSUE DAMAGE MAY OCCUR IF THE PROSTHESIS IS TOO SHORT, TOO LONG, OR IS PUSHED FREQUENTLY AGAINST THE ESOPHAGEAL WALL BY A TRACHEAL CANNULA, STOMA BUTTON, OR THE PATIENT¿S FINGER. IN PATIENTS UNDERGOING RADIOTHERAPY THIS MAY HAPPEN MORE EASILY. SOLUTION: IF THE PATIENT EXPERIENCES SORENESS OR PAIN IN THE AREA AROUND THE VOICE PROSTHESIS, INSPECT THE TISSUE AROUND THE TE PUNCTURE BY ENDOSCOPY TO AVOID SEVERE DAMAGE.¿ RISK ASSESSMENT RH028-HAZ ED 04, H09: HARM TO PUNCTURE OR ESOPHAGUS OR USED BY PATIENT WITH BLEEDING DISORDER, LEADING TO BLEEDING/HEMORRHAGE (SEVERE BLEEDING 15-30% OF BLOOD VOLUME). CLINICAL SEVERITY = 6, SERIOUS. RH028-HAZ ED 04, H13: CONTAMINATION BY MICROORGANISM LEADING TO INFECTION. CLINICAL SEVERITY = 6. HOWEVER, IN PRODUCT RISK ASSESSMENT (UDFMECA) RH028 PROVOX2 ED08, ITEMS 9.15+9.20 ADDRESS USAGE WITH VP IN COMBINATION WITH LARYTUBE/BUTTON. POINTING TO H15: WRONG SIZE INSERTED (TOO SHORT) LEADING TO NECROSIS. CLINICAL SEVERITY = 6. THE RISK IS CONSIDERED TO BE COVERED IN THE PRODUCT RISK ASSESSMENT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER HAS BEEN NOTIFIED REGARDING THE PASSING OF A PATIENT WHO WAS USING A REP DREAMSTATION AUTO CPAP DEVICE. THE PATIENT'S MOTHER REPORTED THAT HER DAUGHTER PASSED AWAY ON (B)(6) 2021, AND HAS REQUESTED TO BE REMOVED FROM ALL FURTHER COMMUNICATION. THE MOTHER ALSO INDICATED THAT SHE DOES NOT HAVE THE DEVICE AVAILABLE FOR RETURN. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE HAS BEEN NO INDICATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
MEDIHONEY GEL IN TUBE, 1.5 FL OZ FLIPCAP - STERILE | K101793 | A CUSTOMER REPORTED THAT HER HUSBAND PASSED AWAY ON (B)(6) 2024, DUE TO AN INFECTION WHILE USING THE MEDIHONEY PRODUCT. THE PRODUCT WAS APPLIED TO A SURGICAL SITE FOLLOWING A PROCEDURE RELATED TO DIABETES, DURING WHICH BONE WAS REMOVED. THE WOUND TREATED BY CLEANING THE WOUND AND CHANGE THE BANDAGE. THE INFECTION BEGAN APPROXIMATELY FIVE AND HALF WEEKS BEFORE THE PATIENT'S PASSING. THE CUSTOMER ONLY BECAME AWARE OF THE INFECTION AFTER HER HUSBAND WAS HOSPITALIZED, THE PHYSICIAN INFORMED HER ABOUT IT. PRIOR TO THE HOSPITALIZATION, SHE HAD NOTICED THAT HER HUSBAND WAS GETTING WEAKER, WHICH MIGHT HAVE CAUSED A FALL. ACCORDING TO INFORMATION PROVIDED BY THE CUSTOMER, THE MEDICAL RECORDS AND AUTOPSY REPORT ARE NOT AVAILABLE. | UPDATED FIELDS: D4, D9, G3, G6, H2, H3, H6, H11. THE MEDIHONEY GEL (ID 31815) WAS NOT RETURNED FOR EVALUATION AND LOT NUMBER INFORMATION HAS NOT BEEN PROVIDED; THEREFORE, AN EVALUATION OF THE DEVICE COULD NOT BE PERFORMED, AND DEVICE HISTORY RECORD (DHR) COULD NOT BE REVIEWED. THE ROOT CAUSE(S) OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE IN THE FUTURE, THIS COMPLAINT WILL BE REOPENED, AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR ISSUES. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED. THE CUSTOMER INDICATED 3 POSSIBLE LOT NUMBERS FOR PRODUCT ID 31815: LOTS 2328, 2332, 2333. THE REVIEW FOCUSED ON IDENTIFYING POSSIBLE DEVIATIONS OR ANOMALIES RELATED TO POSSIBLE CAUSES OF FLIP CAP - STERILE. NO DEVIATIONS (NCS) OR DOCUMENTATION ERRORS WERE IDENTIFIED IN THE DHR RECORDS. THE INSPECTION AND TESTING RESULTS COMPLIED WITH ESTABLISHED SPECIFICATIONS. ADDITIONAL INFORMATION RECEIVED: PATIENT'S DAUGHTER (B)(6) STATED: "DAD WAS ADMITTED TO THE HOSPITAL IN OCTOBER WITH A BROKEN HIP. HE HAD AN INFECTION IN HIS FOOT THAT HE WAS TREATING WITH MEDI HONEY, AND HE ENDED UP HAVING A METATARSAL BONE REMOVED DUE TO THE INFECTION. HE THEN WENT IN FOR HIP SURGERY AND PASSED AWAY."
CAPIOX® FX ADVANCE HOLLOW FIBER OXYGENATOR W INTEGRATED AF W HARDSHELL RES. | K151791 | ADDITIONAL INFORMATION FROM THE CLINICAL SPECIALIST STATES, THAT THE REVIEW OF THE PUMP RECORD REVEALED THAT THE PARTIAL PRESSURE OXYGEN (PO2) VALUES WERE IN THE 200MMHG RANGE, DESPITE INCREASING FRACTION OF INSPIRED OXYGEN (FIO2) SETTINGS FROM 70 PERCENT TO 100 PERCENT. PO2 VALUES IN THE 200S ARE NOT AT CRITICALLY SUPPRESSED LEVELS TO HARM ACCEPTABLE OXYGENATION VALUES. ACTIVATED CLOTTING TIME (ACT) VALUES WERE ABOVE 480 SECONDS, IMPLYING ADEQUATE HEPARINIZATION OF THE PATIENT. THE HISTORY OF THIS PATIENT WAS A 45-YEAR-OLD THAT HAD SEVERE CARDIOMYOPATHY AND A HISTORY OF MITRAL VALVE REPAIR VIA MITRAL CLIP CATHETERIZATION IN ATTEMPT TO IMPROVE HEART FUNCTION. THIS WAS A VERY SICK PATIENT THAT PRESENTED TO THE OPERATING ROOM (OR) ON A VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) FOR INSERTION OF LEFT AND RIGHT VENTRICULAR ASSIST DEVICES (VADS) IN ATTEMPT TO STABILIZE THE PATIENT'S CRITICAL CONDITION. THE CLINICIAN STATED THAT THE PATIENT WAS ADEQUATELY OXYGENATED, AND THIS REVIEW OF THE PUMP RECORD AGREES WITH THAT ASSESSMENT. | THIS FOLLOW-UP REPORT IS SUBMITTED TO FDA IN ACCORD WITH APPLICABLE REGULATIONS ¿ AND AS INDICATED BY TERUMO CARDIOVASCULAR SYSTEMS IN THE INITIAL REPORT SUBMITTED TO THE FDA ON FEBRUARY 21, 2025. UPON FURTHER INVESTIGATION OF THE REPORTED EVENT, THE FOLLOWING INFORMATION IS NEW AND/OR CHANGED: D9 (DEVICE AVAILABILITY - ADDED DATE RETURNED TO MANUFACTURER), G3 (DATE RECEIVED BY MANUFACTURER), G6 (INDICATION THAT THIS IS A FOLLOW-UP REPORT), H2 (FOLLOW-UP DUE TO ADDITIONAL INFORMATION), H3 (EVALUATION IS IN PROGRESS, BUT NOT YET CONCLUDED). A SECOND FOLLOW-UP WILL BE SUBMITTED UPON COMPLETION OF THE INVESTIGATION AND/OR SUBMISSION OF NEW INFORMATION, THUS TCVS REFERENCES CONCLUSIONS CODE 11. ALL AVAILABLE INFORMATION HAS BEEN PLACED ON FILE IN QUALITY MANAGEMENT FOR APPROPRIATE TRACKING, TRENDING, AND FOLLOW-UP.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PIEZOLITH 3000 TRIPLE FOCUS LITHOTRIPTER | K072538 | ON (B)(6) 2025, WE WERE NOTIFIED THAT THE PATIENT PASSED AWAY AFTER HAVING AN EXTRACORPOREAL SHOCKWAVE LITHOTRIPSY TREATMENT USING THE PIEZOLITH 3000 TRIPLE FOCUS MACHINE. THE PATIENT WAS TREATED ON (B)(6) 2024 AND PASSED AWAY ON (B)(6) 2024. AN INVESTIGATION INTO THIS EVENT IS UNDERWAY. IT IS UNCLEAR WHETHER THE EVENT IS RELATED TO THE DEVICE. A SUPPLEMENTAL MDR WILL BE SUBMITTED AS NEEDED AS THE INVESTIGATION EVOLVES. | AFTER NOTIFICATION OF THIS INCIDENT, THE MAINTENANCE RECORDS FOR THE MACHINE WERE REVIEWED. ON NOVEMBER 11, 2024, THE MACHINE RECEIVED ITS REGULARLY SCHEDULED PREVENTATIVE MAINTENANCE WHERE NO MECHANICAL ISSUES WERE NOTED. THE PATIENT WAS THE THIRD OF EIGHT TREATMENTS THAT DAY. FIVE TREATMENTS WERE PERFORMED AFTER THE PATIENT'S TREATMENT USING THE SAME MACHINE WITH NO KNOWN COMPLICATIONS. SINCE THE PATIENT'S TREATMENT, THERE HAVE BEEN 188 TREATMENTS PERFORMED ON THIS MACHINE WITH NO KNOWN COMPLICATIONS. AN INVESTIGATION INTO THIS EVENT IS UNDERWAY. IT IS UNCLEAR WHETHER THE EVENT IS RELATED TO THE DEVICE. A SUPPLEMENTAL MDR WILL BE SUBMITTED AS NEEDED AS THE INVESTIGATION EVOLVES.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION DUE TO DEATH HOWEVER NO OTHER MEDICAL OR CLINICAL INFORMATION WERE REPORTED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
CARDIOSAVE HYBRID W/ E/F PLUG | K112372 | N/A | UPDATED FIELDS - B4, G3, G6, H2, H3, H6 ( TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSION), H11 A GETINGE FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO EVALUATE THE IABP UNIT AND WAS ABLE TO CONFIRM THE REPORTED ISSUE. THE INTERNAL BLOCKS ARE ADJUSTED (THE PIM AND DRIVE MANIFOLD BLOCKS WERE RELOCATED TO ACHIEVE BETTER SEALING, AS WELL AS THE INTERNAL CONNECTORS) THE FIBER OPTIC CABLE WAS DETECTED TO BE PINCHED, IT IS REPLACED BY A NEW CABLE. ALL CHECKS ARE PERFORMED ACCORDING TO PROTOCOL. THE CORRECT OPERATION OF THE EQUIPMENT IS VERIFIED. EQUIPMENT SUITABLE FOR USE.
PRLNE BLU 36IN 2-0 D/A V-7 MP | K133356 | IT WAS REPORTED THAT A PATIENT UNDERWENT A DISSECTION CARDIAC SURGERY ON THE NIGHT OF (B)(6) 2025 AND PURSE-STRING SUTURE WAS USED. THE SUTURE BROKE, RESULTING IN DEATH ON (B)(6) 2025. ADDITIONAL INFORMATION WAS REQUESTED. | PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. H6 COMPONENT CODE: G07002 DEVICE NOT RETURNED. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. ATTEMPTS HAVE BEEN MADE TO RETRIEVE THE DEVICE. TO DATE THE DEVICE HAS NOT BEEN RETURNED. IF THE DEVICE OR FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WOULD THE SURGEON LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? WHAT IS THE CAUSE OF DEATH? PLEASE PROVIDE THE PATIENT'S DEMOGRAPHIC INFORMATION INCLUDING AGE, GENDER, WEIGHT, BMI AT THE TIME OF INDEX PROCEDURE NAME OF INDEX SURGICAL PROCEDURE? THE DIAGNOSIS AND INDICATION FOR THE INDEX SURGICAL PROCEDURE? WHAT WAS THE INITIAL APPROACH FOR THE INDEX SURGICAL PROCEDURE? (OPEN, LAPAROSCOPIC OR OTHER)? ON WHAT TISSUE WAS THE SUTURE USED? WHAT WAS THE TISSUE CONDITION (NORMAL, THIN, CALCIFIED, FRAGILE, DISEASED)? HOW WAS THE SUTURE PLACED (INTERRUPTED OR CONTINUOUS)? HOW WAS THE SUTURE ORIGINALLY TIED (MULTIPLE KNOTS, SQUARE KNOT, ETC.)? WHAT INSTRUMENTS WERE USED IN THIS PROCEDURE TO HANDLE THE SUTURE? PLEASE DESCRIBE ANY MEDICAL/SURGICAL INTERVENTION REQUIRED FOR THIS SUTURE EVENT INCLUDING DATES AND RESULTS. DID THE OPERATING SURGEON OBSERVE ANY SUTURE DEFICIENCY OR ANOMALY BEFORE, DURING, AFTER THE SUTURE PLACEMENT OR DURING ANY RE-OPERATION? WHAT SYMPTOMS DID THE PATIENT EXPERIENCE FOLLOWING THE INDEX SURGICAL PROCEDURE? ONSET DATE? OTHER RELEVANT PATIENT HISTORY/CONCOMITANT MEDICATIONS? WHAT IS THE PHYSICIAN¿S OPINION AS TO THE ETIOLOGY OF OR CONTRIBUTING FACTORS TO THIS EVENT? LOT NUMBER? WILL PRODUCT BE RETURNED? IF SO, PLEASE PROVIDE THE RETURN DATE AND TRACKING INFORMATION. SURGEON¿S NAME?
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A PERITONEAL DIALYSIS (PD) PATIENT CONTACT REPORTED TO FRESENIUS THAT THE PATIENT EXPIRED WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS REPORTED THAT THE PATIENT'S DEATH WAS NOT PD RELATED. THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. ADDITIONAL INFORMATION WAS OBTAINED DURING FOLLOW-UP WITH THE PATIENT'S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). PER THE PDRN, THE PATIENT EXPIRED ON (B)(6) 2025. THE PATIENT HAD COMPLETED TREATMENT AND WAS STILL ATTACHED TO THE CYCLER. IT WAS EXPLAINED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF CARDIOVASCULAR DISEASE WITH PREVIOUS STROKE (UNRELATED TO PD THERAPY) AND WAS DECLINING IN OVERALL HEALTH. THE NURSE INDICATED THAT THE PATIENT EXPIRED FROM CARDIAC ARREST WHICH WAS ATTRIBUTED TO THE PATIENT¿S PRE-EXISTING CARDIOVASCULAR COMORBID CONDITIONS. THE PDRN STATED THERE WERE NO ISSUES WITH THE PD TREATMENT AND STATED SEVERAL TIMES THAT THE DEATH WAS NOT RELATED TO PRODUCT OR PD THERAPY. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXITS BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH FROM CARDIAC ARREST. HOWEVER, THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER OR ANY ISSUES WITH PD TREATMENT. THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED PRE-EXISTING CARDIAC COMORBIDITIES AND STROKE (UNRELATED TO DIALYSIS) WITH REPORT OF DECLINE IN OVERALL HEALTH. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT'S PRE-EXISTING CARDIAC COMORBID CONDITIONS AND STROKE PLACE THAT PATIENT AT HIGHER RISK OF MORTALITY. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
NOVALUNG CONSOLE (LINUX BOARD) | K191407 | A USER FACILITY REPORTED THAT A PATIENT CAME TO THE CARDIAC CATHETERIZATION LAB FOLLOWING A VENTRICULAR TACHYCARDIA AND VENTRICULAR ECTOPIC ARREST FOR INITIATION OF EXTRACORPOREAL CARDIOPULMONARY RESUSCITATION. THE NOVALUNG DEVICE WAS PREPPED TO HAND UP THE PATIENT CIRCUIT INTO THE STERILE FIELD TO START SUPPORT. WHEN HANDED UP TO THE STERILE FIELD, THE P2 PRESSURE CABLE BROKE AND WAS NOT DETECTED ON THE CONSOLE SCREEN. AN ATTEMPT WAS MADE TO PULL THE CORD AND SWITCH TO ANOTHER CABLE; HOWEVER, THE CABLE WAS STUCK IN THE DEVICE. WITHIN A FEW MINUTES OF ECMO SUPPORT THE P3 PRESSURE CABLE ALSO WAS NOT READING A PRESSURE, AND IT WAS FOUND TO BE IN THE XLUNG KIT, BUT NOT ABLE TO READ A PRESSURE. THIS CABLE ALSO REMAINED IN THE DEVICE. ECMO SUPPORT WAS NOT STOPPED DUE TO THE URGENCY OF SUPPORT AT THAT TIME. THE PATIENT WAS HOOKED UP TO BLOOD PRESSURE AND EKG MONITORS TO ENSURE BEST CARE PRACTICES WERE FOLLOWED. ALSO, THE NOVALUNG CONSOLE WAS READING A TOTAL FLOW TO PATIENT WITH RPMS. THERE WAS NO COLOR CHANGE THAT WAS NOTED ON ARTERIAL LIMB OF CIRCUIT UNTIL THE OXYGEN LEVEL WAS INCREASED TO MAXIMUM SUPPORT OF 100% OXYGEN. UPON ARRIVAL TO THE ICU, THE PATIENT PRESSURE CABLES WERE SWITCHED, AND ONCE CHANGED, THE CLINICIANS WERE ABLE TO NOTE THE NEGATIVE PRESSURE (P1) WAS -400 MMHG. ULTIMATELY, THE PATIENT EXPIRED WITHIN 5 HOURS OF ARRIVAL. THE DISPOSABLE WAS SENT BACK TO FRESENIUS DUE TO THE OXYGENATOR NOT FUNCTIONING (THE DATE AND DESTINATION OF THE SHIPPED PRODUCT WERE NOT REPORTED RESULTING IN NO RECORD OF RETURN OR EVALUATION). ARTERIAL BLOOD GASES WERE DRAWN ON THE POST-OXYGENATOR SIDE WITH AN FIO2 OF 100% AND THE HIGHEST YIELD WAS AN ARTERIAL OXYGEN OF 87 MMHG. AN ARTERIAL OXYGEN DIRECTLY FROM THE PATIENT WAS 89 MMHG WHICH INDICATED THE PATIENT'S LUNGS WERE DOING THE WORK AS IT WAS HIGHER THAN THE DEVICE. NO FURTHER INFORMATION WAS PROVIDED BY THE USER FACILITY. | ADDITIONAL INFORMATION: B5, D9 PLANT INVESTIGATION: NO PARTS WERE RETURNED FOR EVALUATION, NO MACHINE LOG FILES PROVIDED, AND NO SERIAL NUMBERS WERE PROVIDED BY THE REPORTING FACILITY. FOUR SIMILAR COMPLAINTS REPORTING A FUNCTIONALITY ISSUE HAVE BEEN REPORTED IN THE LAST 24 MONTHS. THE REASON FOR THE ABSENCE OR INCORRECT VALUES BEING DISPLAYED ON THE CONSOLE MAY BE ATTRIBUTED TO A DEFECTIVE CABLE, A FAULTY CONNECTION BETWEEN THE CABLE AND SENSOR OR A DEFECT IN THE SENSOR ITSELF.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED INFORMATION INDICATING A PATIENT DIED WHILE ON VENTILATION THERAPY WITH A TRILOGY EVO VENTILATOR. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY AN UNINTENTIONAL USER ERROR BASED ON THE REPORTED STATEMENT THAT THE PATIENT CAME HOME FROM THE HOSPITAL AND WAS ON THE UNIT AND THE FAMILY SAID THEY HAD NO IDEA HOW TO USE OR WORK THE DEVICE (VENTILATOR). THE PATIENT PASSED AWAY. DEVICE EVALUATION AND LOG ATTACHMENTS ARE CURRENTLY PENDING. THEREFORE, DEVICE CAUSE OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH WHILE ON THE DEVICE CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE INFORMATION PROVIDED. THE CAUSE OF THE PATIENT¿S DEATH IS CURRENTLY UNKNOWN. INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION HAS BEEN COMPLETED.
EMERGE | K113220, | IT WAS REPORTED THAT PERORATION AND DEATH OCCURRED. ON (B)(6) 2025, THE PATIENT PRESENTED WITH STABLE ANGINA. THE SAME DAY THE PATIENT UNDERWENT PERCUTANEOUS CORONARY INTERVENTION USING CUTTING BALLOON ANGIOPLASTY IN THE PROXIMAL LEFT ANTERIOR DESCENDING (LAD). THE LESION WAS PREDILATED WITH A 15 MM X 3.00 MM WOLVERINE CORONARY CUTTING BALLOON AT THE LAD. INTRAVASCULAR ULTRASOUND (IVUS) WAS PERFORMED, AND ADDITIONAL DILATATION WAS PERFORMED WITH A 3.50 MM X 15 MM EMERGE BALLOON AND A 4.00 MM X 8 MM EMERGE BALLOON. AN ELLIS CLASS 3 PERFORATION IN THE PROXIMAL LAD OCCURRED WHICH WAS CONSIDERED TO HAVE BEEN CAUSED BY THE 3.50 X 15MM EMERGE BALLOON. THE PERFORATION WAS COVERED INITIALLY WITH A 4.00 X 20 MM NON-BOSTON SCIENTIFIC (NON-BSC) STENT AND PERICARDIOCENTESIS PERFORMED. A 3.00 X 20 MM NON-BSC STENT WAS IMPLANTED DISTAL TO THE INITIAL COVER STENT IN VIEW OF PERSISTENT EXTRAVASATION. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED DUE TO PROFOUND HYPOTENSION. THE PATIENT WAS INTUBATED, AND VENO-ARTERIAL EXTRACORPOREAL MEMBRANE OXYGENATION (VA ECMO) WAS INSERTED. THE LEFT MAIN (LM) WAS RE-ENGAGED VIA THE LEFT FEMORAL ARTERY (LFA) USING AN 8FR NON-BSC GUIDING CATHETER. IT WAS SEEN THAT THE PERFORATION WAS SEALED, AND THE FIRST DIAGONAL WAS LOST BUT THERE WAS SATISFACTORY IVUS AND ANGIOGRAPHIC RESULT. TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE) SHOWED RESTRICTION AND A CALCIFIED AORTIC VALVE. BALLOON VALVULOPLASTY WAS PERFORMED WITH A 20 MM NON-BSC BALLOON. A REPERFUSION SHEATH 6FR WAS INSERTED TO THE RIGHT SUPERFICIAL FEMORAL ARTERY (SFA). THE PATIENT WAS ON ECMO WITH LOW FLOWS AND CHUGGING. THE PERICARDIAL DRAIN HAD BEEN REMOVED. THERE WAS CONCERN FOR RIGHT-SIDED CREPITUS AND A CHEST X RAY WAS OBTAINED. THE PATIENT HAD WORSENING PRESSOR REQUIREMENTS STATUS POST 7 LITERS OF INTRAVENOUS FLUID PROMPTING TRANSFUSION OF MULTIPLE BLOOD PRODUCTS. THE PATIENT SUFFERED A CARDIAC ARREST AND RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED AFTER 1 ROUND OF CPR, EPI AND BICARB. BEDSIDE CARDIAC ULTRASOUND SHOWED MODERATE PERICARDIAL EFFUSION. THE NEXT DAY, THE PATIENT CONDITION WAS FURTHER COMPLICATED BY ECMO CHUGGING AND ELEVATED CENTRAL VENOUS PRESSURE (CVP). ADDITIONALLY, THE PATIENT EXPERIENCED AN ACUTE ISCHEMIC STROKE OF THE LEFT OCCIPITAL AND CEREBELLAR LOBES WITHOUT THROMBOLYSIS OR THROMBECTOMY, AND DEVELOPED THROMBOCYTOPENIA, LEADING TO THE DISCONTINUATION OF HEPARIN. ON (B)(6) 2025, THE PATIENT SUBSEQUENTLY UNDERWENT REPEAT PERICARDIAL DRAINAGE WITH PIGTAIL PLACEMENT AND SWAN GANZ INSERTION IN CATHETERIZATION LAB. ON (B)(6) 2025, THE PATIENT NOTED TO HAVE WORSENING PRESSOR REQUIREMENTS, WORSENING LACTATE, WORSENING RENAL FUNCTION, AND DEVELOPED SUSTAINED VENTRICULAR TACHYCARDIA. THE PATIENT WAS STARTED ON AMIODARONE. THE PATIENT WAS WEANED OFF SEDATION WITH NO NEUROLOGIC RESPONSE NOTED OVER COURSE OF 24 HOURS. ON (B)(6) 2025, THE PATIENT WAS NOTED TO BE ASYSTOLE ON MONITOR. THE ECMO CIRCUIT WAS CLAMPED, AND THE VENTILATOR STOPPED. THE SUBJECT WAS PRONOUNCED AS DECEASED. | G4 PMA/510(K): K163174.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED INFORMATION INDICATING A PATIENT DIED WHILE ON VENTILATION THERAPY WITH A TRILOGY EVO VENTILATOR. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY AN UNINTENTIONAL USER ERROR BASED ON THE REPORTED STATEMENT THAT THE PATIENT CAME HOME FROM THE HOSPITAL AND WAS ON THE UNIT AND THE FAMILY SAID THEY HAD NO IDEA HOW TO USE OR WORK THE DEVICE (VENTILATOR). THE PATIENT PASSED AWAY. DEVICE EVALUATION AND LOG ATTACHMENTS ARE CURRENTLY PENDING. THEREFORE, DEVICE CAUSE OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH WHILE ON THE DEVICE CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE INFORMATION PROVIDED. THE CAUSE OF THE PATIENT¿S DEATH IS CURRENTLY UNKNOWN. INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER¿S INVESTIGATION HAS BEEN COMPLETED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT A PATIENT INJURY OCCURRED WHILE PORTS WERE BEING PLACED FOR A DA VINCI-ASSISTED SACROCOLPOPEXY PROCEDURE. ALL PORTS WERE PLACED AND VISUALIZATION WAS CONFIRMED. BEFORE THE DA VINCI SYSTEM WAS DOCKED TO THE PATIENT, THE ANESTHESIOLOGIST OBSERVED THAT THE PATIENT'S BLOOD PRESSURE WAS DROPPING. THE SURGEON DID NOT VISUALIZE ANY BLEEDING AT THE TIME, BUT CONVERTED THE PROCEDURE TO OPEN AND REQUESTED MULTIPLE GENERAL SURGEONS TO ASSIST. THE ROBOTICS COORDINATOR REPORTED THAT THE LOCATION OF THE BLEEDING WAS NEVER FOUND. NO FURTHER INFORMATION WAS PROVIDED REGARDING WHAT OCCURRED AFTER CONVERTING TO OPEN. THE ROBOTICS COORDINATOR STATED THE THEORY WAS THAT THE VERESS NEEDLE USED DURING TROCAR PLACEMENT DAMAGED RETROPERITONEAL ANATOMY. ATTEMPTS WERE MADE TO SPEAK WITH THE SURGEON, HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED. | NO SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE AS THE DA VINCI SYSTEM WAS NEVER DOCKED / USED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT ACCORDING TO THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, THE PATIENT IN THIS REPORT IS SPECULATED TO HAVE DIED POSSIBLY AS A RESULT OF A VERESS NEEDLE INJURY WHICH IS A THIRD PARTY DEVICE. THE INTUITIVE SURGICAL DA VINCI ROBOT WAS NEVER DOCKED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED AS SAME DAY OF SURGICAL PROCEDURE.
EXETER V40(TM). FEM. STEM 44MM | K011623 | REVISION THA WAS PERFORMED ON (B)(6) 2025 TO REPLACE THE IMPLANT DEVICES (COMPETITOR PRODUCTS) WITH NEW ONES (STRYKER'S PRODUCTS) BECAUSE OF THE NECK FRACTURE OF THE COMPETITOR STEM. THE CUP WAS REPLACED SUCCESSFULLY, AND THE SURGEON BEGAN TO REPLACE THE STEM AS A NEXT STEP. HOWEVER, DURING STEM REPLACEMENT, THE PATIENT'S ELECTROCARDIOGRAM SUDDENLY SHOWED ABNORMAL WAVEFORMS AND HER CONDITION WORSENED. ALTHOUGH IT WAS NOT CONFIRMED PROPERLY, THE ANESTHESIOLOGIST WHO PARTICIPATED IN THE OPERATION STATED "IT'S A PULMONARY EMBOLISM" DURING A LIFESAVING PROCEDURE OF THE SUDDEN CHANGE IN THE PATIENT'S CONDITION, SO THERE IS A POSSIBILITY THAT THE PATIENT WENT INTO CARDIAC ARREST DUE TO A PULMONARY EMBOLISM. AS THE PATIENT WENT INTO CARDIAC ARREST, THE SURGERY DISCONTINUED, AND THE PATIENT MANAGED TO REGAIN CONSCIOUSNESS ON THE SAME DAY THROUGH 30 MINUTES OF CARDIAC MASSAGE AND ECMO. THE PATIENT'S FINAL OUTCOME AFTER REGAINING CONSCIOUSNESS IS UNKNOWN. IT WAS REPORTED THAT THE PATIENT HAD NO ABNORMALITIES SUCH AS PRIMARY DISEASE AT THE TIME OF SURGERY OTHER THAN STEM NECK FRACTURE. UPDATE: PATIENT DIED ON (B)(6) 2025 DUE TO A PULMONARY EMBOLUS. | CORRECTED DATA: DEVICE DETAILS (CAT#, GTIN, PRODUCT CODE, COMMON DEVICE NAME). AN EVENT REGARDING PATIENT FACTORS (PULMONARY EMBOLISM) INVOLVING A EXETER STEM WAS REPORTED. THE EVENT WAS NOT CONFIRMED. METHOD & RESULTS: PRODUCT EVALUATION AND RESULTS: MATERIAL ANALYSIS, VISUAL, FUNCTIONAL AND DIMENSIONAL INSPECTIONS COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED. CLINICIAN REVIEW: NO MEDICAL RECORDS WERE RECEIVED FOR REVIEW WITH CLINICAL CONSULTANT. PRODUCT HISTORY REVIEW: REVIEW OF THE DEVICE HISTORY RECORDS INDICATE DEVICES WERE MANUFACTURED AND ACCEPTED INTO FINAL STOCK WITH NO RELEVANT REPORTED DISCREPANCIES. COMPLAINT HISTORY REVIEW: THERE HAVE BEEN NO OTHER SIMILAR EVENTS FOR THE LOT REFERENCED. CONCLUSIONS: IT WAS REPORTED THAT DURING THE SURGERY, AFTER THE CUP PLACEMENT WAS COMPLETED AND THE STEM ON THE FEMORAL SIDE WAS BEING RE-REPLACED (THE SALES COMMENT WAS, "THE SURGICAL FIELD DID NOT LOOK CLEAR, SO THIS IS JUST SPECULATION, BUT WE THINK IT WAS WHEN WE WERE REMOVING THE SOFT TISSUE"), THE PATIENT'S ELECTROCARDIOGRAM SUDDENLY SHOWED ABNORMAL WAVEFORMS AND HER CONDITION SUDDENLY WORSENED. THE EXACT CAUSE COULD NOT BE DETERMINED AS INSUFFICIENT INFORMATION WAS PROVIDED. NO FURTHER INVESTIGATION FOR THIS EVENT IS POSSIBLE AT THIS TIME. IF DEVICES AND / OR ADDITIONAL INFORMATION BECOME AVAILABLE TO INDICATE FURTHER EVALUATION IS WARRANTED, THIS RECORD WILL BE REOPENED.
NXSTAGE SYSTEM ONE | K050525 | A REPORT WAS RECEIVED ON 03 FEB 2025 FROM THE NURSE OF A MALE PATIENT WITH A MEDICAL HISTORY INCLUDING CHRONIC HYPOTENSION AND END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ADDITIONAL INFORMATION WAS RECEIVED 05-07 FEB 2025 FROM THE HOME THERAPY MANAGER (HTM) WHO STATED THE PATIENT WAS FOUND BY THE CAREGIVER SURROUNDED BY AN UNSPECIFIED AMOUNT OF BLOOD (SOURCE OF BLOOD LOSS NOT SPECIFIED). THE CAREGIVER PERFORMED CARDIOPULMONARY RESUSCITATION (CPR) AND CALLED EMERGENCY MEDICAL SERVICES (EMS). UPON EMS ARRIVAL, RESUSCITATION EFFORTS WERE CONTINUED BUT UNSUCCESSFUL, AND THE PATIENT EXPIRED, TIME NOT PROVIDED. | THE CYCLER WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
INTELLIVUE MX40 802.11A/B/G | K113125 | IT WAS REPORTED THAT WHILE MONITORING THE PATIENT THE READINGS OF THE SP02 WERE INCONSISTENT. THE SOURCE NOTES INDICATED THAT THE PATIENT HAS PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED. ADDITIONAL INFORMATION IS REQUESTED FOR FURTHER CLARIFICATION AND ASSESSMENT OF THE CUSTOMER¿S ALLEGED HARM. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
CARESCAPE B450 | K213363 | IT WAS REPORTED THAT NO AUDIBLE ALARMS WERE PROVIDED WHEN THE PATIENT EXPIRED. | LEGAL MANUFACTURER: (B)(4). A1-2, 4-6: NOT PROVIDED BY THE CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
EFFICIA DFM100 | K110825 | PHILIPS RECEIVED A COMPLAINT ON THE EFFICIA DFM100 DEFIBRILLATOR MONITOR INDICATING THE CUSTOMER WAS USING ECG ON A DECEASED PATIENT AND DOTTED LINES AND JUMPING WAVE PATTERNS APPEARED ON THE SCREEN. | REPORTING INSTITUTION PHONE # AND REPORTER PHONE # UNKNOWN.
BIPAP SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER WAS CONTACTED ABOUT A BIPAP AUTO, DOM DEVICE. THE PATIENT WAS REACHED OUT TO ABOUT RETURNING HIS RECALLED DEVICE AFTER HE RECEIVED HIS REPLACEMENT. HE MENTIONED THAT HE HAS BEEN USING HIS MOTHER'S CPAP DEVICE SINCE SHE PASSED AWAY AND HAS ONLY RECENTLY OPENED HIS REPLACEMENT DEVICE. HE RECEIVED A NOTIFICATION THAT HE NEEDED TO RETURN HIS RECALLED DEVICE AND HAS RECEIVED THE SHIPPING LABEL. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE MOTHER'S DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
ION | K182188 | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE. THE PHYSICIAN OBTAINED BIOPSY SAMPLES FROM THE RIGHT LOWER LOBE TARGET LESION. THERE WERE NO ISSUES WITH THE ION SYSTEM. AFTER COMPLETING THE ION PORTION OF THE PROCEDURE, THE ION SCOPE WAS REMOVED AND MANUAL BRONCHOSCOPY WAS PERFORMED. BLEEDING WAS NOTED AT THE BIOPSY SITE WHILE OBTAINING A BRONCHOALVEOLAR LAVAGE SAMPLE. THE ESTIMATED BLOOD LOSS WAS 60ML. THE BLOOD ENTERED THE ENDOTRACHEAL TUBE (ETT) AND THE PATIENT SUFFERED A CARDIAC ARREST, PROMPTING ADVANCED CARDIAC LIFE SUPPORT INTERVENTIONS. THE PATIENT WAS REVIVED AND HEMOSTASIS WAS ACHIEVED. THE ETT WAS REPLACED WITH A DOUBLE LUMEN ETT AND INTRAVENOUS TRANEXAMIC ACID AND AN EPINEPHRINE DRIP WERE ADMINISTERED. THE PATIENT EXPERIENCED A SECOND CARDIAC ARREST 2 DAYS POST-PROCEDURE AND EXPIRED. THE PHYSICIAN REPORTED THAT THE BLEEDING AND PATIENT OUTCOME WERE DUE TO SIGNIFICANT RISKS ASSOCIATED WITH THE BIOPSY PROCEDURE AND COMPOUNDED BY THE PATIENT'S SEVERE PREEXISTING COMORBID CONDITIONS, INCLUDING HEART FAILURE WITH REDUCED EJECTION FRACTION, PULMONARY HYPERTENSION, AND SEVERE COPD. | REVIEW OF THE ION SYSTEM LOGS FOUND NO RELATED ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY; THE ION PORTION OF THE CASE WAS COMPLETED, THE ION SCOPE WAS REMOVED AND MANUAL BRONCHOSCOPY WAS PERFORMED. BLEEDING WAS NOTED WHILE OBTAINING A BRONCHOALVEOLAR LAVAGE SAMPLE. THIS WAS COMPLICATED BY A CARDIAC ARREST. RETURN OF SPONTANEOUS CIRCULATION (ROSC) WAS ACHIEVED WITH RESUSCITATIVE EFFORTS AND THE PATIENT WAS HOSPITALIZED IN THE CCU. THE PATIENT SUFFERED ANOTHER ARREST 2 DAYS AFTER THE PROCEDURE AND DIED. THERE IS NO ALLEGATION THAT A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT, OR ACCESSORY OCCURRED. GIVEN THE AVAILABLE DATA THE REPORTED EVENT WAS LIKELY PROCEDURE-RELATED AND NOT DEVICE-RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW-RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE-CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF RELATIVELY SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
HAMILTON G5 | K193228 | HAMILTON MEDICAL AG RECEIVED THE FOLLOWING EVENT DESCRIPTION: QUOTATION FROM THE REPORTER: " HAMILTON G5 SN (B)(6) WAS ON A PATIENT WHO EXPIRED, UNKNOWN OF ANY PROBLEMS OR ISSUES THAT LED TO THIS. " "TECH DID INSPECT MACHINE AFTER INCIDENT AND MENTIONED THAT EXP. VALVE MAY BE STICKY. " ACCORDING TO THE INFORMATION RECEIVED IT IS UNKNOWN IF THE PATIENT WAS HARMED BY MACHINE. A MEDICAL INTERVENTION WAS NECESSARY, BUT THE SPECIFIC TYPE OF INTERVENTION REQUIRED WAS NOT SPECIFIED. MOREOVER, THE MANUFACTURER WAS INFORMED THAT THE USER REPORTED THIS EVENT TO LOCAL AUTHORITIES, ALTHOUGH NO REFERENCE NUMBER WAS PROVIDED. FURTHER INQUIRIES HAVE BEEN SENT TO THE CUSTOMER TO OBTAIN ADDITIONAL DETAILS ABOUT THE EVENT. THE CAUSALITY OF THE ADVERSE EVENT ASSOCIATED WITH THE HAMILTON-G5 HAS NOT YET BEEN DETERMINED AND IS CURRENTLY UNDER INVESTIGATION | HAMILTON MEDICAL AG COMPLAINT NUMBER: (B)(4). INVESTIGATION ONGOING.
RING ANNULOPLSTY TRICUSP CONTOUR 3D 28MM | K101212 | MEDTRONIC RECEIVED INFORMATION FROM A PATIENT REPRESENTATIVE THAT 1 DAY POST IMPLANT OF THIS 28MM ANNULOPLASTY RING, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE RING COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. | SELECT PATIENT INFORMATION CANNOT BE INCLUDED IN THE REGULATORY REPORT DUE TO REGIONAL PRIVACY REGULATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
WATCHMAN TRUSTEER ACCESS SYSTEM | K240018 | IT WAS REPORTED THAT ST ELEVATION, BRAIN INJURY, AND DEATH OCCURRED. A CONCOMITANT PROCEDURE WAS BEING PERFORMED WITH INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING UNDER DEEP SEDATION WITH NO INTUBATION PERFORMED. A PULSE FIELD ABLATION (PFA) PROCEDURE WAS PERFORMED SUCCESSFULLY FIRST USING FARAPULSE PFA SYSTEM. SOMETIME BETWEEN THE CAVOTRICUSPID ISTHMUS LINE (CTI) LINE ABLATION PORTION AND THE INSERTION OF THE WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) FOR THE INITIATION OF THE LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE, THE PATIENT HAD SUDDEN ST ELEVATION AND HYPOTENSION. AIR WAS SUSPECTED, SO NITRO MEDICATION WAS ADMINISTERED VIA INTRACARDIAC DELIVERY. A 31 MM WATCHMAN FLX PRO DELIVERY SYSTEM AND CLOSURE DEVICE (WDS) WAS QUICKLY INSERTED THROUGH THE WAS AND DEPLOYED SUCCESSFULLY, JUST IN CASE AIR WAS PRESENT IN THE DISTAL LAA. THE INTERVENTIONAL CARDIOLOGIST PERFORMED A DIAGNOSTIC INTERVENTION TO LOOK FOR PRESENT AIR IN THE CORONARY ARTERIES, BUT NO AIR WAS EVER VISUALIZED IN THE PATIENT. THE ST ELEVATIONS SELF-RESOLVED AND THE PATIENT REGAINED NORMAL HEMODYNAMICS. THE PROCEDURE WAS CONCLUDED. THE PATIENT WAS HOSPITALIZED YET NEVER REGAINED CONSCIOUSNESS POST INDEX PROCEDURE. MAGNETIC RESONANCE IMAGING (MRI) SCAN AND MULTIPLE COMPUTED TOMOGRAPHY (CT) SCANS WERE PERFORMED AND REVEALED NO AREA OF CONCERN, AND NO AIR WAS FOUND. HOSPICE CARE WAS INITIATED AND THE PATIENT DIED SEVERAL DAYS POST INDEX PROCEDURE ON AN UNKNOWN DATE. THE OFFICIAL CAUSE OF DEATH IS ANOXIC BRAIN INJURY DUE TO GLOBAL HYPOXIA. THE PHYSICIAN STATED THE DEATH HAD NO OFFICIAL RELATIONSHIP TO THE DEVICES USED OR THE PROCEDURE ITSELF AND IT WAS UNKNOWN WHAT CAUSED THE ANOXIC BRAIN INJURY OTHER THAN IT WAS POSSIBLE THE NON-INTUBATED PATIENT HAD TAKEN A DEEP BREATH WHILE THE WAS WAS IN THE GROIN. | B3: DATE OF DEATH IS UNKNOWN.
REP DREAMSTATION AUTO BIPAP, DOM | K131982 | IT WAS REPORTED THAT THE REP DREAMSTATION AUTO BIPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | THIS REPORT IS BEING SUBMITTED TO PROVIDE THE DEVICE'S EVALUATION RESULTS, UPDATED CODING, AND TO INCLUDE THE PRODUCT UDI. THE REP DREAMSTATION AUTO BIPAP DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE DEVICE WAS VISUALLY INSPECTED DURING THE EVALUATION, AND THE COMPLAINT COULD NOT BE REPRODUCED. THE EVALUATION FOUND THE DEVICE WAS MISSING FILTERS. NO OTHER FAILURES WERE OBSERVED. THE DEVICE WAS SCRAPPED DUE TO THE CUSTOMER'S REQUEST. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION AUTO CPAP | K131982 | IT WAS REPORTED THAT THE PATIENT OF THE DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE DEVICE WILL NOT BE RETURNED FOR EVALUATION AND THE DEVICE IS CURRENTLY REGISTERED TO A NEW USER. AT THIS TIME, THE DEVICE WILL NOT BE RETURNING, AND NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
MICRA | K132030 | IT WAS FURTHER REPORTED THAT THE AFTER THE LEADLESS IPG WAS RECOVERED BY THE FIRST SNARE IT WAS RELEASED AGAIN IN THE RIGHT ATRIUM, VERY MOBILE AND SNARE WAS RENEWED, BUT ONLY ON THE TINES. THEREFORE, A SECOND SNARE WAS USED AND THE LEADLESS IPG WAS SUCCESSFULLY RECOVERED. A LARGE THROMBI WAS OBSERVED ON THE IPG AND THE SNARE. AIRLOCK OF THE DELIVERY SYSTEM WAS RINSED IN ORDER TO REMOVE THE THROMBI. MEDICATION WAS THEN ADMINISTERED THROUGH THE AIRLOCK. "A FEW MINUTES LATER, THE PATIENT PRESENTS WITH ACUTE DYSPNEA AND THEN CARDIAC ARREST WITH ASYSTOLE". | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
ECHELON FLEX 60 POWERED PLUS ENDOPATH STAPLER ARTICULATING ENDOSCOPIC LINEAR CUT | K110385 | IT WAS REPORTED VIA JOURNAL ARTICLE: TITLE: PROPOSAL OF THE NOVEL PREDICTIVE MODEL FOR POSTOPERATIVE PANCREATIC FISTULA IN DISTAL PANCREATECTOMY FOR PANCREATIC TUMOR BASED ON PREOPERATIVE IMAGING PARAMETERS AND STAPLER HANDLING. AUTHORS: KOSHIRO MORINO I KENZO NAKANO I TOSHIHIKO GOTO I TAKAFUMI MACHIMOTO. CITATION: WORLD J SURG. 2024;1-11. DOI: 10.1002/WJS.12113. THE PURPOSE OF THIS SINGLE-CENTER RETROSPECTIVE STUDY WAS TO INVESTIGATE PREOPERATIVE RISK FACTORS AND ESTABLISH THE PREDICTIVE MODEL FOR POPF. FURTHER, WE ESTIMATED THE ASSOCIATION BETWEEN SURGICAL VARIABLES AND POPF, THEN EXPLORED SAFE AND FAVORABLE STAPLING TECHNIQUES FOR PANCREATIC TRANSECTION. BETWEEN 2014 AND 2022, A TOTAL OF 81 PATIENTS (37 FEMALES AND 44 MALES) INVOLVED IN THIS STUDY. TRANSECTION OF THE PANCREAS WAS PERFORMED USING BLUE/GOLD CARTRIDGE OF ENDOPATH¿ STAPLER OR GST¿ SYSTEM OF THE ECHELON FLEX¿ (ETHICON ENDOSURGERY) AND PURPLE/BLACK CARTRIDGE OF TRI-STAPLE TM TECHNOLOGY OF ENDO GIA¿ (COVIDIEN MEDTRONIC). 24 REINFORCEMENT USING POLYGLYCOLIC ACID MESH (NEOVEIL® GUNZE LTD.) WAS USED, ACCORDING TO SURGEONS' PREFERENCE, WITHOUT ANY SET INDICATIONS. AT LEAST ONE DRAIN WAS PLACED NEAR THE STUMP OF THE REMNANT PANCREAS. THE D-AMY WAS MEASURED ON POSTOPERATIVE DAY (POD) 1, 2, 3 ROUTINELY AND LATER DEPENDING ON THE CLINICAL SITUATION. THE FOLLOW-UP COMPUTED TOMOGRAPHY (CT) WAS PERFORMED POD 3-5 ROUTINELY. THE DRAIN WAS REMOVED WHEN THE CT SHOWED NO RESIDUAL FLUID AND DRAINAGE FLUID AMYLASE (D-AMY) WAS LESS THAN ONE THOUSAND ON POD 3-5. SOMATOSTATIN ANALOGS WERE NOT ADMINISTERED POSTOPERATIVELY TO PREVENT POPF. THE FOLLOW-UP DATA RELATED TO THESE PATIENTS WAS UPDATED IN OCTOBER 2023. REPORTED COMPLICATIONS INCLUDED. ECHELON FLEX¿ (ETHICON ENDOSURGERY). ECHELON FLEX BLUE/GOLD CARTRIDGE (ETHICON ENDOSURGERY). POPF (POSTOPERATIVE PANCREATIC FISTULA) (N=?). TREATMENT: DRAIN PLACEMENT, MONITORING D-AMY LEVELS, CT SCAN FOR FOLLOW-UP, AND REMOVAL OF THE DRAIN WHEN CRITERIA WERE MET MORBIDITY AS DEFINED BY CLAVIEN-DINDO CLASSIFICATION VERSION 2.0 (N=?). TREATMENT: NECESSITY FOR REPEAT SURGERY OR RE-ADMISSION. MORTALITY (N=?). TREATMENT: NOT REPORTED. IN CONCLUSION, THE NOVEL PREOPERATIVE MODEL CONSISTING OF IMAGING PARAMETERS EXHIBITED EXCELLENT PREDICTABILITY, AND FOUND THICK CARTRIDGE USE AND LONG PRE-FIRING COMPRESSION TO BE PROTECTIVE FACTORS AGAINST POPF. THIS MODEL MAY FACILITATE THE DEVELOPMENT OF PREVENTIVE SURGICAL STRATEGIES TO REDUCE POPF. | (B)(4). DATE SENT: 3/4/2025. D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. THIS REPORT IS RELATED TO A JOURNAL ARTICLE, THEREFORE NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT/BATCH NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: DOES THE AUTHOR/SURGEON BELIEVE THAT THE ETHICON DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT COMPLICATIONS MENTIONED IN THE ARTICLE? IF YES, PLEASE EXPLAIN. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
TEMPUS LS | K200849 | THIS REPORT IS BASED ON INFORMATION PROVIDED BY REMOTE SERVICE ENGINEER RSE AND IS BEING INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON THE TEMPUS LS MANUAL INDICATING SHOWED THE DEVICE DISPLAYED AN ORANGE BANNER STATING, "CHECK ECG LEADS", EVEN THOUGH THEY WERE CONNECTED TO THE DEVICE. MULTIPLE KNOWN WORKING ECG LEADS WERE ATTEMPTED, BUT NO CHANGE. THE CUSTOMER STATED THEY WERE TRANSITIONING FROM MOTORING THE PATIENT, 18-YEAR-OLD FEMALE, TO ATTEMPTING TO PACE AND WERE "UNABLE TO DO SO BECAUSE DEVICE WOULD NOT READ ECG LEAD 2 OR 3". THE PATIENT WAS IN BRADYCARDIC RHYTHM BEFORE SWITCHING TO PACING. THE PATIENT WENT INTO CARDIAC ARREST AND SUBSEQUENTLY DIED.
TOUCHSOFT COAGULATOR | K093079 | IT WAS REPORTED THAT HOSPITAL "HAD A PATIENT THAT HAD 2 DUODENAL AVMS, FOR WHICH THE TOUCHSOFT PROBE WAS USED. THE SETTINGS WERE APPROPRIATE, WITH IT BEING ON SOFT COAG. AND AT 30-40 WATTS. THE COAGULATION APPEARED TO BE APPROPRIATE. THE PATIENT WAS DISCHARGED, AND A DAY LATER WAS HAVING ABDOMINAL PAIN. THE PATIENT HAD A DUODENAL PERFORATION REQUIRING SURGERY." UPON FOLLOW-UP, THE HOSPITAL EXPLAINED, TWO DAYS AFTER THE SURGERY, THE PATIENT CALLED THE DOCTOR COMPLAINING ABOUT PAIN. THE PATIENT RETURNED TO THE HOSPITAL, AND THE DOCTOR FOUND AN INFECTION IN THE DUODENUM. THE PATIENT WAS ADMITTED TO THE HOSPITAL WHERE A CT SCAN WAS PERFORMED, AND A PERFORATION WAS FOUND. THE PATIENT WAS TAKEN TO SURGERY AND THEN WAS ADMITTED TO THE ICU, WHERE THE PATIENT REMAINED FOR SIX WEEKS UNTIL PASSING AWAY. IN ADDITION, THE HOSPITAL REPRESENTATIVE STATED THAT THE PATIENT HAD SEVERAL OTHER CO-MORBIDITIES AND DID NOT KNOW EXACTLY WHAT CAUSED THIS EVENT TO HAPPEN. THE HOSPITAL REPRESENTATIVE STATED THAT THE DOCTOR DID NOT THINK THAT THE PROBE CONTRIBUTED TO THE PATIENT'S PASSING BUT CANNOT RULE IT OUT. AVMS WERE 2 SMALL, NONBLEEDING ANGIODYSPLASIAS LOCATED IN THE SECOND PORTION OF THE DUODENUM. TOUCHSOFT PROBE SETTINGS USED WERE 30, INCREASED TO 40 DURING TREATMENT. NO ABNORMALITIES WERE NOTED DURING TREATMENT. PATIENT ALSO RECEIVED BIOPSY OF 10MM NODULE IN DUODENAL BULB. PERFORATION WAS LOCATED IN THIRD PORTION OF THE DUDENUM PER OPERATIVE REPORT. | MANUFACTURING RECORDS COULD NOT BE REVIEWED AS THE FACILITY WAS NOT ABLE TO PROVIDE LOT INFORMATION FOR THE DEVICE INVOLVED IN THE EVENT. PER THE DESCRIPTION OF THE PROCEDURE PROVIDED BY THE FACILITY, GENERATOR SETTINGS USED WERE WITHIN THE PARAMETERS RECOMMENDED IN THE IFU.
RENTAL, DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A RENTAL, DREAMSTATION ST30. THE PATIENT WAS PASSED AWAY AND THAT WAS THE REASON FOR RETURNING THE DEVICE. THERE WAS NO ALLEGATION THAT THE PATIENT'S DEATH WAS RELATED TO THE DEVICE. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE WAS RETURNED TO MANUFACTURER INVESTIGATION LABORATORY. THE DEVICE PASSED ALL THE TEST. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. THE MANUFACTURE'S INVESTIGATION IS ONGOING. A FOLLOW UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURE'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT WAS PASSED AWAY. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN BOX H: EVALUATION METHOD CODE, EVALUATION RESULT CODE, COMPONENT CODE AND CONCLUSION CODE HAS BEEN UPDATED.
FREESTYLE LIBRE 2 | K193371 | ABBOTT DIABETES CARE (ADC) RECEIVED A MEDWATCH REPORT WHICH REPORTED THE FOLLOWING INFORMATION: A HEALTHCARE PROFESSIONAL (HCP) REPORTED THAT THE CUSTOMER, WHO WAS USING A FREESTYLE LIBRE 2 SENSOR HAD "PASSED AWAY". AN OFFICIAL CAUSE OF DEATH HAS NOT BEEN PROVIDED AT THIS TIME AND NO ALLEGATION OF DEVICE DEFICIENCY HAS BEEN MADE. THE REPORT INDICATED THAT THE HCP DID NOT CONSENT TO FOLLOW UP. ADC IS THEREFORE UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. THEREFORE, BASED ON THE LIMITED INFORMATION PROVIDED, IT CANNOT BE REASONABLY CONCLUDED THAT AN ADC PRODUCT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO THE EVENT. | THIS ISSUE DOES NOT MEET REPORTABILITY CRITERIA; HOWEVER, IT IS BEING REPORTED TO THE FDA AS THE COMPLAINT WAS RECEIVED VIA USER REPORT. A FOLLOW-UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS OBTAINED. THE DATE OF THE INCIDENT IS UNKNOWN. THE DATE ENTERED IN SECTION B3 IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. THE DEVICE MFG DATE IS UNKNOWN. THE DATE ENTERED IN SECTION H4 IS THE DATE ABBOTT DIABETES CARE BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH OF THE PATIENT. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS INSPECTED AND THE UNIT PASSED THE FUNCTIONAL TEST. THE DEVICE'S EVENT LOG WAS REVIEWED BY THE SERVICE CENTER AND ERROR CODE FOUND. ADDITIONALLY, CUSTOMER COMPLAINT WAS NOT REPRODUCED, AND THE DEVICE WAS SCRAPPED. THERE WAS AN ALLEGATION OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY. AT THIS TIME MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS NOTIFIED THAT A PATIENT WHO HAD BEEN USING A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER, STATING THAT HER HUSBAND HAD PASSED AWAY AND THAT SHE NO LONGER REQUIRED THE DEVICE. SINCE THE DEVICE HAD BEEN USED, THE RESPIRATORY THERAPIST INFORMED THE CALLER THAT IT WAS NOT NECESSARY TO RETURN THE DEVICE. THE CALLER INQUIRED ABOUT THE POSSIBILITY OF FINANCIAL COMPENSATION, AS SHE HAD HEARD ABOUT A SETTLEMENT. UPON REVIEWING THE X PRIORITY NOTIFICATION, THE RESPIRATORY THERAPIST CONFIRMED THAT THE AFFECTED DEVICE HAD NOT BEEN RETURNED. THE CALLER INDICATED THAT SHE DID NOT HAVE ANY TRACKING INFORMATION FOR THE DEVICE, EXPLAINING THAT THE DURABLE MEDICAL EQUIPMENT (DME) PROVIDER HAD TAKEN POSSESSION OF IT. THE CALLER WAS SUBSEQUENTLY INFORMED THAT PHILIPS' ONLY REMEDIATION FOR THE AFFECTED DEVICE WAS A REPLACEMENT UNIT, WHICH THE CALLER ACKNOWLEDGED AND UNDERSTOOD. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION OR ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS NOTIFIED THAT A PATIENT WHO HAD BEEN USING A DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CONTACTED THE MANUFACTURER, STATING THAT HER HUSBAND HAD PASSED AND REQUESTED A REFUND FOR THE DEVICE. IN RESPONSE, THE MANUFACTURER CLARIFIED THAT DEVICES ARE BEING PROVIDED TO PATIENTS THROUGH THE REPAIR AND REPLACEMENT PROGRAM, AND THAT FINANCIAL COMPENSATION, REIMBURSEMENT, OR PAYMENT IS NOT BEING OFFERED. ADDITIONALLY, THE WIFE WAS PROVIDED WITH THE CONTACT INFORMATION (PHONE NUMBER AND EMAIL ADDRESS) OF A THIRD-PARTY SETTLEMENT GROUP, WHICH HAS BEEN RETAINED TO ADMINISTER AN ECONOMIC LOSS SETTLEMENT, SHOULD SHE HAVE ANY FURTHER QUESTIONS. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE IS NO INDICATION OR ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
INQUIRY¿ STEERABLE DIAGNOSTIC CATHETER | K961924 | DURING AN ATRIAL FLUTTER PROCEDURE, THE PATIENT BECAME BRADYCARDIC, WENT INTO CARDIAC ARREST, AND EXPIRED. THE PROCEDURE BEGAN BY POSITIONING THE INQUIRY CATHETER IN THE CORONARY SINUS AND USING A NON-ABBOTT INTRACARDIAC ECHO PROBE TO MAP THE RIGHT ATRIUM. THE MAP AND GEOMETRY WERE COLLECTED WITH A NON-ABBOTT CIRCULAR CATHETER. THE TYPICAL ATRIAL FLUTTER ABLATION WAS PERFORMED WITH A NON-ABBOTT 8MM RF ABLATION CATHETER. AFTER MAPPING AND ABLATION IN THE RIGHT ATRIUM WERE COMPLETED, THE NON-ABBOTT CIRCULAR CATHETER WAS MOVED TO THE LEFT ATRIUM VIA TRANSSEPTAL PUNCTURE USING SWARTZ AND AGILIS SHEATHS AND A BRK NEEDLE. GEOMETRY OF THE LEFT ATRIUM AND A VOLTAGE MAP WERE COLLECTED AND SHOWED AN ELECTRICAL SIGNAL IN THE PULMONARY VEINS. ISOLATION OF THE PULMONARY VEINS BEGAN WITH A NON-ABBOTT (FARAPULSE, BOSTON SCIENTIFIC) CATHETER AND NON-ABBOTT (BOSTON SCIENTIFIC) EQUIPMENT AND TECHNICAL SUPPORT. IMMEDIATELY AFTER ABLATION OF THE PULMONARY VEINS, A NEW MAP OF THE LEFT ATRIUM WAS COLLECTED USING THE NON-ABBOTT CIRCULAR CATHETER TO CHECK WHETHER THE VEINS WERE ISOLATED. AT THAT MOMENT, AFTER THE GEOMETRY COLLECTION BEGAN ON THE POSTERIOR WALL OF THE LEFT ATRIUM, THE PATIENT BECAME BRADYCARDIC AND WENT INTO CARDIAC ARREST. THE TEAM PERFORMED RESUSCITATION MANEUVERS, BUT AFTER A PERIOD OF ATTEMPTS, THE PATIENT DIED. | ADDITIONAL INFORMATION: G3, H2, H3, H6, H11. THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED CARDIAC ARREST AND DEATH COULD NOT BE CONCLUSIVELY DETERMINED.
DREAMSTATION BIPAP AVAPS | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) REPORTED THE PATIENT EXPIRED PEACEFULLY IN HIS SLEEP DUE TO A CARDIAC ARREST ON (B)(6) 2025. THE PDRN STATED THE PATIENT'S TREATMENT WAS COMPLETED WITHOUT ANY KNOWN ALARMS/ISSUES (VERIFIED THROUGH TREATMENT DATA). THE PDRN CONFIRMED THE PATIENT'S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST, SECONDARY TO THE PATIENT'S ADVANCED AGE AND CARDIAC COMORBID CONDITIONS. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT'S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE DEFINITIVE CAUSE OF THE PATIENT¿S CARDIAC ARREST IS UNKNOWN; THEREFORE, CAUSALITY CANNOT FIRMLY BE ESTABLISHED. HOWEVER, THE PATIENT¿S PDRN REPORTED THE PATIENT WAS ELDERLY AND POSSESSED A SIGNIFICANT CARDIAC HISTORY. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO THE PATIENT¿S UTILIZATION OF ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S). BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING | K201571 | ON 15-APR-2025, A DEVICE EVALUATION WAS COMPLETED BY SOLVENTUM QUALITY ENGINEERING. ON 30-DEC-2024, THE DEVICE WAS TESTED PER QUALITY CONTROL PROCEDURE BY THE SOLVENTUM SERVICE CENTER, AND THE DEVICE PASSED AND MET SPECIFICATION PRIOR TO PATIENT PLACEMENT. ON (B)(6) 2025, THE DEVICE WAS PLACED WITH THE PATIENT. ON 15-APR-2025, THE DEVICE WAS TESTED PER QUALITY CONTROL PROCEDURE BY SOLVENTUM QUALITY ENGINEERING AND NO FAILURES WERE FOUND WITH THE DEVICE RELATED TO THIS EVENT. | BASED ON THE ADDITIONAL INFORMATION PROVIDED REGARDING THE DEVICE, THE ASSESSMENT REMAINS THE SAME; IT CANNOT BE DETERMINED THAT THE ALLEGED PATIENT EXPIRATION IS RELATED TO THE ACTIV.A.C.¿ THERAPY SYSTEM WITH ION PROGRESS¿ REMOTE THERAPY MONITORING. THE DEVICE MET SPECIFICATIONS BEFORE PATIENT PLACEMENT AND NO FAILURES WERE FOUND WITH THE DEVICE RELATED TO THIS EVENT POST-PLACEMENT.
MICRA | K132030 | IT WAS FURTHER REPORTED THAT THE AIR EMBOLISM RELATED TO THE SHEATH INSERTION WAS A CONTRIBUTING FACTOR TO THE DEATH. ACCORDING TO THE PHYSICIAN NO BLOOD RETURN WAS NOTED WHEN ASPIRATION WAS ATTEMPTED. A FLUOROSCOPY X-RAY IMAGE THEN DEMONSTRATED APPARENT AIR IN THE RIGHT ATRIUM AND RIGHT VENTRICLE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CENTRIMAG BLOOD PUMP | K020271 | THE ECMO SPECIALIST AND REGISTERED NURSE (RN) WERE BATHING THE PATIENT, AFTER THE BATH, THE RN WALKED PAST THE MACHINE AND BUMPED IT WITH THEIR HIP. THE MACHINE THEN BEGAN ALARMING LOW FLOW AND THE ECMO SPECIALIST RESPONDED AND NOTICED EXSANGUINATION COMING FROM THE TUBING. REGARDING THE SECOND CIRCUIT THAT COULD NOT BE PLACED ON THE PATIENT, IT WAS UNABLE TO BE IDENTIFIED HOW THERE WAS AIR IN IT. IT BEGAN ALARMING WHEN THERE WERE BUBBLES AFTER PERFORMING THE CIRCUIT EXCHANGE. | CORRECTED DATA SECTION H10: MEDWATCH NUMBER ADDED. NO FURTHER INFO WAS PROVIDED. A SUPPLEMENTAL REPORT WILL BE REPORTED ONCE THE MANUFACTURER¿S INVESTIGATION IS COMPLETED.
AIRSEAL | K190303 | THE SALES REPRESENTATIVE REPORTED ON BEHALF OF THE CUSTOMER THAT THE AS-IFS1, AIRSEAL IFS, 120V WAS BEING USED ON 20JAN25 AND ¿WHILE THE CLAMP WAS PLACED, THEY ENCOUNTERED A SIGNIFICANT AMOUNT OF BLEEDING. BLOOD ENTERED THE AIRSEAL TUBE SET, AND THE FLUID INGRESS WARNING APPEARED. THE TUBE SET WAS NOT CHANGED AND THE UNIT SHUT DOWN. SURGEON CONVERTED TO OPEN.¿. IT WAS FURTHER REPORTED THAT THE BLOOD ENTERED THE AIRSEAL TUBE SET AS "...THEY GOT INTO HEAVY BLEEDING AND WERE USING EXCESSIVE SUCTION.". WHEN THE AIRSEAL FLUID INGRESS WARNING APPEARED THE TUBESET WAS NOT CHANGED OUT "...AS THEY STARTED RECEIVING ERROR CODES, THEY WERE ON CLAMP AND TRYING TO MANAGE BLEEDING SO POSSIBLY DUE TO MULTIPLE THINGS GOING ON. THEY ULTIMATELY CONVERTED TO OPEN INSTEAD OF SWAPPING THE TUBE SET.". THE IAS12-120LPI AND ASM-EVAC1 WILL BE LISTED AS CONCOMITANT DEVICES AND THERE WERE NO ALLEGATIONS AGAINST THEM. THIS REPORT IS BEING RAISED ON THE BASIS OF DEATH RELATED TO A DEVICE BEING USED DURING A PROCEDURE WITHOUT ANY REPORT OF A MALFUNCTION. | THE REPORTED DEVICE HAS BEEN RETURNED TO CONMED; HOWEVER, THE INVESTIGATION OF THE DEVICE HAS NOT BEEN COMPLETED. A SUPPLEMENTAL AND FINAL REPORT WILL BE FILED FOLLOWING THE COMPLETION OF THE DEVICE COMPLAINT INVESTIGATION. WE WILL CONTINUE TO MONITOR PER REGULATORY REQUIREMENTS TO HELP ENSURE PATIENT SAFETY.
CAPIOX | K022115 | THE USER FACILITY REPORTED THAT THE INVOLVED OXYGENATOR DID NOT FUNCTION PROPERLY. THE EVENT OCCURRED DURING CARDIOPULMONARY BYPASS. THE EVENT RESULTED IN DELAY IN BEGINNING OR CONTINUING THE SURGICAL PROCEDURE. AFTER 3 HOURS AND 40 MINUTES THE DEVICE'S ABILITY TO OXYGENATE WAS NOT FUNCTIONING. SWEEP WAS AT 1.2 LPM AND 60% FIO2, PO2 WAS 295 THEN NEXT ENTRY FELL TO 95 AND CONTINUED FALLING DESPITE INCREASES IN FIO2 AND SWEEP. A SECOND OXYGENATOR WAS ADDED TO THE RECIRCULATION LINE TO FACILITATE OXYGENATION. THEY DID NOT CHANGE OUT THE OXYGENATOR THEY ADDED ANOTHER OXYGENATOR TO THE CIRCUIT TO FINISH THE CASE. THE SURGERY WAS COMPLETED SUCCESSFULLY. NO BLOOD LOSS WAS REPORTED. A DEATH OCCURRED. THE PATIENT DID NOT DIE IN SURGERY, THEY DIED IN INTENSE CARE UNIT (ICU). | D6A: IMPLANTED DATE: DEVICE WAS NOT IMPLANTED. D6B: EXPLANTED DATE: DEVICE WAS NOT EXPLANTED. E3: OCCUPATION: PERFUSION SUPERVISOR. THE ACTUAL SAMPLE WAS NOT RETURNED. THE MANUFACTURING RECORD AND THE SHIPPING INSPECTION RECORD OF THE ACTUAL PRODUCT FOUND NO ANOMALY. NO OTHER SIMILAR REPORT OF THE PRODUCT WITH THE INVOLVED PRODUCT CODE/LOT NUMBER WAS FOUND. BASED ON THE INVESTIGATION RESULT, NO ANOMALY WAS FOUND IN THE MANUFACTURING RECORDS. AS A CAUSE OF DECREASING IN PAO2, BASED ON OUR PAST EXPERIENCE, FOLLOWING FACTORS WERE INFERRED. HOWEVER, SINCE THE ACTUAL SAMPLE COULD NOT BE CONFIRMED, IT WAS NOT POSSIBLE TO CLARIFY THE CAUSE OF OCCURRENCE. REGARDING FIO2 DURING CIRCULATION, THE AMOUNT OF OXYGEN SUPPLY WAS INSUFFICIENT COMPARED TO THE PATIENT'S OXYGEN CONSUMPTION. GAS EXCHANGE WAS HINDERED BY SOME FACTOR (E.G., BLOOD CLOT, WET LUNG). AS THE PATIENT'S METABOLISM BECAME MORE ACTIVE DUE TO REWARMING, THE OXYGEN SUPPLY WAS INSUFFICIENT COMPARED TO THE PATIENT'S OXYGEN CONSUMPTION, CAUSING A DECREASE IN SVO2. THE INSTRUCTIONS FOR USE (IFU) (CAPIOX RX05) HAS THE FOLLOWING WARNINGS, METHOD OF OPERATION AND PRECAUTION REGARDING GAS EXCHANGE FAILURE: "START GAS SUPPLY WITH V/Q=1, AND FIO2=100%, THEN MAKE ADJUSTMENTS BASED ON BLOOD GAS MEASUREMENTS." "MEASURE BLOOD GASES AND MAKE NECESSARY ADJUSTMENTS AS FOLLOWS. A. CONTROL PAO2 BY CHANGING CONCENTRATION OF OXYGEN IN VENTILATING GAS USING GAS BLENDER. TO DECREASE PAO2, DECREASE FIO2. TO INCREASE PAO2, INCREASE FIO2. B. CONTROL PACO2 BY CHANGING THE TOTAL GAS FLOW. TO DECREASE PACO2, INCREASE TOTAL GAS FLOW. TO INCREASE PACO2, DECREASE TOTAL GAS FLOW." "UPON PATIENT REWARMING, ADJUST O2 CONCENTRATION, GAS FLOW RATE AND BLOOD FLOW RATE BY INCREASING THEM AS NEEDED BASED ON AN INCREASE IN PATIENTS METABOLISM. FAILURE TO ADJUST THE GAS SUPPLY AND THE BLOOD FLOW RATE APPROPRIATELY MAY CAUSE INSUFFICIENT O2 SUPPLY NEEDED OR THE AMOUNT OF THE PATIENT'S GASEOUS METABOLISM." "A PHENOMENON CALLED WET LUNG MAY OCCUR WHEN WATER CONDENSATION OCCURS INSIDE FIBERS OF MICROPOROUS MEMBRANE OXYGENATORS WITH BLOOD FLOWING EXTERIOR TO THE FIBERS. THIS MAY OCCUR WHEN OXYGENATORS ARE USED FOR A LONGER PERIOD OF TIME. IF WATER CONDENSATION AND/OR A DECREASE IN PAO2 AND/OR AN INCREASE IN PACO2 IS NOTED DURING EXTENDED OXYGENATOR USE, BRIEFLY INCREASING THE GAS FLOW RATE MAY IMPROVE THE PERFORMANCE. INCREASE GAS FLOW RATE, TO 5 L/MIN FOR 10 SECONDS. DO NOT REPEAT THIS FLUSHING TECHNIQUE, EVEN IF OXYGENATOR PERFORMANCE IS NOT IMPROVED." TERUMO MEDICAL CORPORATION (TMC) (IMPORTER) REGISTRATION NO. (B)(4) IS SUBMITTING THIS REPORT ON BEHALF OF ASHITAKA FACTORY OF TERUMO CORPORATION (MANUFACTURER) REGISTRATION NO. (B)(4).
CONFIDA BRECKER CURVE GUIDEWIRE | K132623 | IT WAS REPORTED THAT DURING A TRANSCATHETER AORTIC VALVE IMPLANT PROCEDURE, RIGHT FEMORAL ACCESS, AN 18 FRENCH INTRODUCER (SENTRANT), AND NON-MEDTRONIC CLOSURE DEVICES (PROGLIDE) WERE USED. THE PRE-SHAPED GUIDEWIRE (CONFIDA) WAS ADVANCED INTO THE LEFT VENTRICLE, AND ARTERIAL PRESSURE CHANGES (HYPOTENSION) WERE OBSERVED. AN URGENT CARDIOLOGY CONSULTATION AND CONTROL ECHOCARDIOGRAM WERE PERFORMED, RULING OUT PERICARDIAL EFFUSION. CONTRAST INJECTIONS AT THE FEMORAL ACCESS SITES AND ABDOMINAL AORTA TO ASSESS FOR BLEEDING OR DISSECTIONS WERE NEGATIVE, AND AN AORTOGRAM SHOWED NO EVIDENCE OF BLEEDING OR ABNORMALITIES. DUE TO THE PATIENT'S DETERIORATING CONDITION, CARDIAC ARREST OCCURRED. RESUSCITATION MANEUVERS WERE INITIATED. ONCE STABILIZED, A PRE-IMPLANT BALLOON DILATION WAS PERFORMED WITH A 18X40 MM SEMI-COMPLIANT BALLOON. URGENT VALVE DEPLOYMENT PROCEEDED UNDER FLUOROSCOPIC GUIDANCE BY ADVANCING THE DELIVERY CATHETER SYSTEM (DCS) WITH THE PORT POSITIONED AT 3 O¿CLOCK. DEPLOYMENT OF THE VALVE WAS INITIATED IN THE IMPLANTATION PROJECTION. THE POSITIONING WAS ADJUSTED, NOTING THAT THE THREE CUSPS WERE VISIBLE UNDER CONTRAST INJECTION. UPON REACHING THE THIRD NODE, PACING WAS INITIATED VIA THE GUIDEWIRE, AND THE VALVE WAS OPENED TO 80%. THE PACEMAKER WAS THEN DISCONNECTED, AND CONTRAST INJECTIONS CONFIRMED VALVE POSITIONING, WITH A 2 MM DEPTH IN BOTH THE NON-CORONARY AND LEFT CORONARY SINUSES. HOWEVER, THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE, LEADING THE SPECIALIST TO PROCEED WITH FULL VALVE DEPLOYMENT. THE PATIENT SUBSEQUENTLY EXPERIENCED ANOTHER CARDIAC ARREST, REQUIRING FURTHER RESUSCITATION EFFORTS. CARDIOPULMONARY RESUSCITATION WAS ADMINISTERED FOR 37 MINUTES. A CONTROL INJECTION REVEALED THAT THE VALVE HAD MIGRATED TOWARD THE NON-CORONARY SIDE, WITH EVIDENCE OF RIGHT CORONARY ARTERY OCCLUSION. AN ATTEMPT WAS MADE TO CANNULATE THE RIGHT CORONARY ARTERY THROUGH THE VALVE, BUT ALL ATTEMPTS WERE UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED. | CONTINUATION OF D10: SECTION D INFORMATION REFERENCES THE MAIN COMPONENT OF THE SYSTEM. OTHER RELEVANT DEVICE(S) ARE: PRODUCT ID EVOLUTFX-23 (SERIAL: (B)(6)); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE (B)(6) 2025; EXPLANT DATE N/A PRODUCT ID D-EVOLUTFX-2329 (LOT: 0012483145); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID L-EVOLUTFX-2329 (LOT: 0012526083); PRODUCT TYPE: 0195-HEART VALVES; IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID SENTRANT SHEATH (LOT: UNKNOWN); PRODUCT TYPE: 0199-INTRODUCER; IMPLANT DATE N/A; EXPLANT DATE N/A MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REP DREAMSTATION AUTO BIPAP, DOM | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S REPRESENTATIVE THAT THE PATIENT HAD PASSED AWAY AND THEY WOULD LIKE TO DONATE THE PATIENT'S DEVICE. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH OF THE PATIENT. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K183387 | PHILIPS RECEIVED A COMPLAINT ON THE PATIENT INFORMATION CENTER IX INDICATING THAT THE CUSTOMER REQUESTED INFORMATION ON WHETHER THE DEVICE HAD AN 'OPTIONAL FUNCTION' FOR EGC ANALYSIS AND ALARMS DURING POSTOPERATIVE CARE AND CARDIAC ARREST. PRIOR TO THE ARREST, THE PATIENT WAS PACED AT 80 BPM WITH AN EXTERNAL PACEMAKER VIA THE JUGULAR VEIN WITH A PLANNED INSERTION OF AN IMPLANTABLE PACEMAKER. THE PATIENT HAD AN UNWITNESSED CARDIAC ARREST, WAS FOUND UNCONSCIOUS, AND CPR WAS INITIATED. THE DEFIBRILLATOR DISPLAYED THE PACED RHYTHM BUT THERE WAS NO PULSE. AN ULTRASOUND WAS PERFORMED, WHICH REVEALED CARDIAC STANDSTILL, AND THE PATIENT WAS PRONOUNCED DEAD. IT WAS CLARIFIED THAT THE DEVICE FUNCTIONED AS INTENDED AND THAT THERE WAS AN ASYSTOLE ALARM GENERATED ON THE PIC IX WHICH NOTIFIED THE NURSING STAFF. THE ALARM WAS SILENCED AT THE BEDSIDE 2 MINUTES LATER. THE RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL, AND THE PATIENT PASSED AWAY. THE DEATH WAS UNRELATED TO THE DEVICE AND THERE WAS NO DEVICE MALFUNCTION. | THIS COMPLAINT IS A SUPPLEMENTAL TO (9610816-2024-001005) DUE TO INCORRECT REGISTRATION NUMBER USED. THE COMPLAINT WAS ESCALATED FOR TECHNICAL INVESTIGATION, AND A CLINICAL APPLICATION SPECIALIST REVIEWED THE EVENT, REVEALING THE ECG RHYTHM AT 1826 IS PRESENT BUT THEN BECOMES VERY CHAOTIC OVER THE NEXT FEW SECONDS, AND TRANSITIONS TO ECG LEADS OFF WITH NO DATA STORED AFTER THAT TIME. THE LOGS ALSO INDICATE THE PATIENT'S DATA WAS ACCESSED BY A USER FIVE TIMES BETWEEN 1830 AND 1900. THE DEVICE WAS FOUND TO BE OPERATING PER SPECIFICATIONS, INCLUDING ALARMS AND CAREEVENT NOTIFICATIONS. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE DEVICE WAS FUNCTIONING AS INTENDED, AND THERE WAS NO TROUBLE FOUND WITH THE DEVICE. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. SECTION E REPORTING INSTITUTION PHONE #: (B)(6), SECTION E REPORTER PHONE #: (B)(6).
DA VINCI ENERGY | K191280 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED MEDIASTINAL MASS RESECTION, THE BRACHIOCEPHALIC ARTERY WAS INJURED AND THE PROCEDURE WAS CONVERTED TO OPEN. THE PATIENT EXPIRED ON POST-OPERATIVE DAY ONE. THE SURGEON REPORTED THAT THE ARTERY WAS INJURED WHEN USING THE SYNCHROSEAL INSTRUMENT TO DISSECT LYMPH NODES. THE ARTERY WAS ADHERED TO TISSUE AND TO THE LYMPH NODES, RESULTING IN THE ARTERY NOT BEING VISIBLE TO THE SURGEON. THE ARTERY BEGAN TO BLEED AFTER IT WAS GRASPED WITH THE SYNCHROSEAL INSTRUMENT. THE PROCEDURE WAS CONVERTED TO THORACOTOMY. THERE WAS NO MALFUNCTION OF THE SYNCHROSEAL INSTRUMENT OR ANY OTHER DA VINCI PRODUCTS USED DURING THIS PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, CURVED BIPOLAR DISSECTOR, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, MEDIUM-LARGE CLIP APPLIER, SYNCHROSEAL. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A CATASTROPHIC INJURY TO THE INNOMINATE ARTERY WHILE ATTEMPTING TO REMOVE A MEDIASTINAL MASS AND ADHERENT LYMPH NODES IN THE VICINITY OF THIS LARGE ARTERY. THE SURGEON COMMENTED THEY DID NOT RECOGNIZE THE ARTERY AS THEY WERE ATTEMPTING TO DISSECT THE NEARBY LYMPH NODES AND INADVERTENTLY GRASPED THIS LARGE VESSEL. THIS SYNCHROSEAL IS NOT INTENDED FOR MAJOR ARTERIES AND THIS IS A SURGICAL TECHNIQUE ERROR. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
DAVINCI X | K171294 | IT WAS REPORTED IN A LOCAL NEWSPAPER ARTICLE THAT A PATIENT EXPIRED AFTER A DA VINCI-ASSISTED HYSTERECTOMY FOR METASTATIC UTERINE CANCER. THE PROCEDURE WAS CONVERTED TO OPEN DUE TO UNSPECIFIED COMPLICATIONS. DURING THE DA VINCI-ASSISTED PORTION OF THE PROCEDURE, LYMPH NODES WERE REMOVED FROM THE PATIENT'S LEFT PELVIS. THE DA VINCI SYSTEM WAS REPOSITIONED TO REMOVE LYMPH NODES ALONG THE ABDOMINAL AORTA, WHICH REQUIRED PLACEMENT OF THREE ADDITIONAL TROCARS THAT WERE INSERTED UNDER VISUAL GUIDANCE. AFTER THE DA VINCI SYSTEM WAS REPOSITIONED, INSUFFLATION WAS LOST AND THE ABDOMINAL CAVITY COLLAPSED. EFFORTS TO RE-INSUFFLATE THE ABDOMEN FAILED. THE PROCEDURE WAS CONVERTED TO OPEN AND LASTED FOR 11 HOURS AND 21 MINUTES. THE ARTICLE STATES THAT THE PATIENT'S SMALL INTESTINE WAS INJURED AND THE LEFT URETER WAS SEVERED. DETAILS OF HOW OR WHEN THESE INJURIES OCCURRED IS NOT PROVIDED. THE ARTICLE STATED INJURIES REPORTEDLY CAUSED LEAKS LEADING TO PERITONITIS AND SEVERE SEPSIS, RESULTING IN A TOTAL OF 20 ADDITIONAL UNSPECIFIED PROCEDURES PERFORMED. THE PATIENT EXPIRED ON POST-OPERATIVE DAY 47. THE NEWSPAPER ARTICLE DID NOT INDICATE WHETHER THE COMPLICATIONS OCCURRED DURING THE ROBOTIC PORTION OR THE OPEN PORTION OF THE PROCEDURE, AS THE INJURIES WERE NOT DISCOVERED UNTIL OVER A DAY LATER. THE SURGEON IDENTIFIED IN THE ARTICLE WAS CONTACTED BUT DECLINED TO PROVIDE ANY INFORMATION. | REVIEW OF THE SYSTEM LOGS FOUND NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, SMALL GRASPING RETRACTOR, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THE PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A ROBOTIC HYSTERECTOMY FOR UTERINE CANCER AND THE PROCEDURE WAS EVENTUALLY CONVERTED TO OPEN. INJURIES TO THE BOWEL OCCURRED BUT THE DETAILS OF HOW OR WHEN THESE INJURIES OCCURRED IS NOT PROVIDED. THE PATIENT SUBSEQUENTLY UNDERWENT MANY ADDITIONAL PROCEDURES AND ULTIMATELY EXPIRED 47 DAYS AFTER THE INDEX OPERATION. THE REASONS FOR THESE ADDITIONAL PROCEDURES, FINDINGS, AND OTHER DETAILS ARE NOT PROVIDED. HOW THE PATIENT DEVELOPED SEPSIS, HOW THEY DIED, AND THE DETAILS OF EVENTS THAT LED TO THEIR DEATH IS NOT PROVIDED. NO ALLEGATION WAS MADE REGARDING ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | K211575 | AN MHRA USER REPORT HAS BEEN RECEIVED, REPORTED BY A HCP "PATIENT HAD USED THE OMNIPOD PUMP TO DELIVER MULTIPLE INSULIN BOLUSES FOR HYPERGLYCEMIA CORRECTION, HOWEVER IT DID NOT APPEAR THAT THE INSULIN BOLUSES WORKED TO REDUCE GLUCOSE LEVELS". PATIENT HEALTH IMPACT IS DEATH. THE REPORTING HCP HAS BEEN IN CONTACT WITH INSULET ON (B)(6)2025 TO ADVISE THAT THE PATIENT WAS FOUND DEAD IN HIS RESIDENCE IN (B)(6) 2024 AND THE CORONER HAS ADVISED THAT THE CAUSE OF DEATH IS DIABETIC KETOACIDOSIS. THE HCP RECEIVED THE DEVICE ON (B)(6)2025. THE PATIENT'S BLOOD GLUCOSE LEVELS WERE REPORTED TO BE ABOVE 27.8 MMOL/L (500.4 MG/DL) BEFORE THEY DECEASED. | THE DEVICE HAS NOT BEEN RETURNED TO INSULET. THE DEVICE IS WITH THE REPORTING HEALTHCARE PROFESSIONAL. IF THE DEVICE IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE SUBMITTED WITH THE INVESTIGATION RESULTS. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED DEATH. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. THE INVESTIGATION IS ONGOING.
MEDIHONEY GEL IN TUBE, .5 FL OZ TWISTOFF - STERILE | K101793 | N/A. | FAILURE ANALYSIS/ROOT CAUSE - THIS COMPLAINT DESCRIBES USE OF MEDIHONEY FOR TREATMENT OF A WOUNDED TOE AND STATES THAT IT DID NOT WORK AND "MADE IT WORSE". IN A FOLLOW UP COMMUNICATION, THE CUSTOMER PROVIDED THE SKU 31805, HOWEVER, THEY DID NOT PROVIDE ANY CLARIFICATION REGARDING "MADE IT WORSE." AS THE REPORTER DESCRIBED THE PATIENT'S CONDITION IN THE INITIAL COMPLAINT AS "AMPUTATION WAS IMMINENT". A MEDICAL ASSESSMENT WAS REQUESTED TO INTEGRA'S MEDICAL AFFAIRS DIRECTOR AND THE MEDICAL ASSESSMENT DETERMINED THE ADVERSE EVENT WAS POSSIBLY INFECTION. THEY CONCLUDED THAT A CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND AN ADVERSE EVENT OF INFECTION IS HIGHLY UNLIKELY. FURTHERMORE, THE REPORTER DID NOT PROVIDE ANY INFORMATION RELATED TO CHANGES IN TEXTURE, COLOR, SMELL, OR CONSISTENCY OF THE PRODUCT, NOR PRESENCE OF FUZZY PATCHES OR FOREIGN PARTICULATE, WHICH, IF PRESENT, COULD INDICATE MICROBIAL PROLIFERATION. IN AN ADDITIONAL COMMUNICATION FROM THE CUSTOMER, THEY STATE THEY DO NOT HAVE ACCESS TO ADDITIONAL MEDICAL RECORDS AND THAT THE PRODUCT HAS BEEN DISCARDED. DUE TO THE LACK OF FURTHER INFORMATION REGARDING THE PRODUCT, SUCH AS PHOTOS OR RETURNED PRODUCT, AND FURTHER INFORMATION REGARDING THE PATIENT, SUCH AS MEDICAL HISTORY, THIS COMPLAINT IS UNCONFIRMED. A COMPLAINT ADJUDICATION REPORT WAS INITIATED TO PROVIDE 3RD PARTY MEDICAL ASSESSMENT FROM BOARD-CERTIFIED INFECTION DISEASE SPECIALIST (INITIALS (B)(6)). THIS REVIEW INDICATED THAT BASED ON THE INFORMATION PROVIDED, AND IN THE ABSENCE OF DETAILED CLINICAL DOCUMENTATION, IT IS THE PROFESSIONAL OPINION OF THE OUTSIDE EXPERT THAT EVEN IF THE MEDIHONEY WAS APPLIED FROM A POTENTIALLY COMPROMISED TUBE, IT DID NOT CONTRIBUTE TO ANY OF THE ADVERSE OUTCOME REPORTED IN THE COMPLAINT AS THE PATIENT PRESENTED WITH SIGNIFICANT COMORBIDITIES THAT OFFER MORE PLAUSIBLE EXPLANATIONS FOR THEIR CLINICAL DETERIORATION OR DEATH. POST THE COMPLETION OF THIS ASSESSMENT, FURTHER INFORMATION PROVIDED BY THE CUSTOMER INDICATES THAT THIS PRODUCT COULD NOT BE PART OF THE RECALL, FOR WHICH THE EXPERT MADE THIS ASSESSMENT. THE CUSTOMER HAD PROVIDED ADDITIONAL INFORMATION ON THIS COMPLAINT WITH THE ORIGINAL ORDER AND SHIPPING INFORMATION FOR THE PRODUCT, PURCHASED ON AMAZON IN MARCH 2023. THE CUSTOMER, UPON THE INITIATION OF THE COMPLAINT, HAD THE UNDERSTANDING THAT THE PRODUCT IN QUESTION WAS PART OF THE RECALL FOR SKU 31815 (LOTS 2328, 2332 AND 2333). THE LOT INFORMATION WAS NOT PROVIDED FOR THIS COMPLAINT HOWEVER, THE EARLIEST LOT OF THE THREE IN THE SCOPE OF THE RECALL, LOT 2328, WAS RELEASED ON 18AUG2023. THIS IS 5 MONTHS AFTER THE PRODUCT IN QUESTION WAS PURCHASED, MEANING THE COMPLAINT PRODUCT COULD NOT HAVE BEEN IN SCOPE OF THE RECALL LOTS, EVEN WITHOUT CONFIRMATION OF WHICH LOT NUMBER THE PRODUCT IS. THE PHOTOGRAPH OF THE PRODUCT PURCHASED ALSO INDICATES THAT IT IS SKU 31805, WHICH IS ALSO OUTSIDE THE SCOPE OF THE RECALL. ADDITIONALLY, THE CUSTOMER PROVIDED PURCHASE DETAILS SHOWING THE PRODUCT WAS PURCHASED FROM AMAZON, SELLER PREFERRED PHARMACY PLUS. FEDERAL LAW REQUIRES THAT THE PRODUCT BE SOLD AND USED ONLY UNDER THE SUPERVISION OF A PHYSICIAN PER THE IFU. THERE ARE NO LEGAL SALES OF ANY MEDIHONEY PRODUCTS ON AMAZON AND INTEGRA HAS CONFIRMED THAT AMAZON IS SELLING COUNTERFEIT PRODUCTS UNDER THE PRODUCT NAME OF MEDIHONEY. WITHOUT THE PRODUCT RETURNED FOR EVALUATION OR PHOTOGRAPHS PROVIDED, THERE IS NO WAY TO CONFIRM IF THE PRODUCT IS A GENUINE INTEGRA DEVICE. INTEGRA REGULATORY HAS BEEN IN COMMUNICATION WITH THE FDA ABOUT THE INAPPROPRIATE SALES OF MEDIHONEY. THERE IS NO CONFIRMATION IF THE CUSTOMER FOLLOWED THE INSTRUCTIONS FOR USE ON THE INDICATIONS, PRECAUTIONS, OR INSTRUCTIONS. IF THE CUSTOMER DID NOT FOLLOW GUIDANCE FROM THEIR DOCTOR, OR INSTRUCTIONS FOR USE OF THE PRODUCT, INFECTION MAY BE CAUSED BY ANY NUMBER OF SOURCES UNRELATED TO THE USE OF A MEDIHONEY PRODUCT. A CORRECTIVE AND PREVENTATIVE ACTION (CAPA) HAS BEEN CREATED FOR PRESCRIPTION/OTC DISTRIBUTION ISSUE TO ADDRESS HOW MEDIHONEY WAS DISTRIBUTED FOR SALE WITHOUT A PRESCRIPTION. ADDITIONAL INFORMATION RECEIVED: THE PATIENT'S DAUGHTER STATED: "MY FATHER WAS 82 YEARS OLD. I DON'T HAVE ANY OF MY FATHER'S MEDICAL INFORMATION AVAILABLE. HE WAS A DIABETIC AND HAD ASTHMA. NO OTHER PROBLEMS THAT I KNOW OF. I CAN'T PROVIDE MEDICAL RECORDS OR ANY OF THAT INFORMATION."
PARAPAC PLUS | K123957 | IT WAS REPORTED THAT THE FILTER ON THE DEVICE WHERE THE CIRCUIT WAS LINKED TO WAS NOT WORKING AND NOT ALARMING. THE PATIENT WAS REPORTED TO HAVE CODED AND EXPIRED DUE TO ACUTE RESPIRATORY FAILURE. THE DATE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED ON THE CAUSE OF DEATH. TO DATE NO FURTHER INFORMATION HAS BEEN PROVIDED. A CLINICAL REVIEW IS IN PROGRESS. | B2 DATE OF DEATH IS UNKNOWN. INVESTIGATION INCLUDING ROOT CAUSE ANALYSIS IS IN PROGRESS. A SUPPLEMENTAL MDR WILL BE FILED AS NECESSARY IN ACCORDANCE WITH 21 CFR 803.56 WHEN ADDITIONAL REPORTABLE INFORMATION BECOMES AVAILABLE.
HOBBS MEDICAL TOUCHSOFT COAGULATOR | K093079 | THE USER FACILITY REPORTED THAT A PATIENT REPORTED ABDOMINAL PAIN AND RETURNED TO THE HOSPITAL ONE DAY AFTER A PROCEDURE WHICH INCLUDED USE OF THE TOUCHSOFT COAGULATOR FOR TREATMENT OF TWO SMALL ARTERIOVENOUS MALFORMATIONS (AVMS) IN THE SECOND PORTION OF THE DUODENUM. UPON RETURN TO THE HOSPITAL, THE PATIENT WAS FOUND TO HAVE AN INFECTION IN THE DUODENUM AND A PERFORATION IN THE THIRD PORTION OF THE DUODENUM WHICH REQUIRED SURGERY. AFTER THE SURGERY, THE PATIENT REMAINED IN THE ICU FOR APPROXIMATELY SIX WEEKS AND THEN PASSED AWAY. | US ENDOSCOPY IS THE COMPLAINT FILE ESTABLISHMENT AND REPACKAGER/RELABELER FOR THE DEVICE SUBJECT OF THIS EVENT. THE DEVICE IS MANUFACTURED BY HOBBS MEDICAL, INC. AND DISTRIBUTED BY US ENDOSCOPY. ON FOLLOW-UP WITH THE USER FACILITY, US ENDOSCOPY LEARNED THAT THE PATIENT'S MEDICAL HISTORY AND COMORBIDITIES INCLUDED GASTRIC CARCINOID TUMOR, COLONIC POLYPS, RENAL CELL CANCER, HYPERTENSION, AND HYPERLIPIDEMIA. UNRELATED TO USE OF THE TOUCHSOFT COAGULATOR FOR AVM TREATMENT, THE PROCEDURE SUBJECT OF THIS REPORTED EVENT INCLUDED BIOPSY OF A 10MM NODULE IN THE DUODENAL BULB. TWO SMALL NONBLEEDING ANGIODYSPLASIAS WERE ALSO NOTED DURING THE PROCEDURE. THE USER FACILITY STATED THAT THEY DO NOT KNOW EXACTLY WHAT CAUSED THIS EVENT TO HAPPEN; THE DOCTOR STATED HE DID NOT THINK THAT THE PROBE CONTRIBUTED TO THE PATIENT'S PASSING BUT COULD NOT RULE IT OUT. ADDITIONALLY, US ENDOSCOPY LEARNED THAT THE USER FACILITY USED A THIRD PARTY ERBE GENERATOR SET AT 30-40 WATTS WITH THE TOUCHSOFT COAGULATOR DURING THE PROCEDURE. THE SETTING WAS CONTRARY TO THE TOUCHSOFT COAGULATOR INSTRUCTIONS FOR USE (IFU) WHICH STATES, "THE TOUCHSOFT COAGULATOR CAN BE USED WITH...GENERATOR MODEL ERBE, RANGE IN WATTS 45-60." THE IFU STATES, "CAUTION: DO NOT APPLY EXCESSIVE HEAT, FORCE OR PROLONGED APPLICATION TIME, AS THIS MAY CAUSE DAMAGE TO THE DEVICE, STICKING OF THE DEVICE TO THE TISSUE OR OTHER UNDESIRABLE EFFECTS, INCLUDING POSSIBLE PERFORATION." THE IFU FURTHER STATES, "ALWAYS CHOOSE THE LOWEST POWER SETTING AND SHORTEST TIME DURATION (3 SECONDS OR LESS) THAT WILL ACHIEVE THE DESIRED CLINICAL RESULT." THE DEVICE SUBJECT OF THE EVENT WAS DISCARDED BY USER FACILITY PERSONNEL AND IS NOT AVAILABLE FOR EVALUATION. THE USER FACILITY WAS UNABLE TO PROVIDE A LOT NUMBER FOR THE DEVICE SUBJECT OF THE EVENT. A REVIEW OF DISTRIBUTION RECORDS INDICATES THAT THE USER FACILITY HAS ORDERED 5 LOTS OF THIS PRODUCT SINCE 2022. THERE HAVE BEEN NO OTHER COMPLAINTS ASSOCIATED WITH ANY OF THESE 5 LOTS. THE IFU INCLUDES THE FOLLOWING WARNINGS AND CAUTIONS, "WARNING: POTENTIAL ADVERSE EFFECTS INCLUDE, BUT ARE NOT LIMITED TO: RISK OF PERFORATION, HEMORRHAGE, INFECTION, THERMAL BURNS, AND EXPLOSION. CAUTION: ONLY THE USER CAN EVALUATE THE CLINICAL FACTORS INVOLVED WITH EACH PATIENT TO DETERMINE IF THE USE OF THIS DEVICE IS INDICATED. THE USER MUST THEN DECIDE ON THE SPECIFIC TECHNIQUE AND PROCEDURE THAT WILL BEST ACCOMPLISH THE DESIRED CLINICAL EFFECT. DO NOT USE THIS DEVICE FOR ANY PURPOSE OTHER THAN THE STATED INDICATION FOR USE. WARNING: THE USER MUST THOROUGHLY UNDERSTAND THE TECHNICAL PRINCIPLES, CLINICAL TECHNIQUES AND RISKS ASSOCIATED WITH THE USE OF HIGH FREQUENCY ELECTROSURGICAL ENERGY IN GENERAL AND THE CHARACTERISTICS OF THEIR CHOSEN ELECTROSURGICAL GENERATOR IN PARTICULAR BEFORE USING THIS PRODUCT." A US ENDOSCOPY ACCOUNT MANAGER OFFERED THE USER FACILITY IN-SERVICE TRAINING ON THE PROPER USE AND OPERATION OF THE TOUCHSOFT COAGULATOR; HOWEVER, THE USER FACILITY DECLINED. NO ADDITIONAL ISSUES HAVE BEEN REPORTED.
SABLE | K192115 | IT WAS REPORTED A WAS SABLE SPACER WAS REMOVED FROM A PATIENT WHO THEN LATER DIED. | THE DEVICE WAS NOT AVAILABLE FOR EVALUATION. IN THE INITIAL SURGERY, THE PATIENT HAD A PRIOR KYPHOPLASTY WITH CEMENT AT T12, WHICH WAS THE SUPERIOR VERTEBRAL BODY OF THE SITE OF INTENDED INTERBODY FUSION. IT IS UNKNOWN WHAT SCREW SYSTEM WAS PRESENT IN THE PATIENT DUE TO THE KYPHOPLASTY. THE INITIAL PROCEDURE WAS ON THURSDAY, (B)(6) 2025. A SABLE SPACER WAS PLACED AT T12-L1 WITH SUPPLEMENTAL FIXATION USING THE REVERE STABILIZATION SYSTEM. IN THE DAYS AFTER SURGERY, THE PATIENT COMPLAINED OF PAIN. A REVISION CASE WAS PERFORMED THE FOLLOWING MONDAY, (B)(6) 2025, TO REMOVE THE IMPLANT. THE IMPLANT WAS REMOVED WITHOUT ISSUE, HOWEVER IN THE FOLLOWING DAYS AFTER SURGERY THE PATIENT PASSED AWAY. IMAGING OF THE SURGICAL SITE WAS PROVIDED. IN THE IMAGING, THE SABLE SPACER APPEARS TO BE IN AN EXPANDED STATE. THE IMPLANT APPEARS TO BE SUBSIDED INTO THE T12 VERTEBRA WITH THE DISTAL TIP OF THE IMPLANT ANGLED CAUDALLY, WITH IMPLANT ENDPLATES NOT PARALLEL TO THE T12 OR L1 VERTEBRAL ENDPLATES. IT APPEARS THAT THE T12 VERTEBRA MAY HAVE FRACTURED, BUT THIS CANNOT BE DEFINITIVELY DETERMINED BASED ON THE PROVIDED IMAGING. NO DETERMINATIONS COULD BE MADE AS TO THE CAUSE OF THE REPORTED ISSUE. THE FOLLOWING SECTIONS HAVE BEEN UPDATED: B4; D6B; G6; H2; H3; H6; H10.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED INFORMATION THAT A TRILOGY EVO DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. LIMITED INFORMATION WAS PROVIDED BY THE CUSTOMER; HOWEVER THE CUSTOMER DID INDICATE THAT THE PATIENT EXPIRED DUE TO A NON-SPECIFIC NATURAL DISEASE PROGRESSION. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION AND WAS FOUND TO PASS TESTING AND TO BE IN OVERALL GOOD CONDITION.
LEAD EXTRACTION EVOLUTION RL CONTROLLED-ROTATION DILATOR SHEATH SET | K141148 | AS REPORTED TO CUSTOMER RELATIONS: CUSTOMER REPORTS DURING THE USE OF A LEAD EXTRACTION EVOLUTION RL CONTROLLED-RL CONTROLLED-ROTATION DILATOR SHEATH SET (G23747), THE EVOLUTION CHEWED THROUGH THE OUTER SHEATH WHILE BEING ACTIVATED INSIDE OF SAID SHEATH. THIS CAUSED THE SYSTEM TO NOT BE COAXIAL WITH THE LEAD AND CAUSED A SUPERIOR VENA CAVA (SVC) TEAR WHICH RESULTED IN A PATIENT DEATH. | BLANK FIELDS ON THIS FORM INDICATED THE INFORMATION IS UNKNOWN, UNCHANGED, OR UNAVAILABLE. THE UDI INFORMATION PER 21CFR 803.52(C)(4) AND 21 CFR 803.50(B) THAT IS REQUIRED IS REASONABLY UNKNOWN DUE TO THE DEVICE ITSELF WAS DISCARDED/NOT RETURNED FOR INVESTIGATION AFTER MANUFACTURER REQUESTED FOR THE RETURN AND THE LOT OF THE DEVICE WAS UNKNOWN. . THE DEVICE WAS NOT RETURNED FOR THE COMPLAINT; THEREFORE, A PHYSICAL INVESTIGATION COULD NOT BE PERFORMED, AND THE CUSTOMER'S COMPLAINT IS ACKNOWLEDGED, BUT COULD NOT BE CONFIRMED, OTHER THAN BY THE CUSTOMER'S TESTIMONY. THE COMPLAINT/EVENT THAT WAS ENTERED AND REPORTED WITHIN TRACKWISE: "THE EVOLUTION CHEWED THROUGH THE OUTER SHEATH WHILE BEING ACTIVATED INSIDE OF SAID SHEATH. THIS CAUSED THE SYSTEM TO NOT BE COAXIAL WITH THE LEAD AND CAUSED AN SVC TEAR WHICH RESULTED IN A DEATH." THE DEVICE HISTORY RECORD (DHR) WAS UNABLE TO BE VIEWED DUE TO THE LOT NUMBER SPECIFIC TO THIS COMPLAINT WAS UNKNOWN. THIS COMPLAINT MODE IS BEING MONITORED, TRACKED AND TRENDED PER THE CVI POST MARKET SURVEILLANCE AND COMPLAINT HANDLING PROCESSES. A RISK ASSESSMENT WILL BE PERFORMED AND DOCUMENTED IN TRACKWISE. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803. THIS REPORT IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT MADE IN IT IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED; THAT A DEATH OR SERIOUS INJURY OCCURRED; OR THAT ANY COOK DEVICE CAUSED OR CONTRIBUTED TO; OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
NOXBOXI NITRIC OXIDE DELIVERY SYSTEM | K231823 | SUMMERLIN HOSPITAL REPORTED NITRIC OXIDE (NO) DOSE FLUCTUATIONS ON A NOXBOXI DEVICE DURING INHALED NITRIC OXIDE (INO) THERAPY. AS REPORTED, THE DEVICE WAS USED OFF-LABEL ON A <34-WEEK NEONATE PATIENT AND WAS ISSUING HIGH AND LOW DOSE ALARMS. HOSPITAL STAFF ATTEMPTED TO RESOLVE THE ISSUE BY REPLACING THE DEVICE WITH A DIFFERENT NOXBOXI UNIT; HOWEVER, THE DOSE FLUCTUATIONS PERSISTED. THREE HOURS AFTER THE INITIAL REPORT WAS MADE TO LINDE, THE HOSPITAL NOTIFIED LINDE THAT THE PATIENT HAD EXPIRED. THERE WAS NO REPORTED RELATIONSHIP BETWEEN THE DEVICE AND PATIENT'S OUTCOME. THE PATIENT WAS ON A BUNNELL JET LIFEPULSE VENTILATOR USING HIGH FREQUENCY JET VENTILATION. | THIS REPORT CONCERNS TWO NOXBOXI DEVICES (SERIAL NUMBERS (B)(6)) AND A REPORTED ADVERSE EVENT ON FEBRUARY 7, 2025. DESPITE A REQUEST FOR DEVICE RETURN, THE HOPITAL DECLINED TO PROVIDE THE DEVICES FOR PHYSICAL EVALUATION. LOG FILE DATA, PROVIDED BY THE HOSPITAL, INDICATES NO RECORDED DEVICE ACTIVITY: (B)(6) FROM (B)(6) 2025, AND (B)(6) FROM (B)(6) 2024 TO (B)(6) 2025. THEREFORE, NEITHER DEVICE WAS POWERED ON FEBRUARY 7, 2025, THE DATE OF THE REPORTED EVENT. A FORMAL REQUEST FOR DEVICE RETURN AND MORE INFORMATION ABOUT THE EVENT HAS BEEN SUBMITTED. HOWEVER, WE HAVE NOT RECEIVED FURTHER INFORMATION REGARDING THE PATIENT'S CONDITION OR CAUSE OF DEATH. CONSEQUENTLY, BASED ON THE AVAILABLE DATA, A CAUSAL RELATIONSHIP BETWEEN DEVICE FUNCTION AND THE REPORTED EVENT CANNOT BE ESTABLISHED.
DAVINCI X | K171294 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | THE FOLLOWING FIELDS WERE UPDATED: B2,B3, H11. ALL MULTI-USE INSTRUMENTS USED IN THE CASE HAVE BEEN USED IN SUBSEQUENT PROCEDURES WITH THE EXCEPTION OF THE SINGLE-USE INSTRUMENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, LARGE NEEDLE DRIVER, TWO CADIERE FORCEPS INSTRUMENTS, AND A SYNCHROSEAL.
LINEAR 7.5FR. 40CC IAB WITH ACCESSORIES, US ONLY | K041281 | IT WAS REPORTED THAT THE INTRA-AORTIC BALLOON (IAB) CATHETER EXPERIENCED A RUPTURE. UNFORTUNATELY, THE PATIENT DID NOT SURVIVE. | UPDATED FIELDS: B4, B5, D8, D9, G3, G6, H2, H3, H6 (MEDICAL DEVICE ¿ PROBLEM CODE, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS AND INVESTIGATION CONCLUSIONS), H8, H11 CORRECTED FIELDS: D7A. THE IAB WAS RETURNED WITH THE MEMBRANE COMPLETELY UNFOLDED AND BLOOD ON THE INTERIOR AND EXTERIOR OF THE CATHETER. THE SHEATH HUB WAS OVER THE MEMBRANE. THE SHEATH TUBING APPEARS TO BE BROKEN OFF FROM THE SHEATH AND NOT RETURNED. THE EXTENDER TUBING WAS ALSO RETURNED. AN UNDERWATER LEAK TEST OF THE BALLOON, CATHETER, Y-FITTING, EXTRACORPOREAL AND EXTENDER TUBING WAS PERFORMED AND A LEAK WAS DETECTED ON THE MEMBRANE APPROXIMATELY 1.3CM FROM THE REAR SEAL MEASURING 0.013CM IN LENGTH. THE REPORTED PROBLEM WAS MOST LIKELY TRIGGERED BY A LEAK WHICH WAS FOUND ON THE MEMBRANE. UNDER MAGNIFICATION, A WHITISH PATCH WAS OBSERVED AROUND THE LEAK. THIS WHITISH PATCH IS THE TYPICAL APPEARANCE OF AN ABRASION MARK WHICH IS CAUSED BY CALCIFIED PLAQUE IN THE AORTA. A LOT HISTORY RECORD REVIEW WAS COMPLETED FOR THE REPORTED PRODUCT. NO NONCONFORMANCES WERE FOUND THAT ARE CONSIDERED TO BE RELATED TO THE EVENT. COMPLAINT RECORD ID #(B)(4).
WATCHMAN TRUSTEER ACCESS SYSTEM | K240018 | IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A 31MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LEFT ATRIUM (LA) OF THE PATIENT. WHILE PREPARING TO INSERT THE PIGTAIL CATHETER THE PATIENT COUGHED, AND AIR WAS SUCTIONED INTO THE LA THROUGH THE WAS. AN AIR EMBOLISM WAS SEEN IN THE LA AND THE PHYSICIAN ASPIRATED THE AIR BUBBLES OUT OF THE PATIENT. THE PHYSICIAN CONTINUED THE PROCEDURE WITH THE HOPE THAT ANY REMAINING AIR COULD BE TRAPPED BY IMPLANTING THE CLOSURE DEVICE. THE WDS WAS INSERTED, AND THE CLOSURE DEVICE WAS SUCCESSFULLY IMPLANTED IN THE LAA OF THE PATIENT. SHORTLY AFTER THE CLOSURE DEVICE HAD BEEN RELEASED THE PATIENT BEGAN TO EXPERIENCE ST SEGMENT ELEVATION WHICH LED TO COMPLETE HEART BLOCK. THE PATIENT RECEIVED CARDIOPULMONARY RESUSCITATION AND IV MEDICATION BEFORE BECOMING STABLE. THE PATIENT WAS STILL RECOVERING FROM THIS EVENT THE NEXT DAY IN THE INTENSIVE CARE UNIT. THE PHYSICIAN BELIEVES THAT THE AIR EMBOLISM REACHED THE PATIENT'S BRAIN.
WATCHMAN FXD CURVE ACCESS SYSTEM | K212228 | IT WAS REPORTED THAT AN AIR EMBOLISM OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN FXD CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS AND THE DILATOR IN THE LEFT ATRIUM (LA) OF THE PATIENT. WHEN THE PHYSICIAN REMOVED THE DILATOR, THE PATIENT COUGHED DEEPLY RESULTING IN AIR BEING PULLED IN THROUGH THE WAS INTO THE LA OF THE PATIENT. THE PHYSICIAN ATTEMPTED TO ASPIRATE THE AIR OUT OF THE PATIENT. THE PATIENT BECAME HEMODYNAMICALLY UNSTABLE WITH A DROP IN BLOOD PRESSURE, OXYGEN LEVELS, ST SEGMENT ELEVATIONS AND BRIEF EPISODES OF VENTRICULAR TACHYCARDIA. THE PATIENT WAS PACED AND INTUBATED AT THIS TIME. ALL THE DEVICES WERE REMOVED FROM THE PATIENT, AND THEY STABILIZED AFTER ALL THE INTERVENTIONS. THE PROCEDURE WAS ENDED WITH NO CLOSURE DEVICE ATTEMPTED TO BE IMPLANTED. IMAGING SHOWED NO FURTHER AIR IN THE LEFT ATRIUM OR ANY OTHER COMPLICATIONS. THE PATIENT WAS EXTUBATED AND HEMODYNAMICALLY STABLE. FOLLOWING MONITORING OF THE PATIENT THE PHYSICIAN BELIEVED THE PATIENT MAY HAVE EXPERIENCED A CEREBRAL VASCULAR ACCIDENT DUE TO THE PATIENT'S SYMPTOMS OF FACIAL DROOP AND SLURRED SPEECH FOLLOWING EXTUBATION. THE PATIENT UNDERWENT COMPUTED TOMOGRAPHY (CT) IMAGING FOR EVALUATION. THE CT OF THE HEAD AND CT ANGIOGRAM WERE ALL NEGATIVE. THE PATIENT WAS KEPT OVERNIGHT FOR OBSERVATION AND TO UNDERGO A MAGNETIC RESONANCE IMAGING SCAN. IT WAS FURTHER REPORTED THAT NO CEREBRAL VASCULAR ACCIDENT WAS CONFIRMED. CT AND MRI IMAGING WERE ALL NEGATIVE. THE PATIENTS SYMPTOMS RESOLVED FROM THIS EVENT. OTHER CHALLENGES AROSE DURING THE HOSPITALIZATION AND THE PATIENT REMAINED IN THE HOSPITAL. TWO (2) WEEKS POST PROCEDURE THE PATIENT DIED DUE TO THE COMPLICATIONS THAT RESULTED DURING HOSPITALIZATION. THE PATIENTS FAMILY CHOSE TO PROVIDE COMFORT CARE TO THE PATIENT BEFORE THE PATIENT DIED.
TEMPUS LS | K200849 | THIS REPORT IS BASED ON INFORMATION PROVIDED BY REMOTE SERVICE ENGINEER RSE AND IS CURRENTLY UNDER INVESTIGATION BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON THE TEMPUS LS MANUAL INDICATING THE DEVICE, DURING A CARDIAC ARREST ON THE MORNING OF (B)(6) 2025, ENCOUNTERED A "CHECK DEFIBRILLATION PADS" ERROR MESSAGE. THE CREWS PLACED DEFIBRILLATOR PADS ON THE PATIENT (35YO MALE), THEN PLUGGED THE PADS INTO THE TEMPUS LS. THE MONITOR SHOWED A RHYTHM DURING COMPRESSIONS FOR 2-3 SECONDS, THEN NO WAVEFORM AND THE SCREEN DISPLAYED THE MESSAGE "CHECK DEFIBRILLATION PADS." PADS WERE CONFIRMED TO BE IN PLACE AND APPROPRIATELY ADHERED TO THE PATIENT. WHEN THE CORDS WERE UNPLUGGED THEN PLUGGED BACK IN, THE MONITOR RECOGNIZED THAT SOMETHING HAD BEEN ATTACHED TO IT, THEN WENT BACK TO THE ERROR MESSAGE. ONCE OR TWICE, THERE WERE 1-2 SECONDS OF A RHYTHM SHOWING ON THE MONITOR, BUT FOR THE MAJORITY OF THE TIME THE CREW SAW THE "CHECK DEFIBRILLATION PADS" ERROR MESSAGE WITH NO RHYTHM SHOWN. THE ISSUE WAS RESOLVED WHEN WE REMOVED THE FIRST SET OF PADS AND REPLACED THEM WITH ANOTHER SET OF PADS. | EVENT DATE AND DATE OF DEATH UPDATED AFTER RECEIVING GFE RESPONSE THAT WAS ORIGINALLY PROVIDED ON MFR REPORT NUMBER 3003832357-2025-000181.
VIVO 45LS | K193586 | JANUARY 9, 2025: BREAS MEDICAL RECEIVED A CUSTOMER REPORT THAT A PATIENT HAD EXPIRED WHILE USING A BREAS VIVO 45LS. THIS WAS PROVIDED BY THE CUSTOMER ROTECH INC.: 'THIS PT DAUGHTER SENT US A TEXT STATING THE MOM PASSED AWAY AND STATED THE MACHINE PLAYED A FACTOR IN HER DEATH I HAVE BEEN TRYING TO GET THIS SENT OFF AND I HAVE NOT HEARD ANYTHING FROM THE FAMILY SINCE THIS DATE ' BREAS RECEIVED THE DEVICE ON JANUARY 16, 2025. THE INVESTIGATION IS ONGOING. LOG FILES HAVE BEEN REVIEWED BUT THERE IS NO DATA IN THE FILES FROM THE REPORTED DATE OF THE EVENT. NOTHING IN THE LOG FILES INDICATES THAT THE VENTILATOR STOPPED WORKING. BREAS REACHED OUT TO THE CUSTOMER VIA EMAIL ON JANUARY 27, 2025 TO REQUEST MORE INFORMATION.
REFLEX CATHETER | K110055 | YAO, J., ZHENG, B., SUN, Q., ZHANG, F., JI, Z., WANG, C., WU, P., & SHI, H. (2024). HYBRID SURGERY FOR SYMPTOMATIC CHRONIC INTERNAL CAROTID ARTERY OCCLUSION: A SINGLE-CENTER EXPERIENCE. ACTA NEUROCHIRURGICA: THE EUROPEAN JOURNAL OF NEUROSURGERY, 166(1). HTTP S://DOI.ORG/10.1007/S00701-024-06250-4. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A RETROSPECTIVE CASE SERIES WAS CONDUCTED AT A SINGLE CENTER REVIEWING CASES OF 27 PATIENTS WHO UNDERWENT HYBRID SURGERY FOR THE TREATMENT OF SYMPTOMATIC CHRONIC INTERNAL CAROTID ARTERY OCCLUSION (ICAO). IT WAS NOTED THAT PATIENT WITH ICAO ARE AT HIGH RISK FOR REOCCLUSION. THE HYBRID SURGICAL TECHNIQUE COMBINES CAROTID ENDARTERECTOMY (CEA) AND ENDOVASCULAR INTERVENTION (EI). THIS PROCEDURE INVOLVES PERFORMING CEA AT THE PROXIMAL END OF THE CAROTID ARTERY TO REMOVE ATHEROSCLEROTIC PLAQUES AND THROMBI. SUBSEQUENTLY, EI TARGETS DISTAL STENOSIS OR OCCLUSION IN THE INTERNAL CAROTID ARTERY THAT CEA CANNOT REACH, THUS SUCCESSFULLY RECANALIZING THE CHRONIC ICAO. MULTIPLE MANUFACTURERS' DEVICES WERE USED IN THE TREATMENT PROCEDURES. IN SOME CASES, THE INTERMEDIATE CATHETER WAS A NAVIEN 6F 115CM CATHETER. HOWEVER, OTHER CATHETERS WERE USED AND IT WAS NOT SPECIFIED WHICH DEVICES WERE USED IN EACH PROCEDURE. NO DEVICE MALFUNCTION OR DEFICIENCY OF THE CATHETER WAS REPORTED. IT WAS NOTED THAT SUCCESSFUL REVASCULARIZATION WAS ACHIEVED IN 100% OF CASES. THE FOLLOWING PERI-PROCEDURE ADVERSE EVENTS WERE REPORTED: ONE PATIENT DEATH OCCURRED. THE PATIENT DEVELOPED ACUTE RENAL FAILURE POST-SURGERY MARKED BY ANURIA, AND UNFORTUNATELY PASSED AWAY ON THE THIRD DAY AFTER THE OPERATION. TWO PATIENTS EXPERIENCED ISCHEMIC COMPLICATIONS. ONE OF THESE PATIENTS HAD INTRAOPERATIVE STENT THROMBOSIS, AS CONFIRMED BY DIGITAL SUBTRACTION ANGIOGRAPHY (DSA) SHOWING THROMBI IN THE DISTAL ICA AND M2 SEGMENT OF THE MIDDLE CEREBRAL ARTERY. THIS PATIENT UNDERWENT A SUCCESSFUL MECHANICAL THROMBECTOMY AND HAD A FAVORABLE RECOVERY, EXHIBITING NO NEUROLOGICAL ISCHEMIC SYMPTOMS AND WAS DISCHARGED WITH AN MRS SCORE OF 0. THE OTHER PATIENT PRESENTED WITH LIMB WEAKNESS IN THE VASCULAR SUPPLY AREA OF THE SURGERY, CONFIRMED BY POSTOPERATIVE DIFFUSION-WEIGHTED IMAGING (DWI) AS A SMALL FOCAL ACUTE CEREBRAL INFARCTION, LIKELY DUE TO EMBOLIZATION OF DEBRIS. SYMPTOMS PARTIALLY ALLEVIATED FOLLOWING IMPROVED CIRCULATION AND REHABILITATION INTERVENTIONS. ONE PATIENT EXPERIENCED A HEMORRHAGIC COMPLICATION THAT WAS ATTRIBUTED TO THE ANTIPLATELET THERAPY. AS THERE WAS NO DEVICE MALFUNCTION AND THE EVENT WAS NOT ASSOCIATED WITH THE DEVICE OR PROCEDURE, THE HEMORRHAGIC EVENT IS NOT CONSIDERED A COMPLAINT. | **SERIOUS ADVERSE EVENTS/INJURIES REPORTED IN THE ARTICLE WILL BE REPORTED SEPARATELY. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
PHOTODYNAMIC BONE STABILIZATION SYSTEM | K181228 | A PATIENT PASSED AWAY DURING THE CLOSING OF AN ILLUMINOSS PROCEDURE, AFTER THE CURING PROCESS WAS COMPLETED. THEY WERE NOT ABLE TO REVIVE THE PATIENT. THE ANATOMY TREATED WAS THE PELVIS, BOTH SIDES BEING TREATED. THE PATIENT WAS "THOUGHT TO HAVE HAD AN EMBOLIC EVENT" PER THE USER. THE PATIENT HAD COPD. | AT THE TIME OF THE INITIAL MDR SUBMISSION, THE INVESTIGATION WAS STILL UNDERWAY. THIS UPDATE IS BEING PROVIDED WITH THE CONCLUSION FROM THE INVESTIGATION. INVESTIGATION: A CALL WAS HELD WITH THE TREATING PHYSICIAN TO DISCUSS THE EVENT. THE 73 YEAR OLD PATIENT HAD A FRAGILITY FRACTURE IN PELVIS WITH A HISTORY OF SMOKING WITH COMORBIDITIES WHICH INCLUDED COPD, EMPHYSEMA, OSTEOPOROSIS, AND RESPIRATORY DISEASES. THE PHYSICIAN TREATED THE SACRUM AND BOTH ILIA, USED 4 BALLOONS BILATERALLY. AT THE COMPLETION OF THE PROCEDURE DURING CLOSING, THE PATIENT BECAME HYPOXIA AND HYPOTENSIVE. THEY COVERED THE WOUNDS AND RAN A CODE FOR A PRESUMED DIAGNOSIS OF A FAT EMBOLISM. THERE WERE DIFFICULTIES DUE TO THE PATIENT'S RESPIRATORY COMORBIDITIES AND THEY WERE NOT ABLE TO STABILIZE THE PATIENT. AFTER 40 MINUTES OF THE CODE, THE CARDIAC PHYSICIAN DETERMINED THAT GIVEN THE LENGTH OF TIME OF HYPOXIA THE PATIENT'S PROGNOSIS WAS NOT GOOD, AND THE PATIENT WAS PRONOUNCED DEAD. THE USER ATTRIBUTED THE CAUSE OF THE DECLINE AND DEATH TO FAT EMBOLISM WHICH WAS SEEN IN THE ECHOCARDIOGRAM. THE PHYSICIAN IN THIS CASE, WAS A VERY EXPERIENCED ILLUMINOSS USER SO IT IS UNLIKELY THAT ANY USER ERROR CONTRIBUTED TO THE EVENT. A MEDICAL OVERSIGHT REVIEW MEETING WAS HELD WITH A MEDICAL REVIEWER, REGULATORY, QUALITY, AND ENGINEERING. THE INFORMATION FROM THE COMPLAINT WAS REVIEWED AND THE MEDICAL REVIEWER STATED THAT GIVEN THE TIMING AND REPORT FROM THE ECHOCARDIOGRAM OF THE FAT EMBOLISM, AND THAT FROM THE LITERATURE A FAT EMBOLISM IS A KNOWN RISK OF THE ILLUMINOSS PROCEDURE, THEY AGREE WITH THE USES CONCLUSIONS. THE MEDICAL REVIEWER ALSO OBSERVED THAT THE PATIENT HAD MANY RESPIRATORY COMPLICATIONS WHICH WOULD INCREASE THE RISK OF EXPERIENCING A FAT EMBOLISM. DHR REVIEW: THE DHRS FOR THE TWO LOTS OF IMPLANTS USED WERE REVIEWED AND FOUND TO BE IN SPECIFICATION AT THE TIME OF MANUFACTURE AND RELEASE. IFU REVIEW: IFU 900356_X STATES THAT RISKS INCLUDE THROMBOEMBOLIC EVENT OR FAT EMBOLISM (BLOOD CLOT, FAT, OR OTHER MATERIAL THAT COULD RESULT IN ORGAN DAMAGE OR FAILURE). POTENTIAL FOR USER ERROR: THE USER IS AN EXPERIENCED ILLUMINOSS USER, SO IT IS UNLIKELY THAT ANY USER ERROR CONTRIBUTED TO THE EVENT. THE USER PERFORMED REAMING AS INSTRUCTED IN THE SURGICAL TECHNIQUE GUIDE 900598_A FOR THE PELVIS. CONCLUSION: THE CAUSE OF PATIENT DEATH WAS A FAT EMBOLISM. THIS PATIENT HAD SEVERAL COMORBIDITIES WHICH INCREASE THE RISK OF FAT EMBOLISM AND THE RISK OF FAT EMBOLISM IS INHERENT TO AN ORTHOPEDIC SURGERY INVOLVING ANESTHESIA.
TRINAV INFUSION SYSTEM | K180677 | IT WAS REPORTED THAT A TRINAV CATHETER WAS USED IN A ROUTINE THYROID GOITER EMBOLIZATION PROCEDURE EMPLOYING RADIAL APPROACH VIA THE PERIPHERAL VASCULATURE/INFERIOR THYROID ARTERY. EMBOLIZATION WAS SUCCESSFUL AND CARRIED TO STASIS. POST-EMBOLIZATION ANGIOGRAPHY WAS PERFORMED VIA THE TRINAV CATHETER INSTEAD OF THE THROUGH THE BASE CATHETER IN THE SUBCLAVIAN ARTERY. POST-EMBOLIZATION CONTRAST INJECTION VOLUMES AND FORCES WERE HIGH AND THERE WAS NO FORWARD FLOW/ROOM, DUE TO PRESENCE OF SPHERES, RESULTING IN REFLUX. ARTERY WAS SUPERIMPOSED OVER RIBS AND SPINE MAKING VISUALIZATION OF REFLUX DIFFICULT. ORIGIN OF THE LEFT VERTEBRAL ARTERY WAS DOMINANT (LARGER THAN THE RIGHT) AND IMMEDIATELY PROXIMAL TO THE INFERIOR THYROID ARTERY, RESULTING IN HIGH FLOW TO THE POSTERIOR PORTION OF THE BRAIN. PATIENT EXPERIENCED MULTIPLE SMALL STROKES IDENTIFIED VIA POST PROCEDURE MRI, STARTED ON DUAL-ANTIPLATELET MEDICATIONS (ASPIRIN AND PLAVIX, LOADING DOSES), AND MONITORED FOR 24 HOURS. CT SCAN SHOWED A LARGE AMOUNT OF HEMORRHAGE INTO THE LEFT LATERAL VENTRICLE THAT SPREAD INTO ADJACENT VENTRICLES AND RESULTED IN LOSS OF DIFFERENTIATION OF GREY AND WHITE MATTER. THE PATIENT WAS TRANSITIONED TO HOSPICE CARE THAT DAY AND DIED SHORTLY THEREAFTER. | NO DEVICE MALFUNCTION OR PERFORMANCE ISSUE ASSERTED OR SUSPECTED BASED ON USER FACILITY REPORT. SUBJECT DEVICE NOT AVAILABLE FOR INVESTIGATION, SCRAPPED BY USER FACILITY AT TIME OF PROCEDURE. DEVICE VARIANT (MODEL NUMBER) AND/OR LOT NUMBER NOT REPORTED BY USER FACILITY.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A PERITONEAL DIALYSIS (PD) PATIENT'S DAUGHTER REPORTED THAT THIS PATIENT ON PD EXPIRED WHILE STILL CONNECTED TO THE FRESENIUS CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. HOWEVER, THERE WERE NO REPORTED ALLEGATIONS THAT THE DEATH WAS RELATED TO ANY PRODUCT RELATED ISSUES. IN ADDITIONAL FOLLOW-UP, THE PATIENT¿S PD NURSE STATED THAT THE PATIENT HAD PRE-EXISTING COMORBID CONDITIONS OF END STAGE RENAL DISEASE (ESRD) WITH ADDITIONAL CARDIAC COMORBIDITIES INCLUDED AORTIC VALVE STENOSIS AND HYPERTENSION. PER THE NURSE, THE PATIENT EXPIRED ON (B)(6) 2025. THE NURSE STATED THAT THE PATIENT¿S DAUGHTER REPORTED THAT THE PATIENT WAS ATTACHED TO THE FRESENIUS CYCLER. HOWEVER, NO TREATMENT DATA WAS AVAILABLE AND IT WAS UNKNOWN IF THE PATIENT WAS IN ACTIVE TREATMENT WHEN SHE EXPIRED. THE NURSE STATED THE PATIENT WAS AWAKE AND WELL AT 8AM. HOWEVER, SUBSEQUENTLY, AT APPROXIMATELY 10AM THE PATIENT WAS FOUND UNRESPONSIVE (BY DAUGHTER). THE NURSE STATED THAT EMERGENCY MEDICAL SERVICES WAS CALLED. THE PATIENT WAS A DNR AND NOT BROUGHT TO THE HOSPITAL; PRONOUNCED DECEASED IN THE HOME. THE NURSE STATED THERE WERE NO REPORTED ISSUES WITH THE PATIENT¿S PD TREATMENT OR PRODUCT IN RELATION TO THE DEATH. THE NURSE STATED THAT THE PATIENT¿S NEPHROLOGIST SUSPECTS THE CAUSE OF DEATH WAS FROM CARDIAC DUE TO NATURAL CAUSES FROM PRE-EXISTING COMORBID CONDITIONS. NO FURTHER INFORMATION ABOUT THE DEATH WAS AVAILABLE. THE FRESENIUS CYCLER WAS PENDING RETURN TO MANUFACTURER AT THE TIME FOLLOW-UP WAS COMPLETED. | CLINICAL REVIEW: TEMPORAL RELATIONSHIP EXISTS BETWEEN PD THERAPY WITH THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. HOWEVER, THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE THAT THE PATIENT¿S DEATH WAS RELATED TO A PRODUCT ISSUE OR MALFUNCTION WITH THE FRESENIUS CYCLER OR ANY ISSUES WITH PD TREATMENT. IT WAS REPORTED THAT THE SUSPECTED CAUSE OF DEATH WAS CARDIAC ARREST. THE PATIENT HAD A MEDICAL HISTORY THAT INCLUDED ESRD WITH PRE-EXISTING CARDIAC COMORBIDITIES THAT INCLUDED AORTIC VALVE STENOSIS AND HYPERTENSION. PATIENTS WITH END STAGE RENAL DISEASE (ESRD) ON DIALYSIS HAVE A SIGNIFICANTLY HIGHER MORTALITY RISK THAN THE GENERAL POPULATION. FURTHERMORE, THE PATIENT¿S PRE-EXISTING CARDIAC COMORBID CONDITIONS INCREASE RISK OF MORTALITY. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT A PATIENT WITH A RENTAL DREAMSTATION ST30 GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINIC REACHED OUT TO THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. A FOLLOW-UP REPORT WILL BE FILED IF ANY FURTHER RELEVANT INFORMATION BECOMES AVAILABLE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. | THE MANUFACTURER PREVIOUSLY REPOTED THAT THEY RECEIVED INFORMATION ALLEGING THAT A PATIENT WITH A RENTAL DREAMSTATION ST30 GB DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CLINIC REACHED OUT TO THE MANUFACTURER TO INQUIRE ABOUT RETURNING THE DEVICE. THE DEVICE UNDERWENT A REPAIR EVALUATION BY A THIRD-PARTY SERVICE CENTER, WHERE IT WAS VISIBLY INSPECTED, TESTED, AND SUBSEQUENTLY PASSED, WITH NO FAULTS FOUND. THE UNIT WAS THEN RETURNED TO LOAN STOCK . NO FURTHER INVESTIGATION IS NECESSARY. IF WE RECEIVE ANY RELEVANT INFORMATION, A FOLLOW UP REPORT WILL BE SUBMITTED.
PROWATERFLEX | K163426 | IT WAS REPORTED THAT THE PATIENT WITH A 100%-OCCLUDED LEFT MAIN TRUNK (LMT) WAS IN CARDIOGENIC SHOCK WHEN HE WAS BROUGHT IN THE CATHETERIZATION LABORATORY. PERCUTANEOUS CORONARY INTERVENTION (PCI) WAS PERFORMED TO TREAT A 100%-OCCLUDED, HEAVILY CALCIFIED LESION OF 30MM IN LENGTH IN THE LMT. AN ASAHI PROWATERFLEX GUIDE WIRE WAS USED WITH AN UNSPECIFIED BALLOON CATHETER TO DEPLOY A 3.5MM-IN-DIAMETER RESOLUTE ONYX STENT (MEDTRONIC) TO SUCCESSFULLY RE-CANALIZE THE LMT. AT THE END OF THE PROCEDURE, THE SUBJECT PROWATERFLEX WAS BEING REMOVED BUT SLIGHT RESISTANCE WAS MET. AS THE GUIDE WIRE WAS CONTINUED TO BE PULLED BACK, THE DISTAL SEGMENT OF THE GUIDE WIRE WAS FOUND BROKEN OFF. THE DEPLOYED STENT WAS IMMEDIATELY CRUSHED IN THE LMT, RESULTING IN UNSUCCESSFUL RE-CANALIZATION. IT WAS INFORMED THAT THE PATIENT EVENTUALLY DIED. THE PHYSICIAN COMMENTED THAT THE SUBJECT PROWATERFLEX MET RESISTANCE PROBABLY BECAUSE THE GUIDE WIRE MIGHT HAVE BEEN TRAPPED BY A PART OF THE STENT STRUT. AS HE KEPT PULLING BACK THE GUIDE WIRE AGAINST RESISTANCE FOR REMOVAL, THE STENT WENT CRUSHED. CONSEQUENTLY, THE LMT WAS COMPLETELY CLOSED OFF, PREVENTING THE PHYSICIAN FROM ENTERING ANY DEVICE IN THE LMT AND CAUSING CARDIAC ARREST. NOTHING WAS ABLE TO RE-CANALIZE THE LESION AT THAT POINT. | MANUFACTURING SITE: ASAHI INTECC (THAILAND) CO., LTD., PATHUM THANI, THAILAND, REGISTRATION NUMBER: (B)(4). THE REPORTED PROWATERFLEX GUIDE WIRE WAS RETURNED FOR INVESTIGATION. THE RETURNED PROWATERLFLEX GUIDE WIRE WAS FOUND THREE-DIMENSIONALLY DEFORMED FOR APPROXIMATELY 30MM FROM THE DISTAL MID SOLDER (SET AT 25MM FROM THE WIRE TIP TO FIX THE OUTER COIL ONTO THE CORE WIRE). THE OUTER COIL OF THE GUIDE WIRE WAS FOUND ELONGATED FOR APPROXIMATELY 45MM FROM THE DISTAL MID SOLDER AND FRACTURED. THE BALL TIP WAS NOT FOUND AT THE VERY DISTAL END OF THE OUTER COIL. THE CORE WIRE WAS COMING OUT OF THE DISTAL MID SOLDER FOR APPROXIMATELY 10MM. THE OUTER COIL WAS FOUND DISARRANGED AT APPROXIMATELY 10MM AND 14MM PROXIMAL TO THE DISTAL MID SOLDER. OBSERVATION BY MICROSCOPE AND SCANNING ELECTRON MICROSCOPE (SEM) OF THE CORE WIRE FOUND NECKING OF THE FRACTURE END AND DIMPLES ON A RELATIVELY FLAT FRACTURE SURFACE, INDICATING THAT TENSIONAL STRESS HAD CONTRIBUTED TO THE CORE WIRE FRACTURE. OBSERVATION BY MICROSCOPE AND SEM OF THE OUTER COIL FRACTURE END FOUND A RELATIVELY FLAT FRACTURE SURFACE, INDICATING THAT TORSION WAS GENERATED AS THE COIL STRUCTURE HAD BEEN STRAIGHTENED BY TENSION. TWISTING OF THE OUTER COIL WAS ALSO OBSERVED AT THE FRACTURE END. MEASUREMENT OF THE RETURNED PROWATERFLEX GUIDE WIRE SUGGESTED THAT THE CORE WIRE WAS FRACTURED AT APPROXIMATELY 15MM FROM THE WIRE TIP AND THE OUTER COIL WAS ELONGATED AND DETACHED TOGETHER WITH THE BALL TIP AT APPROXIMATELY 20MM FROM THE WIRE TIP. LOT HISTORY REVIEW REVEALED NO ANOMALY RELATING TO THE REPORTED EVENT. NO OTHER SIMILAR PRODUCT EXPERIENCE REPORT WAS RECEIVED FROM THIS LOT. BASED ON THE OBTAINED INFORMATION AND THE INVESTIGATION OUTCOME, IT WAS PRESUMED THAT TENSIONAL STRESS GENERATED WITH GUIDE WIRE MANIPULATION DURING WITHDRAWAL ATTEMPTS MIGHT HAVE BEEN LOCALLY ACCUMULATED ON THE DISTAL SEGMENT OF THE SUBJECT PROWATERFLEX GUIDE WIRE WHILE ITS DISTAL SEGMENT HAD BEEN CAUGHT BY THE DEPLOYED STENT STRUT. CONSEQUENTLY, THE GUIDE WIRE WAS FRACTURED. IT WAS CONCLUDED THAT THIS EVENT WAS NOT ATTRIBUTED TO PRODUCT QUALITY. CAPA: NO CAPA WILL BE TAKEN. THE INSTRUCTIONS FOR USE (IFU) STATES: [WARNINGS] OBSERVE MOVEMENT OF THIS GUIDE WIRE IN THE VESSELS. BEFORE THIS GUIDE WIRE IS MOVED OR TORQUED, THE TIP MOVEMENT SHOULD BE EXAMINED AND MONITORED UNDER FLUOROSCOPY. DO NOT MOVE OR TORQUE THE GUIDE WIRE WITHOUT OBSERVING CORRESPONDING MOVEMENT OF THE TIP; OTHERWISE, THE GUIDE WIRE MAY BE DAMAGED AND/OR TRAUMA MAY OCCUR. IN ADDITION, ENSURE THAT THE DISTAL TIP OF THIS GUIDE WIRE AND ITS LOCATION IN THE VESSEL ARE VISIBLE DURING MANIPULATIONS OF THE GUIDE WIRE. NEVER PUSH, AUGER, WITHDRAW, OR TORQUE THIS GUIDE WIRE THAT MEETS RESISTANCE. TORQUING OR PUSHING THIS GUIDE WIRE AGAINST RESISTANCE MAY CAUSE DAMAGE AND/OR TIP SEPARATION OF THIS GUIDE WIRE OR DIRECT DAMAGE TO A VESSEL. RESISTANCE MAY BE FELT AND/OR OBSERVED UNDER FLUOROSCOPY BY NOTING ANY BUCKLING OF THE GUIDE WIRE. IF THE PROLAPSE OF THE GUIDE WIRE TIP IS OBSERVED, DO NOT ALLOW THE TIP TO REMAIN IN A PROLAPSED POSITION; OTHERWISE DAMAGE TO THE GUIDE WIRE MAY OCCUR. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE ANY NECESSARY REMEDIAL ACTION. IF RESISTANCE IS FELT DUE TO SPASM, BENDING OF THE GUIDE WIRE, OR DUE TO TRAP WHILE OPERATING THIS GUIDE WIRE IN THE BLOOD VESSEL OR REMOVING IT, DO NOT TORQUE AND/OR PULL THE GUIDE WIRE ITSELF. STOP THE PROCEDURE. DETERMINE THE CAUSE OF RESISTANCE UNDER FLUOROSCOPY AND TAKE APPROPRIATE REMEDIAL ACTION. IF THE GUIDE WIRE IS MOVED EXCESSIVELY, IT MAY BREAK OR BECOME DAMAGED, WHICH MAY CAUSE BLOOD VESSEL INJURY OR RESULT IN FRAGMENTS BEING LEFT INSIDE THE VESSEL. [MALFUNCTION AND ADVERSE EFFECTS] SEPARATION OF THE GUIDE WIRE.
ALARIS SYSTEM | K211218 | IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY?...THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. " | OMIT : B17 - DEVICE NOT RETURNED, C20 - NO FINDINGS AVAILABLE, ADDITIONAL INFORMATION : DEVICE AVAILABLE FOR EVAL?, RETURNED TO MANUFACTURER ON, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX A, G, B, C, D CODES AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE CUSTOMER SOUGHT INFORMATION AROUND THE MEDICATION INFUSION DETAILS AND WHETHER THERE WAS ANY TIME WHERE THE MEDICATIONS WERE NOT INFUSING. REVIEW OF THE DEVICE LOGS IDENTIFIED INSTANCES OF OCCLUDED PATIENT SIDE ALARMS, BUT THERE WERE NO INDICATIONS OF DEVICE MALFUNCTIONS OR INTERRUPTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE PATIENT¿S CONDITION. INTERRUPTIONS IN INFUSION WERE NOT OBSERVED OR DURING DEVICE TESTING. INSPECTION AND TESTING OF THE SUSPECT PUMP MODULES AND PCU DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULES WERE INFUSING WITHIN SPECIFICATIONS. THE PCU PASSED PREVENTATIVE MAINTENANCE TESTING AND WAS OPERATING AS INTENDED. REVIEW OF THE SUSPECT PCU BATTERY LOGS DID NOT SHOW ANY EVENTS OR MALFUNCTIONS THAT COULD HAVE CONTRIBUTED TO THE REPORTED INCIDENT. THERE WERE NO INSTANCES OF THE SYSTEM POWERING OFF DURING THE REPORTED INCIDENT TIME. THE INITIAL INFUSION PROGRAMMING OF NOREPINEPHRINE (DRUGID = 308), AT A CONCENTRATION OF 4 MG / 250 ML, WAS RECORDED ON 7FEB2025 AT 1:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6). THE INFUSION RATE WAS SET AT 41.25 ML, THE VOLUME TO BE INFUSED (VTBI) WAS 100 ML. A BASIC INFUSION WAS PROGRAMMED ON (B)(6) 2025 AT 3:13 PM ON THE SUSPECT PUMP MODULE SN (B)(6) AT A RATE OF 100 ML/HR AND A VTBI OF 950 ML. A THIRD PUMP MODULE (SN (B)(6), NOT RECEIVED) WAS IDENTIFIED TO BE IN USE DURING THE REPORTED INCIDENT TIME. A BASIC INFUSION WAS PROGRAMMED AND STARTED ON (B)(6) 2025 AT 1:47 PM, WITH A RATE OF 999 ML/HR AND A VTBI OF 950 ML. AT 2:40 PM, THE CONCOMITANT PUMP MODULE SN (B)(6) WAS CHANNELED OFF. THEN AT 2:59 PM AN INFUSION FOR THE DRUG CEFTRIAXONE (DRUGID = 131), AT A CONCENTRATION OF 1000 MG / 50 ML, WAS PROGRAMMED AND STARTED. THE INFUSION RATE WAS 100 ML/HR AND THE VTBI WAS 50 ML. THE DOSE FOR THE INFUSION OF NOREPINEPHRINE ON THE PUMP MODULE SN (B)(6) WAS CHANGED TO 0.4 MCG/KG/HR AT 3:18 PM AND THEN CHANGED AGAIN TO 0.22 MCG/KG/HR AT 3:20 PM. THE DOSE WAS CHANGED TO 0.25 MCG/KG/HR AT 4:28 PM. AN AIR-IN-LINE ALARM WAS OBSERVED ON THE CONCOMITANT PUMP MODULE SN (B)(6), LASTING 2 MINUTES BEFORE THE UNIT WAS CHANNELED OFF. THE TOTAL VOLUME INFUSED (TVI) WAS 902.015 ML. FIVE (5) OCCLUDED PATIENT SIDE ALARMS WERE OBSERVED BETWEEN 1:13 PM AND 4:27 PM ON THE PUMP MODULE SN (B)(6). THE ALARM AT 2:00 PM LASTED 8 MINUTES BEFORE THE INFUSION WAS RESTARTED. THEN, AT 3:15 PM, THE ALARM LASTED 2 MINUTES AND 48 SECONDS BEFORE THE INFUSION WAS RESUMED. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 3:46 PM, BUT WAS THEN POWERED ON AGAIN AT 4:24 PM AND AN INFUSION OF NOREPINEPHRINE AT A DOSE OF 0.22 MCG/KG/HR AND A VTBI OF 200 ML WAS PROGRAMMED AND STARTED AT 4:26 PM. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:29 PM. THE TVI WAS RECORDED AT 101.491 ML. THE SUSPECT PUMP MODULE SN (B)(6) WAS CHANNELED OFF AT 4:30 PM. THE TVI WAS RECORDED AT 127.003 ML. IT IS NOT POSSIBLE TO DETERMINE WHAT THE ACTUAL VOLUME IS WITHIN AN IV CONTAINER AT THE START AND ENDING OF AN INFUSION FROM A REVIEW OF THE PCU EVENT LOG. IT IS ALSO NOT POSSIBLE TO DETERMINE WHAT THE FLUID IS THAT IS CONTAINED WITHIN THE IV CONTAINER. THIS IS NOT A FUNCTION OF THE PCU EVENT LOG; IT ONLY CAN REFLECT THE INPUTTED INFORMATION IT RECEIVES EITHER MANUALLY OR REMOTELY WITH RESPECT TO VOLUME TO BE DELIVERED AND FLUID SELECTED. THE DEVICE WAS REPORTEDLY IN USE FOR TREATMENT PURPOSES AS INTENDED PER 21 CFR 820.198(D)(2). ROOT CAUSE: THE ROOT CAUSE OF THE REPORTED INTERRUPTION OF INFUSIONS IS DUE TO THE USER POWERING OFF THE PUMP MODULES. IT WAS NOTED THAT THE SUSPECT PUMP MODULES SN (B)(6) AND SN (B)(6) WERE POWERED OFF ON (B)(6) 2025 AT 4:29 PM AND 4:30 PM, RESPECTIVELY. THE CONCOMITANT PUMP MODULE SN (B)(6) WAS POWERED OFF AT 3:28 PM. NO FUNCTIONALITY ISSUES OR ANOMALIES WERE OBSERVED WITH THE DEVICES THAT WOULD CONTRIBUTE TO THE REPORTED EVENT. THE RETURNED PUMP MODULES WERE FOUND TO BE INFUSING WITHIN SPECIFICATION. NOTE THAT THIS REPORT LISTS IMDRF ANNEX A1801, G04037, C0601 CODES NOT ASSOCIATED WITH THE REPORTED EVENT BUT IDENTIFIED AS REPORTABLE MALFUNCTIONS OBSERVED ON THE DEVICE DURING INVESTIGATION ARE UNRELATED TO THE REPORTED ISSUE. THESE OTHER FAILURES PRESUMPTIVELY DID NOT CAUSE OR CONTRIBUTE TO THE REPORTED ISSUE.
ALARIS SYSTEM | K211218 | IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR (B)(6) 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED. " | THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
OMNIPOD DASH INSULIN MANAGEMENT SYSTEM | K211575 | IT HAS BEEN REPORTED THAT THE PATIENT HAS PASSED AWAY ON (B)(6) 2024. THE PATIENT REPORTEDLY DECEASED DUE TO DIABETIC KETOACIDOSIS (DKA). IT WAS ALSO REPORTED THAT INSULIN DELIVERY STOPPED ON (B)(6) 2024, WITH THE REASON FOR THIS BEING UNKNOWN. INFORMATION ON ANY PRODUCT ISSUES THE CUSTOMER EXPERIENCED WAS NOT PROVIDED. ACCORDING TO THE HOSPITAL, THE PATIENT'S BLOOD OXYGEN LEVELS WERE ACIDOTIC AND THERE WERE HIGH LEVELS OF TRAMADOL PRESENT. IT WAS ALSO REPORTED THAT THE PATIENT HAD A HISTORY OF DIABETES RELATED COMPLICATIONS THAT INCLUDED HIGH HBA1C (HEMOGLOBIN A1C), GASTROPARESIS, UVEITIS (TREATED WITH HIGH DOSE STEROIDS) AND RETINOPATHY. | ACCORDING TO THE COMPLAINANT THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION. WE ARE UNABLE TO DETERMINE IF ANY PRODUCT CONDITION COULD HAVE CONTRIBUTED TO THE REPORTED DEATH AND DIABETIC KETOACIDOSIS. NO LOT RELEASE RECORDS WERE REVIEWED, AS THE PRODUCT LOT NUMBER WAS NOT PROVIDED. HIGH BLOOD GLUCOSE IS A COMMON SYMPTOM FOR PEOPLE WITH DIABETES (GLUCOSE MONITORING DATA FROM PEOPLE WITH DIABETES INDICATE THAT ON AVERAGE, THEY CAN EXPERIENCE BLOOD GLUCOSE LEVELS ABOVE 250 MG/DL FOR 14-25% OF THE TIME[1][2][3]..), AND IT WOULD BE CHALLENGING TO SPECULATE ON A CAUSE FOR THE COMPLAINTS WITHOUT RECEIVING THE DEVICES BACK FOR AN ENGINEERING INVESTIGATION. [1] BECK RW, BERGENSTAL RM, CHENG P, KOLLMAN C, CARLSON AL, JOHNSON ML, RODBARD D. THE RELATIONSHIPS BETWEEN TIME IN RANGE, HYPERGLYCEMIA METRICS, AND HBA1C. J DIABETES SCI TECHNOL 2019;13:614-626. [1] WELSH JB, DERDZINSKI M, PARKER AS, PUHR S, JIMENEZ A, WALKER T. REAL-TIME SHARING AND FOLLOWING OF CONTINUOUS GLUCOSE MONITORING DATA IN YOUTH. DIABETES THER 2019; 10:751-755. [1] PUHR S, DERDZINSKI M, WELSH JB, PARKER AS, WALKER T, PRICE DA. REAL-WORLD HYPOGLYCEMIA AVOIDANCE WITH A CONTINUOUS GLUCOSE MONITORING SYSTEM'S PREDICTIVE LOW GLUCOSE ALERT. DIABETES TECHNOL THER 2019; 21:155-158.
FEEDING TUBE | K060944 | GASTRIC PERFORATION WHEN USED FOR GASTRIC FEEDING. THE TUBE WAS PLACED IN A 27 WEEK PREMATURE BABY ON (B)(6) 2024, AND ON (B)(6) THEY DIAGNOSED NEONATAL GASTRIC PERFORATION (NGP). THE BABY WAS TRANSFERRED TO (B)(6) HOSPITAL, A SPECIALIST PAEDIATRIC HOSPITAL WHERE HE DIED. THE BABY WAS A TWIN, AND THE OTHER TWIN ALSO EXPERIENCED A NGP BUT HE SURVIVED. FOR ADMINISTRATION OF ENTERAL NUTRITION. | THE INVOLVED DEVICE IS NOT AVAILABLE. THE CUSTOMER DID NOT PROVIDE THE INVOLVED BATCH NUMBER. THEREFORE, VYGON IE HAS PROVIDED THE BATCHES SUPPLIED TO THIS CUSTOMER IN PREVIOUS 4 MONTHS: BATCH NUMBER 200324PA, 210224PA, 041223PA. THE BATCH REVIEW OF THESE 3 BATCHES CONFIRMS THAT THE FEEDING TUBES ARE COMPLIANT, NO ANOMALIES AND NO NON-CONFORMITIES WERE RECORDED. THESE FEEDING TUBE HAVE ALL A ROUNDED DISTAL END TO AVOID ANY TRAUMA TO PATIENT. THE MANUFACTURING PROCESS OF THESE NUTRISAFE FEEDING TUBES INCLUDES A 100% LEAK AND OBSTRUCTION TESTS. THERE ARE ALSO QUALITY CONTROL CHECKS ARE AS VISUAL ASPECT, DISTAL END ROUNDING CONTROL, DISTANCE BETWEEN THE FIRST MARKING AND THE DISTAL END CENTIMER'S MARKING, USEFUL LENGTH. IN THE INSTRUCTION OF USE THE INFORMATION ARE DESCRIBED AS FOLLOWS: PLACEMENT OF THE FEEDING TUBE: FOR ORAL OR NASAL PLACEMENT, MEASURE THE LENGTH OF TUBE TO BE INSERTED ACCORDING TO STANDARD PROTOCOLS. LUBRICATE THE FEEDING TUBE. INTRODUCE THE FEEDING TUBE IN ONE OF THE NOSTRILS AND MOVE THE TUBE CAREFULLY TOWARDS THE POSTERIOR NASOPHARYNX, ORIENTED TOWARDS THE OESOPHAGUS. PROCEED SLOWLY TO AVOID THE RISK OF THE PATIENT VOMITING. BEND THE HEAD FORWARD TO CLOSE THE TRACHEA AND OPEN THE ESOPHAGUS. ADVANCE THE TUBE INTO THE OESOPHAGUS. ONCE THE MARK IS REACHED, STOP THE ADVANCEMENT OF THE TUBE. FIXATION OF THE FEEDING TUBE: AFTER PREPARING THE SKIN, DRESS ADHESIVE STRIP AROUND THE TUBE AND FIX IT ON THE NOSE. VERIFICATION: CHECK FEEDING TUBE POSITIONING PREFERABLY BY X-RAY. IF USING A FEEDING TUBE WITH STYLET, REMOVE GENTLY THE STYLET FROM THE FEEDING TUBE. -PRECAUTION FOR USE: THE PLACEMENT OF THE TUBE IS VERIFIED ONCE A DAY AND SYSTEMATICALLY BEFORE EACH USE (IN ANY CASE FOLLOW THE PROTOCOLS OF YOUR INSTITUTION). NUTRISAFE2 AND NUTRIFIT FEEDING TUBES ARE CONTRAINDICATED FOR PATIENTS WITH CONGENITAL ABNORMALITIES OF THE GASTROINTESTINAL TRACT. THE INFORMATION WE RECEIVED ARE POOR. WE DO NOT KNOW IF THE PLACEMENT HAS BEEN CHECKED AFTER INSERTION, DAILY AND THEN BEFORE EACH USE, IF THE PATIENT HAD NO ABNORMAL ANATOMY. FURTHERMORE, CASES OF NEONATAL GASTRIC PERFORATION ARE VERY RARE, BUT THEY ARE NEVERTHELESS WELL REPORTED IN THE LITERATURE. IN THESE CASES, THE STOMACH WALLS ARE NOT MATURE AND ARE OFTEN VERY THIN AND WEAK AND NOT RESISTANT TO EVEN MINIMAL PRESSURE. PREMATURITY AND LOW BIRTH WEIGHT ARE THE MAIN FACTORS ASSOCIATED WITH NGP WITH HYPOTENSION, SEPSIS, HIGH GASTRIC ACIDITY AS SIGNIFICANT CONTRIBUTORY FACTORS. MECHANICAL VENTILATION, SPECIFICALLY NASAL VENTILATION ARE ALSO SUSPECTED OF BEING CONTRIBUTORY FACTORS. THE PRESENCE OF A NG TUBE IS OBVIOUSLY ANOTHER RISK FACTOR BUT THERE ARE EVEN CASES OF SPONTANEOUS NGP REPORTED WHEN THERE IS NO NG TUBE IN SITU. THE MOST LIKELY CAUSE IS A RARE COMPLICATION ON A PREMATURE BABY. IT SEEMS NOT TRACED TO THE DEVICE. FROM HISTORICAL DATA ANALYSIS OF THIS NUTRISAFE2 ENTERAL TUBES PVC RANGE CODES 36X.XXX OF SIMILAR INCIDENT (GASTRIC PERFORATION) THESE 3 LAST YEARS, WE CAN CONCLUDE THAT THIS IS THE VERY FIRST CASE OF PERFORATION.
ALARIS SYSTEM | K211218 | IT WAS INITIALLY REPORTED A REQUEST TO REVIEW THE DEVICE LOGS FOR 07 FEBRUARY 2025 BETWEEN THE HOURS OF 1320 AND 1640. NO FURTHER INFORMATION WAS PROVIDED AT THE TIME. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE PATIENT WAS TRANSPORTED FROM EMERGENCY DEPARTMENT (ED) TO RADIOLOGY FOR APPROXIMATELY 15 MINUTES. LEVOPHED AND BICARB WAS INFUSING VIA PUMP AT THE TIME. UPON RETURN TO THE ED, THE PUMP WAS FOUND "NOT INFUSING." PER REPORT, THE PATIENT DIED AND HAD A DNR (DO NOT RESUSCITATE) CODE STATUS "WITH FAMILY DECISION OF NO FURTHER CARE." WHEN ASKED IF THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH, THE HOSPITAL'S PATIENT SAFETY MANAGER RN STATED THAT "CLARIFYING THE CONCERN IS IF THE PUMP SHUT DOWN DURING THE INFUSION POTENTIALLY DUE TO BATTERY LIFE AS IT WAS UNPLUGGED, WAS THE PUMP CONTINUOUSLY INFUSING AND THE DISPLAY WAS NOT VISIBLE. REALLY JUST SEEKING IF THE PUMP WAS CONTINUOUSLY INFUSING OR NOT AND IF NOT, WHY? THE PUMP NOR THE INFUSION(S) LED NOR CONTRIBUTED TO THE DEMISE OF THE PATIENT. WE ARE REVIEWING FOR PROCESS FAILURES INCLUDING TEAM MEMBER HANDOFF IN COMMUNICATION." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE LOGS, SPECIFICALLY IDENTIFYING "THE MEDICATION INFUSION DETAILS OF ML/HR. AND IF THERE WAS ANY TIME, THE MEDICATIONS WERE NOT INFUSING AS WE ARE REVIEWING FOR OUR INVESTIGATION OF PROCESSES. THE PATIENT WAS A DNR, WITH FAMILY PRESENTING THE DNR DOCUMENT UPON THEIR ARRIVAL TO THE ED." | THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DAVINCI 5 | K232610 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | ADDITIONAL INFORMATION WAS RECEIVED FROM THE SITE RISK MANAGER. SECTION B5 INFORMATION: IT WAS REPORTED THAT THE PROCEDURE WAS A HYSTERECTOMY FOR A RIGHT OVARIAN MASS. THERE WERE NO INTRA-OPERATIVE COMPLICATIONS DURING THE DA VINCI-ASSISTED PROCEDURE. POST-OPERATIVELY, THE PATIENT WAS INITIALLY DOING WELL IN POST-ANESTHESIA CARE (PACU);HOWEVER, WAS LATER FOUND TO BE UNRESPONSIVE. THE PATIENT WAS RETURNED TO THE OR WHERE TWO AREAS OF RUPTURED ABDOMINAL WALL VARICES WERE FOUND AND LIGATED. THE ESTIMATED BLOOD LOSS WAS 6.7 LITERS. THE PATENT WAS TRANSFUSED WITH MULTIPLE BLOOD PRODUCTS. THE ABDOMEN WAS PACKED, ABTHERA WOUND DRESSING WAS PLACED, AND THE PATIENT WAS TRANSFERRED TO THE SURGICAL INTENSIVE CARE UNIT. THE PATIENT WAS LATER RETURNED TO THE OR AGAIN FOR BLEEDING. NO LARGE VESSELS WERE VISIBLE, THE BLEEDING APPEARED TO BE COAGULOPATHIC AND WAS NOTED FROM THE PERIHEPATIC REGION, AND FROM THE EDGES OF THE MIDLINE LAPAROTOMY. THE ABDOMEN WAS PACKED WITH LAP PADS AND THE BLEEDING SLOWED. DURING THE PROCEDURE, THE PATIENT'S PULSE WAS LOST; RESUSCITATION WAS SUCCESSFUL WITH RETURN OF SPONTANEOUS CIRCULATION. THE VASCULAR TEAM PERFORMED SUCTION EMBOLECTOMY OF CLOT FROM THE RIGHT ATRIUM AND RIGHT VENTRICLE; HOWEVER, SIGNIFICANT CLOT BURDEN REFORMED AS THE CLOTS WERE MECHANICALLY EXTRACTED. THE PATIENT'S FAMILY ELECTED TO WITHDRAW CARE AND THE PATIENT EXPIRED. THE CUSTOMER STATED THAT NO DA VINCI PRODUCTS CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CUSTOMER ATTRIBUTED THE PATIENT OUTCOME TO THEIR HISTORY OF MELANOMA, ESOPHAGEAL VARICES, GASTROESOPHAGEAL REFLUX DISEASE (GERD), CIRRHOSIS, SEIZURES, AND HYPERTENSION. RE-REVIEW OF THE EVENT BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD SIGNIFICANT COMORBIDITIES PRIOR TO THE OPERATION, SPECIFICALLY ALCOHOLIC LIVER DISEASE. THIS DIAGNOSIS CAN LEAD TO A PATIENT HAVING SIGNIFICANT ABNORMAL VENOUS COLLATERALIZATION BECAUSE OF PORTAL HYPERTENSION INCLUDING ABDOMINAL WALL AND GASTRIC VARICES AND COAGULOPATHY AS RESULT OF POOR HEPATIC FUNCTION. THIS PATIENT UNDERWENT A HYSTERECTOMY BUT LATER BLED FROM THE VARICES AND EVENTUALLY DIED AS A RESULT OF A CASCADE OF COMPLICATIONS FROM THESE VARICES. NO ALLEGATIONS HAVE BEEN MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. THE POSTOPERATIVE COMPLICATIONS SUGGEST THIS DEATH WAS MOST LIKELY A RESULT OF THE PATIENT PRE-EXISTING DISEASE CONTRIBUTION. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
TRILOGY EVO | K181166 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY WHILE USING A VENTILATOR. THE INFORMATION INDICATES EMPORIUM IS UNABLE TO RECOVER THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED THAT PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PDRN. THE PATIENT WAS IN PD TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER ON (B)(6) 2025. THE PATIENT EXPERIENCED SHORTNESS OF BREATH. THE PATIENT¿S PARTNER CALLED 911 AND DISCONNECTED THE PATIENT FROM THE TREATMENT. EMERGENCY MEDICAL SERVICES (EMS) ARRIVED TO FIND THE PATIENT IN CARDIAC ARREST. CARDIOPULMONARY RESUSCITATION (CPR) WAS ADMINISTERED. RESUSCITATIVE EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. NO AUTOPSY WAS CONDUCTED. THE CAUSE OF DEATH IS DOCUMENTED AS CARDIAC ARREST, CAUSE UNKNOWN. NO ISSUES WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT WERE REPORTED PRIOR TO THE PATIENT¿S DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. | PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST ¿ PASSED. SYSTEM AIR LEAK TEST ¿ PASSED. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
INTRI24 INTRODUCER SHEATH | K233646 | AN 85-YEAR-OLD FEMALE ARRIVED AT THE HOSPITAL WITH SHORTNESS OF BREATH AND CHEST PAIN. THE PATIENT HAD A HISTORY OF CARDIAC ISSUES AND WAS LISTED AS "DO NOT RESUSCITATE". THE PATIENT WAS DIAGNOSED WITH MASSIVE PULMONARY EMBOLISM AND A THROMBECTOMY WAS PERFORMED WITH INARI DEVICES. THE PATIENT WAS INTUBATED AND ON TWO PRESSORS AT THE BEGINNING OF THE THROMBECTOMY. THE PROCEDURE BEGAN WITHOUT ISSUES AND A LARGE AMOUNT OF CLOT WAS REMOVED FROM THE PATIENT'S RIGHT SIDE THROUGH FIVE ASPIRATIONS WITH THE TRIEVER20 (T20). THE PATIENT'S OXYGEN (O2) SATURATION ROSE, AND IMAGING REVEALED THAT THE RIGHT SIDE WAS CLEAR. THE PHYSICIAN RE-POSITIONED TO THE LEFT AND WAS IN PROCESS OF OBTAINING WIRE POSITIONING WHEN A NURSE NOTICED A LARGE VOLUME OF BLOOD ON THE GROUND. UPON FURTHER INVESTIGATION, THE STOPCOCK OF THE LARGE BORE SYRINGE SIDE PORT WAS NOT CLOSED AFTER IMAGING WAS PERFORMED. IT WAS ESTIMATED THAT THERE WAS 500 CC OF BLOOD LOST FROM THE SIDE PORT BEING LEFT OPEN. THIS WAS IN ADDITION TO 300 CC OF BLOOD LOSS FROM ASPIRATIONS LEADING TO A TOTAL OF APPROXIMATELY 800 CC OF BLOOD LOSS. THE PATIENT'S BLOOD PRESSURE AND O2 DECLINED. THE PHYSICIAN MANAGED TO RE-POSITION TO THE LEFT SIDE WHILE AN ORDER OF BLOOD WAS PROCESSING. TWO FURTHER ASPIRATIONS WERE PERFORMED WHICH CLEARED SOME CLOT. FLOWSAVER WAS USED TO RETURN BLOOD. FLUID AND BLOOD WERE ADMINISTERED TO THE PATIENT. THE PHYSICIAN PERFORMED TWO ADDITIONAL ASPIRATIONS BUT WAS UNABLE TO FULLY CLEAR THE LEFT SIDE. AT THE TIME THE PROCEDURE WAS CONCLUDED, AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT. THE PATIENT PASSED AWAY IN THE MIDDLE OF THE NIGHT. | THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE NOT PROVIDED, THE LOT HISTORY RECORDS OF ALL POTENTIAL LOTS SHIPPED TO THE FACILITY WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH WAS THE RESULT OF SEVERE DISEASE STATE AND USER ERROR (BLOOD LOSS DUE TO FAILURE TO CLOSE THE SIDE PORT STOPCOCK). BLOOD LOSS, BLEEDING, AND DEATH ARE LISTED AS POTENTIAL COMPLICATIONS/ADVERSE EVENTS. MANUFACTURER REFERENCE #: (B)(4).
HARMONIC ACE 5MM SHEAR | K120729 | TITLE: A COMPARATIVE STUDY BETWEEN LAPAROSCOPIC SPLENECTOMY VERSUS OPEN SPLENECTOMY. AUTHOR:1JAGDISH PRASAD, 2PARVINDER SINGH SANDHU. CITATION: ONLINE ISSN: 2250-3137. PRINT ISSN: 2977-0122. THE STUDY WAS A RANDOMIZED COMPARATIVE STUDY CONDUCTED IN THE DEPARTMENT OF GENERAL SURGERY OF SAWAI MAN SINGH HOSPITAL, JAIPUR. INITIALLY, SPLENECTOMY WAS DONE BY OPEN TECHNIQUE, INDICATIONS BEING MAINLY HAEMATOLOGICAL DISORDERS. CHRONIC ITP. WITH GROWING EXPERIENCE AND ADVANCEMENT IN TECHNOLOGY THE LAPAROSCOPIC APPROACH EXTENDED TO CYSTS,TUMOURS, STORAGE DISEASES AND EVEN TRAUMA IT WAS ALSO USED FOR HAEMATOLOGICAL DISORDERS MAINLY CHRONIC ITP. WITH GROWING EXPERIENCE AND ADVANCEMENT IN TECHNOLOGY THE LAPAROSCOPIC APPROACH EXTENDED TO CYSTS, TUMORS, STORAGE DISEASES AND EVEN TRAUMA. BETWEEN JUNE 2011 TO JUNE 2012. TOTAL 25 PATIENTS WERE INCLUDED IN STUDY PROSPECTIVELY INTO TWO GROUPS, GROUP A & GROUP B. GROUP A(N= 12 AND N=13) CASES IN GROUP B RESPECTIVELY. THE PATIENTS IN GROUP A UNDERWENT LAPAROSCOPIC SPLENECTOMY. WHILE GROUP B UNDERWENT OPEN SPLENECTOMY. THE MEAN AGE IN GROUP A WAS 35.41 YEARS WHILE IN GROUP B WAS 29.53 YEARS. SPLENIC SPAN OF 11-20 CM WAS FOUND IN 33% PATIENTS IN GROUP A AND 46% OF PATIENTS IN GROUP B .MINIMUM PREOPERATIVE PLATELET COUNT IN GROUP A WAS 7000/L IN A CASE OF I.T.P. AND THAT IN GROUP B WAS 30000/L. ACCESSORY SPLEEN WAS DETECTED IN 1 OF THE PATIENT (8%) IN LAPAROSCOPIC GROUP WHEREAS NONE OF THE PATIENT WAS FOUND TO BE HAVING ACCESSORY SPLEEN IN OPEN GROUP. GROUP A DEVELOPED STEROID TOXICITY IN THE FORM OF HYPERTENSION, DIABETES AND PAN GASTRITIS .THE OPERATIVE TIME IN GROUP A WAS IN THE RANGE OF 90-360 MINUTES WITH MEAN OPERATING TIME OF 18. MEAN INTRA OPERATIVE BLOOD LOSS IN GROUP A WAS 60 ML ±38 ML AND FOUND TO BE LESS THAN GROUP B MEAN DRAIN OUTPUT WAS 134ML± 145ML IN GROUP A WITH RANGE OF 50-550ML IN GROUP-A MEAN FOLLOW UP PERIOD WAS 9 MONTHS WITH A RANGE OF 4-18 MONTHS. IN GROUP-B MEAN FOLLOW UP PERIOD WAS 7 MONTHS WITH A RANGE OF 1-18 MONTHS. COMPLICATION DURING THE PERIOD OF FOLLOW UP WAS FOUND IN 17% OF PATIENTS IN GROUP-A AND 15% IN GROUP B. ENDOSTAPLER (ENDOLINEAR CUTTER ATS FLEX 45MM; ECHELON 60M.) WAS USED IN GROUP-A HILAR CONTROL WAS ACHIEVED. ETHILON 3-0 SUTURES WAS USED WHEN THE SKIN SUTURES WERE APPLIED. HARMONIC SCALPEL USED WHEN THE PERITONEUM OVERLYING THE SPLENIC VESSELS WAS DIVIDED REPORTED COMPLICATIONS: HARMONIC DEVICE (ETHICON ENDO-SURGERY). DEATH ON SECOND POST-OPERATIVE DAY DUE TO INTRA ABDOMINAL HEMORRHAGE LEADING TO SHOCK WHICH WAS SEEN DURING RE-EXPLORATION AS DIFFUSE PERITONEAL BLEEDING (N=1). IN CONCLUSION, LAPAROSCOPIC SPLENECTOMY IS SAFE AND EFFICACIOUS FOR MILD TO MODERATE SIZE SPLEENS. THE OPERATIVE TIME FOR LAPAROSCOPIC SPLENECTOMY IS LONGER THAN OPEN SPLENECTOMY BUT AFTER ACHIEVING LEARNING CURVE IT CAN BE SHORTENED SIGNIFICANTLY. | (B)(4). DATE SENT: 3/12/2025. B3: PUBLICATION YEAR OF 2024. D4: BATCH # UNK. THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE MANUFACTURING RECORDS CANNOT BE REVIEWED AS THE LOT/BATCH NUMBER HAS NOT BEEN PROVIDED. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE AUTHOR/SURGEON BELIEVE THAT THE ETHICON DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT COMPLICATIONS AND/OR DEATH MENTIONED IN THE ARTICLE? IF YES, PLEASE EXPLAIN IS THE AUTHOR/SURGEON INTERESTED IN SPEAKING WITH AN ETHICON MEDICAL PROFESSIONAL? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
INTELLIVUE MP70 | K151681 | IT WAS REPORTED THE MONITOR IN THE NICU DID NOT GENERATE A DESATURATION ALARM. THE CUSTOMER INDICATED THE PATIENT WAS FOUND WITH AN SPO2 OF 14% AND NO ALARM WAS GENERATED. AFTER CHECKING THE PATIENT DATA, IT WAS DETERMINED THE PATIENT HAD BEEN DESATURATING FOR 30 MINUTES WITH NO ALARM, WITH ALARM LIMITS SET AT 85% FOR YELLOW AND 80% FOR RED ALARMS FOR PREMATURE PATIENTS. THE PATIENT WAS FOUND WITH A GRAY COMPLEXION AND BRADYCARDIA. THE PATIENT WAS STIMULATED AND UNTIL THEY RECOVERED THEIR OXYGEN SATURATION. THE PATIENT DIED THE NEXT DAY DUE TO SEPSIS SECONDARY TO ENTEROCOLITIS. | PATIENT INVOLVEMENT WAS REPORTED, THE PATIENT WAS FOUND WITH A GRAY COMPLEXION AND BRADYCARDIA. THE PATIENT WAS STIMULATED UNTIL THEY RECOVERED THEIR OXYGEN SATURATION. THE PATIENT DIED THE NEXT DAY DUE TO SEPSIS SECONDARY TO ENTEROCOLITIS. THE RSE REVIEWED THE LOGS PROVIDED AND DETERMINED THAT THE PATIENT ENTERED THE ALARM CONDITION AT 2:49 A.M (B)(6) 2025, FOR WHICH THE DEVICE GENERATED AN ALARM WHICH WAS ACKNOWLEDGED. AS THE ALARM WAS REPEATED, ACKNOWLEDGEMENTS OCCURRED. THE PATIENT DID NOT GET OUT OF THE ALARM CONDITION. THIS CAN BE SEEN ON THE TREND CURVES. THIS IS WHY THE ALARM REVIEW DOES NOT INDICATE NEW ALARMS BUT INDICATES THE ALARM REPETITION ACKNOWLEDGEMENTS. IT IS STILL THE SAME ALARM. NO NEW ALARMS ARE CAPTURED IN THE PATIENT DATA AS THE ORIGINAL ALARM CONDITION NEVER RESOLVED. THEREFORE, THE DEVICE FUNCTIONED AS INTENDED. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED THE INFORMATION AND LOGS PROVIDED. THE PSE STATED THAT THE INFORMATION CONFIRMS THE DETERMINATION BY THE RSE. A REVIEW OF THE LOGS CONFIRMED THE DEVICE FUNCTIONED AS INTENDED. THE DEVICE GENERATED AN ALARM WHICH WAS ACKNOWLEDGED. THE ALARM WAS REPEATED AND ACKNOWLEDGEMENTS OCCURRED. ATTEMPTS WERE MADE TO CLARIFY SOME ADDITIONAL CLINICAL INFORMATION, BUT THE CUSTOMER HAS NOT RESPONDED TO ANY REQUESTS.
DEVILBISS 525 OXYGEN CONCENTRATOR | K071397 | DEVILBISS HEALTHCARE WAS NOTIFIED OF AN INCIDENT INVOLVING AN OXYGEN CONCENTRATOR BY THE LOCAL AUTHORITIES REGARDING A FIRE THAT OCCURRED IN AN APARTMENT OF A NURSING HOME. THE RESIDENT WAS FOUND TO HAVE DIED IN THE FIRE. THE END USER WAS A KNOWN SMOKER. DRIVE DEVILBISS HEALTHCARE IS CURRENTLY INVESTIGATING THE INCIDENT, INCLUDING ATTEMPTING TO INSPECT THE DEVICE IN COOPERATION WITH THE LOCAL AUTHORITIES.
INTRAOSSEOUS, INTRAOSSEOUS NEEDLE KITS | K203193 | IT WAS REPORTED BY THE CUSTOMER THAT WHILE USING THE 55MM NEEDLE IT BENT. THE PATIENT WAS A FULL CODE AND REQUIRED A SECONDARY SITE AT THE HUMORAL HEAD THAT WAS UTILIZED WITH SUCCESS. AFTER RESUSCITATION EFFORTS, PT WAS CONFIRMED DEAD ON SCENE FOLLOWING OUR MEDICAL DIRECTOR OFF LINES. NO FURTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. EACH EVENT REPORTED TO BD IS EVALUATED AND INVESTIGATED IN ACCORDANCE WITH OUR COMPLAINT INVESTIGATION PROCEDURES. THE INVESTIGATION PROCESS INCLUDES, BUT IS NOT LIMITED TO, EVALUATION OF THE EVENT DETAILS PROVIDED BY THE COMPLAINT FACILITY, A REVIEW OF COMPLAINT HISTORY, MANUFACTURING RECORDS AND RISK DOCUMENT WHERE APPLICABLE, AND AN EVALUATION OF THE SUBJECT DEVICE WHEN AVAILABLE TO IDENTIFY POTENTIAL CONTRIBUTING FACTORS. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION.
NC TREK | K110134 | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE LEFT ANTERIOR DESCENDING (LAD) CORONARY ARTERY WITH MILD CALCIFICATION, MILD TORTUOSITY AND 80% STENOSIS. THE 4.0X15 MM NC TREK WAS USED FOR POST-DILATATION AFTER STENT IMPLANTATION. THE BALLOON WAS INFLATED TWICE TO 10-12 ATMOSPHERES BUT WOULD NOT DEFLATE. NEGATIVE PRESSURE WAS HELD FOR 2 MINUTES BUT THE BALLOON COMPLETELY FAILED TO DEFLATE. SEVERAL GUIDE WIRES AND A MICROCATHETER WERE ADVANCED TO ATTEMPT TO REMOVE THE DEVICE AS WELL AS AN ATTEMPT TO ADVANCE A 1X10 MM BALLOON. EVENTUALLY THE BALLOON WAS REMOVED STILL INFLATED TOGETHER WITH THE GUIDING CATHETER AND FORCE WAS APPLIED BUT THE DISTAL END (TIP) OF THE DEVICE AND ITS SHEATH DETACHED AND REMAINED IN THE ANATOMY. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR BYPASS SURGERY AND THE PATIENT ULTIMATELY EXPIRED ON 02/18/2025. NO ADDITIONAL INFORMATION WAS PROVIDED. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS; HOWEVER, A PICTURE WAS PROVIDED BY THE ACCOUNT. THE PICTURE SHOWS A SEPARATION ON THE DISTAL SHAFT, CONFIRMING THE REPORTED SEPARATION. PRODUCTION RECORD AND CORRECTIVE AND PREVENTATIVE ACTIONS (CAPA) REVIEWS WERE PERFORMED AND REVEALED NO INDICATION OF A PRODUCT QUALITY ISSUE. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO OTHER SIMILAR COMPLAINTS FROM THIS LOT. A MEDICAL REVIEW WAS CONDUCTED BY AN ABBOTT VASCULAR CLINICAL SPECIALIST. THIS WAS A CASE OF A PATIENT WHO UNDERWENT AN INTERVENTIONAL PROCEDURE ON (B)(6) 2025. AFTER STENT IMPLANTATION IN THE PROXIMAL LAD, A 4.00*15 NC TREK BALLOON WAS INFLATED TO A NORMAL PRESSURE OF 12 ATM. UPON DEFLATION, THE BALLOON COULD NOT BE RE- FOLDED. INCREASING THE PRESSURE AND ATTEMPTING DEFLATION AGAIN, THE BALLOON STILL COULD NOT BE RETRACTED. EVENTUALLY, THE BALLOON WAS WITHDRAWN TOGETHER WITH THE GUIDING CATHETER (GC), BUT THE DISTAL END OF THE BALLOON AND ITS SHEATH REMAINED IN THE PATIENT'S BODY. THE PATIENT WAS TRANSFERRED TO ANOTHER HOSPITAL FOR OPEN-HEART SURGERY AND UNFORTUNATELY EXPIRED 1-DAY POST-PROCEDURE. DURING THE PROCEDURE, THE PRESSURE PUMP OF THE BALLOON EXPERIENCED PRESSURE DECAY, AND THE BALLOON COULD NOT BE RETRACTED. A NEW GUIDE WIRE (A TOTAL OF 4) WAS IMMEDIATELY SENT UNDER THE SUPPORT OF A MICROCATHETER. ATTEMPTS TO RETRACT THE BALLOON WITH THE GUIDE WIRES WERE UNSUCCESSFUL. AN ATTEMPT TO PASS A 1.0*10MM PRE-DILATION BALLOON ALSO FAILED. AFTER SEVERAL REPORTEDLY FORCEFUL ATTEMPTS TO REMOVE THE BALLOON CATHETER WERE MADE, THE ENTIRE SYSTEM WAS REMOVED AS A SINGLE UNIT. HOWEVER, IT APPEARS THAT THE BALLOON FRACTURED AND SEPARATED AT SOME POINT DURING THE RETRIEVAL ATTEMPTS. UPON REMOVAL OF THE ENTIRE SYSTEM TOGETHER, THE DISTAL TIP AND THE SHAFT OF THE BALLOON REPORTEDLY REMAINED IN THE LM ARTERY. FORCE WAS APPLIED UPON REMOVAL CAUSING SEPARATION OF THE NC TREK AND PORTION OF THE BALLOON WAS RETAINED IN THE CORONARY ARTERY. A BALLOON THAT REMAINS INFLATED IN A CORONARY ARTERY WILL COMPLETELY OCCLUDE THE VESSEL AND THIS CAN SOON LEAD TO CRITICAL COMPLICATIONS, SO BALLOON ENTRAPMENT REQUIRES EMERGENCY BAILOUT, UNLIKE OTHER ENTRAPMENT OF OTHER DEVICES. ENTRAPMENT OF BALLOON WITH A FULLY OR PARTIALLY DEPLOYED BALLOON CAN ALSO RESULT IN ACUTE OCCLUSION OF CORONARY ARTERY. THESE CATHETER REMNANTS CAN BE READILY RECOVERED FROM THE CORONARY ARTERIES WITH THE USE OF VARIOUS INTERVENTIONAL DEVICES WITH A HIGH SUCCESS RATE. IN THE EVENT OF FAILED INTERVENTIONAL RETRIEVAL AND PERSISTENT SIGNS OF ISCHEMIA, PATIENTS ARE URGENTLY REFERRED TO EMERGENT CARDIAC SURGERY. THE EMERGENCY BYPASS SURGICAL PROCEDURE IS DETERMINED PRIMARILY BY THE CLINICAL STATUS OF THE PATIENT BEFORE THE OPERATION. FOR PATIENTS IN CARDIOGENIC SHOCK, THE RISK OF DEATH REMAINS HIGH AROUND 20% TO 50%. ADDITIONAL RISK FACTORS THAT CAN LEAD TO COMPLICATIONS DEPEND UPON THE DURATION, EXTENT, AND LOCATION OF ISCHEMIA. WITHOUT THE CINE IMAGES, THE NC TREK DEVICE, OR ADDITIONAL INFORMATION AVAILABLE, IT CANNOT BE DEFINITIVELY STATED THAT THE NC TREK WAS THE DIRECT CAUSE OF DEATH IN THIS PATIENT. HOWEVER, WITH THE INFORMATION PRESENTED, IT APPEARS THAT IT MOST LIKELY CONTRIBUTED TO THE PATIENT'S DEATH. IT WAS REPORTED BY THE ACCOUNT THAT THE BALLOON WAS REMOVED FULLY INFLATED. IT SHOULD BE NOTED THE CORONARY DILATATION CATHETERS (CDC), NC TREK RX, GLOBAL, INSTRUCTION FOR USE (IFU, INSTRUCTIONS FOR USE SECTION) STATES: WITHDRAW THE DEFLATED DILATATION CATHETER AND GUIDE WIRE FROM THE GUIDING CATHETER THROUGH THE HEMOSTATIC VALVE. IN THIS CASE, IT IS LIKELY THAT THE IFU VIOLATION CONTRIBUTED TO THE REPORTED DIFFICULTY REMOVING THE DEVICE. ADDITIONALLY, THE ACCOUNT REPORTED THAT FORCE WAS USED DURING REMOVAL, AS DIFFICULTY WAS EXPERIENCED. IT SHOULD BE NOTED THAT THE CORONARY DILATATION CATHETERS (CDC), NC TREK RX, GLOBAL, INSTRUCTION FOR USE (IFU, WARNING SECTION) STATED: IF RESISTANCE IS FELT, DETERMINE THE CAUSE BEFORE PROCEEDING. CONTINUING TO ADVANCE OR RETRACT THE CATHETER WHILE UNDER RESISTANCE MAY RESULT IN DAMAGE TO THE VESSELS AND / OR DAMAGE / SEPARATION OF THE CATHETER. IN THIS CASE, IT IS LIKELY THAT THE COMBINATION OF THE IFU DEVIATION AND THE REPORTED DIFFICULTY REMOVING THE DEVICE CONTRIBUTED TO THE REPORTED SEPARATION. THE REPORTED PATIENT EFFECT OF DEATH IS LISTED IN THE CORONARY DILATATION CATHETERS (CDC), NC TREK RX, GLOBAL, INSTRUCTIONS FOR USE AS A KNOWN PATIENT EFFECT. BASED ON THE INFORMATION PROVIDED, A DEFINITIVE CAUSE FOR THE REPORTED FAILURE TO DEFLATE COULD NOT BE DETERMINED. POSSIBLE FACTORS THAT MAY CONTRIBUTE TO FAILURE TO DEFLATE THE BALLOON INCLUDE, BUT ARE NOT LIMITED TO, DEFLATION TECHNIQUE, LOOSE CONNECTION WITH THE INDEFLATOR, CONTAMINATION IN THE INFLATION LUMEN OR DAMAGE TO THE GUIDE WIRE AND/OR INFLATION LUMEN. THE REPORTED DIFFICULTY REMOVING THE DEVICE AND SEPARATION APPEAR TO BE RELATED TO OPERATIONAL CONTEXT. IN THIS CASE, IT IS LIKELY THAT SINCE THE BALLOON WOULD NOT DEFLATE, THE BDC BECAME STUCK, RESULTING IN THE REPORTED DIFFICULTY REMOVING THE DEVICE FROM THE GUIDING CATHETER. ADDITIONAL TREATMENT WITH SEVERAL GUIDE WIRES AND A MICROCATHETER WAS ATTEMPTED TO REMOVE THE INFLATED BALLOON BUT WERE UNSUCCESSFUL. UPON FURTHER MANIPULATION OF THE DEVICE, AGAINST RESISTANCE AND WITH FORCE, THE DEVICE ULTIMATELY SEPARATED AND REMAINED IN THE PATIENT ANATOMY. THE REPORTED UNEXPECTED MEDICAL INTERVENTION, FOREIGN BODY IN PATIENT, SURGICAL INTERVENTION AND HOSPITALIZATION APPEAR TO BE RELATED TO CIRCUMSTANCES OF THE PROCEDURE. IT CANNOT BE DEFINITIVELY STATED THAT THE NC TREK WAS THE DIRECT CAUSE OF DEATH IN THIS PATIENT; HOWEVER, WITH THE INFORMATION PRESENTED, IT APPEARS THAT IT MOST LIKELY CONTRIBUTED TO THE PATIENT'S DEATH. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO THE DESIGN, MANUFACTURE, OR LABELING OF THE DEVICE. H6: MEDICAL DEVICE PROBLEM CODE 2017: EXCESSIVE FORCE. H6: MEDICAL DEVICE PROBLEM CODE 2017: FAILURE TO FOLLOW STEPS / INSTRUCTIONS.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION DEVICE. THERE IS AN ALLEGATION THAT THE DEVICE WAS NOT WORKING CORRECTLY, AND THE PATIENT WAS HAVING PROBLEMS BREATHING. THE PATIENT WAS TAKEN TO AN EMERGENCY ROOM WHERE THEY PASSED FROM HYPERTENSION AND CARDIAC ARREST. THE PATIENT¿S FAMILY ALSO REPORTED THAT THE UNIT WAS BEING CLEANED WITH AN OZONE-BASED DISINFECTION DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
MINICAP WITH POVIDONE-IODINE SOLUTION | K152129 | IT WAS REPORTED THAT THE PATIENT EXPERIENCED AN UNSPECIFIED INFECTION AFTER USING BAXTER MINICAPS AND SUBSEQUENTLY PASSED AWAY. THE CAUSE OF THE UNSPECIFIED INFECTION WAS NOT REPORTED. HOSPITALIZATION AND TREATMENT FOR THE UNSPECIFIED INFECTION WAS NOT REPORTED. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. | SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
FILMARRAY® MENINGITIS/ENCEPHALITIS (ME) PANEL | K160462 | SUMMARY: (B)(6) HÔPITAL (QUEBEC, CANADA) REPORTED A POTENTIAL FALSE NEGATIVE CRYPTOCOCCUS NEOFORMANS/GATTII RESULT ON THE FILMARRAY MENINGITIS/ENCEPHALITIS (ME) PANEL AFTER TESTING A PATIENT'S CEREBROSPINAL FLUID (CSF) SAMPLE. DUE TO THE FILMARRAY ME PANEL RESULT, THE PATIENT'S DIAGNOSIS AND TREATMENT WAS DELAYED. THE CUSTOMER REPORTED THAT THE PATIENT PASSED AWAY AND THEIR DEATH APPEARED TO BE MULTIFACTORIAL, BUT THE DIAGNOSTIC DELAY POSSIBLY CONTRIBUTED TO THE UNFAVORABLE OUTCOME FOR THE PATIENT. UPON INVESTIGATION, NO MALFUNCTION OCCURRED AND THE FILMARRAY ME PANEL WAS WORKING WITHIN SPECIFICATION. THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT C. NEOFORMANS/GATTII RESULT WAS SENSITIVITY/SPECIFICITY DIFFERENCES BETWEEN THE FILMARRAY ME PANEL AND TRADITIONAL LABORATORY METHODS. | INVESTIGATION: THE PATIENT WAS A 71-YEAR-OLD MALE WHO WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2024, WITH NONSPECIFIC SYMPTOMS. ON (B)(6) 2025, MORE PRONOUNCED NEUROLOGICAL SYMPTOMS APPEARED, LEADING TO CSF COLLECTION BY LUMBAR PUNCTURE. ON (B)(6) 2025, THE CSF SAMPLE WAS TESTED ON THE FILMARRAY ME PANEL. THE FILMARRAY ME PANEL REPORTED ALL ANALYTES AS NOT DETECTED. ON (B)(6) 2025, THE CSF CULTURE WAS POSITIVE FOR CRYPTOCOCCUS, WHICH PROMPTED THE CUSTOMER TO PERFORM CRYPTOCOCCAL ANTIGEN (CRAG) TEST ON THE CSF SAMPLE, CONFIRMING CRYPTOCOCCUS. TREATMENT WAS INITIATED FOLLOWING THE POSITIVE CRYPTOCOCCUS RESULT. THE CUSTOMER REPORTED THAT THIS RESULTED IN A PROCESSING DELAY OF 9 DAYS ((B)(6) 2024 TO (B)(6) 2025). PLEASE NOTE THAT THE DELAY ASSOCIATED WITH THE FILMARRAY ME PANEL RESULT WAS 4 DAYS AS THE TEST WAS PERFORMED ON (B)(6) 2025. ACCORDING TO THE CUSTOMER, CRYPTOCOCCUS WAS INITIALLY RULED OUT DUE TO THE NEGATIVE FILMARRAY ME PANEL RESULT, THE ABSENCE OF RISK FACTORS, AND A NON-SUGGESTIVE CLINICAL PICTURE. THE CUSTOMER REPORTED THAT DUE TO THE COMPLEXITY OF THE CASE, THE PATIENT WAS TRANSFERRED TO A TERTIARY CENTER ON (B)(6) 2025, AND DIED ON (B)(6) 2025. WHEN ASKED IF THE PATIENT WAS AFFECTED BY THE FILMARRAY ME PANEL RESULT, THE CUSTOMER STATED THAT IT WAS DIFFICULT TO ASSESS WITH CERTAINTY WHETHER THE PATIENT WAS DIRECTLY HARMED BY THE DIAGNOSTIC DELAY, HOWEVER, A DELAY IN TREATMENT IN A SERIOUS INFECTION SUCH AS CRYPTOCOCCOSIS, CLINICAL COURSE CAN BE INFLUENCED. THEY STATED, "IN THIS CASE, THE PATIENT'S DEATH APPEARS TO BE MULTIFACTORIAL, BUT THIS DIAGNOSTIC DELAY POSSIBLY CONTRIBUTED TO THE UNFAVORABLE OUTCOME." BIOFIRE MEDICAL AFFAIRS ALSO PROVIDED AN ASSESSMENT OF THE CASE. MEDICAL AFFAIRS STATED THAT THE GOLD STANDARD FOR DETECTING CRYPTOCOCCUS IS CULTURE AND CRYPTOCOCCAL ANTIGEN (CRAG). MOLECULAR METHODOLOGIES CAN PROVIDE VALUE IN ESTABLISHING THE ETIOLOGY OF DISEASE; HOWEVER, PERFORMANCE IS INSUFFICIENT TO RULE OUT INFECTION. A GLOBAL GUIDELINE FOR THE DIAGNOSIS OF CRYPTOCOCCOSIS RECOMMENDS THE DIAGNOSTIC WORKUP FOR A PATIENT WITH SUSPECTED INFECTION TO REQUIRE ASSESSMENT FOR CENTRAL NERVOUS SYSTEM (CNS), PULMONARY, AND OTHER SYSTEM INVOLVEMENT. SPECIFIC TO CNS INFECTION, THE RECOMMENDATIONS ARE FOR THE MEASUREMENT OF CSF OPENING PRESSURE, GLUCOSE, PROTEIN, CELL COUNTS, MICROSCOPY, CULTURE, AND QUANTIFICATION OF CSF CRAG.1 THE MAJOR CONTRIBUTING FACTORS FOR THE DIAGNOSTIC DELAY WAS: DEVIATION FROM CLINICAL GUIDELINES FOR EVALUATING PATIENTS SUSPECTED FOR CRYPTOCOCCAL INFECTION CUSTOMER MISCONCEPTION OF THE PERFORMANCE OF MOLECULAR METHODS FOR CRYPTOCOCCUS DETECTION. REFERENCE: 1. GLOBAL GUIDELINE FOR THE DIAGNOSIS AND MANAGEMENT OF CRYPTOCOCCOSIS: AN INITIATIVE OF THE ECMM AND ISHAM IN COOPERATION WITH THE ASM - THE LANCET INFECTIOUS DISEASES. ACCESSED (B)(6) 2025. HTTPS://WWW.THELANCET.COM/JOURNALS/LANINF/ARTICLE/PIIS1473-3099(23)00731-4/FULLTEXT. IN ADDITION TO THE CUSTOMER INFORMATION AND MEDICAL AFFAIRS ASSESSMENT, ANALYSIS OF THE FILMARRAY ME PANEL RUN FILE SHOWED NO AMPLIFICATION OR MELT SIGNATURE FOR THE C. NEOFORMANS/GATTII ASSAY. QUALITY CONTROL (QC) RECORDS FOR POUCH LOT# 3FBH24 (KIT LOT# 1625324) WERE REVIEWED. THIS POUCH LOT PASSED QC CRITERIA AND WAS FOUND WITHIN SPECIFICATIONS. NO RUN MALFUNCTION OCCURRED AT THE CUSTOMER SITE AND THE FILMARRAY INSTRUMENT (SERIAL NUMBER# (B)(6)) WAS WORKING WITHIN DESIGNED SPECIFICATIONS. CONCLUSION: THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT C. NEOFORMANS/GATTII RESULT WAS SENSITIVITY/SPECIFICITY DIFFERENCES BETWEEN THE FILMARRAY ME PANEL AND TRADITIONAL LABORATORY METHODS. ACCORDING TO THE FILMARRAY ME PANEL INSTRUCTIONS BOOKLET (WWW.ONLINE-IFU.COM/ITI0035), PATIENTS WITH A SUSPICION OF CRYPTOCOCCAL MENINGITIS AND A NEGATIVE CRYPTOCOCCAL PCR RESULT, SUCH AS BY THE FILMARRAY ME PANEL, SHOULD BE TESTED FOR CRYPTOCOCCAL ANTIGEN (CRAG). ACCORDING TO THE "CRYPTOCOCCUS DETECTION BY THE BIOFIRE® FILMARRAY® MENINGITIS/ENCEPHALITIS (ME) PANEL" TECHNICAL NOTE [FLM1-PRT-0278] , THERE HAVE BEEN REPORTS OF NEGATIVE RESULTS BY THE FILMARRAY ME PANEL IN PATIENTS WITH NEWLY DIAGNOSED CRYPTOCOCCAL MENINGITIS AND POSITIVE CRAG AND/OR CULTURE. ADDITIONALLY, DISCORDANT CRAG AND FILMARRAY ME PANEL RESULTS HAVE BEEN OBSERVED IN PERSONS EITHER ON ANTIFUNGAL TREATMENT AT THE TIME OF TESTING OR WITH A PAST HISTORY OF TREATED DISEASE IN WHICH THE CRAG REMAINS POSITIVE. THE FILMARRAY ME PANEL DETECTS NUCLEIC ACID WITHIN THE CSF SAMPLE, WHILE CRAG TEST DETECTS THE CRYPTOCOCCUS ANTIGEN IN THE CSF SAMPLE. PCR-BASED DIAGNOSIS OF CRYPTOCOCCAL MENINGITIS HAS NOT BEEN WIDELY DEVELOPED GIVEN THE HIGH SENSITIVITY, WIDE AVAILABILITY, AND LOW COST OF CRAG TESTING. A NEGATIVE FILMARRAY ME PANEL RESULT DOES NOT EXCLUDE THE POSSIBILITY OF CNS INFECTION AND SHOULD NOT BE USED AS THE SOLE BASIS FOR DIAGNOSIS, TREATMENT, OR OTHER MANAGEMENT DECISIONS. CLINICAL PERFORMANCE: ACCORDING TO TABLE 9. FILMARRAY ME PROSPECTIVE CLINICAL PERFORMANCE SUMMARY OF THE FILMARRAY ME PANEL INSTRUCTION BOOKLET (WWW.ONLINE-IFU.COM/ITI0035), THE PERFORMANCE CLAIM FOR THE C. NEOFORMANS/GATTII ASSAY COMPARED TO PCR WITH BI-DIRECTIONAL SEQUENCING SHOWED AN OVERALL POSITIVE PERCENT AGREEMENT (PPA) OF 100% (1/1) AND AN OVERALL NEGATIVE PERCENT AGREEMENT (NPA) OF 99.7% (95% CI 99.3-99.9%) C. NEOFORMANS/GATTII WAS DETECTED IN 2/4 FALSE POSITIVE SPECIMENS USING A COMMERCIALLY AVAILABLE ANTIGEN TEST. ONE FALSE POSITIVE SPECIMEN WAS POSITIVE BY STANDARD CULTURE. TO SUPPLEMENT THE RESULTS OF THE PROSPECTIVE CLINICAL STUDY, AN EVALUATION OF 235 PRESELECTED ARCHIVED SPECIMENS (OF WHICH 25 WERE NEGATIVE) WAS PERFORMED. ACCORDING TO TABLE 14. FILMARRAY ME PANEL ARCHIVED SPECIMEN PERFORMANCE DATA SUMMARY OF THE FILMARRAY ME PANEL INSTRUCTION BOOKLET, THE C. NEOFORMANS/GATTII ASSAY COMPARED TO A CONFIRMATORY MOLECULAR TEST (E.G., PCR WITH BI-DIRECTIONAL SEQUENCING) SHOWED A PPA OF 100% (95% CI 83.2-100%) AND AN NPA OF 100% (95% CI 97.8-100%). ONE SPECIMEN WAS SEQUENCED AND IDENTIFIED AS C. GATTII AND 18 WERE C. NEOFORMANS.
LIFE SCOPE PT | K220976 | THE CUSTOMER REPORTED THAT A PATIENT PASSED AWAY. | THE CUSTOMER REPORTED THAT A PATIENT PASSED AWAY. THE CUSTOMER DISCOVERED A 20-MINUTE GAP WITH NO CLINICAL ALARMS TRIGGER TO INDICATE THAT THE PATIENT WAS CONTINUALLY DESATURATING OR THAT THE EKG LEADS WERE OFF FOR NURSING TO RESPOND. ACCORDING TO THE CUSTOMER, THE PATIENT HAD ALREADY PASSED AWAY DURING THOSE 20 MINUTES. IT IS POSSIBLE THAT THE PATIENT WAS NEVER DIRECTLY ADMITTED TO THE BEDSIDE MONITOR. PER THE CUSTOMER, IN THE ER IT'S COMMON FOR PATIENTS NOT TO BE FORMALLY ADMITTED TO THE BEDSIDE MONITOR AND THE MONITOR MAY BE RUNNING CONTINUOUSLY. THE CUSTOMER LATER CONFIRMED THAT THE PATIENT WAS NOT ADMITTED BY NAME INTO THE SYSTEM, THE PATIENT WAS JUST BEING MONITORED WITHOUT A NAME IN THE ER. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICE(S) WERE USED IN CONJUNCTION WITH THE BSM: CNS: MODEL #:CU-152R SERIAL #: (B)(6). DEVICE MANUFACTURER DATA: 11/26/2021 UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NO MODEL #:PU-681RA SERIAL #: (B)(6). DEVICE MANUFACTURER DATA: 02/04/2022. UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NO.
PENDITURE LAA EXCLUSION SYSTEM | K232295 | MEDTRONIC RECEIVED INFORMATION THROUGH THE CLIP-IT CLINICAL POST-MARKET STUDY, THAT FOLLOWING A PROCEDURE INVOLVING USE OF THE PENDITURE CLIP ON (B)(6) 2025, IT WAS REPORTED THAT THE PATIENT HAD DIED ON (B)(6) 2025. THE ADVERSE EVENT WAS DEEMED BY THE SITE AS POSSIBLY RELATED TO THE CLIP AND AS NOT RELATED TO THE PENDITURE DELIVERY SYSTEM DEVICE. | MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE ADVERSE EVENT IS POSSIBLY RELATED TO THE PENDITURE CLIP AND POSSIBLY RELATED TO THE PENDITURE DELIVERY SYSTEM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT PASSED AWAY WHILE USING A TRILOGY EVO DEVICE. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ALTHOUGH THE CAUSE OF THE DEATH WAS DETERMINED TO BE PROGRESSION OF ALS, IT IS UNKNOWN ABOUT THE CAUSAL RELATIONSHIP WITH THIS DEVICE. THE PATIENT USED THE DEVICE FOR APPROXIMATELY 8 MONTHS. PRELIMINARY EVALUATION FROM THE SALES REPRESENTATIVE REVIEW OF CARE ORCHESTRATOR, SHOWS "THE LAST USE DATE/TIME WAS FROM THE EVENING OF (B)(6) 2025 TO THE MORNING OF (B)(6) (2025). (THE SALES REPRESENTATIVE'S OPINION IS) PROBABLY, THE DEVICE REMAINED STANDBY, AND THEY DID NOT PRESS START THERAPY". THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A PATIENT PASSED AWAY WHILE USING A TRILOGY EVO DEVICE. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ALTHOUGH THE CAUSE OF THE DEATH WAS DETERMINED TO BE PROGRESSION OF ALS, IT IS UNKNOWN ABOUT THE CAUSAL RELATIONSHIP WITH THIS DEVICE. THE PATIENT USED THE DEVICE FOR APPROXIMATELY 8 MONTHS. PRELIMINARY EVALUATION FROM THE SALES REPRESENTATIVE REVIEW OF CARE ORCHESTRATOR, SHOWS "THE LAST USE DATE/TIME WAS FROM THE EVENING OF (B)(6) TO THE MORNING OF (B)(6) (2025). (THE SALES REPRESENTATIVE'S OPINION IS) PROBABLY, THE DEVICE REMAINED STANDBY, AND THEY DID NOT PRESS START THERAPY." THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH." UPON EVALUATION OF THE DEVICE THE EVENT LOG CONFIRMED THAT THE DEVICE WAS TURNED OFF AT 08:13 ON (B)(6) AND REMAINED INACTIVE UNTIL IT ENTERED STANDBY MODE AT 03:36 ON (B)(6). AN OPERATIONAL CHECK WAS PERFORMED USING A BREATHING CIRCUIT, BUT NO FAILURES WERE CONFIRMED. FUNCTIONAL TESTS OF THE PRESSURE SENSOR AND FLOW SENSOR WERE PERFORMED, AND NO FAILURES WERE CONFIRMED. THE DEVICE WAS DISASSEMBLED, AND AN INTERNAL INSPECTION WAS CARRIED OUT, BUT NO FAILURES WERE CONFIRMED. BASED ON THE ABOVE RESULTS, NO FAILURES WERE CONFIRMED WITH THE DEVICE, AND IT HAS BEEN DETERMINED TO BE OPERATING NORMALLY. PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY. SOFTWARE VERSION WAS CHECKED (VER.1.06.10.00).
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A DREAMSTATION AUTO CPAP. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE CONTRIBUTED TO OR CAUSED THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
CATHERA | K151638 | MADJIDYAR, J., KELLER, E., WINKLHOFER, S., TOTH, D., BARNAURE, I., SCHUBERT, T., THURNER, P., FIERSTRA, J., WILLMS, J. F., REGLI, L., & KULCSAR, Z. (2023). SINGLE-ANTIPLATELET REGIMEN IN RUPTURED CEREBRAL BLOOD BLISTER AND DISSECTING ANEURYSMS TREATED WITH F LOW-DIVERTER STENT RECONSTRUCTION. JOURNAL OF NEUROINTERVENTIONAL SURGERY, 15(10), 953¿957. HTTPS://DOI.ORG/10.1136/JNIS-2022-019361. MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A SINGLE-CENTER CASE SERIES OF NINE PATIENTS TREATED FOR RUPTURED BLOOD BLISTER ANEURYSMS (BBA) OR DISSECTING ANEURYSMS WITH ASSOCIATED SUBARACHNOID HEMORRHAGE (SAH) USING PIPELINE FLEX WITH SHIELD TECHNOLOGY FLOW DIVERSION STENTS FROM 2019 TO 2022. THE CASES WERE OFF-LABEL USE FOR THE PIPELINE EMBOLIZATION DEVICES DUE TO USE OF SINGLE ANTIPLATELET TREATMENT AS DUAL ANTIPLATELET TREATMENT IS RECOMMENDED PER THE INSTRUCTIONS FOR USE (IFU). FURTHERMORE, PER THE IFU, THE PIPELINE¿ FLEX EMBOLIZATION DEVICE WITH SHIELD TECHNOLOGY¿ SHOULD NOT BE USED ALONE AS SOLE THERAPY FOR ACUTELY RUPTURED ANEURYSMS. THE PURPOSE OF THE ARTICLE WAS TO INVESTIGATE THE EFFICACY OF A SINGLE ANTIPLATELET REGIMEN TO PREVENT ANTIPLATELET TREATMENT RELATED HEMORRHAGIC COMPLICATIONS. ALL PIPELINES WERE IMPLANTED WITHOUT PROCEDURAL ADVERSE EVENTS AND NO DEVICE MALFUNCTION WAS REPORTED IN THE ARTICLE. TWO PATIENTS ALSO HAD COIL EMBOLIZATION IN ADDITIONAL TO THE STENT IMPLANTATION. ONE PATIENT HAD TWO PIPELINES IMPLANTED TELESCOPING IMPLANTED IN THE INITIAL PROCEDURE FOR FULLY AFFECTIVE RECONSTRUCTION OF THE CAROTID SEGMENT. IT SHOULD BE NOTED THAT THE PHENOM 27 MICROCATHETER WAS USED IN THE FLOW DIVERSION STENTING PROCEDURES. ONE PATIENT DEATH WAS REPORTED: PATIENT 7 WAS TREATED FOR A FUSIFORM DISSECTION ANEURYSM OF THE RIGHT VERTEBRAL ARTERY AND SUFFERED A LETHAL RE-RUPTURE SIX DAYS AFTER THE INDEX PROCEDURE. THE FOLLOWING NON-LETHAL ADVERSE EVENTS WERE REPORTED: IN PATIENT 3, A SECOND PIPELINE WAS IMPLANTED 14 DAYS AFTER THE INITIAL TREATMENT PROCEDURE DUE TO INCREASING SIZE OF THE ANEURYSM. THREE PATIENTS HAD NEW NEUROLOGICAL DISABILITY REPORTED DURING FOLLOW-UP. TWO PATIENTS HAD A NEW MODERATE DISABILITY (MRS 3 AT 6¿10 MONTHS) AND ONE PATIENT HAD A NEW MODERATELY SEVERE DISABILITY (MRS 4 AT 3 MONTHS). THE CAUSE OF THE NEW NEUROLOGICAL DEFICIT WAS NOT SPECIFIED IN THE ARTICLE. TWO PATIENT WERE REPORTED TO HAVE SUBTOTAL ANEURYSM OCCLUSION AT 3 AND 6 MONTHS WITH NO LATER FOLLOW UP AVAILABLE AND NO ADDITIONAL TREATMENT, INTERVENTION, OR ASSOCIATED SYMPTOMS REPORTED IN THE ARTICLE. IT SHOULD BE NOTED THAT THERE WERE ADVERSE EVENTS REPORTED IN THE ARTICLE WHICH WERE NON-COMPLAINT EVENTS PER 027-WI185, DUE TO CAUSE BEING ATTRIBUTED TO THE PATIENTS' PRESENTING CONDITION OF SAH, NOT WITH THE MEDTRONIC DEVICE OR THE PROCEDURE. THE NON-COMPLAINT EVENTS INCLUDED: SEVERE VASOSPASM ASSOCIATED WITH THE SAH OCCURRED IN TWO PATIENTS, LEADING TO DELAYED CEREBRAL ISCHEMIA. FOUR PATIENTS NEEDED VENTRICULOPERITONEAL SHUNTING DUE TO SAH-RELATED HYDROCEPHALUS. | SEPARATE REPORTS WILL BE SUBMITTED FOR THE SERIOUS INJURIES REPORTED IN THE ARTICLE THAT DID NOT HAVE AN OUTCOME OF MORTALITY. A2. REPORTED PATIENT AGE IS REPRESENTATIVE OF THE MEAN AGE OF ALL PATIENTS INCLUDED IN THE STUDY GROUP. A3. REPORTED PATIENT SEX IS REPRESENTATIVE OF THE MAJORITY (5/9) PATIENTS INCLUDED IN THE STUDY GROUP. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
NXSTAGE SYSTEM ONE | K122051 | A REPORT WAS RECEIVED ON 19 FEB 2025 FROM THE HOME THERAPY NURSE (HTN) OF A 55-YEAR-OLD MALE PATIENT WITH A MEDICAL HISTORY INCLUDING END STAGE RENAL DISEASE, WHO STATED THE PATIENT EXPIRED DURING A HOME HEMODIALYSIS TREATMENT ON (B)(6) 2025. ALTHOUGH REQUESTED, ADDITIONAL INFORMATION WAS NOT PROVIDED. | THE DEVICE WAS NOT RECEIVED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD (DHR) WAS CONDUCTED WHICH CONFIRMED THAT THE DEVICE MET ALL QUALITY CRITERIA AND MANUFACTURING SPECIFICATIONS PRIOR TO RELEASE. THERE WAS NO INDICATION OF A DEVICE MALFUNCTION FROM THE AVAILABLE INFORMATION. THE NXSTAGE SYSTEM ONE USER GUIDE OUTLINES RISKS ASSOCIATED WITH PERFORMING HEMODIALYSIS THERAPY AND WARNS ALL TREATMENTS MUST BE ADMINISTERED UNDER A PHYSICIAN'S PRESCRIPTION AND MUST BE PERFORMED BY A TRAINED AND QUALIFIED PERSON, CONSIDERED TO BE COMPETENT IN THE USE OF THIS DEVICE BY THE PRESCRIBING PHYSICIAN. UDI: (B)(4).
RESPIRONICS | K102985 | PHILIPS RECEIVED A COMPLAINT FROM THE CUSTOMER, REPORTING THAT THE V60 VENTILATOR DID NOT ALARM WHEN THE DEVICE WAS DISCONNECTED. THE DEVICE WAS IN CLINICAL USE WHEN THE ISSUE OCCURRED. IT WAS REPORTED THAT THE EVENT WAS FATAL TO THE PATIENT. A FOLLOW UP WAS PERFORMED WITH THE KEY MARKET (KM), AND THE RESPONDER WAS UNABLE TO PROVIDE ANY DETAILS REGARDING THE PATIENT STATUS. THE KM RESPONDER REPORTED THAT THEY HAVE REACHED OUT TO THE HOSPITAL SEVERAL TIMES; HOWEVER, THE HOSPITAL STAFF HAVE NOT PROVIDED ANY DETAILS. CLINICAL ASSESSMENT HAS FOUND NO CAUSAL RELATIONSHIP OF THE DEVICE TO THE PATIENT EXPIRATION, HOWEVER CONTRIBUTION OF THE DEVICE AND ITS USE TO THE PATIENT EXPIRATION HAS BEEN CONFIRMED: FORESEEABLE UNINTENTIONAL USE ERROR - ADDITION OF RESISTIVE COMPONENTS TO THE PATIENT SIDE OF THE PROXIMAL PRESSURE LINE RESULTING IN COMPROMISED PATIENT DISCONNECTION ALARM FUNCTIONALITY. THE USER REPORTED THAT THE PATIENT WAS CONNECTED TO THE V60 VENTILATOR, BUT AFTER THE FOLLOWING MORNING, THE USER FOUND THAT THE VENTILATOR HAD BEEN DISCONNECTED FROM THE PATIENT, AND THAT THE DEVICE DID NOT ALARM (PATIENT DISCONNECT). AFTER SEVERAL INQUIRIES, AND A VIDEO FOOTAGE OF THE DEVICE, IT WAS OBSERVED THAT USER HAD USED TWO BACTERIA FILTER (MAKE/MODEL UNSPECIFIED) CONNECTED TO THE DEVICE. PER THE INSTRUCTIONS FOR USE (IFU), IT IS ADVISED TO "AVOID ADDING ANY COMPONENTS TO THE PATIENT CIRCUIT THAT ARE NOT ABSOLUTELY NECESSARY. ADDITIONAL COMPONENTS INSTALLED IN THE PATIENT CIRCUIT CAN CHANGE THE PRESSURE GRADIENT ACROSS THE VENTILATOR BREATHING SYSTEM, INCREASE THE DEAD SPACE, AND ADVERSELY AFFECT THE VENTILATOR PERFORMANCE." IN ADDITION, THE IFU STATES, "AVOID ADDING RESISTIVE CIRCUIT COMPONENTS ON THE PATIENT SIDE OF THE PROXIMAL PRESSURE LINE. SUCH COMPONENTS MAY DEFEAT THE DISCONNECT ALARM." THE DEVICE WAS THEN EVALUATED BY PHILIPS SERVICE ENGINEER (SE), AND IT WAS REPORTED THE RESPIRATORY CIRCUIT WAS CHECKED, AND IT WAS CONFIRMED THAT A FILTER BETWEEN THE PATIENT AND EXHALATION PORT, AND A NEBULIZATION DEVICE BETWEEN THE FILTER AND EXHALATION PORT. WHEN THE SE DISCONNECTED THE CIRCUIT CONSUMABLE, THE PATIENT DISCONNECT ALARM IS RE-ESTABLISHED. THIS INVESTIGATION IS ONGOING. | BASED ON THE INFORMATION PROVIDED, THE EVENT HAS SHOWN THAT THE DEVICE PERFORMED AS INTENDED WITH NO FAILURE OR MALFUNCTION TO PERFORM TO MANUFACTURER DECLARE SPECIFICATIONS. THE ADDITIONAL COMPONENTS ADDED TO THE PATIENT SIDE OF THE PROXIMAL PRESSURE LINE RESULTED IN THE DEVICE PATIENT DISCONNECTION ALARM BEING DEFEATED. FURTHER GOOD FAITH EFFORTS YIELDED NO ADDITIONAL INFORMATION REGARDING THE PATIENT STATUS. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED, THE COMPLAINT FILE WILL BE REOPENED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
BIPAP AUTOSV ADVANCED SYSTEM ONE 60 SERIES | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO AN INFORMATION REQUEST CONCERNING DISPOSING OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
CARDIOSAVE HYBRID, TYPE B PLUG | K112372 | N/A | UPDATED FIELDS - B4, E1(POSTAL CODE), G3, G6, H2, H11. CORRECTED FIELDS - D10.
CARDIOSAVE HYBRID, TYPE B PLUG | K112372 | N/A. | UPDATED : B4, G3, G6, H2, H6(MEDICAL DEVICE ¿ PROBLEM CODE), H11. CORRECTED FIELDS - D10.
SENSATION PLUS 7.5FR. 40CC IAB & ACCESSORIES | K122628 | IT WAS REPORTED THAT THE INTRA-AORTIC BALLOON (IAB) CATHETER EXHIBITED BLOOD BACKFLOW ON UNIT (B)(6). DESPITE THIS OBSERVATION, THE SAME CATHETER WAS SUBSEQUENTLY USED ON UNIT (B)(6). AFTER THE PATIENT WAS DECLARED DECEASED, ATTEMPTS WERE MADE TO REMOVE THE BALLOON; HOWEVER, REMOVAL WAS UNSUCCESSFUL DUE TO CLOT FORMATION. AS A RESULT, THE CATHETER HAD TO BE CUT FOR REMOVAL. | UPDATED FIELDS: B4, G3, G6, H2, H6 (TYPE OF INVESTIGATION, MEDICAL DEVICE ¿ PROBLEM CODE, INVESTIGATION FINDINGS, COMPONENT CODES, INVESTIGATION CONCLUSIONS), H11. CORRECTED FIELDS: D7A, D10. NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. THE FAILURE MODE IS ADDRESSED IN THE RISK FILE AND IS OPERATING WITHIN ITS RISK PROFILE. THE IFU ADDRESSES THE REPORTED FAILURE. THE INVESTIGATION DOES NOT INDICATE THAT THE DEVICE WAS INADVERTENTLY RELEASED AS NON-CONFORMING OR AN ADULTERATED PRODUCT OR WAS A COUNTERFEIT. EACH IAB MANUFACTURED IS 100% INSPECTED AND THE CONTROLS PUT IN PLACE DURING THE MANUFACTURING OF THE IAB WOULD CATCH A DEFECT AND NOT ALLOW NON-CONFORMING DEVICE TO BE RELEASED. THE TREND REVIEW DID NOT IDENTIFY AN ADVERSE TREND (INCREASE IN NUMBER OF COMPLAINTS OVER PAST THREE (3) MONTHS). BASED ON THE RATIONAL PROVIDED ABOVE, NO ESCALATION TO THE CAPA PROCESS IS REQUIRED. COMPLAINT ID: (B)(4).
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION ST30 DEVICE. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE OPERATED AS DESIGNED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
HEMOSTAR | K051748 | IT WAS REPORTED THAT THE PATIENT STARTED HEMODIALYSIS IN THE LEFT FOREARM UNDER THE INTERNAL FISTULA PUNCTURE, ULTRAFILTRATION FOR ABOUT TWO HOURS. THE NURSE WAS READY TO GIVE ROUTINE CARE TO THE PATIENT'S INTRACERVICAL DIALYSIS CATHETER, AND IMMEDIATELY FOUND THAT THE PATIENT'S RESPIRATION WAS SHORT OF BREATH. THE NURSE REPORTED TO THE DOCTOR IMMEDIATELY, PLACED THE PATIENT IN LYING POSITION, AND STOP ULTRAFILTRATION. IT WAS IDENTIFIED THAT THE LONG-TERM TUBE ARTERIAL END OF THE CLAMP, 1CM BELOW THE CATHETER WAS RUPTURED, THE CATHETER CLAMP INTACT, WAS IN CLAMPED CLOSED STATE. REPORTEDLY, THE PATIENT EXPERIENCED CONFUSION, CALLED OUT, BILATERAL PUPIL DILATION, RESPIRATION 29-35 TIMES / MIN, HEART RATE 120-136 TIMES / MIN, OXYGEN SATURATION 65-69%, AND HEMODIALYSIS TREATMENT WAS TERMINATED. THE HCP ACTIVELY RESUSCITATED THE CAROTID ARTERY AFTER THE CTA REPORT SUGGESTS THAT SOME OF THE INTRACEREBRAL VEINS AND FASCIA GAS SHADOW IN THE SINUS, CONSIDERED THE POSSIBILITY OF AN AIR EMBOLISM. AFTER THAT, THE PATIENT WAS TRANSFERRED TO ICU FOR TREATMENT. FURTHERMORE, THE TREATMENT WAS STILL INEFFECTIVE, AND THE PATIENT WAS PRONOUNCED DEAD. | H10: MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED, AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: ONE ELECTRONIC PHOTO WAS REVIEWED. THE PHOTOGRAPH SHOWED AN INSERTED DIALYSIS CATHETER WITH EXTENSION LEGS AND SOME OF THE BIFURCATION SIGNIFICANTLY YELLOWED, AS WELL AS THE MARKER INSERTS ON THE CLAMPS. THE TWO HUBS WERE BEING HELD IN A GLOVED HAND, BEING SURROUNDED LOOSELY BY GAUZE. BOTH CLAMPS APPEARED TO BE CLOSED. THE RED (ARTERIAL) EXTENSION LEG OF THE CATHETER HAD A LARGE, EVEN, CLEAN PARTIAL DISCONTINUITY ON THE PROXIMAL SIDE OF THE CLOSED CLAMP THAT APPEARED TO BE A CUT, BASED ON ITS EVENNESS AND APPARENT SMOOTHNESS. THE CUT WAS ROUGHLY TRANSVERSE, THOUGH AT A SHALLOW ANGLE RELATIVE TO THE TUBING AXIS. NO BLEEDING/LEAKAGE FROM THE CATHETER WAS EVIDENT, BUT AIR WAS VISIBLE WITHIN THE EXTENSION LEG ON THE SIDE OF THE CUT. THEREFORE, BASED ON THE PHOTO/VIDEO REVIEW, THE REPORTED RUPTURE WILL BE MORE SPECIFICALLY DESCRIBED AS A CUT, AND CONFIRMED, AND THE PROBLEMS OF DISCOLORATION AND AIR WITHIN THE DEVICE CAN BE CONFIRMED ALSO BASED ON THE PROVIDED PHOTOS/VIDEO. A DEFINITIVE ROOT CAUSE FOR THE DISCOLORATION, CUT, AND AIR LEAK COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE A LABELING REVIEW IS NOT REQUIRED. THE INFORMATION PROVIDED BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
TUBING SET | K112360 | COMPLAINT ID# (B)(4). | RECEIVED A REQUEST FROM THE FDA ON APRIL 7, 2025 STATING-"THIS IS TO BRING TO YOUR ATTENTION THAT WE ARE IN RECEIPT OF REPORT NUMBER 3013876692-2025-00019-1 WITH THE REPORT DATE 03/20/2025. UPON REVIEW OF THE INFORMATION INCLUDED FOR THE REPORT, IT REFERENCES RELATED MEDICAL DEVICE REPORT (MDR) (B)(4) IDENTIFIED BY YOUR FIRM FOR THE SAME EVENT. HOWEVER, THIS INFORMATION WAS NOT INCLUDED IN THE CORRESPONDING BLOCK H10, RELATED REPORT NUMBER FIELD, OF MEDWATCH FORM 3500A." THIS SUBMISSION SERVES TO ADDRESS THIS ERROR NOTED ABOVE.
AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | K163000 | IT WAS REPORTED THAT ON (B)(6) 2025, AN AMPLATZER AMULET LEFT ATRIAL APPENDAGE (LAA) OCCLUDER WAS CHOSEN FOR IMPLANTATION UTILIZING A 14F TORQVUE 45X45 DELIVERY SHEATH. DURING ADVANCEMENT OF THE DELIVERY SYSTEM, RESISTANCE WAS ENCOUNTERED AND ADVANCEMENT WAS STOPPED. ON FLUOROSCOPY, THE SHEATH WAS WRAPPED AROUND ITSELF AND KINKED SO THE DELIVERY SYSTEM WAS REMOVED AND REPLACED. A 18F GROIN SHEATH WAS PLACED AND THE REPLACEMENT 14F TORQVUE 45X45 DELIVERY SYSTEM WAS ADVANCED. THE DELIVERY SYSTEM EXPERIENCED RESISTANCE WHEN TRYING TO CROSS THE SEPTUM. THE DELIVERY SYSTEM WAS REMOVED AND THE SEPTUM WAS DILATED. THE DELIVERY SYSTEM WAS THEN ABLE TO CROSS THE SEPTUM AND ADVANCED TO THE LAA. BLOOD PRESSURE BECAME UNSTABLE SO THE CASE WAS ABORTED. NO EFFUSION OR GROIN COMPLICATIONS WERE OBSERVED. ONCE PATIENT WAS BROUGHT TO THE RECOVERY AREA BLOOD PRESSURE DROPPED. A HEMATOMA WAS NOTED IN GROIN AREA. MANUAL COMPRESSION WAS PERFORMED, AND THE PATIENT CODED AND DIED. IT IS BELIEVED THAT THE PATIENT HAD A BLEED AT THE GROIN SITE AND ENTERED CARDIOGENIC SHOCK DUE TO A LOW EJECTION FRACTION. THE CAUSE OF THE BLEED IS UNKNOWN AS AN ARTERIAL LINE HAD ALSO BEEN INSERTED. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF DIFFICULT TO ADVANCE THROUGH PATIENT ANATOMY WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION, THE CAUSE OF THE REPORTED DIFFICULT TO ADVANCE THROUGH PATIENT ANATOMY COULD NOT BE DETERMINED. THE PATIENT EFFECTS OF CARDIAC ARREST, HEMORRHAGE, HEMATOMA, AND HYPOTENSION COULD NOT DETERMINED. DEATH WAS RELATED TO PATIENT'S COMORBIDITIES. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING.
BENEVISION N17 PATIENT MONITOR | K213799 | DURING AN INVESTIGATION FOLLOW-UP ON APRIL 7, 2025, THE CUSTOMER CLARIFIED THE INVESTIGATION REQUEST. NO ALLEGED MALFUNCTION WAS REPORTED. THE CUSTOMER IS REQUESTING THE KEYSTROKE FOR THE A7 ANESTHESIA MACHINE (SERIAL NUMBER (B)(6)) AND THE BENEVISION N17 PATIENT MONITOR (SERIAL NUMBER (B)(6)). THE REQUEST IS FOR RAW DATA FROM 09:00 AM TO 11:00 AM ON (B)(6) 2025. SOME VENTILATION CHANGES WERE MADE TO THE ANESTHESIA MACHINE AT 10:28. THEREFORE, THE CUSTOMER REQUESTED TO PROVIDE DETAILS ON THE KEYSTROKE ACTIVITY ON THE ANESTHESIA MACHINE AND BEDSIDE MONITOR 10 MINUTES BEFORE AND 10 MINUTES AFTER THAT TIME. THE PATIENT EXPIRED DAYS AFTER (B)(6) 2025 (THE EXACT DATE IS UNKNOWN). | A MINDRAY FIELD SERVICE REPRESENTATIVE COLLECTED THE PATIENT DATABASE AND THE BENEVISION N17 MONITOR LOGS FOR REVIEW. MINDRAY PROVIDED THE BENEVISION N17 KEYSTROKE ACTIVITY FOR THE REQUIRED PERIOD. A MINDRAY SERVICE REPRESENTATIVE, UPON CLARIFYING THE REQUEST, RETURNED TO THE SITE TO COLLECT LOGS FROM THE ANESTHESIA UNIT. HOWEVER, SINCE THE ATTEMPT TO COLLECT DATA WAS ON (B)(6) 2025, THE DATA FOR THE OCCURRENCE DATE ((B)(6) 2025) IS NO LONGER AVAILABLE ON THE A-SERIES. NO DEVICE MALFUNCTION WAS REPORTED WITH THE MINDRAY DEVICES. UNITS ARE MONITORED PER SPECIFICATIONS.
INTRACEPT RF PROBE | K180369 | IT WAS REPORTED THAT THE PATIENT HAD AN INTRACEPT PROCEDURE AND THE PROCEDURE WAS COMPLETED SUCCESSFULLY. HOWEVER, THE PATIENT HAD AN ISSUE WITH THE ANESTHESIA. THE PATIENT WAS TRANSPORTED TO THE EMERGENCY ROOM (ER) POST-PROCEDURE. THE PATIENT PASSED AWAY AFTER SOME ANESTHESIA CHALLENGES POST-PROCEDURE. THE PHYSICIAN ACKNOWLEDGED THAT THE PATIENT'S DEATH WAS NOT A PROCEDURE OR DEVICE-RELATED COMPLICATION. | CORRECTION TO INITIAL EMDR IN BLOCK D4.
MERGE CARDIO | K192276 | MERGE CARDIO IS A SYSTEM INTENDED TO BE USED TO ACQUIRE, STORE, PRINT, TRANSFER, AND ARCHIVE CLINICAL INFORMATION FROM MERGE HEALTHCARE AND OTHER VENDORS SYSTEMS INCLUDING IMAGES, HEMODYNAMIC STUDIES AND REPORTS, MEASUREMENTS (VIA IMPORT FROM DICOM STRUCTURED REPORTING, TEXT FILES OR OPTICAL CHARACTER RECOGNITION OF MEASUREMENTS CAPTURED ON IMAGES) AND CARDIOLOGY SIGNAL (WAVEFORM) DATA. MERGE CARDIO IS INTENDED TO ALLOW USERS TO REVIEW DIAGNOSTIC AND NON-DIAGNOSTIC QUALITY IMAGES, ANNOTATE STUDIES, PERFORM DIGITAL SUBTRACTION ON IMAGES, TO PERFORM QUANTITATIVE MEASUREMENTS ON IMAGES (INCLUDING BUT NOT LIMITED TO QUANTITATIVE CORONARY ANALYSIS, LEFT VENTRICULAR ANALYSIS, TIME, AREA, LENGTH, VELOCITY, ANGLE, VOLUME, AND VELOCITY-TIME INTEGRALS), TO GENERATE PHYSICIAN GENERATED CLINICAL REPORTS (VIA STRUCTURE REPORTING AND TEMPLATE BASED TOOLS), AND TO STORE THIS INFORMATION IN A DATABASE. ON 02/19/2025, A CUSTOMER CONTACTED MERGE HEALTHCARE FOR ASSISTANCE WITH MERGE CARDIO'S STUDY LIST FILTERS. MERGE TECHNICAL SUPPORT REMOTELY CONNECTED TO THE CUSTOMER'S SYSTEM TO ASSIST WITH FILTER QUESTIONS. THE CUSTOMER WAS ATTEMPTING TO LOCATE A ECHOCARDIOGRAM PATIENT STUDY PERFORMED ON (B)(6) 2025 AND WAS UNABLE TO CREATE A NEW FILTER TO LOCATE THE IMAGING STUDY. TECHNICAL SUPPORT PROVIDED ASSISTANCE TO ENABLE THE CUSTOMER TO CREATE THE STUDY LIST FILTER. MERGE TECHNICAL SUPPORT ASSISTED THE CUSTOMER WITH FINDING THE STUDY AND CONFIRMED THAT THE NEW STUDY LIST FILTER WORKED AS EXPECTED. THE CUSTOMER REPRESENTATIVE MENTIONED TO TECHNICAL SUPPORT THAT THE PATIENT HAD PASSED AWAY ON (B)(6) 2025. MERGE HEALTHCARE FOLLOWED UP WITH THE CUSTOMER TO GATHER MORE INFORMATION AND DETERMINED THAT THE PATIENT OUTCOME WAS DUE TO A PREEXISTING CONDITION. BASED ON THE AVAILABLE INFORMATION, MERGE CARDIO IS PERFORMING AS DESIGNED AND DID NOT CAUSE OR CONTRIBUTE TO THE PATIENT OUTCOME. IF NEW INFORMATION BECOMES AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A PHILIPS CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
BIPOLAR VENTRICULAR ENDOCARDIAL LEAD | K900550 | VOLUNTARY MEDWATCH FORMS MW5166910, MW5166911 RECEIVED STATE: IT WAS REPORTED THAT THE PATIENT IMPLANTED WITH THIS (B)(6) PACEMAKER AND FOUR NON-(B)(6) LEADS (TWO CAPPED/ABANDONED AND TWO ACTIVE) UNDERWENT A LEAD EXTRACTION PROCEDURE DUE TO VEGETATION ON THE LEADS. UPON REMOVAL OF THE ABANDONED RIGHT ATRIAL (RA) LEAD MODEL 1018T, THE PATIENT'S BLOOD PRESSURE DROPPED AND AN EMERGENCY THORACOTOMY WAS PERFORMED. THE PATIENT HAD A TEAR IN THE RIGHT ATRIUM AND SUPERIOR VENA CAVA (SVC). THE PATIENT DEVELOPED DISSEMINATED INTRAVASCULAR COAGULATION (DIC), THE DAMAGED AREAS OF THE HEART COULD NOT BE REPAIRED, AND THE PATIENT DIED ON THE OPERATING TABLE. THE EXPLANTED PACEMAKER WAS PLANNED TO BE RETURNED FOR DISPOSAL. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). | CORRECTION: SECTION E2 - "HEALTH PROFESSIONAL?" SHOULD HAVE BEEN "NO INFORMATION" CORRECTION: CUSTOMER SECTION SHOULD HAVE BEEN LEFT BLANK AS THE ORIGINAL REPORT SOURCE IS NOT KNOWN AND THE INFORMATION WAS RECEIVED VIA THE FDA. NOTE: HOSPITAL FACILITY, PATIENT DETAILS AND SPECIFIC PRODUCT DETAILS WERE NOT PROVIDED. FURTHER INFORMATION WAS NOT AVAILABLE.
RADIOGRAPHY 7300 C | K233662 | IT WAS REPORTED THAT DURING A PATIENT EXAMINATION, AN OLD AGE PATIENT FELL OFF FROM THE TABLE WHILE THE OPERATOR WAS IN THE CONTROL ROOM. SIX DAYS LATER, THE PATIENT PASSED AWAY IN THE HOSPITAL, THOUGH IT'S UNCLEAR WHETHER THE FALL CONTRIBUTED TO THE DEATH. THE FALL OCCURRED ONLY ONCE AND WAS PROMPTLY FOLLOWED UP WITH A CT SCAN FOR INJURIES. THERE WAS NO IMPACT TO THE USER. THE PATIENT IMPACT IS DEATH. | THE RADIOGRAPHY 7300 C IS USED IN GENERAL RADIOGRAPHIC EXAMINATIONS AND APPLICATIONS. IT IS SUITABLE FOR ALL ROUTINE RADIOGRAPHY EXAMINATIONS, INCLUDING SPECIALIST AREAS LIKE INTENSIVE CARE, TRAUMA OR PEDIATRIC WORK, EXCLUDING FLUOROSCOPY, ANGIOGRAPHY AND MAMMOGRAPHY. STATEMENT OF THE REPORTED PROBLEM: PHILIPS RECEIVED FEEDBACK THROUGH CFMS (CUSTOMER FEEDBACK MANAGEMENT SYSTEM) THAT A PATIENT FELL FROM THE TABLE OF THE RADIOGRAPHY 7300 C SYSTEM DIED 6 DAYS LATER. THE SEQUENCE OF EVENTS IS PROVIDED BELOW. ACCORDING TO THE HEALTH CARE PROVIDERS AT THE CUSTOMER LOCATION, ON (B)(6) 2025, 01:15 PM A FEMALE PATIENT, AGE 89 YEARS, UNDERWENT RADIOLOGICAL EXAMINATION WITH THE RADIOGRAPHY 7300 C SYSTEM. RADIOGRAPHS OF THE LUMBAR SPINE WERE REQUESTED FOR VERTEBRAL FRACTURES AS THE PATIENT WAS EXPERIENCING INCREASED PAIN IN THE LUMBAR SPINE AND PELVIS. PATIENT WAS STABLE BUT LIMITED IN COOPERATION. A ROLL BOARD WAS USED TO TRANSFER THE PATIENT FROM THE BED ONTO THE X-RAY TABLE, WITH NO CONFIRMATION OF ANY RESTRAINTS OR HAND GRIPS BEING USED. JUST AS THE OPERATOR STEPPED INTO THE CONTROL ROOM TO START THE SCAN, THE PATIENT TURNED WHILE ON THE TABLE AND FELL OFF THE TABLE ONTO THE FLOOR. SHORTLY AFTER THE FALL, THE PATIENT WAS STABLE AND CONSCIOUS. THE PATIENT WAS EXAMINED FOR INJURIES DUE TO THE FALL, AND A CT SCAN OF THE HEAD WAS PERFORMED THAT SHOWED GENERALIZED INVOLUTION AND PATCHY, INHOMOGENEOUS MARGIN DEGENERATION, AND PREVIOUSLY KNOWN FALX CALCIFICATIONS. THE CT SHOWED NO EVIDENCE OF A CALVARIAL FRACTURE, NO BLEEDING, NO CONTUSION, AND NO ACUTE CSF CIRCULATION DISTURBANCE. AFTER THE FALL, THE CASE WAS ALSO PRESENTED TO THE SURGEON ON DUTY AND THE PATIENT UNDERWENT ASSESSMENT IN THE EMERGENCY ROOM, BUT NO FURTHER DETAILS REGARDING THE TREATMENT WERE PROVIDED BY THE HEALTH CARE PROVIDERS AT THE CUSTOMER LOCATION. AFTER SIX DAYS, ON (B)(6) 2025, THE PATIENT DIED. THE DEATH CERTIFICATE AND OFFICIAL CAUSE OF DEATH WERE NOT PROVIDED DUE TO DATA PROTECTION REGULATIONS IN GERMANY. ANALYSIS PERFORMED: A SERVICEMAX CASE WAS NOT CREATED FOR THIS CASE AS THERE WAS NO ALLEGATION OF PRODUCT MALFUNCTION FROM THE CUSTOMER. NO FIELD SERVICE ENGINEER VISITED THE SITE TO EVALUATE THE SYSTEM. HOWEVER, COMPLAINT WAS ESCALATED FOR A TECHNICAL INVESTIGATION TO EVALUATE THE SYSTEM PERFORMANCE AT THE TIME OF THE INCIDENT ((B)(6) 2025). THE SYSTEM LOG FILES FROM THE DATE OF INCIDENT (B)(6) 2025 WERE RETRIEVED AND EVALUATED BY SYSTEM ENGINEER. THE RESULTS OF LOG FILE ANALYSIS ARE DOCUMENTED IN THE TECHNICAL INVESTIGATION CHILD RECORD- (B)(4). SUMMARY OF ANALYSIS OF LOG FILE RECEIVED (I.E. ST.VINZENZ-KRANKENHAUSHANAU_RADIOGRAPHYDR7300C_RAUMB_20250417_1345_SYSLOG.ZIP): 1) AROUND THE EVENT TIME I.E. 25TH OF FEBRUARY 2025, 01:15 PM, IT WAS OBSERVED THERE ARE NO ERROR LOGS REPORTED ON THE TABLE OR ANY GEOMETRY COMPONENTS. 2) THERE IS NO TABLE MOVEMENT INITIATED DURING THIS TIME. 3) THERE IS NO EVIDENCE OF ANY EXPOSURE INITIATED BY THE OPERATOR DURING THIS TIME. THE INSTRUCTIONS FOR USE (IFU, 3000 140 63981) FOR THE RADIOGRAPHY 7300 C DEVICE CONTAINS INFORMATION ABOUT THE USE OF HAND GRIPS (P286) FOR PATIENT SUPPORT, SIDE BARS (P291) AND ADJUSTABLE STRAPS (P293) FOR PATIENT SUPPORT. TWO TABLE TYPES CAN BE USED WITH THE RADIOGRAPHY 7300 C DEVICE, THE HEIGHT ADJUSTABLE TABLE (TH2) AND THE SINGLE SIDE SUSPENDED TABLE (TH-S). ADJUSTABLE STRAPS AND A SIDE BAR CAN BE USED WITH THE TH-S, WHILE HAND GRIPS CAN BE USED WITH THE TH2 TO HELP STABILIZE THE PATIENT. THE INSTRUCTIONS FOR USE (IFU, 3000 140 63981) FOR THE RADIOGRAPHY 7300 C DEVICE WARNS THAT THE PATIENT SHOULD BE SECURED AND UNDER CONSTANT OBSERVATION WHILE ON THE TABLE (P143) AND ADVISES THAT THE PATIENT USE HAND GRIPS AND/OR STRAPS IF NECESSARY (P144). FURTHER INVESTIGATION REVEALED THAT THE CUSTOMER DID NOT PURCHASE HAND GRIPS FOR TABLE TH-2 FROM PHILIPS OR FROM OTHER VENDORS. CUSTOMER DID NOT USE ANY MATTRESS AT THE TIME OF THE INCIDENT AS CONFIRMED BY THE FIELD SERVICE ENGINEER ON 02 MAY 2025. AS NO DEVICE FAULTS WERE REPORTED BY THE CUSTOMER AND NO DEVICE ERRORS WERE FOUND DURING SUBSEQUENT LOG FILE ANALYSIS, EVIDENCE DOES NOT EXIST TO SUGGEST THAT A DEVICE MALFUNCTION CAUSED OR CONTRIBUTED TO INJURY OR DEATH OF THE PATIENT.
INTELLIVUE X3 | K171801 | IT WAS REPORTED AFTER A RECENT NEW DEVICE INSTALLATION, THERE WAS AN INTENSIVE CARE UNIT (ICU) PATIENT WHO PASSED AWAY. THE PATIENT WAS IN BRADYCARDIA AND LAYING DOWN IN THEIR BED AWAITING TRANSFER TO THE WARD BUT STAFF LATER FOUND THEM DECEASED. THE CUSTOMER INDICATED THERE ARE MORE ALARMS WITH THE NEW MONITORS, ALARMS WERE MISSED OR IGNORED BY STAFF, WHICH MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. IN ADDITION, THE CUSTOMER INDICATED THERE WERE ISSUES WITH THE WAY OF WORKING IN THE ICU AS WELL. THE CUSTOMER INVESTIGATED THE LOGS, REVEALING TECHNICAL INOP ALARMS FOR ECG PATCH AND SPO2 CONNECTIONS; HOWEVER, IT WAS FOUND THE DEVICES BEHAVED PER SPECIFICATION. | REPORTING INSTITUTION PHONE#: (B)(6). REPORTER PHONE#: (B)(6). A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS NOTIFIED THAT A PATIENT WAS IN BRADYCARDIA AND LAYING DOWN IN THEIR BED AWAITING TRANSFER TO THE WARD BUT STAFF LATER FOUND THEM DECEASED. THE CUSTOMER INDICATED THERE ARE MORE ALARMS WITH THE NEW MONITORS, ALARMS WERE MISSED OR IGNORED BY STAFF, WHICH MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. IN ADDITION, THE CUSTOMER INDICATED THERE WERE ISSUES WITH THE WAY OF WORKING IN THE INTENSIVE CARE UNIT (ICU) AS WELL. THE CUSTOMER INVESTIGATED THE LOGS, REVEALING TECHNICAL INOP ALARMS FOR ECG PATCH AND SPO2 CONNECTIONS; HOWEVER, IT WAS FOUND THE DEVICES BEHAVED PER SPECIFICATION. ADDITIONALLY, SOME PATIENT DATA WAS NOT AVAILABLE TO SEND TO PHILIPS AS THE PATIENT WAS NOT ADMITTED TO THE SYSTEM. ADMISSION TO THE SYSTEM IS A NEW WORKFLOW FOR THE CUSTOMER WITH THE NEW MONITORS. A PHILIPS REMOTE SERVICE ENGINEER (RSE) SUPPORTED THE CUSTOMER TO COLLECT THE LOGS. THE EVENT TOOK PLACE IN ROOM 9 ON (B)(6) 2025 AT 1930-2255 HOURS. THE CUSTOMER VERIFIED THAT A 3-LEAD ECG CABLE WAS IN USE. THE DATA WAS SENT TO A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL SPECIALIST (PCS) FOR REVIEW AND A SUMMARY OF THE FINDINGS IS DOCUMENTED BELOW. THE LOGS SHOW THAT THE MONITOR WAS TAKEN OUT OF STANDBY AT 19:35:56 (B)(6) 2025 19:35:56 (B)(6), RUM 9 STANDBY AVSLUTAS MT34463 EXITING STANDBY. AT 19:35:52, ECG LEADS OFF IS ANNOUNCED. (B)(6) 2025 19:35:57 (B)(6), RUM 9 EKG LÖSA ELEKTR SKAPADES 19:35:52. PIC IX: KGAIXAKUTC1 LOGGED INOP. AT 19:37:00, HEART RATE ALARMS ARE SWITCHED OFF AND PULSE BECOMES THE ALARM SOURCE. THE ECG/ARRHYTHMIA ALARMS ARE CHANGED FROM ON TO OFF. FROM THE IFU, ¿SELECTING PULSE AS THE ACTIVE ALARM SOURCE FOR HR/PULSE SWITCHES OFF THE ARRHYTHMIA ALARMS, INCLUDING ASYSTOLE, VFIB AND VTACH ALARMS, AND THE HEART RATE ALARMS. THIS IS INDICATED BY THE MESSAGE ECG/ARRH ALARMSOFF (UNLESS THIS HAS BEEN CONFIGURED OFF FOR YOUR MONITOR), AND THE CROSSED-OUT ALARM SYMBOL BESIDE THE ECG HEART RATE NUMERIC. THE MESSAGE ECG/ARRH ALARMSOFF CAN BE CONFIGURED OFF, OR TO SWITCH TO A YELLOW (MEDIUM SEVERITY) INOP AFTER A FIXED NUMBER OF HOURS. HIGH AND LOW PULSE RATE AND EXTREME BRADYCARDIA AND EXTREME TACHYCARDIA ALARMS FROM PULSE ARE ACTIVE.¿ IF THE ALARM SOURCE IS SET TO AUTO, THIS SWITCHING OFF MAY BE AUTOMATICALLY FOLLOWING A MANUAL REQUEST TO TURN ON. WHEN THE ALARM SOURCE IS SET TO AUTO, THE MONITOR WILL USE THE HEART RATE FROM THE ECG MEASUREMENT AS THE ALARM SOURCE WHENEVER THE ECG MEASUREMENT IS SWITCHED ON AND AT LEAST ONE ECG LEAD CAN BE MEASURED WITHOUT AN INOP CONDITION. SEE DATA SET B FOR MANY EXAMPLES OF ALARM ON/OFF ENTRIES. AT 19:37:07. **SPO2 86<90 WAS GENERATED. AT 19:37:16, **PULSE 45<50 WAS GENERATED. 19:37:27, ALARMS WERE ACKNOWLEDGED AT MT34463 BEDSIDE MONITOR. **SPO2 86<90 ENDED AT 19:37:57 AND THEN GENERATED AGAIN AT 19:38:22. **PULSE 40<50 ENDED AT 19:38:02. ** PULSE 35<50 GENERATED AT 19:38:21. THESE ALARMS ENDED AT 19:38:29 WHEN THE ***BRADY/PULSE ALARM 26<40 GENERATED. THE ***BRADY/PULSE ALARM ENDED WITH A MAXIMUM VIOLATION OF 25<40 AT 19:38:38. THE ECG/ARRHYTHMIA ALARMS WERE SWITCHED ON AND OFF SEVERAL TIMES BETWEEN 19:37 AND 22:46:30. AT 19:41:32, **PULS 46<50 GENERATED AND ENDED AT 19:42:21 WITH A MAXIMUM OF VIOLATION OF 35. **PULSE 46<50 GENERATED AT 19:43:09 AND ENDED AT 19:44:02. AT 19:44:24, **PULS 42<50 GENERATED. AT 19:44:40, ECG LEAD OFF ENDED AND THE ECG/ARRHYTHMIA ALARM SWITCHED BACK ON. THIS RESULTED IN THE **PULSE ALARM ENDING AND THE SOUNDING OF THE *HR 48<50 ALARM AT 19:44:46. THE AUDIT LOG CONTINUES TO SHOW *HR ALARMS GENERATING AND ENDING UNTIL 20:10:31 WHEN THE ***XBRADY 39<40 GENERATED. THE ALARM ENDED AT 20:10:47 WITH A MAXIMUM VIOLATION OF 36 <40. SPO2 PULSE ? AND SPO2 NO PULSE INOPS ARE NOTED. THE PATIENT CONTINUED TO HAVE LOW HR AND XBRADY ALARMS AT 20:46:06, CANNOT ANALYSIS ECG GENERATED. AT 21:24:59, !!ECG LEADS OFF INOP GENERATED. A DEVICE PROBLEM IS NOT CONFIRMED. THE INVESTIGATION DETERMINED THAT THE DEVICE OPERATED AS EXPECTED. A CUSTOMER RESPONSE LETTER IS BEING PROVIDED TO THE CUSTOMER TO RESOLVE THE ISSUE.
DS2ADV AUTO CPAP W/HUMID+HT CELL/BT | K131982 | THE MANUFACTURER RECEIVED A CALL ON DREANSTATION REPLACEMENT DEVICE, FROM THE PATIENT'S PARTNER CALLED ON HIS BEHALF. THE PATIENT PASSED AWAY AND SHE WAS NOT SURE WHAT TO DO WITH HIS REPLACEMENT DEVICE. SHE WANTED TO RETURN THE DEVICE IF SHE COULD. SHE DID NOT HAVE EMAIL ACCESS. THERE WAS NO REPORT OF PATIENT HARM OR INJURY. THERE WAS NO REPORT OF MEDICAL INTERVENTION. REPORTABLE, ALTHOUGH THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO THE DEATH, CAUSE OF DEATH IS UNKNOWN. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER.
REMSTAR AUTO A-FLEX | K091319 | THE MANUFACTURER HAS BEEN INFORMED OF THE RECENT PASSING OF A PATIENT WHO WAS USING A REMSTAR AUTO A-FLEX DEVICE. ACCORDING TO THE PATIENT'S WIFE, HER HUSBAND RECENTLY PASSED AWAY, AND SHE WAS UNABLE TO LOCATE THE DEVICE. SHE HAS INDICATED THAT SHE WILL RETURN IT IF SHE COMES ACROSS IT WHILE SORTING THROUGH HIS BELONGINGS. NO FURTHER DETAILS HAVE BEEN PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE HAD ANY ROLE IN OR CONTRIBUTED TO THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER HAS BEEN NOTIFIED OF THE RECENT PASSING OF A PATIENT WHO WAS USING A DREAMSTATION CPAP PRO DEVICE. THE PATIENT'S DAUGHTER HAS REQUESTED INFORMATION REGARDING THE RECALL AND WHETHER THE RECALLED DEVICE COULD BE LINKED TO HER MOTHER'S CANCER DIAGNOSIS. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | IT WAS REPORTED THAT THE RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE IS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THERE WERE NO REPORTS OF MEDICAL INTERVENTION. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING EVALUATION, NO ERROR WAS FOUND IN THE DEVICE LOG, AND THE DEVICE WAS AWAITING SCRAPPING CONFIRMATION. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
ZIO AT | K240029 | THE PATIENT EXPIRED DURING THEIR PRESCRIBED ZIO AT WEAR PERIOD. IRHYTHM ATTEMPTED TO GATHER MORE INFORMATION ABOUT THE CAUSE OF DEATH FROM THE ACCOUNT, BUT NO FURTHER DETAILS WERE PROVIDED. | THE INVESTIGATION IS STILL ONGOING, AND A FOLLOW-UP REPORT WILL BE SUBMITTED UPON ITS COMPLETION.
BIOFIRE® BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL | K193519 | SUMMARY: (B)(6) HEALTHCARE REPORTED A POTENTIAL NEGATIVE METHICILLIN-RESISTANT STAPHYLOCOCCUS AUREUS (MRSA) RESULT ON THE BIOFIRE BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL AFTER TESTING A PATIENT'S BLOOD CULTURE SAMPLE. THE CUSTOMER REPORTED THAT DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S TREATMENT WAS IMPACTED, CONTRIBUTING TO THEIR DEATH. UPON INVESTIGATION, NO MALFUNCTION OCCURRED, AND THE BIOFIRE BCID2 PANEL WAS WORKING WITHIN SPECIFICATION. THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT MRSA WAS DUE TO AN ORGANISM STRAIN THAT HAS REDUCED REACTIVITY WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. | INVESTIGATION: THE PATIENT WAS A 9-YEAR-OLD MALE. THE PATIENT PRESENTED WITH SIGNS AND SYMPTOMS OF TACHYPNEA AND TACHYCARDIA. THE CUSTOMER REPORTED THERE WAS ALSO CONCERN FOR SEPSIS WITH ELEVATED C-REACTIVE PROTEIN (CRP) WHEN BLOOD CULTURES WERE DRAWN. THE PATIENT WAS ALREADY PAN-CYTOPENIC AND NEUTROPENIC. WHEN ASKED WHAT THE UNDERLYING CAUSE OF THE PATIENT'S IMMUNOCOMPROMISED STATUS (PAN-CYTOPENIC AND NEUTROPENIC) WAS, THE CUSTOMER COMMUNICATED THE FOLLOWING: "MULTIFACTORIAL. THE PATIENT WAS PREVIOUSLY HEATHY BUT PRESENTED WITH [FEBRILE INFECTION-RELATED EPILEPSY SYNDROME] FIRES AND ALSO DEVELOPED PANCYTOPENIA (OF UNKNOWN ETIOLOGY DESPITE HEME/ONC INVOLVEMENT) AND [HEMOPHAGOCYTIC LYMPHOHISTIOCYTOSIS] HLH DURING THE PATIENT'S PROLONGED COURSE OF ILLNESS. THE PATIENT RECEIVED CYTOKINE DIRECTED THERAPY WITH BOTH IL-6 BLOCKER AND ANAKINRA AS WELL AS HIGH-DOSE STEROIDS FOR FIRES. ALSO RECEIVED IMMUNOSUPPRESSIVE THERAPY FOR HLH." WHEN ASKED IF THE PATIENT WAS DIAGNOSED WITH ANY ORGAN FAILURE AT THE TIME OF TESTING, SUCH AS SEPTIC SHOCK, THE NEED FOR MECHANICAL VENTILATION, OR ANY OTHER FORM OF ORGAN SUPPORT, THE CUSTOMER REPORTED THAT THE PATIENT WAS ALREADY VENTILATED DUE TO NEUROLOGICAL ISSUE AND RENAL FAILURES (UNKNOWN TIME POINT FOR VENTILATION AND RENAL FAILURE) AND THE PATIENT DEVELOPED HEPATIC FAILURE DURING ANTIBIOTIC TREATMENT REQUIRING CONTINUOUS RENAL REPLACEMENT THERAPY (CRRT). ON (B)(6) 2025, THE PATIENT'S BLOOD CULTURE SAMPLE WAS TESTED ON THE BIOFIRE BCID2 PANEL. THE BIOFIRE BCID2 PANEL REPORTED STAPHYLOCOCCUS SPP. AND S. AUREUS AS DETECTED. THE BIOFIRE BCID2 PANEL WAS POSITIVE FOR MECA/C BUT NEGATIVE FOR MREJ, RESULTING IN A NOT DETECTED RESULT FOR 'MECA/C AND MREJ (MRSA).' NOTE THAT THE BIOFIRE BCID2 PANEL WILL ONLY REPORT A DETECTED RESULT FOR 'MECA AND MREJ (MRSA)' IF S. AUREUS, MREJ, AND MECA/C ASSAYS ARE ALL POSITIVE. MICROSCAN SUSCEPTIBILITY TESTING SHOWED OXACILLIN (MIC= >2 UG/ML) RESISTANCE AND WAS CEFOXITIN SCREEN POSITIVE (MIC= >4 UG/ML). THE ABBOTT CLEARVIEW PENICILLIN-BINDING PROTEIN 2A (PBP2A) CULTURE COLONY TEST WAS POSITIVE FOR PBP2A. GRAM STAIN SHOWED GRAM-POSITIVE COCCI IN CLUSTERS. THE CUSTOMER REPORTED ANOTHER BLOOD CULTURE SAMPLE WAS TESTED ON THE BIOFIRE BCID2 PANEL AND WAS POSITIVE FOR MSSA. THE RUN FILE INFORMATION FOR THIS TEST WAS NOT PROVIDED. THE CUSTOMER REPORTED DUE TO THE BIOFIRE BCID2 PANEL RESULT, THE PATIENT'S ANTIBIOTIC TREATMENT WAS CHANGED TO MSSA DIRECTED THERAPY. THIS INAPPROPRIATE CHANGE IN TREATMENT OCCURRED FOR TWO DAYS. THE CUSTOMER STATED THAT THE FALSE NEGATIVE MRSA RESULT ON THE BIOFIRE BCID2 PANEL "DIRECTLY IMPACTED THE PATIENT'S TREATMENT, RESULTING IN THEIR DEATH." THE CUSTOMER COMMUNICATED THE FOLLOWING REGARDING THE PATIENT'S TREATMENT CHANGE: "WHEN THE PATIENT DETERIORATED, INITIALLY, THE PATIENT WAS ON VANCOMYCIN. HOWEVER, ONCE WE RECEIVED THE RESULT OF [BIOFIRE BCID2 PANEL] INDICATING MSSA, WE SWITCHED TO OXACILLIN FOR DIRECTED THERAPY (AND ADDED CLINDAMYCIN FOR TOXIN MEDIATED DISEASE AS HE ALSO DEVELOPED ERYTHRODERMA) GIVEN THE PATIENT ALSO DEVELOPED AKI (WANTED TO AVOID VANC) AS WELL AS RESPIRATORY FAILURE (DAPTO WAS NOT OPTIMAL). GIVEN PANCYTOPENIA, LINEZOLID WAS NOT HIGH ON CUSTOMER OPTIONS INITIALLY. ONCE CULTURE SHOWED MRSA, CUSTOMER SWITCHED OXACILLIN TO CEFTAROLINE FOR MRSA." THE ANALYSIS OF THE INITIAL BIOFIRE BCID2 PANEL RUN FILE PROVIDED BY THE CUSTOMER SHOWED ROBUST AMPLIFICATION SIGNATURES FOR BOTH MECA/C AND S. AUREUS ASSAYS, WHILE NO MELT/AMPLIFICATION SIGNATURE WAS SEEN FOR THE MREJ ASSAY. QUALITY CONTROL (QC) RECORDS FOR POUCH LOT# 3H7D24 (KIT LOT# 1977124) WERE ALSO REVIEWED. THIS POUCH LOT PASSED QC CRITERIA AND WAS FOUND WITHIN SPECIFICATIONS. NO RUN MALFUNCTION OCCURRED AT THE CUSTOMER SITE AND THE FILMARRAY INSTRUMENT (SERIAL NUMBER# (B)(6)) WAS WORKING WITHIN DESIGNED SPECIFICATIONS. IN-HOUSE INVESTIGATION: BIOFIRE RECEIVED ONE TUBE OF ISOLATE AND TWO TUBES OF BLOOD CULTURE ALIQUOTS FOR TESTING. THE TWO BLOOD CULTURE ALIQUOTS WERE TESTED ON THE BIOFIRE BCID2 PANEL. THE BIOFIRE BCID2 PANEL DETECTED S. AUREUS, MECA/C AND MREJ (MRSA) FOR BOTH ALIQUOTS. THE MECA/C AND S. AUREUS ASSAYS EACH HAD ROBUST SIGNATURES, WHILE A LATE SIGNATURE WAS OBSERVED FOR THE MREJ ASSAY. THE ISOLATE WAS TESTED TWICE ON THE BIOFIRE BCID2 PANEL AND S. AUREUS, MECA/C AND MREJ (MRSA) WAS DETECTED BOTH TIMES. AGAIN, THE MECA/C AND S. AUREUS ASSAYS EACH HAD ROBUST SIGNATURES, WHILE A LATE SIGNATURE WAS OBSERVED FOR THE MREJ ASSAY. GROWTH OF ONE MORPHOLOGY TYPE WAS SEEN FROM CULTURE. ISOLATES WERE TESTED USING VITEK® MS AND VITEK® 2, WITH BOTH PLATFORMS IDENTIFYING THE ORGANISM AS S. AUREUS, VITEK® MS REPORTING A HIGH CONFIDENCE. VITEK® 2 AST CONFIRMED MRSA WITH A CEFOXITIN POSITIVE CALL AND OXACILLIN RESISTANCE. BASED ON THE BIOFIRE BCID2 PANEL'S AMPLIFICATION SIGNATURE PATTERN OF S. AUREUS, MECA/C, AND MREJ ASSAYS, INITIAL FINDINGS INDICATED IT WAS LIKELY THAT THE PATIENT SAMPLE CONTAINED A MRSA STRAIN CARRYING AN MREJ VARIANT WITH REDUCED REACTIVITY ON THE BIOFIRE BCID2 PANEL. WHOLE GENOME SEQUENCING WAS PERFORMED TO PROVIDE ADDITIONAL INFORMATION ON THE SPECIFIC MREJ VARIANT. THE MRSA ISOLATE WAS GROWN TO A HIGH TITER IN LIQUID CULTURE TO PREPARE A GLYCEROL STOCK FOR WHOLE GENOME SEQUENCING. ANALYSIS OF TARGET SEQUENCES OBSERVED MULTIPLE MISMATCHES UNDER THE MREJ ASSAY'S PRIMER BINDING SITES. THE ANALYSIS FURTHER INDICATED THE ISOLATE RECOVERED FROM THE PATIENT'S BLOOD CULTURE IS LIKELY AN MREJ VARIANT THAT IS EXPECTED TO EITHER HAVE REDUCED OR NO REACTIVITY (NOT DETECTED) WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. HOWEVER, THE BLOOD CULTURE ALIQUOT AND ISOLATE PROVIDED BY THE CUSTOMER BOTH DETECTED MREJ BY THE BIOFIRE BCID2 PANEL DURING INTERNAL TESTING, SUGGESTING THE MREJ VARIANT CAN BE REPORTED AS DETECTED WHEN THE TARGET NUCLEIC ACID IS PRESENT AT A HIGH CONCENTRATION. ALTERNATIVELY, THE MREJ VARIANT MAY BE REPORTED AS NOT DETECTED WHEN THE TARGET IS NOT PRESENT IN THE BLOOD CULTURE AT HIGH ENOUGH CONCENTRATION, WHICH IS LIKELY THE REASON IT WAS NOT DETECTED AT THE CUSTOMER'S SITE. CONCLUSION: THE INVESTIGATION CONCLUDED THAT THE MOST LIKELY CAUSE FOR THE DISCREPANT MRSA RESULT WAS DUE TO A MREJ VARIANT THAT HAS REDUCED REACTIVITY WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. THE BIOFIRE BCID2 PANEL IS A MULTIPLEXED NUCLEIC ACID TEST INTENDED FOR THE SIMULTANEOUS QUALITATIVE DETECTION AND IDENTIFICATION OF MULTIPLE BACTERIAL AND YEAST NUCLEIC ACIDS AND SELECT GENETIC DETERMINANTS ASSOCIATED WITH ANTIMICROBIAL RESISTANCE. EACH AMR GENE RESULT IS ASSOCIATED WITH A SINGLE CORRESPONDING ASSAY EXCEPT FOR THE MECA/C AND MREJ RESULT, WHICH IS DEPENDENT ON BOTH THE MECA/C ASSAY AND THE MREJ ASSAY. IN ORDER FOR MECA/C AND MREJ (MRSA) TO BE REPORTED AS DETECTED, THE FILMARRAY MUST BE POSITIVE FOR ALL THREE ASSAYS: S. AUREUS, MECA/C, AND MREJ. MORE INFORMATION ON MECA/C AND MREJ ASSAY INTERPRETATION CAN BE FOUND IN TABLE 10 OF THE BIOFIRE BCID2 PANEL INSTRUCTION FOR USE (IFU) (WWW.ONLINE-IFU.COM/ITI0048). THE ANALYSIS OF THE INITIAL BIOFIRE BCID2 PANEL RUN FILE PROVIDED BY THE CUSTOMER SHOWED ROBUST AMPLIFICATION SIGNATURES FOR BOTH MECA/C AND S. AUREUS ASSAYS, WHILE NO MELT/AMPLIFICATION SIGNATURE WAS SEEN FOR THE MREJ ASSAY. HOWEVER, THE CUSTOMER'S SUBCULTURE AND SUSCEPTIBILITY TESTING REVEALED MRSA AND GRAM STAIN SHOWED GRAM-POSITIVE COCCI IN CLUSTERS. TO INVESTIGATE FURTHER, IN-HOUSE TESTING AT BIOFIRE WAS PERFORMED, AND THE CUSTOMER'S RESULT OF A NEGATIVE MREJ/NEGATIVE MRSA COULD NOT BE REPRODUCED. ALL BIOFIRE BCID2 PANEL TESTS RAN INTERNALLY WERE POSITIVE RESULTS FOR MECA/C, S. AUREUS, AND MREJ (MRSA). HOWEVER, IT IS NOTEWORTHY THAT THE MREJ SIGNATURES OBSERVED WERE LATER AND LESS ROBUST THAN THOSE SEEN FOR MECA/C AND S. AUREUS. ADDITIONALLY, ANALYSIS OF THE WHOLE GENOME SEQUENCE IDENTIFIED A MREJ VARIANT THAT WOULD BE EXPECTED HAVE REDUCED REACTIVITY WITH THE BIOFIRE BCID2 PANEL'S MREJ ASSAY. ACCORDING TO THE "METHICILLIN RESISTANCE GENES DETECTION BY THE BIOFIRE BLOOD CULTURE IDENTIFICATION 2 (BCID2) PANEL" TECHNICAL NOTE [BFR0002-4828], METHICILLIN RESISTANCE IN STAPHYLOCOCCUS SPECIES IS PRIMARILY CONFERRED WHEN THE MOBILE SCCMEC GENETIC CASSETTE, CARRYING EITHER OF THE RESISTANCE GENES MECA OR MECC, INTEGRATES INTO A SPECIFIC REGION OF THE STAPHYLOCOCCUS GENOME. IN S. AUREUS, THIS STAPHYLOCOCCAL CASSETTE CHROMOSOME MEC (SCCMEC) INSERTION CREATES A JUNCTION BETWEEN THE S. AUREUS GENOME AND THE SCCMEC CASSETTE THAT IS CALLED MREJ (SCCMEC RIGHT-EXTREMITY JUNCTION). THE BIOFIRE BCID2 PANEL'S MREJ ASSAY IS DESIGNED TO DETECT THIS SPECIFIC INTEGRATION EVENT IN THE S. AUREUS GENOME. THERE ARE A VARIETY OF DIFFERENT MREJ TYPES (I-XXI) CLASSIFIED BY THE POINT OF INSERTION OF SCCMEC INTO S. AUREUS. THE MREJ ASSAY IS PREDICTED TO HAVE IMPAIRED REACTIVITY OR BE NON-REACTIVE TO SOME OF THESE MREJ TYPES AND VARIANT SEQUENCES. THESE MREJ TYPES INCLUDE IX, XV, AND XVIII, AS WELL AS TYPES XIX AND XX (ASSOCIATED WITH METHICILLIN-SENSITIVE S. AUREUS; MSSA), AND MREJ SEQUENCES ANNOTATED FROM NON-AUREUS STAPHYLOCOCCUS SPECIES AND NON-STAPYLOCOCCI ORGANISMS. MORE INFORMATION CAN BE FOUND IN TABLE 118. RESULTS FOR MECA/C AND MREJ(MRSA) IN STAPHYLOCOCCUS AUREUS ISOLATES TESTED AND PREDICTED REACTIVITY FOR MREJ TYPES IN THE BIOFIRE BCID2 PANEL'S IFU. OVERALL, EACH AMR GENE ASSAY ON THE BIOFIRE BCID2 PANEL WAS FOUND TO DETECT THE MAJORITY OF AMR GENE TYPES BASED ON TESTING AND ANALYSIS OF AVAILABLE SEQUENCE DATA. HOWEVER, THE MREJ ASSAY MAY BE AFFECTED BY ORGANISM STRAINS THAT CONTAIN SEQUENCE VARIANTS OR REARRANGEMENTS IN THE GENE TARGETS OF THE ASSAY. THE BCID2 PANEL'S MREJ ASSAY IS DESIGNED TO TARGET THE MOST PREVALENT TYPES OF MREJS, BUT NOT ALL KNOWN TYPES. TYPES KNOWN TO HAVE IMPAIRED OR NO REACTIVITY ARE LISTED IN TABLE 118 OF THE BIOFIRE BCID2 PANEL'S IFU, HOWEVER, THE SPECIFIC MREJ VARIANT RECOVERED FROM THIS PATIENT BLOOD CULTURE WAS UNABLE TO BE DEFINITIVELY LINKED TO THOSE LISTED IN TABLE 118. HOWEVER, BASED ON THE SEQUENCE ALIGNMENT, THIS SEQUENCE IS NOT EXPECTED TO BE DETECTED BY THE MREJ ASSAY ON THE BCID2 PANEL, AND MAY BE A NOVEL MREJ TYPE. NOTE: PER THE "LIMITATIONS" SECTION OF THE BIOFIRE BCID2 PANEL IFU, ANTIMICROBIAL RESISTANCE CAN OCCUR VIA MULTIPLE MECHANISMS. A NOT DETECTED RESULT FOR A GENETIC MARKER OF ANTIMICROBIAL RESISTANCE DOES NOT INDICATE SUSCEPTIBILITY TO ASSOCIATED ANTIMICROBIAL DRUGS OR DRUG CLASSES. CULTURE IS REQUIRED TO OBTAIN ISOLATES FOR ANTIMICROBIAL SUSCEPTIBILITY TESTING, AND BIOFIRE BCID2 PANEL RESULTS SHOULD BE USED IN CONJUNCTION WITH CULTURE RESULTS FOR THE DETERMINATION OF SUSCEPTIBILITY OR RESISTANCE. ALL IDENTIFICATION RESULTS PROVIDED BY THE BIOFIRE BCID2 PANEL ARE INTENDED TO BE INTERPRETED IN CONJUNCTION WITH GRAM STAIN RESULTS. THE BIOFIRE BCID2 PANEL RESULTS MUST BE CORRELATED WITH THE CLINICAL HISTORY, EPIDEMIOLOGICAL DATA, AND OTHER DATA AVAILABLE TO THE CLINICIAN EVALUATING THE PATIENT. CLINICAL PERFORMANCE: CLINICAL PERFORMANCE CAN BE FOUND IN TABLE 65. BIOFIRE BCID2 PANEL CLINICAL PERFORMANCE SUMMARY, MECA/C AND MREJ (MRSA) AND TABLE 81. MECA/C AND MREJ (MRSA) PERFORMANCE TABLE (AS COMPARED TO PHENOTYPIC AST METHODS FOR METHICILLIN (OXACILLIN/CEFOXITIN) RESISTANCE ON CULTURED ISOLATE(S) FROM PROSPECTIVE PBC SPECIMENS) OF THE BIOFIRE BCID2 PANEL IFU (WWW.ONLINE-IFU.COM/ITI0048).
LLD EZ LEAD LOCKING DEVICE | K142116 | A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD DUE TO CIED/POCKET INFECTION. A SPECTRANETICS LLD EZ LEAD LOCKING DEVICE (LLD EZ) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 16F GLIDELIGHT LASER SHEATH, PROGRESS WAS MADE TO THE MID INNOMINATE VEIN. THEN, SWITCHING TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH, ADVANCEMENT WAS MADE INTO THE RIGHT VENTRICLE. WHILE PULLING WITH TRACTION, THE LEAD RELEASED, EXCEPT THE PATIENT¿S BLOOD PRESSURE DROPPED, AND A PERICARDIAL EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). RESCUE EFFORTS BEGAN, INCLUDING PERICARDIOCENTESIS AND STERNOTOMY. AN RV PERFORATION WAS DISCOVERED AND REPAIRED, BUT THE PATIENT¿S EJECTION FRACTION (EF) REMAINED LOW, AND WAS PLACED ON EXTRACORPOREAL MEMBRANE OXYGENATION (ECMO) TO ALLOW THE HEART RECOVER. HOWEVER, THE FOLLOWING DAY, THE DECISION WAS MADE TO WITHDRAW LIFE SUPPORT, AND THE PATIENT DID NOT SURVIVE. THIS REPORT CAPTURES THE LLD EZ DEVICE IN USE WHEN THE RV PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. | A2) PATIENT DATE OF BIRTH - NOT AVAILABLE FROM FACILITY. A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. H3) THE LLD WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) CARDIAC PERFORATION AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE LLD. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, RESPIRATORY TRACT IRRITATION, LUNG DISEASE CANCER, AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S HUSBAND INQUIRED ABOUT HOW TO DISPOSE OF THE DEVICE. HE WAS ADVISED TO CONTACT THE MEDICAL COMPANY THAT PROVIDED IT, AS THEY COULD ASSIST IN PROPERLY DISPOSING OF THE DEVICE. WHEN THE HUSBAND ASKED IF HE COULD SIMPLY THROW IT AWAY, HE WAS RECOMMENDED TO REACH OUT TO THE DOCTOR TO OBTAIN THE NAME OF THE DME PROVIDER, WHO COULD HELP WITH THE DISPOSAL PROCESS. THE HUSBAND THEN EXPRESSED CONCERNS ABOUT THE TIME INVOLVED IN DRIVING AROUND AND MENTIONED THAT HE MIGHT JUST DISPOSE OF THE DEVICE HIMSELF. NO ADDITIONAL DETAILS WERE PROVIDED, AND THERE ARE NO ALLEGATIONS SUGGESTING THAT THE DEVICE CONTRIBUTED TO OR CAUSED THE PATIENT'S DEATH. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
ANGIOVAC | K190594 | AN ANGIODYNAMICS TERRITORY MANAGER REPORTED A PATIENT DEATH THAT OCCURRED DURING USE OF AN ANGIOVAC C180 WITH OBTURATOR PG AND CIRCUIT WITH BUBBLE TRAP. THE FOLLOWING WAS REPORTED: "PATIENT HAD HISTORY OF CANCER 2018 (CERVICAL CANCER), ATRIAL FIBRILLATION, COPD, CTEPH, AND DVT. PATIENT HAD BLOOD IN URINE IN (B)(6) AND WAS TAKEN OFF XARELTO. INSTRUCTED TO BEGIN XARELTO AGAIN AFTER 5 DAYS. PATIENT FAILED TO BEGIN TAKING XARELTO AGAIN AS INSTRUCTED. YESTERDAY PATIENT PRESENTED AT ED AT (B)(6) WITH SUSPECTED PULMONARY EMBOLISM. PATIENT WAS TRANSFERRED TO DEACONESS GATEWAY FOR TREATMENT. CT SCAN SHOWED SMALL SADDLE EMBOLI AS WELL AS EMBOLI IN LEFT PULMONARY ARTERY THAT WASN'T DEEMED CLINICALLY SIGNIFICANT PER DR. (B)(6) . TEE SHOWED LARGE RA MASS (GREATER THAN 4CM). MASS APPEARED TO BE BROAD BASED, CONCENTRIC, AND LACKING JAGGED EDGES. SMALLER CLOT (2CM) APPEARED TO BE PRESENT AT THE IVC/RA JUNCTION. PATIENT HAD BEEN INTUBATED DURING THE TEE AND HAD REMAINED INTUBATED SINCE. DR. (B)(6), CT SURGEON, SUSPECTED CARDIAC TUMOR BASED ON CONVERSATIONS IN THE CONTROL ROOM AND HE MENTIONED THE RISK OF EMBOLIZATION AND MASSIVE PE TO DR. (B)(6). PRIOR TO CASE BEGINNING, DR.(B)(6) MET WITH ADAM DAVIS AND I TO DISCUSS SPECIFICS OF THE CASE. WE DISCUSSED RIJ CANNULA/LFV RETURN AND RFV CANNULA/LFV RETURN WITH DR. (B)(6); HE CHOSE RFV CANNULA/LFV RETURN APPROACH FOR THIS CASE. DEBULKING OF THE RIGHT ATRIAL CLOT/MASS WAS THE GOAL OF THE CASE. PATIENT WAS ALREADY INTUBATED AND RIGHT ATRIAL MASS WAS CONFIRMED TO BE PRESENT PRIOR TO CASE. WE GAINED ACCESS USING A 26F GORE DRY SEAL ON THE RIGHT AND A 17F MEDTRONIC ARTERIAL RETURN CATHETER ON THE LEFT. PREP OF THE ANGIOVAC DEVICE WAS WITHOUT EVENT AND PER ANGIODYNAMICS IFU. DR. (B)(6) GAINED ACCESS WITH AN AMPLATZ SUPER STIFF WIRE; WIRE WAS IN PLACE AND ANGIOVAC CANNULA AND OBTURATOR WERE LOADED ONTO THE WIRE AND PLACED IN THE RIGHT ATRIUM. ONCE DR. (B)(6) WAS HAPPY WITH THE POSITION OF THE ANGIOVAC CANNULA, THE WIRE AND OBTURATOR WERE REMOVED. FUNNEL WAS THEN EXPOSED. AT THIS POINT DR. (B)(6) ASKED THE PERFUSION TEAM TO COME ON SUCTION. THEY DID AND FLOW VARIED BETWEEN 2.5 TO 5L/MIN. ON SEVERAL OCCASIONS THE FUNNEL ENGAGED MATERIAL/ATRIAL WALL AND FLOW WAS COMPLETELY BLOCKED. THE CANNULA WOULD BE REPOSITIONED EACH TIME AND FLOW WAS RESTORED. THE RIGHT ATRIAL MASS WAS ENGAGED AND A LARGE AMOUNT OF MATERIAL WAS REMOVED (BRIGHT WHITE MATERIAL) AND LODGED IN THE BUBBLE TRAP. IT WAS NOTICED ON TEE THAT THE MAJORITY OF THE MATERIAL WAS NO LONGER IN THE RA, WITH ONLY A SINGLE PIECE OF MATERIAL REMAINING. THE REMAINING MATERIAL WAS ROUGHLY 3CM LONG AND EXTREMELY PEDUNCULATED AND STALK LIKE. AT THIS POINT A LONG PAUSE FOLLOWED IN CARDIAC FUNCTION, FOLLOWED BY CARDIAC ARREST. HE CALLED TO INITIATE CPR AND RESUSCITATION EFFORTS BEGAN IN THE FORM OF CHEST COMPRESSIONS . THE ANGIOVAC DEVICE AND MEDTRONIC ARTERIAL RETURN SHEATH WERE BOTH REMOVED. CPR WAS MAINTAINED FOR 14 MINUTES WITH BRIEF PAUSES TO CHECK FOR RETURN OF FUNCTION. DR. (B)(6) BELIEVED THAT MATERIAL HAD EMBOLIZED FROM THE RA TO THE RIGHT PULMONARY ARTERY. HE CALLED FOR THE INARI REPRESENTATIVE. A PIGTAIL CATHETER WAS PLACED IN THE MAIN PA AND A FILLING DEFECT WAS NOTICED IN THE LEFT PULMONARY ARTERIES. DR.(B)(6) PLACED A FLOWTRIEVER IN THE LEFT PULMONARY ARTERY AND MADE 7 TO 8 PULLS IN AN ATTEMPT TO TRY AND EVACUATE THE MATERIAL. HE DID NOT ASPIRATE ANY MATERIAL FROM THE PULMONARY ARTERIES AND THE DECISION WAS MADE AFTER 27 MINUTES TO GO SPEAK WITH THE FAMILY AND CEASE RESUSCITATION EFFORTS." DR. (B)(6) IS A VERY EXPERIENCED ANGIOVAC USER AND FELT THIS WAS THE BEST OPTION TO ATTEMPT TO TREAT A VERY, CHRONICALLY ILL PATIENT AND DOES NOT BELIEVE THERE WAS ANY SORT OF DEVICE MALFUNCTION THAT CONTRIBUTED TO THIS EVENT. THIS MEDICAL OPINION WAS DISCUSSED AND AGREED UPON WITH SEVERAL PHYSICIANS WITHIN THIS ACCOUNT. | THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT THROMBOSIS (PULMONARY EMBOLISM - PE) THAT RESULTED IN CARDIAC ARREST CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THE PATIENT SERIOUS ADVERSE EVENT (SAE). NO ANGIOVAC CANNULA/CIRCUIT WAS RETURNED FOR EVALUATION AS THERE WAS NO REPORTED DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING THE PROCEDURE. AS NOTED IN THE EVENT DESCRIPTION: "THE PHYSICIAN IS A VERY EXPERIENCED ANGIOVAC USER AND FELT THIS WAS THE BEST OPTION TO ATTEMPT TO TREAT A VERY, CHRONICALLY ILL PATIENT AND DOES NOT BELIEVE THERE WAS ANY SORT OF DEVICE MALFUNCTION THAT CONTRIBUTED TO THIS EVENT [PATIENT PE, CARDIAC ARREST AND EXPIRATION]. THIS MEDICAL OPINION WAS DISCUSSED AND AGREED UPON WITH SEVERAL PHYSICIANS WITHIN THIS ACCOUNT." A DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/ASSEMBLY LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. THERE HAVE BEEN NO OTHER REPORTED COMPLAINTS FOR THE INDICATED PACKAGING/ASSEMBLY LOTS FOR ANY DEVICE FAILURE MODES OR SIMILAR PATIENT THROMBOSIS/PE ISSUE. LABELING REVIEW: THE ANGIOVAC CANNULA SAMPLE WAS NOT RETURNED FOR EVALUATION SINCE THERE WAS NO REPORTED DEVICE MALFUNCTION DURING THE PROCEDURE. IT CANNOT BE DETERMINED IF DEVICE WAS USED IN ACCORDANCE WITH ITS LABELING. DIRECTIONS FOR USE (DFU) THAT IS PROVIDED WITH THIS DEVICE CONTAINS THE FOLLOWING STATEMENTS: WARNINGS: - SELECTION OF THE PATIENT AS A CANDIDATE FOR USE WITH THIS DEVICE AND FOR SUCH PROCEDURES AS IT IS INTENDED IS THE PHYSICIANS' SOLE RESPONSIBILITY. THE OUTCOME IS DEPENDENT ON MANY VARIABLES INCLUDING, PATIENT PATHOLOGY, SURGICAL PROCEDURE, AND PERFUSION PROCEDURE/TECHNIQUE. THE BENEFITS OF USE OF THIS DEVICE MUST BE WEIGHED AGAINST THE RISKS INCLUDING RISKS OF SYSTEMIC ANTICOAGULATION AND MUST BE ASSESSED BY THE PRESCRIBING PHYSICIAN. - AS WITH ALL MEDICAL DEVICES, THIS DEVICE AND ANCILLARY EQUIPMENT ARE TO BE USED BY TRAINED PHYSICIANS ONLY. SPECIFICALLY, THIS DEVICE IS TO BE USED ONLY BY MEDICAL PERSONNEL EXPERIENCED WITH INITIATING AND MONITORING EXTRACORPOREAL BYPASS AND BY PHYSICIANS TRAINED AND EXPERIENCED USING SURGICAL AND/OR PERCUTANEOUS (SELDINGER) VASCULAR ACCESS TECHNIQUES. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EXTRACORPOREAL BLOOD VESSEL DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE INSTRUCTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, EXTRACORPOREAL CIRCULATION AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: - DISTAL EMBOLIZATION OF THROMBUS. - PULMONARY EMBOLISM. - ARRHYTHMIAS. - DEATH. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT WILL CONTINUE TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
MX40 1.4 GHZ SMART HOPPING | K113125 | IT WAS REPORTED ALARMS WERE NOT GENERATED FOR AN INCIDENT WITH A TELEMETRY PATIENT WHO PASSED AWAY. THE TELEMETRY PATIENT WAS CHECKED BY THE INCOMING NURSE AS THE CENTRAL STATION WAS DISPLAYING A QUESTION MARK IN THE PATIENT SECTOR AND FOUND UNRESPONSIVE WITH NO PULSE. CPR WAS INITIATED. THE STAFF WAS NOT SURE WHAT HAPPENED BUT IT WAS DETERMINED THE PATIENT'S TELEMETRY DEVICE WAS NEVER ASSIGNED AT THE CENTRAL STATION; HOWEVER, THIS WAS NOT THE CASE, AND THE PATIENT WAS ADMITTED. THE CUSTOMER INDICATED THE PATIENT PASSED AWAY DUE TO UNKNOWN CAUSES A FEW DAYS LATER. THE AUDIT LOGS REVEALED ALARMS HAD BEEN GENERATED AND ACKNOWLEDGED. | THE EVENT TOOK PLACE ON FEBRUARY 16, 2025 BETWEEN 12:53 AND 18:45 IN BED (B)(4). DURING THE INVESTIGATION, A TECHNICAL AND CLINICAL ANALYSIS OF THE LOG DATA PROVIDED WAS PERFORMED. THE FINDINGS ARE BELOW. THE INVESTIGATION FOUND THAT ON (B)(6) 2025 AT 22:09:13, THE PATIENT WAS ADMITTED. FROM ADMISSION TO ON (B)(6) 2025 2:06:35 (3 HOURS 57 MIN 22 SECONDS), THERE WAS EITHER NO ECG ALARMING POSSIBLE OR NO PHYSIOLOGICAL MONITORING AT ALL POSSIBLE FOR A TOTAL OF 1 HOUR 34 MIN AND 18 SECONDS, DUE TO: ECG LEADS OFF, LEADSET UNPLUGGED, NO DATA TELE CONDITIONS. DURING THIS PERIOD OF TIME THERE WERE ALSO LA [LEFT ARM] LEAD OFF INOPS WHICH LIMITS ARRHYTHMIA ANALYSIS TO JUST ONE LEAD FOR A TOTAL OF 27 MINUTES AND 33 SECONDS. ANALYSIS AND ALARMS OCCURRED BETWEEN ON (B)(6) 2025 2:06:35 AND ON (B)(6) 2025 12:51:53 (10 HOURS 45 MINUTES 16 SECONDS). STARTING ON (B)(6) 2025 12:51:53 UNTIL ON (B)(6) 2025 18:45:54 (5 HOURS 54 MINUTES AND 1 SECOND), THERE WAS EITHER NO ECG ALARMING POSSIBLE OR NO PHYSIOLOGICAL MONITORING AT ALL POSSIBLE FOR A TOTAL OF 3 MINUTES AND 52 SECONDS, DUE TO: ECG LEADS OFF, LEADSET UNPLUGGED (3 MINUTES 49 SECOND EVENT ¿ THIS LEADS TO AN ARRYTHMIA RELEARN) CONDITIONS. FOLLOWING THESE EVENTS, THERE WERE LA LEAD OFF INOPS WHICH LIMITS ARRHYTHMIA ANALYSIS TO JUST ONE LEAD FOR A TOTAL OF 1 HOUR 51 MINUTES AND 48 SECONDS. THE LA LEAD OFF INOP WAS CORRECTED AT 15:57 AND MONITORING WITH ALL AVAILABLE ECG LEADS CONTINUED UNTIL THE V LEAD OFF ALARM GENERATED AT 18:45:46. STARTING ON (B)(6) 2025 18:51:08, THERE WAS A 2 HOUR 20 MINUTES AND 28 SECONDS ECG LEADS OFF CONDITION (NO ECG ALARMING POSSIBLE). THE ECG LEADS OFF ENDED DUE TO NO DATA TELE AT 21:11:15 AND THEN THE PATIENT WAS TRANSFERRED TO A DIFFERENT ROOM AT 21:18 ON (B)(6) 2025. THEY WOULD HAVE HAD A ¿?¿ IN THE SECTOR FOR HR WITH A FLAT WAVEFORM WHILE THE ECG LEADS OFF WAS PRESENT. ONCE THE DEVICE LOST CONNECTION TO THE PIC IX, THERE IS NO ¿?¿. THE SECTOR IS JUST BLANK. THE INVESTIGATION FOUND THAT THERE WAS EITHER NO ECG ALARMING POSSIBLE, NO PHYSIOLOGICAL MONITORING AT ALL POSSIBLE, OR LIMITED MONITORING DUE TO ONGOING ECG LEAD OFF ISSUES DURING THE REPORTED PERIOD OF TIME. FROM THE AUDIT LOG, ECG MONITORING WAS AVAILABLE FROM 15:57 WHEN THE LA LEAD OFF ENDED UNTIL!! ECG LEADS OFF GENERATED AT 18:50:58. THIS ALARM SOUNDED UNTIL IT WAS ACKNOWLEDGED AT 19:07:58 AND CONTINUED TO REMIND EVERY THREE MINUTES UNTIL 21:08:44. THIS ALARM WAS ACKNOWLEDGED AGAIN AT 19:42:01. THE REPORTED PROBLEM IS NOT CONFIRMED. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE ANY DEVICE MALFUNCTION; HOWEVER, A COMBINATION OF USE ERROR, ALARM MANAGEMENT, AND CLINICAL WORKFLOW DURING USE OF THE DEVICE MAY HAVE BEEN FACTORS IN THE PATIENT'S OUTCOME.
DREAMSTATION | K102465 | THE MANUFACTURER RECEIVED INFORMATION STATING A PATIENT SADLY PASSED AWAY. THE PATIENT WAS USING A BIPAP AVAPS30 DEVICE, ALTHOUGH NO INFORMATION WAS PROVIDED TO VERIFY IF THE DEVICE WAS IN USE WHEN THE PATIENT PASSED. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. ALSO, PER REPORT OF THE EVENT, "THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH." THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE AND DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS YET TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, WE ARE UNABLE TO CONFIRM IF ANY MALFUNCTION OCCURRED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON 06/MAR/2025 REPORTEDLY DUE TO PERITONITIS AND A CLOSTRIDIUM DIFFICILE (C-DIFF) INFECTION. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, HOSPITAL RECORDS, TREATMENT RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, AUTOPSY REPORT, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: IT IS UNKNOWN IF A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF PERITONITIS, A C-DIFF INFECTION, AND DEATH. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. ACCORDING TO THE INITIAL REPORTING, THE PATIENT¿S DEATH WAS ATTRIBUTED TO THE PERITONITIS AND C-DIFF INFECTIONS. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP INFORMATION PRECLUDED A MORE COMPREHENSIVE CLINICAL ASSESSMENT. THOSE INDIVIDUALS UNDERGOING PD FOR RRT (MANUAL OR CYCLER BASED) ARE AT HIGH RISK FOR INFECTIONS OF THE PERITONEUM. ADDITIONALLY, RISK FACTORS ASSOCIATED WITH THE DEVELOPMENT OF C. DIFFICILE INCLUDE LONGER DURATIONS OF ANTIBIOTIC ADMINISTRATION REQUIRED FOR ESRD PATIENTS TO CLEAR THE INFECTION. LASTLY, THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER AND LIBERTY CYCLER SET CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE LACK OF PATIENT DATA, TREATMENT DATA, CLINICAL FOLLOW-UP, DISCHARGE SUMMARY, DEATH CERTIFICATE, ESRD DEATH NOTIFICATION AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
LIBERTY CYCLER SET, SINGLE CONN./EXT. DL | K173718 | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 REPORTEDLY DUE TO PERITONITIS AND A CLOSTRIDIUM DIFFICILE (C-DIFF) INFECTION. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, HOSPITAL RECORDS, TREATMENT RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, AUTOPSY REPORT, DISCHARGE SUMMARY) WERE UNSUCCESSFUL. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: IT IS UNKNOWN IF A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF PERITONITIS, A C-DIFF INFECTION, AND DEATH. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. ACCORDING TO THE INITIAL REPORTING, THE PATIENT¿S DEATH WAS ATTRIBUTED TO THE PERITONITIS AND C-DIFF INFECTIONS. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP INFORMATION PRECLUDED A MORE COMPREHENSIVE CLINICAL ASSESSMENT. THOSE INDIVIDUALS UNDERGOING PD FOR RRT (MANUAL OR CYCLER BASED) ARE AT HIGH RISK FOR INFECTIONS OF THE PERITONEUM. ADDITIONALLY, RISK FACTORS ASSOCIATED WITH THE DEVELOPMENT OF C. DIFFICILE INCLUDE LONGER DURATIONS OF ANTIBIOTIC ADMINISTRATION REQUIRED FOR ESRD PATIENTS TO CLEAR THE INFECTION. LASTLY, THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER AND LIBERTY CYCLER SET CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CAUSAL OR CONTRIBUTORY ROLE IN THE SERIOUS ADVERSE EVENTS. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION CAUSED AND/OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. HOWEVER, GIVEN THE LACK OF PATIENT DATA, TREATMENT DATA, CLINICAL FOLLOW-UP, DISCHARGE SUMMARY, DEATH CERTIFICATE, ESRD DEATH NOTIFICATION AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
ION | K182188 | IT WAS REPORTED THAT A PATIENT UNDERWENT AN UNEVENTFUL ION ENDOBRONCHIAL LUNG BIOPSY WITH MEDIASTINAL STAGING PROCEDURE. THE PHYSICIAN REPORTED THAT THERE WERE NO INTRA-PROCEDURAL COMPLICATIONS OR MALFUNCTIONS WITH THE ION SYSTEM. APPROXIMATELY 30 MINUTES POST-PROCEDURE, THE PATIENT SHOWED ACUTE STROKE SYMPTOMS OF LEFT-SIDED WEAKNESS AND GAZE DEVIATION TREATED WITH TISSUE PLASMINOGEN ACTIVATOR (TPA) THAT RESULTED IN EPISODES OF HEMOPTYSIS. THE PHYSICIAN STATED THAT THE PATIENT SUFFERED AN ISCHEMIC STROKE; UNSPECIFIED HEAD IMAGES DID NOT SHOW ANY ACUTE CHANGES. THE PATIENT EXPIRED SHORTLY THEREAFTER. THE CAUSE OF THE STROKE WAS REPORTED TO BE RELATED TO THE PATIENT¿S MEDICAL HISTORY, INCLUDING SUBSTANTIAL CARDIOVASCULAR DISEASE WITH MULTIPLE STENTS AND MULTIPLE STROKES TREATED WITH ASPIRIN AND PLAVIX THERAPY. THE PLAVIX HAD BEEN HELD FOR 5 DAYS PRIOR TO THE PROCEDURE, BUT THE ASPIRIN THERAPY WAS CONTINUED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT AN ELDERLY PATIENT WITH MULTIPLE MEDICAL CO-MORBIDITIES INCLUDING MULTIPLE PRIOR STROKES, CARDIOVASCULAR DISEASE REQUIRING MULTIPLE VASCULAR STENTS UNDERWENT AN ION LUNG BIOPSY. PLAVIX WAS HELD FOR 5 DAYS IN PREPARATION FOR THE BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS TRANSPORTED TO THE PACU. ABOUT 30 MINUTES LATER THE PATIENT DEVELOPED A LEFT SIDED WEAKNESS AND GAZE DEVIATION PROMPTING TPA TO TREAT AN ACUTE STROKE. WORKUP WAS CONSISTENT WITH AN ISCHEMIC STROKE AND THE PATIENT DIED LATER THE SAME DAY. BASED ON THE AVAILABLE DATA THE EVENT WAS PROCEDURE RELATED IN A PATIENT WITH MULTIPLE PRIOR STROKES WHOSE PLAVIX HAD BEEN STOPPED FOR THE BIOPSY AND WAS NOT DEVICE RELATED. BRONCHOSCOPY AND BIOPSY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW COMPLICATION RATE AND RARELY ASSOCIATED WITH FATAL COMPLICATIONS. IN A MULTICENTER PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES THE TOTAL NUMBER OF ANY COMPLICATIONS WAS REPORTED TO BE 227 (1.08%) WITH 5 (0.02%) TRANSIENT ISCHEMIC ATTACKS AND 4 (0.02%) TOTAL DEATHS BUT NO STROKES. ANOTHER MULTICENTER STUDY REPORTED 13 (2.2%) COMPLICATIONS WITH NO STROKES AND NO DEATHS ASSOCIATED WITH 581 CASES. A RECENT PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 BRONCHOSCOPIES REPORTED 1 ASSOCIATED DEATH (0.08%) AND NO STROKES. ANOTHER PROSPECTIVE SERIES OF 415 ION PROCEDURES REPORTED 1 CVA (0.2%). THERE WAS NO ALLEGATION OF A MALFUNCTION OF THE ION SYSTEM, INSTRUMENT OR ACCESSORIES ASSOCIATED WITH THE REPORTED COMPLICATION. THE RISK OF STROKE ASSOCIATED WITH GENERAL ANESTHESIA HAS BEEN REPORTED TO VARY FROM 0.1-1.9% IN NON-CARDIAC, NON-NEUROLOGIC, AND NON-MAJOR SURGERY AND AS HIGH AS 10% IN THE SETTING OF BRAIN OR HIGH-RISK CARDIOVASCULAR SURGERY. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009;71(1). OST DE, ERNST A, LEI X, ET AL. DIAGNOSTIC YIELD AND COMPLICATIONS OF BRONCHOSCOPY FOR PERIPHERAL LUNG LESIONS. RESULTS OF THE AQUIRE REGISTRY. AM J RESPIR CRIT CARE MED. 2016;193(1):68-77. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BATEMAN BT, SCHUMACHER HC, WANG S, SHAEFI S, BERMAN MF. PERIOPERATIVE ACUTE ISCHEMIC STROKE IN NONCARDIAC AND NONVASCULAR SURGERY: INCIDENCE, RISK FACTORS, AND OUTCOMES. ANESTHESIOLOGY. 2009. SELIM M. PERIOPERATIVE STROKE. NEW ENGLAND JOURNAL OF MEDICINE. 2007. SECTION A2: AGE REPORTED AS "IN THEIR 80'S."
PHILIPS RESPIRONICS E30 WITH HUMIDIFIER | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A "PATIENT DEATH WAS REPORTED SEVERAL DAYS AFTER USING THE E30 VENTILATOR. THE HOSPITAL STATES THAT PHILIPS DID NOT PERFORM SERVICE TO THE SYSTEM NOR PROCEED WITH THE SCHEDULED RECALL AND COLLECTION, AS PER THE OFFICIAL DOCUMENTATION. THE FAMILY HAS SUBMITTED THE CORRESPONDING LEGAL DEATH CERTIFICATE (ACTA DE DEFUNCIÓN / DEATH CERTIFICATE) THAT MIGHT DIRECTLY LINK THE INCIDENT TO THE USE OF THE DEVICE. THE FAMILY IS NOW REPORTING THE LACK OF FOLLOW UP ON THE RECALL, THE SERVICE OR EXCHANGE NOT BEING PERFORMED, AND OF COURSE THE CASE OF THEIR DECEASED PATIENT PLEASE BE ADVISED THAT PHILIPS MEXICO, UNDER THE SIGNATURE OF (B)(6), ISSUED AN OFFICIAL COMMUNICATION INDICATING THAT ALL E30 SYSTEMS WERE TO BE RECOLLECTED BEFORE (B)(6) 2024. AS OF TODAY, THIS RECOLLECTION HAS NOT BEEN EXECUTED, AND ADDITIONAL PATIENT RISK REPORTS CONTINUE TO BE FILED ON DIFFERENT E-30 SYSTEMS". THE COMPLAINT HAS BEEN REVIEWED BY A CLINICAL EXPERT, AND IT HAS BEEN DETERMINED THAT THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. NO SERIAL NUMBER HAS BEEN PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
EVIS EXERA DUODENOVIDEOSCOPE | K954451 | ADDITIONAL INFORMATION WAS RECEIVED FROM THE AUTHOR THAT THERE WAS NO MALFUNCTION OF THE OLYMPUS DEVICES. IT WAS REPORTED THERE WAS NO LINK BETWEEN THE PATIENT'S DEATH AND THE DUODENOSCOPE. THE PATIENT'S DEATH WAS RELATED TO A COMPLICATION OF THE ANESTHESIA. | THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO PROVIDE ADDITIONAL INFORMATION RECEIVED FROM THE AUTHOR. UPDATED FIELDS: B5.
CONFIDA BRECKER CURVE GUIDEWIRE | K132623 | ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS IN EXTREMELY FRAGILE CONDITION WITH STAGE IV PULMONARY ARTERY DISEASE PRIOR TO THE VALVE IMPLANT PROCEDURE. THEY WERE EVALUATED FOR A VALVE DUE TO SEVERE AORTIC STENOSIS. IT WAS NOT POSSIBLE TO USE VIABLE FEMORAL ACCESS, THEREFORE RIGHT CAROTID ACCESS WAS USED AS AN ALTERNATIVE STRATEGY. PRIOR TO THE 23 MILLIMETER (MM) VALVE BEING INSERTED INTO THE PATIENT, AS THE GUIDEWIRE CROSSED INTO THE LEFT VENTRICLE, THE PATIENT SUFFERED CARDIAC ARREST. ADVANCED CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED IMMEDIATELY, WHICH LASTED MORE THAN 20 MINUTES WITHOUT RESTORING SPONTANEOUS CIRCULATION. UP TO THIS POINT, THE VALVE HAD NOT BEEN INSERTED, IMPLANTED OR INTERACTED WITH THE PATIENT IN ANY WAY. IN AN ATTEMPT TO REVERSE THE PATIENT'S CRITICAL CONDITION, THE MEDICAL TEAM DECIDED TO PROCEED WITH THE IMPLANTATION OF THE VALVE AS A LAST RESORT HOPING TO RESTORE CARDIAC FUNCTION. HOWEVER, THE PATIENT NEVER REGAINED EFFECTIVE CIRCULATION AND THEY DIED DURING THE PROCEDURE. THE FATAL OUTCOME WAS NOTED TO BE DUE TO THE PATIENT'S UNDERLYING CRITICAL ILLNESS AND EXTREME INSTABILITY. THEREFORE, THE DEVICE DID NOT CONTRIBUTE TO OR WORSEN THE REPORTED EVENT. DURING THE PROCEDURE AND AFTERWARDS, THE LEAD IMPLANT PHYSICIAN STATED " THE VALVE DID NOT CAUSE ANY HARM TO THE PATIENT. THE DEVICE WAS NOT INVOLVED IN THE CARDIAC ARREST OR THE PATIENT'S HEMODYNAMIC DETERIORATION." | UPDATED DATA: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
NIPRO BLOOD TUBING SET W/TP | K072024 | FEMALE PATIENT WITH CVC LINE AND FOR SOME UNKNOWN REASONS THERE WAS AN ACCIDENTAL DISCONNECTION FROM THE VENOUS BLOODLINE CONNECTED TO THE CVC LINE AND THEREFORE A MASSIVE BLEEDING, WITH THE DEATH OF THE PATIENT, CONFIRMED THE (B)(6) BY THE CUSTOMER TOWARDS THE SALES REP, WHO VISITED THE CENTER ON THE (B)(6). THE COSTUMER DOES NOT BLAME THE MACHINE, OR THE LINES, AS THEY THREW AWAY EVERYTHING AND USED THE MACHINE FOR 4 MORE TREATMENTS AFTERWARDS. THIS IS ALSO THE REASON THE COSTUMER DID NOT REPORT THIS ISSUE TO NIPRO. HOWEVER, BECAUSE THE MACHINE WAS USED AFTERWARDS FOR FOUR MORE TREATMENTS, WE COULD NOT SAVE ANY DATA, LOGS OR PICTURES.
INJECTOMAT AGILIA BR | K121613 | THE FOLLOWING HAS BEEN REPORTED BY CUSTOMER: THE CUSTOMER REPORTS THAT DURING THE ADMINISTRATION OF MEDICATION, A LARGER AMOUNT THAN PROGRAMMED WAS INFUSED, RESULTING IN THE DEATH OF A NEWBORN. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION.
ILET BIONIC PANCREAS | K231485 | ON 2/25/25 A BETA BIONICS CLINICAL DIABETES SPECIALIST (CDS) REPORTED THAT AN ILET USER EXPERIENCED A LOW BLOOD GLUCOSE (BG) EVENT AND WAS HOSPITALIZED. IT WAS LATER NOTED THAT THE USER IS NOW DECEASED. HOWEVER, THE PHYSICIAN IS UNCERTAIN WHETHER THE CAUSE OF DEATH IS RELATED TO DIABETES. AT THE TIME OF THIS REPORT, MULTIPLE ATTEMPTS BY BETA BIONICS TO DETERMINE IF THE USER WAS ON THE ILET AT THE TIME OF THE EVENT AS WELL AS CAUSE OF DEATH HAVE BEEN UNSUCCESSFUL. | NO PRODUCT WAS RETURNED FOR EVALUATION. ATTEMPTS TO DOWNLOAD THE DEVICE ENGINEERING LOGS WERE UNSUCCESSFUL, INDICATING THAT THE LOGS HAVE NOT BEEN SYNCED SINCE DEVICE WAS SHIPPED FROM MANUFACTURING. AS SUCH, A LOG REVIEW WAS UNABLE TO BE CONDUCTED NO PRODUCT PERFORMANCE ISSUES CAN BE IDENTIFIED. IF THE PRODUCT IS RECEIVED AT A LATER DATE, OR THE ILET DATA IS UPLOADED, THE COMPLAINT WILL BE REOPENED AND INVESTIGATED ACCORDINGLY. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. WITHOUT DEVICE DATA OR A DATE OF DEATH, IT CANNOT BE DETERMINED IF THE USER WAS WEARING THE ILET AT THEIR TIME OF DEATH OR IF THE ILET WAS RELATED TO THEIR PASSING. CHART NOTES ON FILE INDICATE A HISTORY OF UNCONTROLLED DIABETES WITH RECURRENT HYPOGLYCEMIA AND AN A1C OF 11.4% PRIOR TO USE OF THE ILET. IT IS POSSIBLE THAT THERE WERE COMORBIDITIES THAT MAY HAVE CONTRIBUTED TO THE DEMISE AND WITHOUT FURTHER INFORMATION, THIS CANNOT BE RULED OUT. THIS CASE WILL BE REVISITED AND INVESTIGATED FURTHER IF ANY NEW INFORMATION BECOMES AVAILABLE.
FILMARRAY® PNEUMONIA PANEL PLUS | K222601 | SUMMARY: (B)(6) FRANCE, REPORTED A POTENTIAL FALSE NEGATIVE KLEBSIELLA OXYTOCA, KLEBSIELLA PNEUMONIAE GROUP, AND CTX-M RESULT ON THE FILMARRAY PNEUMONIA PANEL PLUS AFTER TESTING A PATIENT'S BRONCHOALVEOLAR LAVAGE (BAL) SAMPLE. IT IS UNKNOWN IF THERE WAS ANY IMPACT TO THE PATIENT DUE TO THE FILMARRAY PNEUMONIA PANEL PLUS RESULT. BIOFIRE HAS REQUESTED FURTHER INFORMATION FROM THE CUSTOMER AND IS CURRENTLY INVESTIGATING THIS EVENT. | INVESTIGATION: THE PATIENT WAS A 65-YEAR-OLD MALE. THE CUSTOMER REPORTED THAT THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025, WITH DYSPNEA AND HEMOPTYSIS. AT THE TIME OF TESTING, THE PATIENT WAS INTUBATED AND VENTILATED IN THE ICU AND ON TAZOBACTAM. ON (B)(6) 2025, A PATIENT'S BAL SAMPLE WAS TESTED ON FILMARRAY PNEUMONIA PANEL PLUS. THE FILMARRAY PNEUMONIA PANEL PLUS REPORTED ALL ANALYTES AS NOT DETECTED. ON (B)(6) 2025, THE PATIENT'S SAMPLE WAS RETESTED ON THE FILMARRAY PNEUMONIA PANEL PLUS. THE FILMARRAY PNEUMONIA PANEL PLUS REPORTED K. OXYTOCA, K. PNEUMONIAE GROUP, AND CTX-M AS DETECTED. THE CUSTOMER REPORTED CHEST COMPUTED TOMOGRAPHY (CT) SCAN WAS SUGGESTIVE OF CYTOMEGALOVIRUS INFECTION, CLINICAL WORSENING AND WORSENING OF LESIONS. ON AN UNKNOWN DATE, MEROPENEM FOR K.PNEUMONIAE ESBL WAS INTRODUCED. THE CUSTOMER STATED THAT THE PATIENT RAPIDLY IMPROVED ON (B)(6) 2025, WITH EXTUBATION. THE PATIENT DIED ON (B)(6) 2025. BIOFIRE'S INVESTIGATION INTO THIS EVENT IS ONGOING AND FURTHER INFORMATION HAS BEEN REQUESTED FROM THE CUSTOMER. THE FULL INVESTIGATION AND ASSOCIATED CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. CONCLUSION: INVESTIGATION ONGOING.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A REMSTAR AUTO A-FLEX DEVICE. THE PATIENT HAS PASSED AWAY. THERE IS NO CUSTOMER INFORMATION HENCE WE CANNOT REACH OUT TO THE CUSTOMER AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. BOX H UPDATED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP DEVICE. THE PATIENT'S WIFE HAS ALLEGED HER HUSBAND PASSED AWAY. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN UNABLE TO CONFIRM FOAM PARTICLES IN THE AIR PATH DURING THE DEVICE EVALUATION. THE DEVICE WAS SCRAPPED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND NO VISIBLE FOAM PARTICLES WERE OBSERVED. ADDITIONALLY, NO FAULTS WERE FOUND AND ALL TESTS WERE PASSED. | THE BECOME AWARENESS DATE, COMPONENT CODE GRID AND CONCLUSION CODE GRID HAS BEEN UPDATED/CORRECTED IN THIS REPORT. THE DEVICE INFORMATION HAS ALSO BEEN UPDATED/CORRECTED. IN THIS REPORT, BOX D, G, H HAS BEEN UPDATED/CORRECTED.
DAVINCI X | K171294 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
CARDIOHELP | K133598 | THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE. COMPLAINT# (B)(4). | MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-03-31: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2024-12-02 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" AND THE CONSEQUENCE FOR THE PATIENT WAS "HEMORRHAGE AFTER REMOVING THE DEVICE AT THE INSERTION SITE. THIS RESULTED IN THE PATIENT'S DEATH. RENACER REGISTRY PATIENT (B)(6).¿ NEW INFORMATION WAS RECEIVED ON 2025-04-02 THAT NO MALFUNCTION ON THE CARDIOHELP DEVICE OCCURRED. THE EVENT WAS NOT RELATED TO THE CARDIOHELP DEVICE. THE TREATMENT WITH THE CARDIOHELP MACHINE WAS FINALIZED ON 2024-11-18. ON (B)(6) 2024, AND THE PATIENT SUFFERED A TEAR OF THE SUPERFICIAL FEMORAL ARTERY WITH MASSIVE HEMORRHAGE AND WAS REVIEWED BY THE SURGICAL TEAM. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" COULD NOT BE LINKED TO A DEVICE DEFECT/FAILURE. THEREFORE THE COMPLAINT COULD NOT BE CONFIRMED. THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
NIO ADULT | K142086 | ADDITIONAL INFORMATION FROM CUSTOMER RECEIVED ON MARCH 30, 2025: "PATIENT WAS IN CARDIAC ARREST WITH CPR BEING PERFORMED. CREW ARRIVED, TOOK OVER CPR. IV WAS ATTEMPTED AND FAILED. NIO WAS BROUGHT OUT, SITE LOCATED AND WHEN WE SPUN THE TOP ON THE NIO TO UNLOCK IT, IT WAS TIGHT TO TURN. PRESSURE PLACED WITH DOMINANT HAND, OTHER HAND STABILIZING NIO. NIO DID NOT DEPLOY, SECOND CREWMEMBER ATTEMPTED TO DEPLOY IT AND IT DID NOT WORK. SECOND NIO WAS BROUGHT OUT AND THE SAME THING HAPPENED. TOP WAS TIGHT TO SPIN AND IT DID NOT DEPLOY FOR EITHER THE RN OR AEMT. MANUAL IO WAS PLACED IN THIS PATIENT FIRST ATTEMPT. CPR CONTINUED. PATIENT DIED. ONE OF THE NIO THAT DID NOT DEPLOY WAS BROUGHT BACK TO THE STATION AND IS CURRENTLY BEING SENT BACK WITH THE PROVIDED SHIPPING LABEL." | PLEASE NOTE THAT IN THE INITIAL REPORT, THE BATCH NUMBER PROVIDED BY THE CUSTOMER WAS INCORRECT (2440043). DURING THE INVESTIGATION OF THE DEVICE, IT WAS CLARIFIED THAT THE CORRECT BATCH NUMBER IS: 2440040. AS PART OF THE INVESTIGATION PROCESS, THE FOLLOWING EXAMINATIONS WERE PERFORMED: VISUAL INSPECTION: THE DEVICE WAS VISUALLY EXAMINED, AND IT APPEARED THAT THE TRIGGER WAS UNCLIPPED. X-RAY EXAMINATION ((B)(6) 2025): THE DEVICE WAS X-RAYED, AND ALL METAL COMPONENTS WERE FOUND TO BE IN THEIR EXPECTED POSITIONS. THERE WAS NOTHING IN THE X-RAY THAT WOULD CLEARLY INDICATE A CAUSE FOR THE DEVICE'S FAILURE TO ACTIVATE. POST-X-RAY FUNCTIONAL TEST (04/11/2025): THE DEVICE WAS TESTED TO DETERMINE IF IT COULD BE ACTIVATED. IT WAS NOTED THAT THE SAFETY CAP WAS SIGNIFICANTLY MORE DIFFICULT TO ROTATE THAN EXPECTED. THE DEVICE WAS PLACED ON A TORQUE TABLE TO MEASURE THE FORCE REQUIRED TO ROTATE THE SAFETY CAP FROM THE "LOCKED" TO "UNLOCKED" POSITION. THE TORQUE MEASUREMENTS RANGED FROM 500 N·MM TO OVER 900 N·MM, MORE THAN TWICE THE VALUES MEASURED ON A DEVICE WHERE THE SAFETY CAP ROTATED AS INTENDED. HOWEVER, THE INCREASED TORQUE REQUIRED TO TURN THE SAFETY CAP WOULD NOT IMPACT THE USER'S ABILITY TO FIRE THE DEVICE PROVIDED THEY COULD MOVE IT FROM THE LOCKED TO UNLOCKED POSITION. POST-ACTIVATION INSPECTION: AFTER ACTIVATING THE DEVICE, THE SAFETY CAP WAS REMOVED AND BOTH THE CAP AND BODY WERE EXAMINED. ADDITIONAL SCRATCHES WERE OBSERVED ON THE BODY OF THE DEVICE. THE MOST LIKELY CAUSE WAS DETERMINED TO BE EXCESS FLASH (MATERIAL LEFT ON THE COMPONENT AFTER MANUFACTURING) NEAR THE PRONG AREA OF THE SAFETY CAP. ROOT CAUSE ANALYSIS: A FISHBONE DIAGRAM WAS COMPLETED TO ANALYZE THE FAILURE MODE. THE MOST LIKELY ROOT CAUSE OF THE INCREASED TORQUE REQUIRED TO TURN THE SAFETY CAP IS FLASH FROM THE MANUFACTURING PROCESS. HOWEVER, THIS IS NOT LIKELY TO HAVE LED TO THE DEVICE NOT FIRING SINCE THE SAFETY CAP COULD BE ROTATED TO THE UNLOCKED POSITION. THE MOST LIKELY ROOT CAUSE FOR THE DEVICE NOT FIRING IS EITHER DUE TO THE USER NOT FULLY PRESSING DOWN ON THE DEVICE OR NOT PRESSING STRAIGHT DOWN ON THE DEVICE. IF THE USER PRESSED DOWN AT AN ANGLE AND THEN PULLED THE TRIGGER BACK WITH EXCESSIVE FORCE WITHOUT FIRST UNLOCKING THE TRIGGER, IT MAY HAVE CAUSED DEFORMATION. CONCLUSIONS: THE CUSTOMER COMPLAINT REGARDING THE DIFFICULTY IN ROTATING THE TOP OF THE DEVICE WAS REPLICATED. THE ISSUE WAS MOST LIKELY CAUSED BY FLASH (MATERIAL LEFT ON THE COMPONENT AFTER MANUFACTURING) IN THE SAFETY CAP, OR DEBRIS (BROKEN FLASH) LOCATED BETWEEN THE SAFETY CAP AND THE DEVICE BODY. THE COMPLAINT REGARDING DIFFICULTY IN FIRING THE DEVICE WAS ALSO REPLICATED. THIS WAS MOST LIKELY CAUSED BY NOT FOLLOWING THE INSTRUCTIONS AS OUTLINED IN THE IFU, AND BY SQUEEZING THE TRIGGER WITH EXCESSIVE FORCE EITHER PRIOR TO FULLY UNLOCKING THE TRIGGER, OR BY ONLY PARTIALLY UNLOCKING THE TRIGGER BY PRESSING DOWN AT AN ANGLE. *NOTE- UPDATED B2 TO INCLUDE "OTHER SERIOUS OR IMPORTANT MEDICAL EVENTS". UPDATED B7 TO INCLUDE RELEVANT INFORMATION.
TUBING SET | K112360 | COMPLAINT ID: (B)(4). | THE EVENT OCCURRED IN FRANCE DURING TREATMENT. IT WAS REPORTED THAT THE CIRCUIT WAS PRIMED WITH WATER IN ACCORDANCE WITH STANDARD PROCEDURES TO PREPARE FOR ECMO SUPPORT. THE FLOW RATE WAS OBTAINED WITHOUT DIFFICULTY. A FEW MINUTES LATER THERE WAS A LOSS OF FLOW, A MACHINE ALARM AND VISIBLE BUBBLES WITHIN THE OXYGENATOR. THE SYSTEM COULD NOT BE DE-AIRED. THE CIRCUIT WAS CLAMPED AND THE CIRCUIT REPLACED. THE PATIENT WAS ADMITTED FOR EMERGENCY CARDIAC ARREST. THE CARE WAS DELAYED FOR THIS PATIENT. THE ECMO WAS INITIATED AT 8 PM ON (B)(6) 2025 AND THE PATIENT PASSED AWAY AT 3 AM ON (B)(6) 2025. NEW INFORMATION WAS RECEIVED ON 2025-04-03 THAT THE PATIENT WAS CONNECTED TO THE FIRST HLS SET WHICH CONTAINED AIR IN THE SYSTEM. LATER, IT WAS CONFIRMED THAT THE REPLACEMENT SET HAD NO FAULT AND WAS DISCARDED BY THE CUSTOMER. THE FIRST SET WAS AVAILABLE FOR INVESTIGATION, WHICH ACCORDING TO THE CUSTOMER HAD NO VISIBLE DEFECT NOR ANY LEAKS DURING/AFTER PRIMING. FURTHER INFORMATION WAS RECEIVED THAT THERE WAS NO AIR ENTRY AFTER THE REPLACEMENT OF THE AFFECTED SET. THE USED CANNULAS WERE NOT REPLACED AND WERE THEREFORE WORKING AS EXPECTED. NO INFORMATION WAS AVAILABLE IF THE USED CANNULAS WERE GETINGE CANNULAS. THE CANNULAS WERE DISCARDED BY THE CUSTOMER. AS PATIENT DEATH WAS REPORTED, A REPORT IS REQUIRED. THE AFFECTED PRODUCT WAS INVESTIGATED BY THE GETINGE LABORATORY ON 2025-06-27 WITH FOLLOWING CONCLUSION: THE FAILURE COULD NOT BE CONFIRMED. THERE WERE NO VISIBLE DAMAGES ON THE PRODUCT. THE PRODUCT COULD BE PRIMED WITHOUT ANY PROBLEMS AND THERE WAS NO AIR ENTRY. THE PRODUCT FUNCTIONALITY WAS AS EXPECTED. A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-07-04 WITH FOLLOWING CONCLUSION: "ON (B)(6) 2025, AN INCIDENT OCCURRED INVOLVING AN HLS SET ADVANCED 7.0 USED WITH THE CARDIOHELP-I SYSTEM DURING VA-ECLS. THE PATIENT, A 45-YEAR-OLD FEMALE ADMITTED AFTER CARDIAC ARREST, RECEIVED ECLS SUPPORT BEGINNING AT APPROXIMATELY 20:15. WHILE INITIAL STARTUP WAS UNEVENTFUL AND TARGET FLOW WAS ACHIEVED, A SUDDEN LOSS OF FLOW OCCURRED SHORTLY THEREAFTER, ACCOMPANIED BY BUBBLE DETECTION ALARMS. AIR BUBBLES WERE OBSERVED IN THE OXYGENATOR AND COULD NOT BE REMOVED. NO AIR ENTRY SOURCE WAS IDENTIFIED AT THE TIME. THE CLINICAL TEAM REPLACED THE CIRCUIT, WHICH FUNCTIONED NORMALLY. THE PATIENT EXPIRED AT AROUND 03:00; NO DIRECT CAUSAL LINK TO THE INCIDENT WAS ESTABLISHED, THOUGH TREATMENT WAS DELAYED. THE SUBSEQUENT INVESTIGATION FOUND NO PRODUCT-RELATED DEFECTS. A MINOR LEAK AT THE VENOUS SAMPLING PORT WAS ATTRIBUTED TO DRIED BLOOD RESIDUE AND WAS NO LONGER PRESENT AFTER CLEANING. THE TUBING, INCLUDING THE CRITICAL VENOUS LINE, SHOWED NO DAMAGE OR KINKS. FUNCTIONAL TESTING, INCLUDING A 60-MINUTE LEAK TEST, CONFIRMED SYSTEM INTEGRITY. A PRESSURE SENSOR MALFUNCTION WAS NOTED BUT DETERMINED TO BE RELATED TO TRANSPORT AND NOT CONNECTED TO THE COMPLAINT. IMPORTANTLY, THE VENOUS CANNULA USED DURING ECMO INITIATION - A CRITICAL COMPONENT IN ASSESSING THE ROOT CAUSE OF AIR INGRESS - WAS NOT RETURNED FOR FURTHER EVALUATION. THIS OMISSION LIMITS THE ABILITY TO FULLY RULE OUT A CONNECTION-RELATED ISSUE OR A DEFECT AT THE INSERTION SITE. LITERATURE SUPPORTS THE RISK OF AIR INGRESS VIA THE VENOUS LINE OR CANNULA DURING ECMO, PARTICULARLY IN THE PRESENCE OF LOOSE CONNECTIONS, NEGATIVE PRESSURE, OR INSUFFICIENT PRIMING. IN ADDITION, AIR INGRESS THROUGH A CENTRAL VENOUS CATHETER, IF PRESENT, COULD BE A CONTRIBUTING FACTOR, PARTICULARLY IF CONNECTIONS WERE NOT FULLY SECURED. THE ASSUMPTION OF VENOUS AIR ENTRAINMENT APPEARS TO CONFORM TO THE EVENT PATTERN AS DESCRIBED: LOSS OF FLOW (DUE TO AIR INVASION OF THE CENTRIFUGAL PUMP), AIR IN THE OXYGENATOR, THEN BUBBLE ALARM. IT IS PRESUMED THAT THE REPORTED BUBBLE ALARM WAS ARTERIAL. NO MENTION OF VENOUS BUBBLE SENSOR DEPLOYMENT, HOWEVER, WAS CITED IN THE COMPLAINT NARRATIVE. IN CONCLUSION AND GIVEN THE ABSENCE OF THE VENOUS CANNULA FOR INVESTIGATION AND THE LIMITED PROCEDURAL DETAILS AVAILABLE, NO DEFINITIVE ROOT CAUSE OF THE AIR INGRESS CAN BE CONFIRMED. HOWEVER, POSSIBLE CAUSES INCLUDE A LEAK AT THE VENOUS CANNULA SITE, AIR ENTRY VIA A CENTRAL VENOUS CATHETER, OR A TRANSIENT LEAK AT THE VENOUS SAMPLING PORT, OBSERVED DURING INITIAL FLUSHING AND ATTRIBUTED TO DRIED BLOOD RESIDUES, THOUGH THIS WAS NOT REPRODUCIBLE UNDER CONTROLLED TESTING CONDITIONS." ACCORDING TO THE INSTRUCTION OF USE OF THE HLS SET IT IS MENTIONED THAT INCORRECT POSITIONING OF THE CARDIOHELP CAN CAUSE AIR PERMEATION ON THE BLOOD SIDE OF THE HLS MODULE ADVANCED. THE CARDIOHELP MUST BE POSITIONED AT A LEVER LOWER THAT THE PATIENT AND SECURED CLOSE TO THE PATIENT. THE FLOW/BUBBLE SENSOR MUST ALWAYS BE AFFIXED ON THE ARTERIAL SIDE OF THE SET IF YOU ARE USING THE CARDIOHELP-I. WITH AN ACTIVATED INTERVENTION, THE DETECTION OF BUBBLES BY THE FLOW/BUBBLE SENSOR TRIGGERS A PUMP STOP. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-07-03. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARDS TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARDS TO THE REPORTED FAILURE. BASED ON THE RESULTS THE REPORTED FAILURE "AIR IN SYSTEM" COULD NOT BE CONFIRMED. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. ACCORDING TO THE RISK REVIEW THE REPORTED FAILURE IS BELOW THE ACCEPTANCE THRESHOLD ACCORDING TO THE RISK MANAGEMENT PLAN OF THE HLS SET. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THEIR LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). IT WAS CONFIRMED THE PATIENT WAS IN ACTIVE TREATMENT UTILIZING THE LIBERTY SELECT CYCLER AT THE TIME OF THE EVENT. IT IS UNKNOWN WHAT PHASE AND CYCLE THE PATIENT WAS IN. THE PATIENT CALLED THEIR SPOUSE INTO THE BEDROOM AND STATED THEY WERE NOT FEELING WELL. THE PATIENT WAS SITTING ON THE EDGE OF THE BED AND BECAME QUIET. THE PATIENT THEN BECAME UNRESPONSIVE. 911 WAS CALLED. THE 911 OPERATOR INSTRUCTED THE SPOUSE TO PLACE THE PATIENT ON THE FLOOR. DESPITE THE PATIENT BEING A DO NOT RESUSCITATE (DNR) ORDER, THE SPOUSE BEGAN CARDIOPULMONARY RESUSCITATION (CPR). EMERGENCY MEDICAL SERVICES (EMS) ARRIVED AND CONTINUED TO PERFORM RESUSCITATIVE EFFORTS. PER THE PATIENT¿S SPOUSE, THEY ADMINISTERED CPR AND MEDICATIONS (UNSPECIFIED). RESUSCITATION EFFORTS WERE UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED AT HOME. NO AUTOPSY WAS COMPLETED. THE PATIENT WAS TRANSPORTED TO THE FUNERAL HOME. THE PATIENT¿S CAUSE OF DEATH IS REPORTEDLY CARDIAC ARRHYTHMIA. THE PATIENT DID HAVE A HISTORY OF UNSPECIFIED CARDIAC ARRHYTHMIA. THERE WAS NO REPORTED ISSUE WITH THE PATIENT¿S TREATMENT, OR THE LIBERTY SELECT CYCLER IN RELATION TO THE EVENT. THE LAST REPORTED ISSUES WERE DRAIN COMPLICATIONS WHICH WERE SUBSEQUENTLY FOUND TO BE CAUSED BY SEVERE CONSTIPATION AS A RESULT OF THE PATIENT BEING BEDRIDDEN. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE REPORTED CAUSE OF DEATH WAS CARDIAC ARRHYTHMIA. CARDIAC DISEASE IS THE MAJOR CAUSE OF DEATH, ACCOUNTING FOR APPROXIMATELY 37 PERCENT OF ALL-CAUSE MORTALITY IN PATIENTS RECEIVING EITHER HEMODIALYSIS OR PERITONEAL DIALYSIS WITH THE LARGEST SPECIFIC CAUSE ATTRIBUTED TO ARRHYTHMIC MECHANISMS OR SUDDEN CARDIAC ARREST (SCA). THE PATIENT REPORTEDLY HAD A HISTORY OF UNSPECIFIED CARDIAC ARRHYTHMIA. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
TUBING SET | K112360 | THE EVENT OCCURRED IN FRANCE. IT WAS REPORTED THAT THE CIRCUIT WAS PRIMED WITH WATER IN ACCORDANCE WITH STANDARD PROCEDURES TO PREPARE FOR ECMO SUPPORT. THE FLOW RATE WAS OBTAINED WITHOUT DIFFICULTY. A FEW MINUTES LATER THERE WAS A LOSS OF FLOW, A MACHINE ALARM AND VISIBLE BUBBLES WITHIN THE OXYGENATOR. THE SYSTEM COULD NOT BE DE-AIRED. THE CIRCUIT WAS CLAMPED AND THE CIRCUIT REPLACED. THE PATIENT WAS ADMITTED FOR EMERGENCY CARDIAC ARREST. THE CARE WAS DELAYED FOR THIS PATIENT. THE ECMO WAS INITIATED AT 8 PM AND THE PATIENT PASSED AWAY AT 3 AM. COMPLAINT ID# (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE FRANCE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
CARDIOHELP | K133598 | THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE. COMPLAINT# (B)(4). | MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-03-31: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2024-12-02 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" AND THE CONSEQUENCE FOR THE PATIENT WAS "HEMORRHAGE AFTER REMOVING THE DEVICE AT THE INSERTION SITE. THIS RESULTED IN THE PATIENT'S DEATH. RENACER REGISTRY PATIENT (B)(6).¿ NEW INFORMATION WAS RECEIVED ON 2025-04-02 THAT NO MALFUNCTION ON THE CARDIOHELP DEVICE OCCURRED. THE EVENT WAS NOT RELATED TO THE CARDIOHELP DEVICE. THE TREATMENT WITH THE CARDIOHELP MACHINE WAS FINALIZED ON 2024-11-18. ON (B)(6) 2024, AND THE PATIENT SUFFERED A TEAR OF THE SUPERFICIAL FEMORAL ARTERY WITH MASSIVE HEMORRHAGE AND WAS REVIEWED BY THE SURGICAL TEAM. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "MAJOR BLEEDING IN THE FEMORAL ARTERY" COULD NOT BE LINKED TO A DEVICE DEFECT/FAILURE. THEREFORE, THE COMPLAINT COULD NOT BE CONFIRMED. THE CONFIRMATION WAS RECEIVED THAT THERE WAS NO PRODUCT MALFUNCTION ON A GETINGE PRODUCT, THEREFORE THIS COMPLAINT IS NOT REPORTABLE THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING.
INJECTOR OPVSU NPS | K063503 | THIS INCIDENT WAS REPORTED BY A FACILITY IN (B)(6) ON (B)(6) 2025. "(B)(6) 2025, GUERBET WAS CONTACTED (B)(6) ASKING FOR DOCUMENTATION OF TRAINING FOR THE STAFF ON (B)(6) 2021. SHE REPORTED A FATAL AIR BOLUS USING THE OPTIVANTAGE, SHE DOES NOT BLAME THE INJECTOR BUT CONSIDERS THIS TECHNOLOGY ERROR WHO HAS BEEN USING THE INJECTOR SINCE INSTALL IN 2021. THEY USE GUERBET SELF FILL SYRINGES.". | OVERALL INVESTIGATION SUMMARY: GUERBET WAS CONTACTED BY A CUSTOMER REPORTING A FATAL AIR BOLUS OCCURRED USING OPTIVANTAGE INJECTOR 849002 SERIAL NUMBER (B)(6). THE CUSTOMER DID NOT BLAME THE INJECTOR AND CONSIDERS THIS A TECHNOLOGIST ERROR. ADDITIONAL DETAILS WERE REQUESTED FROM THE CUSTOMER BUT HAVE NOT BEEN RECEIVED AT THE TIME OF COMPLAINT INVESTIGATION CLOSURE. THIS COMPLAINT WILL BE REOPENED WHEN THE ADDITIONAL DETAILS ARE RECEIVED. THERE WAS NO REPORT OF MALFUNCTION OR REQUEST FOR EVALUATION OF THE INJECTOR; THEREFORE, THE INVESTIGATION REVIEWED HISTORICAL RECORDS AND DESIGN RECORDS. OPTIVANTAGE WAS NOT DESIGNED AND IS NOT CAPABLE OF DETECTING OR PREVENTING AIR INJECTIONS; HOWEVER, AS LISTED BELOW, THE IFU INCLUDES INSTRUCTIONS FOR THE PROPER PURGING OF AIR FROM THE SYSTEM PRIOR TO INJECTION, AS WELL AS PRECAUTIONS AND STEPS TO PREVENT AIR INJECTIONS. HISTORICAL COMPLAINTS FOR OPTIVANTAGE AIR INJECTIONS WERE REVIEWED AND ALL INSTANCES WERE DETERMINED TO BE USER-RELATED, INDEPENDENT OF THE INJECTOR. THE RISK ASSESSMENT AND CONTROL DOCUMENTATION FOR THE OPTIVANTAGE WAS REVIEWED AND CONFIRMED THAT THE HAZARDOUS SITUATION OF AIR INJECTIONS LEADING TO AIR EMBOLISMS HAVE BEEN MITIGATED WITHIN THE MANUFACTURER'S CAPACITY. THE MITIGATIONS IN PLACE INCLUDE, BUT ARE NOT LIMITED TO WARNING STATEMENTS IN THE INSTRUCTIONS FOR USE (IFU), SYRINGE ADAPTERS PROVIDING A CLEAR VIEW OF THE SYRINGE, PROMPTS APPEARING ON SCREEN TO ENSURE THE USER FOLLOWS A PROPER FILL/PURGE SEQUENCE ON THE SYRINGES, A "CHECK FOR AIR..." MESSAGE ON THE POWERHEAD AND CONSOLE DISPLAYS, AND THE IFU CONTAINING INSTRUCTIONS ON HOW TO CLEAR AIR BUBBLES FROM THE CONSUMABLES. A REVIEW OF GUERBET'S COMPLAINT TRACKING SYSTEM RELATED TO THIS UNIT WAS CONDUCTED BUT FOUND NO RELATED INCIDENTS. GUERBET TECHNICAL SERVICE TEAM HAS NEVER SERVICED THIS INJECTOR OUTSIDE OF THE INITIAL INSTALLATION AND APPLICATIONS TRAINING. THE HEALTH HAZARD EVALUATION (HHE) (B4) WAS OPENED IN RELATION TO THIS COMPLAINT. THE HHE CONCLUSION INDICATES NO ACTIONS ARE WARRANTED AT THIS TIME. IMPACT ASSESSMENT SUMMARY: THE AIR BOLUS WAS FATAL AND THE PATIENT PASSED. IMDRF CODES = B13, B11; C19, C23; D11. ROOT/PROBABLE CAUSE CODE: PERSONNEL - PERFORMANCE - FAILED TO FOLLOW PROCEDURE. ROOT/ PROBABLE CAUSE SUMMARY: AIR INJECTIONS ARE GENERALLY USER-RELATED AS THE OPTIVANTAGE INJECTOR REQUIRES THE USERS TO MANUALLY PURGE THE AIR FROM THE SYSTEM AND CONFIRM THAT THEY HAVE CHECKED FOR AIR. BASED ON THIS AND THE INVESTIGATION, THERE IS NO NEED FOR A CAPA AT THIS TIME. COMPLAINTS ARE TRENDED AND REPORTED DURING MANAGEMENT AND METRIC REVIEWS TO CONSIDER CORRECTIVE ACTION. QUALITY WILL CONTINUE TO MONITOR AND TREND FOR SIMILAR ISSUES. DISPOSITION SUMMARY: THE UNIT REMAINED IN USE.
CLEARVIEW ARTERIOTOMY SHUNT | K120612 | A LITERATURE REVIEW WAS CONDUCTED ON A PROSPECTIVE STUDY FROM JANUARY 2015 TO DECEMBER 2023, WHICH COLLECTED DATA ON ALL PATIENTS WHO UNDERWENT MINIMALLY INVASIVE TOTAL ARTERIAL CORONARY ARTERY BYPASS GRAFTING, TO EVALUATE THE PROGRESS AND DEVELOPMENT OF THIS TECHNIQUE. A TOTAL OF 279 CONSECUTIVE PATIENTS WERE EVALUATED. THE MEAN AGE WAS 66 +/- 7 YEARS, WITH NEARLY TWO-THIRDS BEING MEN AND ONE-THIRD HAVING DIABETES. A SIMILAR PROPORTION OF PATIENTS HAD TRIPLE VESSEL DISEASE AND LEFT MAIN DISEASE. THE FOLLOWING MEDTRONIC DEVICES WERE USED FOR ALL PROCEDURES: OCTOPUS STABILIZER, ARTERIOTOMY SHUNT, CLEARVIEW BLOWER AND THE QUICKFLOW DISTAL PERFUSION SET (LOT NUMBERS NOT PROVIDED). AMONG ALL PATIENTS, 1 IN HOSPITAL MORTALITY OCCURRED. BASED ON THE AVAILABLE INFORMATION, THE DEATH MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: LOW CARDIAC OUTPUT, EXTRA CORPOREAL LIFE-SUPPORT INTERVENTION, SECONDARY CONVERSION TO STERNOTOMY, PERIOPERATIVE MYOCARDIAL INJURY, UNPLANNED POSTOPERATIVE PERCUTANEOUS CORONARY INTERVENTION, RETHORACOTOMY FOR BLEEDING, NEED FOR POSTOPERATIVE DIALYSIS, WOUND INFECTIONS, HIGH POSTOPERATIVE CREATINE KINASE MB, INCOMPLETE REVASCULARIZATION AND TRACHEOTOMY. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. | B3: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT G2: CITATION: AUTHORS: VEREVKIN, ALEXANDER ET AL. TITLE: MINIMALLY INVASIVE CORONARY ARTERY BYPASS GRAFTING VIA LEFT ANTERIOR MINITH ORACOTOMY: SETUP, RESULTS, AND EVOLUTION OF A NEW SURGICAL PROCEDURE. JTCVS TECHNIQUES, VOLUME 29, 28 ¿ 39. 10.1016/J.XJTC.2024.10.022 NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CLEARVIEW BLOWER/ MISTER | K973485 | A LITERATURE REVIEW WAS CONDUCTED ON A PROSPECTIVE STUDY FROM JANUARY 2015 TO DECEMBER 2023, WHICH COLLECTED DATA ON ALL PATIENTS WHO UNDERWENT MINIMALLY INVASIVE TOTAL ARTERIAL CORONARY ARTERY BYPASS GRAFTING, TO EVALUATE THE PROGRESS AND DEVELOPMENT OF THIS TECHNIQUE. A TOTAL OF 279 CONSECUTIVE PATIENTS WERE EVALUATED. THE MEAN AGE WAS 66 +/- 7 YEARS, WITH NEARLY TWO-THIRDS BEING MEN AND ONE-THIRD HAVING DIABETES. A SIMILAR PROPORTION OF PATIENTS HAD TRIPLE VESSEL DISEASE AND LEFT MAIN DISEASE. THE FOLLOWING MEDTRONIC DEVICES WERE USED FOR ALL PROCEDURES: OCTOPUS STABILIZER, ARTERIOTOMY SHUNT, CLEARVIEW BLOWER AND THE QUICKFLOW DISTAL PERFUSION SET (LOT NUMBERS NOT PROVIDED). AMONG ALL PATIENTS, 1 IN HOSPITAL MORTALITY OCCURRED. BASED ON THE AVAILABLE INFORMATION, THE DEATH MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: LOW CARDIAC OUTPUT, EXTRA CORPOREAL LIFE-SUPPORT INTERVENTION, SECONDARY CONVERSION TO STERNOTOMY, PERIOPERATIVE MYOCARDIAL INJURY, UNPLANNED POSTOPERATIVE PERCUTANEOUS CORONARY INTERVENTION, RETHORACOTOMY FOR BLEEDING, NEED FOR POSTOPERATIVE DIALYSIS, WOUND INFECTIONS, HIGH POSTOPERATIVE CREATINE KINASE MB, INCOMPLETE REVASCULARIZATION AND TRACHEOTOMY. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. | B3: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT G2: CITATION: AUTHORS: VEREVKIN, ALEXANDER ET AL. TITLE: MINIMALLY INVASIVE CORONARY ARTERY BYPASS GRAFTING VIA LEFT ANTERIOR MINITH ORACOTOMY: SETUP, RESULTS, AND EVOLUTION OF A NEW SURGICAL PROCEDURE. JTCVS TECHNIQUES, VOLUME 29, 28 ¿ 39. 10.1016/J.XJTC.2024.10.022 NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DISTAL PERFUSION KIT | K994171 | A LITERATURE REVIEW WAS CONDUCTED ON A PROSPECTIVE STUDY FROM JANUARY 2015 TO DECEMBER 2023, WHICH COLLECTED DATA ON ALL PATIENTS WHO UNDERWENT MINIMALLY INVASIVE TOTAL ARTERIAL CORONARY ARTERY BYPASS GRAFTING, TO EVALUATE THE PROGRESS AND DEVELOPMENT OF THIS TECHNIQUE. A TOTAL OF 279 CONSECUTIVE PATIENTS WERE EVALUATED. THE MEAN AGE WAS 66 +/- 7 YEARS, WITH NEARLY TWO-THIRDS BEING MEN AND ONE-THIRD HAVING DIABETES. A SIMILAR PROPORTION OF PATIENTS HAD TRIPLE VESSEL DISEASE AND LEFT MAIN DISEASE. THE FOLLOWING MEDTRONIC DEVICES WERE USED FOR ALL PROCEDURES: OCTOPUS STABILIZER, ARTERIOTOMY SHUNT, CLEARVIEW BLOWER AND THE QUICKFLOW DISTAL PERFUSION SET (LOT NUMBERS NOT PROVIDED). AMONG ALL PATIENTS, 1 IN HOSPITAL MORTALITY OCCURRED. BASED ON THE AVAILABLE INFORMATION, THE DEATH MAY HAVE BEEN ATTRIBUTED TO MEDTRONIC PRODUCT. AMONG ALL PATIENTS, ADVERSE EVENTS INCLUDED: LOW CARDIAC OUTPUT, EXTRA CORPOREAL LIFE-SUPPORT INTERVENTION, SECONDARY CONVERSION TO STERNOTOMY, PERIOPERATIVE MYOCARDIAL INJURY, UNPLANNED POSTOPERATIVE PERCUTANEOUS CORONARY INTERVENTION, RETHORACOTOMY FOR BLEEDING, NEED FOR POSTOPERATIVE DIALYSIS, WOUND INFECTIONS, HIGH POSTOPERATIVE CREATINE KINASE MB, INCOMPLETE REVASCULARIZATION AND TRACHEOTOMY. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. | B3: EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT G2: CITATION: AUTHORS: VEREVKIN, ALEXANDER ET AL. TITLE: MINIMALLY INVASIVE CORONARY ARTERY BYPASS GRAFTING VIA LEFT ANTERIOR MINITH ORACOTOMY: SETUP, RESULTS, AND EVOLUTION OF A NEW SURGICAL PROCEDURE. JTCVS TECHNIQUES, VOLUME 29, 28 ¿ 39. 10.1016/J.XJTC.2024.10.022 NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
BIPAP SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER BECAME AWARE OF AN ALLEGATION DUE TO A USER OF BIPAP AUTO BIFLEX. DAUGHTER CALLED IN STATING PATIENT PASSED AWAY AND DOES NOT NEED THE REPLACEMENT DEVICE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE DEVICE HAS BEEN RETURNED FROM THE THIRD-PARTY SERVICE CENTER. EVALUATION SHOWS THE UNIT IS FUNCTIONAL. THERE ARE NO VISIBLE FOAM PARTICLES AND HAS CONTAMINATION FINDING WHICH IS DUST. THE DEVICE WAS ROUTED TO SCRAP. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THIS COMPLAINT IS INCLUDED IN THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES.
CRAGG-MCNAMARA | K940634 | TORRES RUIZ, I., OOI, X. Y., HARRY, L., KOKSOY, C., PALLISTER, Z. S., GILANI, R., MILLS, J. L., BAILEY, C. J., <(>&<)> CHUNG, J. (2024). MULTILEVEL THROMBOTIC OR EMBOLIC BURDEN AND ITS ROLE IN SEX-RELATED OUTCOMES IN ACUTE LIMB ISCHEMIA. JOURNAL OF VASCULAR SURGERY, 80(6), 1796¿1803. HTTPS://DOI.ORG/10.1016/J.JVS.2024.06.007 MEDTRONIC REVIEW OF THE LITERATURE ARTICLE FOUND A TWO-CENTER RETROSPECTIVE REVIEW OVER 9 YEARS OF 170 PATIENTS TREATED FOR ACUTE LIMB ISCHEMIA (ALI). THE PURPOSE OF THE STUDY WAS TO QUANTIFY THE EFFECT OF SEX UPON AMPUTATION-FREE SURVIVAL (AFS) AFTER A PERCUTANEOUS-FIRST APPROACH FOR ALI. IT WAS NOTED THAT IN SOME CDT CASES, CRAGG-MCNAMARA CATHETERS WERE USED FOR THROMBOLYTIC INFUSION. 92 PATIENTS RECEIVED CATHETER-DIRECTED THROMBOLYSIS (CDT) ALONE, AND 35 PATIENTS WERE TREATED WITH A COMBINATION OF CDT AND ASPIRATION THROMBECTOMY. FOR PATIENTS WHO EXPERIENCED COMPLICATIONS, IT WAS NOT SPECIFIED WHICH TYPE OF TREATMENT(S) THE PATIENT RECEIVED. THE ARTICLE REPORTS SOME COMPLICATIONS RELATED TO THE PROCEDURES BUT DOES NOT SPECIFICALLY MENTION DEVICE MALFUNCTIONS OR CATHETER-RELATED COMPLICATIONS. THE COMPLICATIONS OBSERVED INCLUDED: 13 BLEEDING EPISODES (8%) 4 CASES OF ATRIAL FIBRILLATION (2%) 3 RE-THROMBOSIS/CLOT EXTENSION EVENTS (1.7%) THE ARTICLE SPECIFIES THAT 11 PATIENTS REQUIRED A MAJOR AMPUTATION FOLLOWING TREATMENT. OF THESE, 2 WERE ABOVE-THE-KNEE AMPUTATIONS (AKAS), AND 9 WERE BELOW-THE-KNEE AMPUTATIONS (BKAS). ADDITIONALLY, THE STUDY REPORTED A 30-DAY MORTALITY RATE OF 8%, WHICH EQUATES TO 13 PATIENTS OUT OF THE 170 WHO PARTICIPATED IN THE STUDY. CAUSE OF DEATH WAS NOT SPECIFIED IN THE ARTICLE. THE ARTICLE NOTED THAT FEMALES WERE MORE LIKELY THAN MALES TO HAVE POOR OUTCOME OR OUTCOME OF MORTALITY. HOWEVER, THE FEMALES IN THE STUDY GROUP ALSO HAD A HIGHER AVERAGE AGE, HAD MORE LIKELY PRESENTED WITH ATRIAL FIBRILLATION, AND MORE OFTEN HAD MULTI-LEVEL OCCLUSIONS PRIOR TO TREATMENT. | A2. THE REPORTED PATIENT AGE IS THE MEDIAN AGE FROM ALL PATIENTS INCLUDED IN THE STUDY. A3A. THOUGH THE MAJORITY (51%) OF PATIENTS IN THE STUDY WERE MALE, IT WAS NOTED THAT FEMALES WERE MORE LIKELY THAN MALES TO HAVE POOR OUTCOME OR OUTCOME OF MORTALITY. THEREFORE, THE REPORTED SEX IS FEMALE. B4. REPORTED EVENT DATE IS DATE THE ARTICLE WAS ACCEPTED FOR PUBLICATION. E1. THE REPORTED FACILITY IS THE FACILITY ASSOCIATED WITH THE COMMUNICATING AUTHOR. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
SUREFORM | K173721 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED SLEEVE GASTRECTOMY SURGICAL PROCEDURE, THE PATIENT WAS RE-HOSPITALIZED ON POST-OPERATIVE DAY (POD) 11 AND ULTIMATELY EXPIRED ON POD 18. THE PATIENT WAS READMITTED ON POD 11 WITH HEMATEMESIS AND SEPSIS WITHOUT SHOCK. THE PATIENT HAD A BMI OF 82 AND WAS INTUBATED FOR AIRWAY PROTECTION. A CT SCAN SHOWED A TINY AMOUNT OF CONTRAST LEAKAGE, BUT IT COULD NOT BE DETERMINED IF THE STAPLE LINE WAS OPEN OR NOT. SOMETIME AFTER THE CT SCAN, THE PATIENT CODED FROM RESPIRATORY FAILURE AND WAS RESUSCITATED. EXPLORATORY SURGERY FOUND A 7CM OPENING AT THE STAPLE LINES. THE SURGEON THOUGHT THAT THE STAPLE LINE WAS NOT OPEN BEFORE THE PATIENT CODED; IT WAS THOUGHT TO BE CAUSED BY, OR BECAME WIDER, DURING THE RESUSCITATIVE EFFORTS. THE SURGEON STATED THAT IF THE PATIENT HAD NOT CODED DUE TO RESPIRATORY FAILURE, THE FINDINGS WOULD HAVE BEEN TREATABLE. THE PATIENT SUFFERED RESPIRATORY FAILURE, CARDIOGENIC SHOCK, RENAL FAILURE, AND AN ANOXIC BRAIN INJURY. THE ULTIMATE CAUSE OF DEATH WAS REPORTED TO BE THE ANOXIC BRAIN INJURY THAT WAS IRREVERSIBLE. DURING THE PROCEDURE, A SUREFORM 60 STAPLER FIRED FOUR BLACK RELOADS WITH BUTTRESS MATERIAL WITHOUT INCIDENT. DURING THE FIFTH BLACK RELOAD FIRE A ¿TISSUE TOO THICK TO CONTINUE¿ MESSAGE APPEARED. THE SURGEON DID NOT THINK THE TISSUE WAS TOO THICK AS THE OTHER FIRINGS WERE UNEVENTFUL. A SIXTH SUREFORM 60 BLACK RELOAD WITHOUT BUTTRESS MATERIAL FIRED SUCCESSFULLY TO FINISH THE STAPLE LINE. AT THE END OF THE PROCEDURE, THE SURGEON INSPECTED THE STOMACH STAPLE LINES VIA ENDOSCOPY AND NO ISSUES WERE FOUND. THE SURGEON STATED THAT THE LEAK WAS LIKELY FROM THE STAPLE LINE THAT HAD THE MISFIRE, BUT THE STAPLE LINES HAD APPEARED IN GOOD CONDITION WHEN INSPECTED DURING THE PROCEDURE; THE LEAK AND THE DEATH ¿ARE TWO SEPARATE ISSUES¿. THE SURGEON ALSO STATED THAT ¿THE STAPLER MISFIRE AND LEAK WERE THE ONLY OBVIOUS ABNORMALITIES AND IT WAS LIKELY RELATED SOMEHOW TO THE PATIENT OUTCOME.¿ | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE 30 DEGREE PLUS, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE ADVANCED STAPLER LOGS SHOWED THAT THE SUREFORM 60 STAPLER WAS USED TO FIRE SIX BLACK RELOADS WITH NO INCOMPLETE CLAMPS OR UNCLAMPS. THE FIFTH BLACK RELOAD FIRE STOPPED AT ABOUT 82%. THE SAME STAPLER INSTRUMENT WAS INSTALLED TO FIRE A SIXTH BLACK RELOAD SUCCESSFULLY. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BARIATRIC SLEEVE GASTRECTOMY PROCEDURE FOR EXTREME MORBID OBESITY AND BMI OF 82. THE INTRAOPERATIVE FINDINGS AT THE TIME OF THE INITIAL PROCEDURE SHOWED NO EVIDENCE OF LEAK. THE PATIENT RETURNED 11 DAYS AFTER SURGERY WITH A SMALL LEAK NOTED ON CT. THE CAUSE AND EXACT LOCATION OF THIS LEAK WAS NOT PROVIDED. THE PATIENT BEGAN HAVING ADDITIONAL COMPLICATIONS INCLUDING RESPIRATORY FAILURE, RENAL FAILURE AND CARDIOGENIC SHOCK. THE PATIENT EVENTUALLY CODED. ON RE-EXPLORATION, THE STAPLE LINE HAD NOW DEHISCED WHICH WAS PRESENT ON THE CONTRAST CT SCAN PREVIOUSLY DONE. THE CAUSE OF THIS DISRUPTION IS UNKNOWN HOWEVER THE SURGEON SPECULATED THIS DISRUPTION MAY HAVE OCCURRED AS A RESULT OF COMPRESSIONS WHEN THE PATIENT CODED. THE PATIENT EVENTUALLY DEVELOPED AN ANOXIC BRAIN INJURY AND DIED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION A4 PATIENT WEIGHT: THE PATIENT BMI OF 82 KG/M2 WAS PROVIDED.
CARDIOSAVE HYBRID W/ E/F PLUG | K112372 | ON (B)(6) 2025, IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) DID NOT SWITCH TO MAINS OPERATION ONCE THE POWER CONNECTION WAS PROPERLY ESTABLISHED. THE DEVICE REMAINED ON BATTERY POWER AND THE CONTROL BLOCK WAS CORRECTLY SECURED IN THE TROLLEY. ON (B)(6) 02025 ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT LATER DIED DUE TO POOR GENERAL CONDITION, NOT BECAUSE THE THERAPY WAS NOT STARTED. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
TRILOGY 100 | K083526 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT WHO HAD BEEN USING A TRILOGY 100 VENTILATOR PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION AGAINST THE DEVICE. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE EVALUATION CONFIRMED THAT THE DEVICE OPERATED AND ALARMED AS DESIGNED. THE DEVICE WAS CLEANED AND PASSED ALL FINAL TESTING.
ENDO ANCHOR SYSTEM - HELI-FX AAA | K140036 | IT WAS REPORTED THAT DURING AN INDEX PROCEDURE INVOLVING THE HELI-FX ENDOSYSTEM AND A N ENDURANT IIS, A RUPTURE WAS SUSPECTED WHEN THE NON-MEDTRONIC WIRE APPEARED TO DEFLECT OUTSIDE THE AORTIC LUMEN UPON REMOVAL OF THE HELI-FX GUIDE. PRE-OPERATIVELY THE DIAMETER OF THE AORTA AT THE RENAL ARTERY MEASURED 26MM WHILE THE LENGTH OF THE PROXIMAL AORTIC NECK WAS 9MM. MODERATE VESSEL TORTUOSITY AND SEVERE VESSEL CALCIFICATION WERE REPORTED. THE PROCEDURE INITIALLY INVOLVED THE SUCCESSFUL PLACEMENT OF A 32MM BIFURCATED DEVICE AND 16X16MM DIAMETER LIMBS, ALONG WITH SIX ENDOANCHORS DUE TO THE PATIENT'S SHORT AND CONICAL NECK ANATOMY. A SELECTIVE ANGIOGRAM WAS PERFORMED, AND NO RUPTURE LEAKS WERE IDENTIFIED AT THAT TIME. HOWEVER, APPROXIMATELY ONE HOUR LATER, THE PATIENT WAS REPORTED TO BE UNSTABLE, PROMPTING ANOTHER ANGIOGRAM WHICH REVEALED A TEAR ABOVE THE LEFT RENAL ARTERY. THE LEFT RENAL ARTERY WAS EMBOLIZED VIA BRACHIAL APPROACH AND COVERED WITH A 32X32X49MM AORTIC CUFF, WHILE THE RIGHT RENAL ARTERY WAS STENTED AND "SNORKELED" BEHIND A 36X36X49MM AORTIC CUFF, MAINTAINING BLOOD FLOW THROUGH THE STENT. THE PATIENT RECEIVED THREE UNITS OF BLOOD AND WAS TRANSFERRED TO CRITICAL CARE. MULTIPLE STROKES WERE CONFIRMED AND THE PATIENT WAS NOT MOVING THE RIGHT SIDE OF THEIR BODY FOR SEVERAL DAYS WHILE INTUBATED. THE PHYSICIAN CONFIRMED THAT THE RUPTURE AND STROKE WERE NOT ATTRIBUTED TO THE HELI-FX GUIDE. IT IS BELIEVED THAT IT WAS THE NON-MEDTRONIC WIRE THAT CAUSED THE RUPTURE AND IT WAS STATED THAT ¿AS THE WIRE WAS IN FRONT OF THE HELI-FX GUIDE AND THE GUIDE WOULD ONLY FOLLOW THE WIRE". THE STROKES WERE BELIEVED TO BE CAUSED FROM THE BRACHIAL ACCESS AND MULTIPLE EXCHANGES OVER THE ARCH DURING THE SECONDARY PROCEDURE TO FIX THE RUPTURE. SEVERAL EMBOLI WERE REPORTED TO AFFECT THE BRAIN STEM. THE PATIENT¿S FAMILY DECIDED TO COMPASSIONATELY REMOVE INTUBATION AND THE PATIENT PASSED AWAY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
XENMATRIX AB | K133223 | PER MEDWATCH REPORT (MW5167617), "ALLEGED EVENT: LIMITED INFORMATION WAS REPORTED THROUGH A LEGAL EVENT THAT A PATIENT WAS IMPLANTED WITH 1ST STRATTICE ON (B)(6) 2016 AND 2ND STRATTICE ON (B)(6) 2017. THE FORM ALSO INDICATES THAT ON AN UNSPECIFIED DATE, THE STRATTICE IMPLANT HAD BEEN REMOVED OR REVISED. NO INFORMATION WAS PROVIDED ON THE INDICATION FOR THE INITIAL SURGERY OR THE REASON FOR THE REMOVAL OR REVISION. PATIENT HAS DIED ON (B)(6) 2022 AND THERE IS NO FURTHER DETAILS ON DEATH CAUSALITY. THIS RECORD CAPTURED 1ST STRATTICE. ON (B)(6) 2025, (B)(6) RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT WAS BORN WITH AN ABDOMINAL WALL DEFECT (GASTROSCHISIS), WHO UNDERWENT A VENTRAL HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT SP100381-101 ON (B)(6) 2016. THE PATIENT UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGICAL MESH ON (B)(6) 2017. THE PATIENT WAS IMPLANTED WITH A SECOND STRATTICE DEVICE LOT SP100537-078 ON THE SAME DATE. THE PATIENT THEN UNDERWENT AN INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WITH A THIRD STRATTICE DEVICE LOT SP100595." THE PATIENT THEN UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH ON (B)(6) 2018 AND IMPLANTED WITH A NONABBVIE DEVICE, XENMATRIX LOT HUCR0748. THE MESH WAS REVISED DUE TO RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WAS A SECOND NON-ABBVIE DEVICE, BARD MESH LOT #HUCQ0919. THE MESH WAS REVISED ON (B)(6) 2022 DUE TO OPEN REPAIR OF RECURRENT VENTRAL INCISIONAL HERNIA WITH BIOLOGICAL MESH, ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION AND IMPLANTED WITH A FOURTH STRATTICE DEVICE LOT SP200312-045." THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC STRATTICE MESH ON OR ABOUT (B)(6) 2016. THE PATIENT ALSO UNDERWENT A OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) ON OR ABOUT(B)(6) 2017. THE PATIENT WAS TREATED FOR INCISIONAL/ VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) BETWEEN (B)(6) 2018. THE PATIENT UNDERWENT OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH (RECURRENT); AT BEDSIDE SEROUS SANGUINOUS OOZING AROUND LEFT JP DRAIN BETWEEN(B)(6) 2018 AND (B)(6) 2018. THE PATIENT WAS TREATED FOR RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON OR ABOUT (B)(6) 2018. THE PATIENT WAS SEEN FOR ABDOMINAL PAIN BETWEEN 2018 AND (B)(6) 2022. THE PATIENT WAS TREATED FOR BOWEL OBSTRUCTION BETWEEN APPROXIMATELY 2017-2018. THE PATIENT RECEIVED TREATMENT FOR ABDOMINAL PAIN/ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR BOWEL OBSTRUCTION, ABDOMINAL PAIN, ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR HERNIA SYMPTOMS FROM APPROXIMATELY 2018 TO (B)(6) 2022. THE PATIENT WAS TREATED OPEN REPAIR OF RECURRENT INCISIONAL HERNIA WITH BIOLOGIC MESH (STRATTICE); ABDOMINAL RE-ENTRY FOR CORRECTION OF REHERNIATION BETWEEN (B)(6) 2022 AND (B)(6) 2022. THE PATIENT RECEIVED POST-OP REHAB BETWEEN 2018-2019. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)". REFERENCE REPORT: MW5167618. LOT #: HUCR0748". | NO CONCLUSIONS CAN BE MADE. AS REPORTED, PATIENT HAD A COMPLEX MEDICAL HISTORY WITH MULTIPLE IMPLANTS AND SURGICAL PROCEDURES. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. REVIEW OF MANUFACTURING RECORDS FOR THE XENMATRIX AB IMPLANTS USED TO TREAT THE PATIENT CONFIRMED THE PRODUCTS WERE MANUFACTURED TO SPECIFICATION. TO DATE, THIS IS THE ONLY REPORTED COMPLAINT FOR THIS MANUFACTURING LOT OF 38 UNITS. NOTE, THE ACTUAL DATE OF DEATH AND THE DATE OF EVENT WERE NOT REPORTED IN THE REPORT SENT TO BD. WE HAVE DOCUMENTED THESE REQUIRED DATE FIELDS WITH A BEST ESTIMATE. THIS MDR REPRESENTS THE XENMATRIX AB (DEVICE #2). AN ADDITIONAL MDR WAS SUBMITTED TO REPRESENT THE XENMATRIX AB (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
XENMATRIX AB | K133223 | PER MEDWATCH REPORT (MW5167617), "ALLEGED EVENT: LIMITED INFORMATION WAS REPORTED THROUGH A LEGAL EVENT THAT A PATIENT WAS IMPLANTED WITH 1ST STRATTICE ON (B)(6) 2016 AND 2ND STRATTICE ON (B)(6) 2017. THE FORM ALSO INDICATES THAT ON AN UNSPECIFIED DATE, THE STRATTICE IMPLANT HAD BEEN REMOVED OR REVISED. NO INFORMATION WAS PROVIDED ON THE INDICATION FOR THE INITIAL SURGERY OR THE REASON FOR THE REMOVAL OR REVISION. PATIENT HAS DIED ON (B)(6) 2022 AND THERAILS OE IS NO FURTHER DETN DEATH CAUSALITY. THIS RECORD CAPTURED 1ST STRATTICE. ON (B)(6) 2025, (B)(6) RECEIVED NOTIFICATION FROM LEGAL THAT THE PLAINTIFF PROFILE FORM WAS RECEIVED ASSOCIATED WITH THIS EVENT. AS PER THE PPF FORM, THE PATIENT WAS BORN WITH AN ABDOMINAL WALL DEFECT (GASTROSCHISIS), WHO UNDERWENT A VENTRAL HERNIA SURGERY AND WAS IMPLANTED WITH A STRATTICE DEVICE LOT SP100381-101 ON (B)(6) 2016. THE PATIENT UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGICAL MESH ON (B)(6) 2017. THE PATIENT WAS IMPLANTED WITH A SECOND STRATTICE DEVICE LOT SP100537-078 ON THE SAME DATE. THE PATIENT THEN UNDERWENT AN INCISIONAL/VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WITH A THIRD STRATTICE DEVICE LOT SP100595." THE PATIENT THEN UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH ON (B)(6) 2018 AND IMPLANTED WITH A NONABBVIE DEVICE, XENMATRIX LOT HUCR0748. THE MESH WAS REVISED DUE TO RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON (B)(6) 2018 AND IMPLANTED WAS A SECOND NON-ABBVIE DEVICE, BARD MESH LOT #HUCQ0919. THE MESH WAS REVISED ON (B)(6) 2022 DUE TO OPEN REPAIR OF RECURRENT VENTRAL INCISIONAL HERNIA WITH BIOLOGICAL MESH, ABDOMINAL RE-ENTRY FOR CORRECTION OF RE-HERNIATION AND IMPLANTED WITH A FOURTH STRATTICE DEVICE LOT SP200312-045." THE FORM LISTS THE CONDITIONS THAT WERE TREATED WERE OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC STRATTICE MESH ON OR ABOUT (B)(6) 2016. THE PATIENT ALSO UNDERWENT AN OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) ON OR ABOUT(B)(6) 2017. THE PATIENT WAS TREATED FOR INCISIONAL/ VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH (STRATTICE) BETWEEN (B)(6) 2018. THE PATIENT UNDERWENT OPEN VENTRAL INCISIONAL HERNIA REPAIR WITH MESH (RECURRENT); AT BEDSIDE SEROUS SANGUINOUS OOZING AROUND LEFT JP DRAIN BETWEEN (B)(6) 2018 AND (B)(6) 2018. THE PATIENT WAS TREATED FOR RECURRENT VENTRAL HERNIA REPAIR WITH BIOLOGIC MESH ON OR ABOUT (B)(6) 2018. THE PATIENT WAS SEEN FOR ABDOMINAL PAIN BETWEEN 2018 AND (B)(6) 2022. THE PATIENT WAS TREATED FOR BOWEL OBSTRUCTION BETWEEN APPROXIMATELY 2017-2018. THE PATIENT RECEIVED TREATMENT FOR ABDOMINAL PAIN/ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR BOWEL OBSTRUCTION, ABDOMINAL PAIN, ADHESIONS BETWEEN (B)(6) 2016 AND (B)(6) 2022. THE PATIENT WAS SEEN FOR HERNIA SYMPTOMS FROM APPROXIMATELY 2018 TO (B)(6) 2022. THE PATIENT WAS TREATED OPEN REPAIR OF RECURRENT INCISIONAL HERNIA WITH BIOLOGIC MESH (STRATTICE); ABDOMINAL RE-ENTRY FOR CORRECTION OF REHERNIATION BETWEEN (B)(6) 2022 AND (B)(6) 2022. THE PATIENT RECEIVED POST-OP REHAB BETWEEN 2018-2019. "THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2)". REFERENCE REPORT: MW5167618. LOT #: HUCR0748". | NO CONCLUSIONS CAN BE MADE. AS REPORTED, PATIENT HAD A COMPLEX MEDICAL HISTORY WITH MULTIPLE IMPLANTS AND SURGICAL PROCEDURES. NO MEDICAL RECORDS, AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED. REVIEW OF MANUFACTURING RECORDS FOR THE XENMATRIX AB IMPLANTS USED TO TREAT THE PATIENT CONFIRMED THE PRODUCTS WERE MANUFACTURED TO SPECIFICATION. TO DATE, THIS IS THE ONLY REPORTED COMPLAINT FOR THIS MANUFACTURING LOT OF 31 UNITS. NOTE, THE ACTUAL DATE OF DEATH AND THE DATE OF EVENT WERE NOT REPORTED IN THE REPORT SENT TO BD. WE HAVE DOCUMENTED THESE REQUIRED DATE FIELDS WITH A BEST ESTIMATE. THIS MDR REPRESENTS THE XENMATRIX AB (DEVICE #1). AN ADDITIONAL MDR WAS SUBMITTED TO REPRESENT THE XENMATRIX AB (DEVICE #2). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
CORTRAK 2 NG/NI FEED, ELECTROMAGNETIC STYLET, ENFIT, CONNECTOR | K220588 | IT WAS REPORTED THE NASO-JEJUNAL (NJ) TUBE WAS PLACED WITH THE NEW MAGNETIC DEVICE. THE CAREGIVERS WERE "CONCERNED AS THE TUBE BECAME STIFF VERY QUICKLY." THE PATIENT ENDED UP HAVING A REPLACEMENT ON (B)(6) 2025 WHICH WAS "ONLY AROUND 3-MONTHS LATER." MOREOVER, WHEN INTERVENTIONAL RADIOLOGIST (IR) WENT TO REPLACE IT, THE TUBE HAD RUPTURED, AND A SEGMENT HAD BROKEN OFF AND WAS LODGED IN THE DUODENUM MAKING IT DIFFICULT FOR A FURTHER NJ TO BE SITED BY IR. THE PATIENT WAS "NOT FIT ENOUGH FOR [A] OGD [OESOPHAGO-GASTRO-DUODENOSCOPY] TO REMOVE IT." THE PATIENT "DIED SHORTLY AFTER THE PROCEDURE FROM ASPIRATION PNEUMONIA. IT WAS NOTED THE PATIENT WAS "ADMITTED WITH AN UGI [UPPER GASTROINTESTINAL] BLEED." | THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD AND UDI NUMBER ARE IN-PROGRESS. ALL INFORMATION REASONABLY KNOWN AS OF 03-APR-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
DREAMSTATION AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THROAT CANCER AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE PATIENT WAS ADMITTED TO HOSPITAL. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
TRILOGY | K093905 | THE MANUFACTURER RECEIVED INFORMATION ON (B)(6) 2025 REGARDING A PATIENT DEATH THAT OCCURRED ON (B)(6) 2025 WHILE THE PATIENT WAS ON VENTILATION THERAPY USING A TRILOGY 202 VENTILATOR. THE REPORTER STATED THE CUSTOMER WOULD LIKE TO KNOW IF THE TRILOGY 202 VENTILATOR ALARMED FOR APNEA DURING USE SINCE THE PATIENT DIED WHILE RECEIVING VENTILATION THERAPY. THE REPORTER STATED THEY ARE NOT SURE IF THE VENTILATOR ALARMED OR NOT. IT WAS REPORTED THE PATIENT WAS ON THE VENTILATOR AND HAD BEEN FOR A NUMBER OF DAYS WITH NO ISSUES. THE REPORTER INDICATED THE ERROR LOGS HAD BEEN REVIEWED BUT PROVIDED NO ADDITIONAL INFORMATION REGARDING THE FINDINGS OF THAT REVIEW. THE REPORTER INDICATED THEY WERE UNAWARE OF ANY OTHER DEVICES USED AT THE TIME OF THE EVENT OTHER THAN THE TRILOGY 202 VENTILATOR. MEDICAL ATTENTION WAS REPORTEDLY SOUGHT AS A DIRECT RESULT OF THE ISSUE, BUT THOSE DETAILS HAVE NOT BEEN PROVIDED. | MDR REPORT 2518422-2025-103669 IS A DUPLICATE OF MDR REPORT 2518422-2025-103169. MDR REPORT 2518422-2025-103669 WAS FILED FOR THE SAME ISSUE REPORTED ON MDR REPORT 2518422-2025-103169. 2518422-2025-103669 B5: THE MANUFACTURER RECEIVED INFORMATION ON 03/14/2025 REGARDING A PATIENT DEATH THAT OCCURRED ON 03/08/2025 WHILE THE PATIENT WAS ON VENTILATION THERAPY USING A TRILOGY 202 VENTILATOR. THE REPORTER STATED THE CUSTOMER WOULD LIKE TO KNOW IF THE TRILOGY 202 VENTILATOR ALARMED FOR APNEA DURING USE SINCE THE PATIENT DIED WHILE RECEIVING VENTILATION THERAPY. THE REPORTER STATED THEY ARE NOT SURE IF THE VENTILATOR ALARMED OR NOT. IT WAS REPORTED THE PATIENT WAS ON THE VENTILATOR AND HAD BEEN FOR A NUMBER OF DAYS WITH NO ISSUES. THE REPORTER INDICATED THE ERROR LOGS HAD BEEN REVIEWED BUT PROVIDED NO ADDITIONAL INFORMATION REGARDING THE FINDINGS OF THAT REVIEW. THE REPORTER INDICATED THEY WERE UNAWARE OF ANY OTHER DEVICES USED AT THE TIME OF THE EVENT OTHER THAN THE TRILOGY 202 VENTILATOR. MEDICAL ATTENTION WAS REPORTEDLY SOUGHT AS A DIRECT RESULT OF THE ISSUE, BUT THOSE DETAILS HAVE NOT BEEN PROVIDED.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION THAT ON (B)(6) 2025 AT 08:31, THE PATIENT'S MOTHER DETACHED THE CATHETER MOUNT TO PERFORM SPUTUM SUCTION. AT THAT TIME, A "CIRCUIT DISCONNECTED" ALARM SOUNDED. THE MOTHER WENT OUT AFTER THAT. AT 08:36 THE SAME MORNING, AN ALARM OCCURRED ON THE MONITOR (THE CATHETER MOUNT MAY HAVE COME OFF DURING THAT TIME). IT WAS REPORTED THAT THE PATIENT'S FATHER DID NOT NOTICE THIS. IT WAS REPORTED THERE WAS NO ALARM LOGGED ON THE TRILOGY EVO VENTILATOR. THE SAME MORNING, AT 09:30, THE PATIENT'S FATHER NOTICED THE CATHETER MOUNT WAS DETACHED BECAUSE A "LOW MINUTE" VENTILATION ALARM 1.8L/MIN. HAD OCCURRED. NO MEDICAL INTERVENTION WAS INITIATED. IT WAS CONFIRMED THE PATIENT HAD CARDIOPULMONARY ARREST. THE PATIENT'S CAUSE OF DEATH HAS BEEN REPORTED AS CARDIAC ARREST. CAUSAL RELATIONSHIP BETWEEN THE MEDICAL DEVICE AND THE PATIENT HARM IN THE OPINION OF THE DOCTOR WAS REPORTED AS: OCCURRENCE RECORD OF AN ALARM (OR THE CIRCUIT DISCONNECT ALARM) WHICH OCCURRED TUESDAY, (B)(6) 2025, IS PRESENT ON THE MONITOR. THE FATHER DID NOT NOTICE IT AT THE TIME. THERE IS A POSSIBILITY THAT THE CATHETER MOUNT CAME OFF DURING THE TIME INDICATED, BUT THERE IS NO ALARM LOGGED ON THE TRILOGY EVO VENTILATOR. THE DOCTOR REPORTEDLY CHECKED THE ALARM LOG ON SITE AND PROVIDED THEIR OPINION AS STATED HERE. ADDITIONAL INFORMATION PROVIDED: ON 01 APRIL 2025, THE PHILIPS SALES REPRESENTATIVE CHECKED THE OPERATING STATUS OF THE TRILOGY EVO VENTILATOR AT THE PATIENT'S HOME TOGETHER WITH THE POLICE. THEY CONFIRMED THAT THE DEVICE OPERATED NORMALLY AND ANNUNCIATED ALARMS. IT WAS REPORTED IT MAY TAKE TIME FOR THE DEVICE TO BE RETRIEVED FOR MANUFACTURER'S INVESTIGATION SINCE THE POLICE MAY COLLECT IT. THE DEVICE WAS NOT AVAILABLE AT THE TIME OF REPORTING. | THE DEVICE WAS RETURNED TO THE MANUFACTURER AND EVALUATED ON 24 APRIL 2025 WITH THE FOLLOWING FINDINGS: WHEN THE CATHETER MOUNT WAS REMOVED USING THE SAME CIRCUIT WITH THE SETTING OF THE ISSUE OCCURRING, IT WAS CONFIRMED THAT "LOW MINUTE VENTILATION" AND "CIRCUIT DISCONNECTED" ALARM OCCURRED. PERFORMED AN OPERATIONAL CHECK AS USUAL, AND THERE WERE NO ABNORMALITIES FOUND. UPON CHECKING THE ERROR LOG, IT WAS CONFIRMED THAT "LOW MINUTE VENTILATION" ALARM OCCURRED FOR 3482 SECONDS AND "CIRCUIT DISCONNECTED" ALARM OCCURRED FOR 3475 SECONDS. THERE WERE NO ERRORS INDICATING ABNORMALITIES OF THE DEVICE. WHEN TESTS FOR THE FLOW OR PRESSURE WERE PERFORMED AS A VERIFICATION, NO ABNORMALITIES WERE OBSERVED. THERE WERE NO ABNORMALITIES FOUND DURING AN INTERNAL INSPECTION WITH DISASSEMBLING THE DEVICE. BASED ON THE ABOVE, IT IS BELIEVED THAT THERE WERE NO ABNORMALITIES IN THE DEVICE, AND THE CATHETER MOUNT WAS REMOVED, THEN "LOW MINUTE VENTILATION" AND "CIRCUIT DISCONNECTED" ALARM KEPT SOUNDING. PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY.
TRILOGY 100 | K083526 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT PASSED AWAY WHILE USING A VENTILATOR. THE INFORMATION INDICATES THAT DUE TO A MISTAKE BY A DOCTOR IN THE TREATMENT FOR REPLACING THE TRACHEOSTOMY STOMA CANNULA EVERY TWO WEEKS. THE UNIT HAS BEEN SEIZED BY THE POLICE. THERE IS NO ALLEGATION THE VENTILATOR MALFUNCTIONED OR FAILED TO DELIVER THERAPY AS DESIGNED. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP, JP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE USER DEVELOPED ASPIRATION PNEUMONIA FROM DEVICE DUE TO VOMIT AND LATER PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THERE WAS REPORT OF PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION 2 ADVANCE AUTO CPAP | K200480 | THE MANUFACTURER BECAME AWARE THAT A USER OF THE DREAMSTATION 2 ADVANCE AUTO CPAP HAD PASSED AWAY BEFORE RECEIVING THE REPLACEMENT MACHINE. IT WAS NOT CONFIRMED WHETHER THE DEVICE HAD CAUSED THE DEATH OR NOT. THE USER'S DAUGHTER WANT TO USE THE MACHINE UNDER HER NAME. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
RENTAL, DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A RENTAL, DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION, THE DEVICE WAS VISUALLY INSPECTED INTERNALLY AND EXTERNALLY, AND NO FAULTS WERE FOUND. IN ADDITION, THE DEVICE PASSED ALL TESTS.
AMPLATZER VASCULAR PLUG II | K071699 | THE ARTICLE, "PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE", WAS REVIEWED. THE ARTICLE PRESENTED A PROSPECTIVE, SINGLE CENTER STUDY TO DETERMINE THE SAFETY AND EFFICACY OF PERCUTANEOUS DEVICE CLOSURE FOR SIGNIFICANT PARAVALVULAR LEAKS (PVLS). DEVICES INCLUDED IN THE STUDY WERE AMPLATZER VASCULAR PLUG II, AMPLATZER VASCULAR PLUG III, AND AMPLATZER DUCT OCCLUDER. THE ARTICLE CONCLUDED THAT TRANSCATHETER PVL CLOSURE IS A SAFE AND EFFECTIVE ALTERNATIVE TO SURGICAL CORRECTION AND SHOULD BE USED IN SELECTED PATIENTS FOR SIGNIFICANT AMELIORATION OF HEART FAILURE SYMPTOMS AS WELL AS HEMOLYTIC ANEMIA WITH CLINICAL AND HEMODYNAMIC IMPROVEMENT. [THE PRIMARY AUTHOR WAS UDAY B KHANOLKAR, NARAYANA HRUDYALAYA, BENGALURU, INDIA. THE CORRESPONDING AUTHOR WAS MEEMANSA KASHYAP BUCH, NARAYANA HRUDYALAYA, BENGALURU, INDIA.] THE TIME FRAME OF THE STUDY WAS FROM 2020 TO 2023. A TOTAL OF 42 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH ALL PATIENTS RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE WAS 50.11 YEARS AND THE MAJORITY GENDER WAS MALE. COMORBIDITIES INCLUDED HEMOLYTIC ANEMIA AND CONGESTIVE CARDIAC FAILURE. | INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. ARTICLE TITLED " PERCUTANEOUS DEVICE CLOSURE FOR PARAVALVULAR LEAKS- A SINGLE CENTRE EXPERIENCE."
HAKIM PROGRAMMABLE VALVE | K221840 | THIS IS 2 OF 2 REPORTS LINKED TO MFG REPORT NUMBERS: 3013886523-2025-00069. A PHYSICIAN REPORTED: ON AN UNSPECIFIED DATE, A CODMAN HAKIM PROGRAMMABLE VALVE (CHPV) WAS IMPLANTED FOR AN UNREPORTED INDICATION. ON AN UNKNOWN DATE, THE PATIENT HAD INEFFICIENCY MANAGING THE CSF (CEREBROSPINAL FLUID) FLOW. THE PATIENT TRAGICALLY PASSED AWAY ON AN UNKNOWN DATE LAST YEAR DUE TO VENTRICULAR COLLAPSE, WHICH WAS ATTRIBUTED TO THE VALVE'S PERFORMANCE BY THE PHYSICIAN. IT WAS UNKNOWN IF AN AUTOPSY WAS PERFORMED. NO PATIENT INFORMATION PROVIDED INCLUDING: AGE, GENDER, CONCURRENT CONDITION, FAMILY HISTORY PAST MEDICATION, CONCOMITANT MEDICATION AND CO-SUSPECT MEDICATIONS, DRUG ALLERGIES. NO ADDITIONAL INFORMATION HAS BEEN PROVIDED AFTER SEVERAL ATTEMPTS. | AN INVESTIGATION HAS BEEN INITIATED BASED ON THE REPORTED INFORMATION. UPON COMPLETION OF THE INVESTIGATION, A FOLLOW-UP REPORT WILL BE SUBMITTED.
HAKIM PROGRAMMABLE VALVE | K221840 | N/A | UPDATED FIELDS: D9, G3, G6, H2, H3, H6, H11. THE HAKIM VALVE (UNKNOWN ID) WAS NOT RETURNED FOR EVALUATION AFTER THREE GOOD FAITH ATTEMPTS (GFES) WERE MADE. LOT NUMBER INFORMATION HAS NOT BEEN PROVIDED; THEREFORE, AN EVALUATION OF THE DEVICE COULD NOT BE PERFORMED, AND DEVICE HISTORY RECORD (DHR) COULD NOT BE REVIEWED. THE ROOT CAUSE(S) OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF ADDITIONAL RELEVANT INFORMATION BECOMES AVAILABLE IN THE FUTURE, THIS COMPLAINT WILL BE REOPENED, AND THE RESPECTIVE EVALUATION PERFORMED. TRENDS WILL BE MONITORED FOR THIS AND SIMILAR ISSUES. AT PRESENT, WE CONSIDER THIS COMPLAINT TO BE CLOSED. A MEDICAL ASSESSMENT WAS REQUESTED TO INTEGRA'S MEDICAL SAFETY TEAM: "THERE IS NO INDICATION THAT THE DEVICES MALFUNCTIONED BUT DO SUGGEST THAT PROCEDURAL DEVIATIONS FROM RECOMMENDED INSTRUCTIONS MAY HAVE CONTRIBUTED TO THE EVENTS, AS DEVICES WERE NOT USED AND MONITORED IN ACCORDANCE WITH ALL INSTRUCTIONS INCLUDED IN THE INSTRUCTIONS FOR USE (IFU). THE INFORMATION GATHERED DOES SUGGEST USE ERROR ONLY. IT APPEARS PROGRAMMING WAS PERFORMED BY UNTRAINED RESIDENTS WHO MAY NOT HAVE POSITIONED THE TRANSMITTER PROPERLY ON THE VALVE NOR RESTRICTED MOVEMENT OF THE COIL DURING PROGRAMING, OR OBSERVED THE PROPER OPERATING TEMPERATURE OF THE PROGRAMMER. ADDITIONALLY, THEY DID NOT TAKE CONFORMATIONAL RADIOGRAPHS POST-PROGRAMMING NOR PRE-PROGRAMMING CONFIRM THE SETTING AND IDENTIFY IF THERE HAD BEEN ANY CHANGE OF VALVE SETTING DUE TO STRONG MAGNETIC FORCE. THE HOSPITAL DOES NOT USE EVDÍS ROUTINELY DURING STABILIZATION OF THEIR TRAUMA PATIENTS BUT INSTEAD USE SHUNTS WHERE FIBRIN CLOTS MAY FORM AND DISRUPT VALVE PROGRAMMING. AS PER A CONFERENCE CALL WITH DR. (B)(6) MEDICAL ADVISOR DR. (B)(6). ON (B)(6) 2025, DR. (B)(6) GROUP WILL BE MEETING WITH DR.(B)(6). TO DISCUSS TRAINING, AND CONFIRM ANY ADDITIONAL COMPLAINTS THAT MAY HAVE OCCURRED AND SHOULD BE REPORTED TO INTEGRA."
SAVVYWIRE | K213854 | EVENT AS DESCRIBED BY OPSENS SALES REPRESENTATIVE: PROCEDURE TYPE: TAVR. VALVE: EDWARDS SAPIEN 26MM. LOSS OF CAPTURE DURING DEPLOYMENT OF THE VALVE. THE EDWARDS VALVE EMBOLIZED. INTERVENTION WAS PERFORMED TO MOVE THE VALVE TO AN APPROPRIATE AREA IN THE AORTA. THE PATIENT DID NOT SURVIVE DESPITE THEIR EFFORTS. PACING RATE 180, OUTPUT OF 20MA. VOO WAS CONFIRMED. BEFORE THE DEPLOYMENT OF THE VALVE TEST PACING WAS CONDUCTED. PACING RATE OF 120, OUTPUT OF 20MA WALKED DOWN UNTIL LOSS OF CAPTURE AT 10MA. | THE SAVVYWIRE WAS DISCARDED AND NOT RETURNED TO THE MANUFACTURER. NO INSPECTION WAS POSSIBLE. NO IMAGING AND NO ADDITIONAL DATA WERE PROVIDED BY THE ACCOUNT AS OF APRIL 4TH, 2025. CHECK OF THE DEVICE HISTORY RECORDS (DHR) OF GUIDEWIRE LOT NUMBER OSW-0323 SHOWED THAT IT WAS RELEASED PER SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION, PACING TEST WAS PERFORMED BEFORE VALVE DEPLOYMENT, AS PRESCRIBED, CONFIRMING VENTRICULAR CAPTURE AND THRESHOLDS. WITHOUT INSPECTING THE WIRE, OPSENS IS NOT ABLE TO FULLY INVESTIGATE THEREFORE NO DEFINITIVE ROOT CAUSE CAN BE CONCLUDED IN THIS REPORT. IF THERE IS ANY FURTHER RELEVANT INFORMATION, A FOLLOW UP REPORT WILL BE SUBMITTED TO THE FDA. THE ADVERSE EVENT RESULT OF THIS INCIDENT IS WELL MENTIONED IN THE SAVVYWIRE INSTRUCTIONS OF USE (IFU): ADVERSE EVENTS THAT MAY RESULT FROM THE USE OF THIS DEVICE INCLUDE, BUT ARE NOT LIMITED TO: ACCESS SITE OR VESSELS COMPLICATIONS, ADDITIONAL SURGICAL PROCEDURE, ALLERGIC REACTIONS, AMPUTATION, ANEURYSM, ANGINA, ARRHYTHMIA, BLEEDING, CARDIAC OR VESSEL PERFORATION/DISSECTION, CORONARY OBSTRUCTION, DEATH, EMBOLISM, FIBRILLATION, FOREIGN BODY/WIRE FRACTURE, HEART BLOCK, HEMATOMA, HYPOTENSION/HYPERTENSION, INFECTION, KIDNEY INJURY/FAILURE, MYOCARDIAL INFARCTION, NEED FOR PERMANENT PACEMAKER, PERICARDIAL EFFUSION, PNEUMOTHORAX, STROKE OR OTHER NEUROLOGIC EVENT, SPASM, TAMPONADE, THROMBUS, VALVE DYSFUNCTION OR COMPLICATIONS, VALVE MALPOSITIONING OR EMBOLIZATION, VASOSPASM, VESSEL OCCLUSION, WIRE ENTRAPMENT/ENTANGLEMENT, X-RAY RADIATION EXPOSURE COMPLICATIONS.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED TO THE CLINIC. NO ADDITIONAL INFORMATION WAS PROVIDED DURING INTAKE, AND SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL CLINICAL INFORMATION (E.G., PATIENT DEMOGRAPHICS, TREATMENT RECORD, HOSPITAL RECORDS, DEATH CERTIFICATE) WERE UNSUCCESSFUL. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. AN EXTERNAL VISUAL INSPECTION WAS PERFORMED ON THE CYCLER AND SHOWED BEZEL DAMAGE. ALTHOUGH THERE WAS EVIDENCE OF DRIED FLUID PRESENT WITHIN THE CASSETTE COMPARTMENT, THERE WERE NO BURRS OR SHARP EDGES IN THE CASSETTE AREA THAT COULD HAVE PUNCTURED THE CASSETTE MEMBRANE. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. AN AS-RECEIVED WITH REDUCED DWELL SIMULATED TREATMENT WAS PERFORMED ON THE CYCLER AND PASSED. THE CYCLER UNDERWENT AND PASSED A SYSTEM AIR LEAK TEST AND VALVE ACTUATION TEST. AN INVESTIGATION OF THE CYCLER MUSHROOM HEADS VERIFIED THAT THE SURFACE CONDITIONS AND ALIGNMENTS WERE WITHIN SPECIFICATION. AN INTERNAL VISUAL INSPECTION IDENTIFIED EVIDENCE OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE DRIED FLUID COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
FREESTYLE LIBRE 2 | K193371 | ABBOTT DIABETES CARE RECEIVED A MEDWATCH REPORT #MW5167182 WHICH REPORTED THE FOLLOWING INFORMATION PERTAINING TO AN ABBOTT DIABETES CARE (ADC) DEVICE: A HEALTHCARE PROFESSIONAL REPORTED ON (B)(6) 2025 THAT A CUSTOMER PASSED AWAY WHILE WEARING AN ADC DEVICE; HOWEVER, NO FURTHER INFORMATION WAS PROVIDED. NO CONTACT INFORMATION WAS PROVIDED TO ADC; THEREFORE, ADC IS UNABLE TO ATTEMPT ANY FOLLOW UP ACTIVITIES RELATED TO THIS EVENT. THE HCP MADE NO ALLEGATION OF DEFICIENCY WITH THE ADC DEVICE; THEREFORE, BASED ON THE LIMITED INFORMATION AVAILABLE, THERE IS NO EVIDENCE TO INDICATE THAT THE ABBOTT DIABETES CARE (ADC) DEVICE CONTRIBUTED TO THE REPORTED INCIDENT. | THIS COMPLAINT WAS RECEIVED VIA USER REPORT # MW5167182. ABBOTT DIABETES CARE (ADC) IS UNABLE TO FURTHER INVESTIGATE THE REPORTED COMPLAINT DUE TO THE ABSENCE OF AN ALLEGED PRODUCT ISSUE. ADDITIONALLY, NO CONTACT INFORMATION WAS PROVIDED IN THE USER REPORT, RENDERING ADC UNABLE TO CONDUCT ANY FOLLOW-UP ACTIVITIES. THE DEVICE MANUFACTURING DATE IS UNKNOWN, AND THE DATE ENTERED IN SECTION H4 REFLECTS THE DATE ADC BECAME AWARE OF THE EVENT. ALL PERTINENT INFORMATION AVAILABLE TO ADC HAS BEEN SUBMITTED.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | CORRECTION TO THE PREVIOUSLY REPORTED INFORMATION: THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS INSPECTED AND NO FAULT FOUND. ADDITIONALLY, ALL TESTS PASSED. THE "PROBLEM CODE GRID" HAS BEEN CORRECTED IN THIS REPORT. IN THIS REPORT, BOX H HAS BEEN UPDATED/CORRECTED.
REACT CATHETER | K182097 | MATSUMOTO, S., SUZUKI, T., EMOTO, S., NAKASE, H., MATSUMURA, Y.; INTERNAL MEDICINE; 2024; 63: 3087-3091; A CEREBRAL EMBOLISM CAUSED BY A MALIGNANT PERIPHERAL NERVE SHEATH TUMOR IN A PATIENT WITH NEUROFIBROMATOSIS TYPE 1; DOI: 10.2169/INTERNALMEDICINE.2996-23 ARTICLE OBJECTIVE: THE DOCUMENT REPORTS A CASE OF A CEREBRAL EMBOLISM CAUSED BY A MALIGNANT PERIPHERAL NERVE SHEATH TUMOR (MPNST) IN A PATIENT WITH NEUROFIBROMATOSIS TYPE 1 (NF-1). THE PATIENT WAS A 31-YEAR-OLD MAN WHO UNDERWENT RESECTION OF THE MPNST ON THE BUTTOCK 3 MONTHS PRIOR TO PRESENTING WITH A HYPERACUTE LEFT MIDDLE CEREBRAL ARTERY EMBOLISM. THE CASE REPORT DETAILS EVENTS THAT OCCURRED OVER A HOSPITALIZATION PERIOD OF 15 DAYS. THE DOCUMENT DOES NOT EXPLICITLY STATE WHETHER MULTIPLE MANUFACTURERS' DEVICES WERE USED. IT PRIMARILY MENTIONS THE USE OF A SPECIFIC MEDTRONIC PRODUCT FOR MECHANICAL THROMBECTOMY. THE DOCUMENT MENTIONS THE USE OF THE ASPIRATION CATHETER REACT-71 FOR MECHANICAL THROMBECTOMY. THE PATIENT DIED ON THE 15TH DAY OF HOSPITALIZATION DUE TO RESPIRATORY FAILURE CAUSED BY LUNG METASTASIS OF THE MPNST. ADVERSE EVENTS/COMPLICATIONS: ON THE 14TH DAY OF ADMISSION, HIS CONSCIOUSNESS LEVEL DECREASED TO GCS E3V2M5, AND NEUROLOGICAL EXAMINATION DEMONSTRATED CONJUGATE DEVIATION OF THE LEFT EYE WITH RIGHT UPPER AND LOWER LIMB PARALYSIS. IN ADDITION, HIS RESPIRATORY CONDITION WORSENED TO SPO2 94% (O2 1 L/MIN). BRAIN COMPUTED TOMOGRAPHY (CT) REVEALED AN INTRATUMORAL HEMORRHAGE IN THE LEFT FRONTAL LOBE NEOPLASM. IN ADDITION, BODY CONTRAST CT REVEALED MULTIPLE TUMORAL LESIONS IN THE LEFT LUNG AND MASSIVE PLEURAL EFFUSION WITH ATELECTASIS IN THE RIGHT LUNG. CT DID NOT REVEAL ARTERIOSCLEROTIC LESIONS IN THE AORTA. ANOTHER TUMOR LESION WAS OBSERVED IN THE RIGHT KIDNEY. AS IT WAS CONSIDERED THAT THE PATIENT WOULD NOT BE ABLE TO TOLERATE SURGERY DUE TO SEVERE ATELECTASIS, CONSERVATIVE TREATMENT FOR THE INTRATUMORAL HEMORRHAGE AND DRAINAGE OF THE PLEURAL FLUID WAS CONDUCTED. ON THE 15TH DAY AFTER ADMISSION, HIS RESPIRATORY STATUS ABRUPTLY WORSENED AND HE DIED OF RESPIRATORY FAILURE. ON AUTOPSY, MULTIPLE NODULES OF THE MPNST OCCUPIED THE RIGHT LUNG. MOREOVER, A LUNG HISTOLOGICAL EXAMINATION OF THE AUTOPSY CONFIRMED TUMOR INVASION OF THE PULMONARY VEINS. THE RESULTS OF ELASTICA VAN GIESON STAINING CONFIRMED THAT THE LESION WAS VENOUS INVASION. MOST TUMOR CELLS HAD SPINDLE-SHAPED AND HYPERCHROMATIC NUCLEI THAT WERE IDENTIFIED AS MPNST. MULTIPLE VASCULAR INVASIONS OF THE MPNST, SUCH AS IN THE KIDNEYS AND RIGHT THYROID GLAND, WERE ALSO OBSERVED. VERRUCAE, CALCIFICATION OF THE VALVES, AND PATENT FORAMEN OVALE WERE NOT OBSERVED IN THE HEART. FURTHERMORE, THERE WAS NO EVIDENCE OF CEREBROVASCULAR ATHEROSCLEROSIS. IN ADDITION, THE HEMORRHAGE IN THE LEFT FRONTAL LOBE WAS HISTOPATHOLOGICALLY IDENTIFIED TO BE THE RESULT OF MPNST. THEREFORE, RESPIRATORY FAILURE DUE TO METASTATIC MPNST IN THE LUNGS AND CEREBRAL HEMORRHAGE CAUSED BY INTRACRANIAL MPNST WERE REGARDED AS THE CAUSE OF DEATH. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER BECAME AWARE OF A REP DREAMSTATION AUTO CPAP DEVICE ALLEGING DEATH OF THE PATIENT. NO MEDICAL INTERVENTION WAS SPECIFIED BY THE PATIENT. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
TI-CRON | K930591 | IT WAS REPORTED THAT DURING AN OPEN EMERGENCY EXPLORATION OF MEDIASTINUM AND EVACUATION OF PERICARDIAL TAMPONADE, THE SUTURE WAS USED WHILE BEING EXPIRED. THE PATIENT DIED. THE PATIENT'S DEATH WAS NOT RELATED TO THE DEVICE. | ADDITIONAL INFORMATION: B5, G3, H6 ADDITIONAL INFORMATION RECEIVED SHOWS THAT THERE IS NO INFORMATION TO REASONABLY SUGGEST THAT THE DEVICE IN THIS REPORT MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH OR SERIOUS INJURY OR THAT THE DEVICE IN THIS REPORT HAS MALFUNCTIONED. THEREFORE, THIS EVENT DOES NOT MEET THE REPORTING REQUIREMENTS IN 21 CFR 803. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGES DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. REPEATED ATTEMPTS TO HAVE THE DEVICE AND COMPONENTS RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. THE MANUFACTURER BELIEVES THEY WILL BE UNABLE TO GATHER ADDITIONAL INFORMATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED. IN THIS REPORT IN BOX G: DATE RECEIVED BY MFG, IN BOX H: EVAL METHOD CODE, EVAL RESULT CODE, CONCLUSION CODE HAS BEEN UPDATED.
MITRACLIP | K190167 | IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT MITRAL REGURGITATION (MR) WITH A GRADE OF UNKNOWN. IT WAS NOTED DURING THE INSERTION OF THE STEERABLE GUIDE CATHETER (SGC), THE DEVICE WAS GETTING HUNG UP ON THE PERCLOSES THAT WERE DEPLOYED PRIOR TO THE GUIDE INSERTION. THE SGC WAS ABLE TO BE ADVANCE INTO THE RIGHT ATRIUM, BLEEDING WAS OBSERVED AROUND THE SGC AND A HEMATOMA WAS FORMING AT THE SITE. IT WAS DECIDED TO IMPLANT ONE MITRACLIP. THE PATIENT'S BLOOD PRESSURE DECREASED ONCE THE CLIP WAS INTRODUCED. THE SGC WAS THEN REMOVED AND THE ACCESS SITE WAS STITCHED UP. PRESSURE WAS APPLIED. HOWEVER, BLEEDING CONTINUED AND FURTHER STITCHING WAS PERFORMED. THE FEMORAL ILIAC VEIN WAS THEN TIED OFF TO PREVENT BLEEDING. BLEEDING CONTINUED. THE PATIENT EXPIRED. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. ADDITIONALLY, A REVIEW OF THE COMPLAINT HISTORY IDENTIFIED NO SIMILAR COMPLAINTS REPORTED FROM THIS LOT. THE INVESTIGATION DETERMINED THE REPORTED PHYSICAL RESISTANCE DURING ADVANCING AND DIFFICULT TO ADVANCE APPEAR TO BE DUE TO THE PROCEDURAL CONDITION DUE TO THE PERCLOSES. A CAUSE FOR THE REPORTED PATIENT EFFECTS OF PERFORATION OF VESSELS (RUPTURE) AND DEATH CANNOT BE DETERMINED. THE BLEEDING/HEMORRHAGE WAS DUE TO THE RUPTURE. THE HEMATOMA AND HYPOTENSION APPEAR TO BE CASCADING EFFECTS OF THE BLEEDING. PERFORATION OF VESSELS (RUPTURE), BLEEDING, HEMATOMA, DEATH, AND HYPOTENSION ARE LISTED IN THE INSTRUCTIONS FOR USE AS KNOWN POSSIBLE COMPLICATIONS ASSOCIATED WITH THE MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGICAL INTERVENTION WERE RESULTS OF CASE SPECIFIC CIRCUMSTANCES AS PRESSURE WAS APPLIED AND THE VASCULAR OPENED THE PATIENTS GROIN AND TIED OFF THE FEMORAL ILIAC VEIN. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN OR LABELING.
RHYTHM ANALYTICS - SOFTWARE (SAMD) | K182344 | COMPLAINT RECEIVED THAT PATIENT PASSED AWAY ON (B)(6) 2025 WHILE ENROLLED IN A HOSPITAL MONITORING PROGRAM ALLEGEDLY DUE TO MYOCARDIAL INFARCTION OR PULMONARY EMBOLISM. | PLEASE NOTE THAT THE INVESTIGATION IS STILL IN PROCESS AT THIS TIME. PLEASE NOTE THAT THE INITIAL ATTEMPT TO FILE THE MEDWATCH FORM 3500A FOR THIS EVENT OCCURED ON 27 MARCH, 2025 IN THE FORM OF AN EMAIL TO THE EMDR AT MDRPOLICY@FDA.HHS.GOV. THE HELP TEAM REDIRECTED THE EMAIL AND INFORMED BIOFOURMIS THAT THE REPORT REQUIRED SUBMISSION THROUGH THE ESG WITHIN THE WEBTRADER SYSTEM. THE BIOFOURMIS TEAM INITIATED THE PROCESS TO APPLY FOR AN ACCOUNT UNDER THE WEBTRADER / ESG SYSTEM BUT WAS UNABLE TO COMPLETE THE ACCOUNT REGISTRATION PROCESS WITHIN THE REQUIRED 30 DAYS FROM BECOMING AWARE OF THE EVENT. ON 1 APRIL, THE NEXT GEN ESG WAS MADE AVAILABLE AND THE LEGACY TEST DATABASE SYSTEM WAS NO LONGER ACTIVE REQUIRING ADDITIONAL STEPS TO COMPLETE THE REGISTRATION PROCESS. EMAILS OF THE EXCHANGE BETWEEN THE ESG AND ESGNG SUPPORT TEAM AND BIOFOURMIS ARE AVAILABLE FOR REVIEW. AT THE TIME OF THIS REPORT THAT WAS GENERATED USING THE E-SUBMITTER AND MOST CURRENT APPLICATIONS UTILIZING THE SAME INFORMATION ORIGINALLY SENT TO THE MDR TEAM ON 27 MAR 2025, BIOFOURMIS WAS STILL AWAITING ASSISTANCE FROM THE ESGNG SUPPORT TEAM TO REGISTER AN ACCOUNT WITHIN THE LEGACY ESG PRODUCTION SYSTEM.
SUREFORM | K183224 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | ADDITIONAL INFORMATION: THE SUREFORM 45 CURVED-TIP STAPLER INSTRUMENT WAS RECEIVED FOR FAILURE ANALYSIS. THE WHITE RELOAD WAS NOT RETURNED FOR INVESTIGATION WITH THE STAPLER. VISUAL INSPECTION DID NOT FIND EVIDENCE OF ANY PROBLEMS. THE INSTRUMENT WAS TESTED ON AN IN-HOUSE SYSTEM, PASSED ALL TESTING AND WAS FULLY FUNCTIONAL. THE FIRING TEST WAS COMPLETED TWO TIMES IN DIFFERENT POSITIONS AND PASSED EACH TIME. NO PRODUCT ISSUE WAS IDENTIFIED.
REMSTAR AUTO A-FLEX W/HUMID, SD CARD | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHENOM CATHETER | K151638 | ADDITIONAL INFORMATION WAS RECEIVED STATING THAT THE PATIENT HAD BEEN HOSPITALIZED PRIOR TO THEIR PASSING AND THEY WERE ADMITTED WITH SPONTANEOUS INTRACEREBRAL HEMORRHAGE (SICH). THE INFORMATION IS CONFIRMED BY THE PHYSICIAN. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DS2ADV AUTO CPAP | K200480 | THIS COMPLAINT HAS BEEN REVIEWED AND IT HAS BEEN DETERMINED THAT THE CUSTOMER HAS ALLEGED PATIENT PASSED AWAY. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
AMPLATZER TORQVUE 45X45 DELIVERY SHEATH | K163000 | SUBSEQUENT TO THE PREVIOUSLY FILED REPORT, ADDITIONAL INFORMATION WAS RECEIVED THAT THE PATIENT WAS ADMINISTERED 6000 UNITS OF HEPARIN FOR PROCEDURE. THE LARGEST DIMENSION OF THE PERICARDIAL EFFUSION WAS UNABLE TO BE OBTAINED DUE TO THE EMERGENCY OF THE SITUATION AND THE PATIENT CODING. | AN EVENT FOR PATIENT EFFECTS OF PERICARDIAL EFFUSION, CARDIAC ARREST, CARDIAC TAMPONADE, AND DEATH WAS REPORTED. A RETURNED DEVICE ASSESSMENT COULD NOT BE PERFORMED AS THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED AND THE PRODUCT MET ALL DEFINED MANUFACTURING SPECIFICATIONS. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED PERICARDIAL EFFUSION CAN LEAD TO CARDIAC TAMPONADE, WHICH CAN CAUSE CARDIAC ARREST, APPEARS TO BE RELATED PROCEDURAL CONDITIONS. THE CAUSE OF THE REPORTED DEATH APPEARS TO BE RELATED TO RIGHT VENTRICLE FAILURE. THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING.
AVEIR¿ INTRODUCER KIT ¿ 50 CM | K211989 | IT WAS REPORTED THE PATIENT PRESENTED FOR IMPLANT. DURING SURGERY, THE INTRODUCER PERFORATED THROUGH THE HEART STRUCTURE. THE PATIENT'S BLOOD PRESSURE DROPPED AND TAMPONADE OCCURRED. THE PATIENT DECEASED. | A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND ALL REQUIRED MANUFACTURING PROCESSES AND INSPECTIONS STEPS WERE CONFIRMED TO BE COMPLETED PER THE REQUIREMENTS. THE DEVICE MET SPECIFICATIONS PRIOR TO LEAVING ABBOTT MANUFACTURING FACILITIES.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A HOME THERAPY PROGRAM MANAGER AND PERITONEAL REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO A CARDIAC ARREST. IT WAS AFFIRMED THE PATIENT WAS CONNECTED TO HER LIBERTY SELECT CYCLER AT THE TIME OF THE CARDIAC ARREST AND HER SUBSEQUENT DEATH. THE PATIENT¿S CARDIAC ARREST WAS ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOPULMONARY DISEASE THAT WAS UNRELATED TO CCPD THERAPY. IT WAS EXPLAINED THAT THE PATIENT WAS FOUND UNRESPONSIVE AND PULSELESS ON THE MORNING OF (B)(6) 2025. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS NOT TRANSPORTED TO THE HOSPITAL AND PRONOUNCED DECEASED IN HER HOME. IT WAS CONFIRMED THE PATIENT¿S CARDIAC ARREST LEADING TO HER DEATH WAS NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | CORRECTION PROVIDED IN G4. ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE VISUAL NO INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. AN (AS-RECEIVED WITH REDUCED DWELL) SIMULATED TREATMENT WAS PERFORMED AND COMPLETED. VALVE ACTUATION TEST PASSED. SYSTEM AIR LEAK TEST PASSED. AN INTERNAL VISUAL INSPECTION OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
POWERED WHEELCHAIR SYSTEM | K182995 | ON MARCH 12, 2025, INVACARE NOTIFIED MOTION CONCEPTS LP OF A INCIDENT INVOLVING ONE OF OUR WHEELCHAIRS, IN WHICH THE END-USER TRAGICALLY PASSED AWAY FROM A HEART ATTACK, AFTER BECOMING PINNED UNDER THE WEIGHT OF THE CHAIR. THE PRODUCT HAD BEEN SOLD TO THE END-USER IN THE UNITED STATES THROUGH THE VAMC - SALISBURY, NC (659/121). AS REPORTED BY INVACARE, THE INCIDENT TOOK PLACE WHILE THE END-USER WAS ATTEMPTING TO DRIVE THE WHEELCHAIR UP A RAMP LEADING INTO HIS RESIDENCE. DURING THE PROCESS, THE WHEELCHAIR REPORTEDLY TIPPED OVER, PINNING THE USER UNDERNEATH. THE END-USER SUFFERED A HEART ATTACK AND SUBSEQUENTLY PASSED AWAY. THE END-USER'S NIECE, WHO WAS PRESENT AT THE TIME OF THE INCIDENT, CALLED EMERGENCY SERVICES. IT WAS FURTHER REPORTED THAT AN AUTOPSY WAS CONDUCTED. AT THE TIME OF THE EVENT, THE SYSTEM HAD BEEN IN USE FOR APPROXIMATELY FIVE MONTHS. | THE MENTIONED SYSTEM WAS SHIPPED FROM MOTION CONCEPTS TO (B)(4) ((B)(4) IMPORTER/DISTRIBUTOR), AS PER DEALER REQUIREMENTS, ON 18-OCT-2024. THE SYSTEM WAS SUBSEQUENTLY SOLD TO THE (B)(6) ON NOVEMBER 19, 2024. A REVIEW OF THE DEVICE HISTORY CONFIRMED THAT NO ISSUES OR CONCERNS HAD BEEN REPORTED BY EITHER THE DEALER OR THE END-USER, PRIOR TO THE REPORTED INCIDENT. MOTION CONCEPTS BECAME AWARE OF THE INCIDENT WHEN AN (B)(6) REPRESENTATIVE VISITED THE END-USER'S RESIDENCE, BY SCHEDULED APPOINTMENT, TO RETURN THE USER'S PREVIOUS WHEELCHAIR. DURING THE VISIT, THE END-USER'S NIECE INFORMED THE REPRESENTATIVE ABOUT THE INCIDENT AND SPECIFICALLY REQUESTED THAT THE SYSTEM INVOLVED IN THE INCIDENT NOT BE HANDLED. DESPITE MOTION CONCEPTS REQUESTING FOR FURTHER DETAILS, NO ADDITIONAL INFORMATION WAS PROVIDED. THIS LACK OF INFORMATION HAS LIMITED OUR ABILITY TO PERFORM AN EVALUATION OR DETERMINE WHETHER THE INCIDENT WAS RELATED TO A SYSTEM MALFUNCTION. THE (B)(6) REPRESENTATIVE CAPTURED SEVERAL PHOTOS OF THE WHEELCHAIR RAMP AT THE RESIDENCE, WHICH WERE SHARED WITH MOTION CONCEPTS WHEN (B)(6) REPORTED THE INCIDENT. UPON A REVIEW OF THE IMAGES, IT APPEARS THAT THE RAMP WAS CONSTRUCTED FROM ORIENTED STRAND BOARD AND WAS LACKING A CURB RAMPS OR CURB CUTS, WHICH IS AN ESSENTIAL DESIGN SAFEGUARD TO HELP TO PREVENT A WHEELCHAIR FROM ACCIDENTALLY TIPPING OVER OR DRIVING OFF THE EDGE. BASED ON THE INFORMATION MADE AVAILABLE TO MOTION CONCEPTS, THERE IS NO INDICATION THAT THERE WAS A SYSTEM MALFUNCTION. HOWEVER, SINCE THE INCIDENT INVOLVED ONE OF OUR WHEELCHAIRS, WE ARE CONSIDERING THIS AS A REPORTABLE EVENT.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | A PERITONEAL DIALYSIS (PD) NURSE REPORTED THAT A PD PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. REPORTEDLY, THE CYCLER WAS NOT ACTIVELY RUNNING. THE PATIENT¿S TREATMENT HAD BEEN COMPLETED. THE PATIENT WAS FOUND BY SPOUSE DECEASED IN THE BEDROOM AT APPROXIMATELY 7AM. IN A FOLLOW UP, ADDITIONAL DETAILS VERIFIED THAT THE PATIENT WAS FOUND DECEASED AT APPROXIMATELY 7AM ON (B)(6) 2025. THE PATIENT¿S SPOUSE FOUND PATIENT COLD AND PULSELESS. THE SPOUSE ADMINISTERED CARDIOPULMONARY RESUSCITATION (CPR) UNTIL THE RESCUE SQUAD ARRIVED. THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME AND TRANSPORTED TO THE FUNERAL HOME. THE PATIENT¿S TREATMENT WAS RECORDED AS COMPLETED AT 5:29AM, HOWEVER; THE PATIENT HAD NOT DISCONNECTED FROM THE LIBERTY SELECT CYCLER. THE EXACT TIME OF DEATH WAS NOT REPORTED. NO ISSUES WERE REPORTED WITH THE CYCLER OR THE PATIENT¿S TREATMENT. THERE WERE 0 DRAIN ALARMS AND 0 FILL ALARMS PER THE TREATMENT SHEET. THE MACHINE HAD A HISTORY OF 5 ALARMS (UNKNOWN TYPE OR DATE). THE PATIENT HAD REPORTED SLOW DRAINS PREVIOUSLY. THE CAUSE OF DEATH WAS DOCUMENTED AS CARDIAC ARREST PER THE DOCTOR. THE DEATH WAS NOT RELATED TO ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY. | ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL ANALYSIS. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT ON THE PUMP. THERE WERE NO VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT ANY FAILURES OR PROBLEMS. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. SYSTEM AIR LEAK TEST PASSED. TEACH PUMP TEST PASSED. VALVE ACTUATION TEST PASSED. AN INTERNAL VISUAL OF THE RETURNED CYCLER WAS PERFORMED. THERE WERE VISUAL INDICATIONS OF DRIED FLUID ON THE BOTTOM COVER BEHIND THE FRONT PANEL ASSEMBLY. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED. THERE WERE NO DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S).UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
SOLITAIRE FR4 | K183022 | LITERATURE WAS REVIEWED REGARDING: "BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE." THE OBJECTIVE IS TO PRESENT A CASE OF MECHANICAL THROMBECTOMY AND SIMULTANEOUS DISTAL VERTEBRAL ARTERY EMBOLIZATION FOR BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 71 CATHETER, SOLITAIRE X, AND PHENOM 17 CATHETER. AMONG PATIENT ADVERSE EVENTS INCLUDED: THE STUDY DESCRIBES A CASE INVOLVING A PATIENT WHO EXPERIENCED A FALL AND SUBSEQUENT MEDICAL INTERVENTIONS. THE PATIENT DEVELOPED ALTERED CONSCIOUSNESS SIX HOURS AFTER THE CERVICAL SPINE TRAUMA AND HAD A RELATIVELY EARLY ONSET OF BASILAR ARTERY OCCLUSION CAUSED BY VERTEBRAL ARTERY INJURY. MECHANICAL THROMBECTOMY WAS PERFORMED USING THE REACT AND SOLITAIRE DEVICE, AND EMBOLIZATION OF THE INJURED VERTEBRAL ARTERY AT THE SOURCE OF THE EMBOLISM USING COILS DELIVERED THROUGH THE PHENOM CATHETER. COMPLETE RECANALIZATION WAS ACHIEVED DURING THE THROMBECTOMY. NO IMPROVEMENT IN NEUROLOGICAL SYMPTOMS WAS OBSERVED AFTER THE SURGERY. DIFFUSION-WEIGHTED IMAGING PERFORMED 13 H AFTER THE ONSET OF ALTERED CONSCIOUSNESS SHOWED INFARCTION IN THE MIDBRAIN, PONS, BOTH OCCIPITAL LOBES, BOTH CEREBELLAR HEMISPHERES, AND THE LEFT THALAMUS. PATENCY OF THE RIGHT VERTEBRAL AND BASILAR ARTERIES AND OCCLUSION OF THE LEFT VERTEBRAL ARTERY WERE CONFIRMED. GIVEN THE EXTENT OF THE BRAIN INFARCTION, A POOR NEUROLOGICAL AND LIFE PROGNOSIS WAS ANTICIPATED. THE PATIENT DIED TWO DAYS AFTER SURGERY DUE TO HYPOTENSION AND CARDIAC ARREST. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | CONTINUATION OF D10: PRODUCT ID REACT-71 (UNKNOWN); PRODUCT TYPE: PRODUCT ID (UNKNOWN); PRODUCT TYPE: ; G2: CITATION: AUTHORS: TANIMURA, M., IKEDA, H., FUJIWARA, T., UEZATO, M., OSUKI, T., KINOSADA, M., KUROSAKI, Y., CHIN, M. BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE. SURGICAL NEUROLOGY INTERNATIONAL 15:12 2024. DOI:10.25259/SNI_948_2023 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REACT CATHETER | K182097 | LITERATURE WAS REVIEWED REGARDING: "BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE." THE OBJECTIVE IS TO PRESENT A CASE OF MECHANICAL THROMBECTOMY AND SIMULTANEOUS DISTAL VERTEBRAL ARTERY EMBOLIZATION FOR BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY. THE FOLLOWING MEDTRONIC DEVICES WERE USED: REACT 71 CATHETER, SOLITAIRE X, AND PHENOM 17 CATHETER. AMONG PATIENT ADVERSE EVENTS INCLUDED: THE STUDY DESCRIBES A CASE INVOLVING A PATIENT WHO EXPERIENCED A FALL AND SUBSEQUENT MEDICAL INTERVENTIONS. THE PATIENT DEVELOPED ALTERED CONSCIOUSNESS SIX HOURS AFTER THE CERVICAL SPINE TRAUMA AND HAD A RELATIVELY EARLY ONSET OF BASILAR ARTERY OCCLUSION CAUSED BY VERTEBRAL ARTERY INJURY. MECHANICAL THROMBECTOMY WAS PERFORMED USING THE REACT AND SOLITAIRE DEVICE, AND EMBOLIZATION OF THE INJURED VERTEBRAL ARTERY AT THE SOURCE OF THE EMBOLISM USING COILS DELIVERED THROUGH THE PHENOM CATHETER. COMPLETE RECANALIZATION WAS ACHIEVED DURING THE THROMBECTOMY. NO IMPROVEMENT IN NEUROLOGICAL SYMPTOMS WAS OBSERVED AFTER THE SURGERY. DIFFUSION-WEIGHTED IMAGING PERFORMED 13 H AFTER THE ONSET OF ALTERED CONSCIOUSNESS SHOWED INFARCTION IN THE MIDBRAIN, PONS, BOTH OCCIPITAL LOBES, BOTH CEREBELLAR HEMISPHERES, AND THE LEFT THALAMUS. PATENCY OF THE RIGHT VERTEBRAL AND BASILAR ARTERIES AND OCCLUSION OF THE LEFT VERTEBRAL ARTERY WERE CONFIRMED. GIVEN THE EXTENT OF THE BRAIN INFARCTION, A POOR NEUROLOGICAL AND LIFE PROGNOSIS WAS ANTICIPATED. THE PATIENT DIED TWO DAYS AFTER SURGERY DUE TO HYPOTENSION AND CARDIAC ARREST. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS. | CONTINUATION OF D10: PRODUCT ID SFR4-4-40-10 (UNKNOWN); IMPLANT DATE N/A; EXPLANT DATE N/A PRODUCT ID (UNKNOWN); G2: CITATION: AUTHORS: TANIMURA, M., IKEDA, H., FUJIWARA, T., UEZATO, M., OSUKI, T., KINOSADA, M., KUROSAKI, Y., & CHIN, M. BASILAR ARTERY OCCLUSION DUE TO VERTEBRAL ARTERY INJURY TREATED WITH THROMBECTOMY AND DISTAL VERTEBRAL ARTERY EMBOLIZATION THROUGH THE UNAFFECTED SIDE. SURGICAL NEUROLOGY INTERNATIONAL 15:12 2024. DOI:10.25259/SNI_948_2023 EARLIEST DATE OF PUBLICATION USED FOR DATE OF EVENT NO UNIQUE DEVICE IDENTIFIER (SERIAL/LOT) NUMBERS WERE PROVIDED; WITHOUT THIS INFORMATION IT COULD NOT BE DETERMINED WHETHER THESE OBSERVATIONS HAVE BEEN PREVIOUSLY REPORTED. WITHOUT RETURN OF THE PRODUCT NO DEFINITIVE CONCLUSION CAN BE MADE REGARDING THE CLINICAL OBSERVATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
QC HEMOSTATIC DRESSING, 2X2, 10X | K123387 | IT WAS REPORTED THAT: "THE PATIENT UNDERWENT SURGERY IN (B)(6) 2024 TO REMOVE A HEMOSTATIC DRESSING THAT HAD BEEN LEFT IN THE BODY DURING ANOTHER SURGERY ON (B)(6) 2022. THE HEMOSTATIC DRESSING HAD CAUSED AN ABSCESS TO FORM AND LED TO A SEVERE INFECTION, AND UNFORTUNATELY, THE PATIENT DIED." ADDITIONAL INFORMATION HAS BEEN REQUESTED AND A FOLLOW UP REPORT WILL BE SUBMITTED IF IT IS RECEIVED. | (B)(4) NO ITEM RETURNED FOR EVALUATION. IFU WARNINGS SECTION CLEARLY STATES:1) PRODUCT IS NOT ABSORBABLE AND MUST BE REMOVED FROM WOUND PRIOR TO WOUND CLOSURE.2) THE QUIKCLOT HEMOSTATIC DRESSING MUST NOT REMAIN IN THE WOUND FOR LONGER THAN 24 HOURS. THIS COMPLAINT IS THE RESULT OF USER ERROR, NOT ADHEARING TO IFU INSTRUCTIONS AND WARNINGS.
FUSION HOLLOW FIBER OXYGENATOR | K230640 | MEDTRONIC RECEIVED INFORMATION THAT DURING USE OF A FUSION OXYGENATOR, IT WAS REPORTED THAT THE PARTIAL PRESSURE OF OXYGEN WAS RECORDED AT 8 WITH 100% FRACTION OF INSPIRED OXYGEN (FI02) BEING DELIVERED THROUGH THE OXYGENATOR. THE CUSTOMER ALSO OBSERVED CLOTS IN THE OXYGENATOR. SEE ATTACHED PHOTOS. THE OXYGENATOR SWAPPED OUT WAS REPLACED WITH ANOTHER OXYGENATOR WHICH WORKED SUFFICIENTLY FOR THE REMAINING BYPASS TIME. THE PATIENT PASSED AWAY. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE PATIENT WAS AT HIGH RISK WITH AN AORTIC DISSECTION. THE PATIENT HAD BEEN IN THE HOSPITAL FOR A FEW DAYS. THE PROCEDURE WAS NOT AN EMERGENCY/CRASH BYPASS. THE PATIENT DID RECEIVE TRASYLOL PRIOR TO THE CASE (9 ANTITHROMBOTIC). THE CUSTOMER MANAGED HIGH AND ACCORDINGLY ACTIVATED CLOTTING TIMES (ACT). THEIR HEMOCHRON ACT DEVICE WAS SUFFICIENT FOR TRASYLOL USE. THE OXYGENATOR REPLACEMENT TOOK 3-4 MINUTES BUT THE CHANGE OUT WAS REQUIRED AT A TIME WHEN THE PATIENT WAS STRUGGLING ON THE TABLE AND VENTILATION HAD TO COMMENCE DURING THE OXYGENATOR SWAP OUT PERIOD. HYPOTENSION WAS EXPERIENCED. THE FIRST BYPASS PERIOD WAS FOR APPROXIMATELY 173 MINUTES. IT WAS UNEVENTFUL FROM AN OXYGENATOR PERFORMANCE PERSPECTIVE. THE PATIENT'S BLOOD GASES, LACTATES, PARTIAL PRESSURE OF OXYGEN (PO2), PARTIAL PRESSURE OF CARBON DIOXIDE (PCO2) AND ACTS WERE ALL SUFFICIENT. AT 15.39PM, PO2 WAS 27MMHG AND ACT WAS 999S. THE FIRST BYPASS ENDED AT 15.44PM. PROTAMINE WAS ADMINISTERED TO THE PATIENT, HOWEVER, ALL SUCTION DEVICES WERE REMOVED FROM THE PATIENT. THERE WAS NO SUCTION BLOOD. THE ARTERIAL AND VENOUS CANNULA REMAINED IN SITU. THE PATIENT WAS TRAN SFUSION WITH VARIOUS DOSES OF 100ML FROM THE PERFUSION CIRCUIT DURING THIS OFF BYPASS STAGE, AVOIDING STASIS IN THE CIRCUIT. THESE 100ML TOP UPS OF PATIENT VOLUME OCCURRED APPROXIMATELY EVERY 30 SECONDS. UPON INSTIGATING A SECOND BYPASS AT 16.12PM, THE CUSTOMER NOTED THAT THE OXYGENATOR PERFORMANCE STARTED TO DETERIORATE. PO2 HAD DROPPED AND INCREASING FIO2 DID NOT RESOLVE THE ISSUE. THE FIO2 VALUES WERE FROM 60 TO 70, TO 80, TO 90, TO 100% WITH PO2¿S FALLING FROM 22.4, 15, 7.9, 7.6MMHG. THE CUSTOMER REPORTED OBSERVING CLOTTING IN THE OXYGENATOR, ALONG WITH PO2 CHALLENGES AND MAX FIO2. THE OXYGENATOR WAS SWAPPED OUT FOR ANOTHER FUSION OXYGENATOR AT APPROXIMATELY AT 17.20PM. SHORTLY AFTER THE OXYGENATOR WAS CHANGED OUT, THE OPERATION WAS STOPPED AS THE PATIENT PASSED AWAY. THE CUSTOMER NOTED THAT THE NEW OXYGENATOR WAS WORKING SUFFICIENTLY. THE CUSTOMER DOES NOT UNDERSTAND WHY THE OXYGENATOR WOULD STRUGGLE ON THE SECOND BYPASS PERIOD AS COMING OFF AND ON BYPASS WAS A COMMON OCCURRENCE. THE CUSTOMER STATED THAT THE OXYGENATOR FAILED AT A T IME IN THE OPERATION WHERE THE PATIENT WAS STRUGGLING DUE TO OPERATION RELATED ISSUES. THE PATIENT'S DEATH MAY BE MULTIFACTORIAL. PRE AND POST OXYGENATOR PRESSURES WERE NOT MEASURED. THE ACT¿S WERE KEPT HIGH THROUGHOUT THE CASE AND THEY WERE ALL ABOVE PROTOCOL. THE CUSTOMER NOTICED CLOTS IN THE OXYGENATOR ON THEIR SECOND BYPASS RUN ONLY. A NORMAL ROLLER PUMP CONFIGURATION WAS USED. | ADDITIONAL INFORMATION B5: MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CUSTOMER STATED THAT MEDTRONIC ARE NOT PRIVILEGED TO THE FINAL INFORMATION REGARDING THE CAUSE OF DEATH, HOWEVER, THEY MENTIONED A RIGHT-SIDE HEART FAILURE BUT THERE WAS NO FORMAL NOTE. THE PRESSURE WAS MEASURED IN THE CIRCUIT POST OXYGENATOR. HEPARIN DOSING WAS MONITORED WITH AN ACT INSTRUMENT TO ACHIEVE OPTIMAL THERAPEUTIC RESPONSE. THE CARDIOTOMY FLOW RATES WERE STANDARD FOR THE SHUNT LINES (E.G. 500ML/MINUTE). THE CUSTOMER USED A FLOW PROBE TO MEASURE THE ACTUAL FLOW TO THE PATIENT. THE SUCKER BLOOD WAS NOT SEQUESTERED TO THE CELL SAVER BEFORE RETURNING TO CARDIOTOMY VENOUS RESERVOIR (CVR). THE LEVEL IN THE CVR WAS NOT UNUSUALLY LOW OR HIGH FOR CVR LEVELS. THERE WAS 500ML OF RINGER¿S SOLUTION, 1000ML OF HARTMANN'S SOLUTION AND 10,000 UNITS OF HEPARIN ADMINISTRATED FOR PRIMING. DURING CARDIOPULMONARY BYPASS (CPB), GELOFUSINE, RINGER'S SOLUTION, BLOOD, CELL SAVER BLOOD, APROTININ (TRASYLOL), MANNITOL (10%), METARAMINOL AND SODIUM BICARBONATE WERE ADMINISTRATED. POST OXYGENATOR PRESSURES WERE 193MMHG, 158MMHG AND 73MMHG. THE CARDIOPLEGIA TEMPERATURE WAS BETWEEN 6.1° TO 7°. THE PATIENT'S CORE TEMPERATURE WAS BETWEEN 36.5° TO 36.8°. THE ARTERIAL TEMPERATURE WAS BETWEEN 22.4° TO 32.7°. AT 17.32PM, IT WAS NOTED THAT THE PULMONARY ARTERY (PA) PRESSURE WAS TOO HIGH. THE PATIENT WAS GIVEN PROSTACYCLIN BEFORE COMING OFF BYPASS. THE PATIENT PASSED AWAY AT 17 .38PM. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT HAD PASSED AWAY DUE TO PULMONARY FIBROSIS. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED IN REFERENCE TO A CUSTOMER REQUESTING INFORMATION ON RETURNING A DEVICE DUE TO A PATIENT PASSING. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT WAS SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION IT WAS NOTED THERE WERE NO ERROR CODES. THE DEVICE WAS TESTED AND PASSED. THE DEVICE WAS RETURNED TO STOCK. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME.
DREAMSTATION AUTO CPAP | K131982 | THE SPOUSE OF THE PATIENT REPORTED THE PATIENT WAS DECEASED. THE REPORTER INQUIRED ABOUT THE RETURN AND/OR DISPOSAL OF THE DEVICE. THE CAUSE OF DEATH WAS NOT PROVIDED BY THE REPORTER. THE REPORTER WAS ADVISED IT WAS SAFE TO DISPOSE OF THE DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
INTELLIPAP AUTO BILEVEL W/HEATED HUMIDIFIER | K091919 | FROM DEVILBISS COMPLAINT RECEIVED BY EMAIL: "I WOULD LIKE TO FILE A COMPLAINT CONCERNING A MEDICAL DEVICE. ON (B)(6) 2025, MY MOTHER WAS PUT ON HOSPICE. SHE WAS DISCHARGED FROM THE HOSPITAL AND WENT HOME AS BIPAP DEPENDENT. MY MOTHER HAD COPD AND THOUGH SHE WAS PUT ON HOSPICE SHE WAS NO WHERE NEAR READY TO PASS. MY MOTHER ATE DINNER ON FRIDAY NIGHT, SHE GOT UP BY HERSELF, USED THE RESTROOM AND WAS FULLY ALLERT. SHE WENT TO SLEEP THAT NIGHT AT 1045PM AND WOKE UP SOMETIME BETWEEN 530 AND 6AM. AT THIS TIME MY MOTHER SAT UP IN HER BED AND WHEN DOING SO HER BIPAP TUBE CAME COMPLETELY OFF OF HER MACHINE. ONCE THIS HAPPENED SHE WAS UNABLE TO REATTACH IT AND I WAS TOO LATE GETTING TO HER. THE MASK WAS STILL FULLY ATTACHED TO HER FACE BUT THE END THAT ATTACHES TO THE BIPAP MACHINE CAME LOOSE. I HAD NO IDEA THAT COULD HAPPEN OR I WOULD HAVE STAYED WHERE I COULD WATCH HER EVERY SECOND OF THE DAY. THE MACHINE HAS NO EJECTION PORT, NO ALARM AND JUST SLIDES RIGHT OFF THE MACHINE. I WAS TOO LATE, BY THE TIME I GOT TO MY MOTHER SHE COULD NOT CATCH HER BREATH AND DIED AT 642AM. MY MOTHER DIED BECAUSE OF THIS MACHINES FAILURE MY MOTHER SUFFERED AND DIED A HORRIBLE DEATH. SOMETHING NEEDS TO BE DONE TO MAKE THESE MACHINES SAFE FOR PEOPLE THAT ARE DEPENDENT UPON THEM. THE SIMPLE SLIDING ON AND OFF IS NOT SAFE AND KILLS PEOPLE WHO ARE DEPENDENT UPON THESE MACHINES. THE MACHINE SHE WAS USING: DEVILBISS INTELLIPAP." SEE DEVILBISS COMPLAINT# (B)(4). MW5168016. | REACT HEALTH HAD PREVIOUSLY SUBMITTED A REPORT STATING THAT A PATIENT HAD DIED WHILE USING A DEVILBISS INTELLIPAP BIPAP. ADDITIONAL INFORMATION WAS RECEIVED ON 4/15/2025 FROM THE HOSPICE PROVIDER ON A PATIENT COMPLAINT FORM. THE HOSPICE PROVIDER ALSO FORWARDED A SERVICE REPORT FROM A THIRD PARTY SERVICE PROVIDER. PER THEIR PAPERWORK, IT WAS NOTED THE STREETS DME LLC MEDICAL DELIVERY TECHNICIAN WHO HAD DELIVERED THE BIPAP ON 2/20/2025 WITH ALL SUPPLIES (MASK, PRESSURE LINE ADAPTOR AND GREY CORRUGATED TUBING) NOTICED THE PATIENT HAD THEIR OWN 5 LITER AND 10 LITER CONCENTRATOR AND ALSO A PHILIPS RESPIRONICS TRILOGY WITH BLUE TUBING. ON 2/24/2025 THE DELIVERY TECHNICIAN PICKED UP ALL EQUIPMENT DUE TO PATIENT PASSING AND NOTICED THEIR BLUE TUBING FROM THE PHILIPS RESPIRONICS TRILOGY WAS ON THE DEVILBISS BIPAP, # (B)(6). THE BIPAP, (B)(6) WAS TESTED BY THE EQUIPMENT PROVIDER ON THE SAME DAY OF THE COMPLAINT WITH NO ISSUES FOUND. THE AIR PRESSURE WAS NORMAL. THE DEVICE WAS THEN SENT TO MEDICAL TECTRONICS, INC. IN HOUSTON, TEXAS FOR A PERFORMANCE EVALUATION. THE TECHNICIAN RAN FULL DIAGNOSTICS ON UNIT UPON INTAKE AND INDICATED NO ISSUES. THE UNIT WAS STATED AS OPERATING ACCORDING TO MANUFACTURER'S SPECIFICATIONS. MULTIPLE ATTEMPTS HAVE BEEN MADE TO HAVE THE DEVICE RETURNED TO REACT HEALTH FOR EVALUATION AND THE DEVICE HAS NOT BEEN RETURNED.
PROGRIP | K140941 | THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG W ITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS. THE PATIENT WAS REPORTED TO HAVE DIED BY SUICIDE IN (B)(6) 2025. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PROGRIP | K140941 | THE PATIENT REPORTED THAT ON (B)(6) 2024 AN OPEN BILATERAL INGUINAL HERNIA CURE WAS PERFORMED USING THE LICHTENSTEIN TECHNIQUE WITH THE PLACEMENT OF TWO RETRO-APONEUROTIC PROSTHESES. THE PROCEDURE INVOLVED A DORSOL DECUBITUS POSITION UNDER GENERAL ANESTHESIA, INCISION IN THE RIGHT LOWER ABDOMINAL FOLD, OPENING OF THE EXTERNAL OBLIQUE FASCIA, DISSECTION OF A LARGE EXTERNAL OBLIQUE HERNIA, SPERM CORD RELEASE, DETACHMENT OF APONEUROTIC MARGINS, AND THE SPREAD OF A MESH PROSTHESIS ATTACHED TO THE PUBIS, CRURAL ARCH, AND CONJOINED TENDON BY SEPARATE POINTS OF A SUTURE. THE WOUND WAS CLOSED WITH APONEUROTIC CLOSURE, CLOSURE OF THE FASCIA CRIBRIFORMIS, NAROPENINE WALL INFILTRATION, AND NON-ABSORBABLE THREADS ON THE SKIN. POST OPERATIVELY, THE PATIENT EXPERIENCED SEVERE PAIN, INTERNAL BURNS ON THE RIGHT AND LEFT, NEEDLE STROKE SENSATION, AND UNPLEASANT DISCOMFORT IN THE TESTICLES WITH PAIN (SYMPTOMATIC BILATERAL GROIN HERNIA). THE PATIENT ALSO REPORTED FEELING AS IF HAVING BRUISES AND AN IMPRESSION OF HAVING CORRUGATED CARDBOARD INSIDE. THESE SYMPTOMS WERE EXACERBATED WHEN STANDING AND SITTING, LEADING TO WAKING UP AT NIGHT DUE TO NEEDLE STROKE AND BURNS, DIFFICULTY GOING BACK TO SLEEP, AND SOMETIMES SLEEPLESS NIGHTS. THE PATIENT FACED GREAT DIFFICULTY IN SQUATTING, STANDING UP, AND LIFTING LOADS, ALONG W ITH LOSS OF TONE AND INTENSE FATIGUE. PSYCHOLOGICAL CONSEQUENCES INCLUDED ALTERATION OF MOOD, ANXIETY, STRESS, DEPRESSION, AND MORBID IDEAS. THE PATIENT WAS REPORTED TO HAVE DIED BY SUICIDE IN (B)(6) 2025. | CORRECTION: B5. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT HAD PASSED AWAY. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
I-NEB CF UK UNIT | K102454 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A I-NEB CF UK UNIT DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE CORRECT B5 SHOULD BE: THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A I-NEB CF UK UNIT DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS BEEN SCRAPPED; THEREFORE, IT CANNOT BE RETRIEVED FOR EVALUATION. IN ADDITION, APPROPRIATE CODE UPDATED/CORRECTED IN THIS FOLLOW-UP REPORT.
LLD EZ LEAD LOCKING DEVICE | K142116 | A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD, RIGHT ATRIAL (RA) LEAD, AND LEFT VENTRICULAR (LV) LEAD DUE TO BACTEREMIA. SPECTRANETICS LLD EZ LEAD LOCKING DEVICES (LLD EZS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 14F GLIDELIGHT LASER SHEATH ON EACH LEAD, ATTEMPTS WERE MADE TO GET THROUGH A BINDING SITE IN THE INNOMINATE. NEXT, SWITCHING TO A SPECTRANETICS 11F TIGHTRAIL SUB-C ROTATING DILATOR SHEATH AND SPECTRANETICS 11F TIGHTRAIL ROTATING DILATOR SHEATH (LONG), ADVANCEMENT WAS MADE INTO THE SUPERIOR VENA CAVA (SVC), ENSURING THE LEADS WERE FREE AND NOT BOUND TO ONE ANOTHER. WITH THE 11F TIGHTRAIL (LONG), THE RV LEAD WAS REMOVED, FOLLOWED BY THE RA LEAD. THEN, WITH THE 11F TIGHTRAIL (LONG) ON THE LV LEAD, PROGRESS WAS MADE DOWN TO THE RA. WHILE APPLYING TRACTION, THE TIGHTRAIL DEVICE WAS ACTUATED, BUT THE LV LEAD WOULD NOT RELEASE, AND THE PATIENT¿S BLOOD PRESSURE BEGAN TO GRADUALLY DROP. A PERFORATION OF THE CORONARY SINUS (CS) WAS CONFIRMED VIA ANGIOGRAM, AND RESCUE EFFORTS BEGAN, INCLUDING CPR, BYPASS, AND STERNOTOMY. THE CS PERFORATION WAS REPAIRED; HOWEVER, THE PATIENT CONTINUED TO HAVE SIGNIFICANT BLOOD LOSS, AND THE LOCATION OF THE BLEEDING WAS NOT IDENTIFIED. FURTHER INTERVENTION WAS CEASED, AND THE PATIENT DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE LLD EZ PROVIDING TRACTION WITHIN THE LV LEAD WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. | A6) PATIENT RACE - NOT AVAILABLE FROM FACILITY. H3) THE LLD EZ WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) PERFORATION OF VESSELS AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE LLD DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
TRILOGY DEVICE | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY AND THE DEATH IS NOT BELIEVED TO BE DUE TO A PRODUCT ISSUE. THE CAUSE OF DEATH IS BELIEVED TO BE ASPHYXIATION, AND THE AUTOPSY RESULTS ARE PENDING. THERE IS A CRIMINAL INVESTIGATION AGAINST THE CAREGIVER DUE TO THE PATIENT'S DEATH AND THE DEVICE DATA HAS BEEN REQUESTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS PASSED AWAY AND THE DEATH IS NOT BELIEVED TO BE DUE TO A PRODUCT ISSUE. THE CAUSE OF DEATH IS BELIEVED TO BE ASPHYXIATION, AND THE AUTOPSY RESULTS ARE PENDING. THERE IS A CRIMINAL INVESTIGATION AGAINST THE CAREGIVER DUE TO THE PATIENT'S DEATH AND THE DEVICE DATA HAS BEEN REQUESTED. ADDITIONAL INFORMATION RECEIVED REPORTS THAT THE CAUSE OF DEATH IS STILL UNKNOWN AND NOT AVAILABLE TO THE HOME SERVICE TO PROVIDE. THE PATIENT HAD RESPIRATOR DEPENDENCE, UNSPECIFIED LEVEL OF CERVICAL SPINAL CORD INJURY, TRACHEOSTOMY, CHRONIC RESPIRATORY FAILURE, HYPOXIA OR HYPERCAPNIA AND GASTROSTOMY. THE PATIENT HAD A HISTORY OF CHRONIC RESPIRATORY FAILURE SECONDARY TO BEING IN AN AUTO PEDESTRIAN MOTOR VEHICLE ACCIDENT AND SUSTAINED A C 1-2 SPINAL CORD INJURY. THE DEVICE IS CURRENTLY UNAVAILABLE FOR RETURN AS IT IS CURRENTLY IN POSSESSION OF LAW ENFORCEMENT UNTIL INVESTIGATION IS COMPLETED.
DREAMSTATION BIPAP AUTOSV | K090539 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION BIPAP SV HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP W/HUMID CELL/BT | K200480 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION 2 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | UDI- NO SN NUMBER.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION DEVICE HAS PASSED DUE TO AN UNKNOWN HEART CONDITION. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP, JP | K131982 | THE MANUFACTURER WAS NOTIFIED ABOUT A PATIENT WHO DIED. THE CLAIM STATES THAT THE PATIENT CHOKED ON VOMIT WHILE USING AN AUTO CPAP A-FLEX DEVICE. IT IS ALLEGED THAT ASPIRATION PNEUMONIA CAUSED THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED.
POWERPORT IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K060812 | ON (B)(6) 2016, PATIENT UNDERWENT BARD TI POWERPORT IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DIAGNOSED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. MANUFACTURING REVIEW: THE DEVICE HISTORY RECORDS HAVE BEEN REVIEWED AND THIS LOT MET ALL RELEASE CRITERIA. INVESTIGATION SUMMARY: THE PHYSICAL DEVICE WAS NOT RETURNED FOR EVALUATION. NO PHOTOS, VIDEOS, OR MEDICAL RECORDS WERE PROVIDED FOR REVIEW. THEREFORE, THE INVESTIGATION IS INCONCLUSIVE FOR THE REPORTED INFECTION AND PATIENT DEATH AS NO OBJECTIVE EVIDENCE WAS PROVIDED FOR REVIEW. A DEFINITIVE ROOT CAUSE FOR THE REPORTED INFECTION AND PATIENT DEATH COULD NOT BE DETERMINED BASED UPON THE PROVIDED INFORMATION. LABELING REVIEW: AS THE REPORTED EVENT DID NOT ALLEGE A LABELING OR USE RELATED ISSUE A LABELING REVIEW IS NOT REQUIRED. D4 (UNIQUE IDENTIFIER (UDI) #), G3, H4, H6 (COMPONENT, METHOD). SECTION A THROUGH F: THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K063377 | ON (B)(6) 2019, PATIENT UNDERWENT BARD POWERPORT CLEARVUE ISP IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DIAGNOSED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K122899 | ON (B)(6) 2018, PATIENT UNDERWENT BARD POWERPORT CLEARVUE ISP IMPLANTABLE PORT PLACEMENT PROCEDURE FOR UNKNOWN MEDICAL CONDITION. ABOUT SOMETIME AFTER PORT PLACEMENT PROCEDURE, PATIENT WAS DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, PATIENT EXPIRED DUE TO SEPTIC SHOCK. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
PENTAX MEDICAL | K180292 | ON 02-APR-2025, PENTAX MEDICAL BECAME AWARE OF A PATIENT DEATH INVOLVING A PENTAX MEDICAL VIDEO GASTROSCOPE, MODEL EG34-I10, SERIAL NUMBER (B)(6). ACCORDING TO THE INITIAL REPORT, THE PHYSICIAN PERFORMED A THERAPEUTIC ENDOSCOPY TO IDENTIFY A GASTROINTESTINAL (GI) BLEEDING SOURCE. THE BLEEDING SOURCE WAS FOUND, HEMOSTASIS WAS ACHIEVED USING ELECTROCOAGULATION, AND CLIPS WERE DEPLOYED TO PREVENT FURTHER BLEEDING. DURING THE PROCEDURE, THE PATIENT'S PHYSICAL CONDITION DETERIORATED. THE ENDOSCOPE WAS REMOVED, CARDIOPULMONARY RESUSCITATION (CPR) WAS PERFORMED, AND A CODE BLUE TEAM WAS ACTIVATED. THE PATIENT WAS PRONOUNCED DEAD ON THE SAME DAY, (B)(6) 2025. THE PENTAX MEDICAL VIDEO GASTROSCOPE REPORTEDLY FUNCTIONED NORMALLY DURING THE PROCEDURE. THERE WAS NO INDICATION OF DEVICE MALFUNCTION, AND THE DEATH WAS NOT ATTRIBUTED TO THE ENDOSCOPE. THE PROCEDURE WAS SCHEDULED IN ADVANCE AND WAS NOT AN EMERGENCY CASE. THE PATIENT HAD A HISTORY OF HEART SURGERY; HOWEVER, DETAILED MEDICAL HISTORY WAS NOT AVAILABLE. THIS EVENT MEETS THE REQUIREMENTS FOR FDA REPORTABILITY; HOWEVER, SUBMISSION OF THIS REPORT DOES NOT CONSTITUTE AN ADMISSION THAT MEDICAL PERSONNEL, USER FACILITY, IMPORTER, MANUFACTURER OR PRODUCT CAUSED OR CONTRIBUTED TO THE EVENT. | H6 CONTINUED: HEALTH EFFECT - CLINICAL CODE: 1762 CARDIAC ARREST, 4461 RESPIRATORY ARREST. HEALTH EFFECT - IMPACT CODE: 1802 DEATH, 4635 UNEXPECTED DETERIORATION. MEDICAL DEVICE PROBLEM CODE: 2993 ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM COMPONENT CODE: 4755. TYPE OF INVESTIGATION: 4110 HISTORICAL DATA ANALYSIS, 3331 ANALYSIS OF PRODUCTION RECORDS, 4111 COMMUNICATION/INTERVIEWS. INVESTIGATION FINDINGS: 213 NO DEVICE PROBLEM FOUND. INVESTIGATION CONCLUSIONS: 4310 CAUSES TRACED TO NON-DEVICE RELATED FACTORS. CORRECTION INFORMATION B4: DATE OF THIS REPORT G6: FOLLOW-UP #: 1 H2: IF FOLLOW-UP, WHAT TYPE? H3: DEVICE EVALUATED BY MANUFACTURE H6: ADVERSE EVENT PROBLEM ADDITIONAL INFORMATION D4: PRIMARY UNIQUE DEVICE IDENTIFIER (UDI) H4: DEVICE MANUFACTURE DATE H11: EVALUATION SUMMARY EVALUATION SUMMARY PENTAX MEDICAL AMERICA AND THE JAPANESE INVESTIGATION TEAM CONDUCTED A GOOD FAITH EFFORT (GFE) TO OBTAIN ADDITIONAL INFORMATION REGARDING THIS COMPLAINT. THE FOLLOWING TWO QUESTIONS WERE SUBMITTED TO THE SITE REPRESENTATIVE MULTIPLE TIMES BETWEEN 11-APR-2025 AND 15-MAY-2025: - HAS THE SUBJECT ENDOSCOPE (EG34-I10_(B)(6)) CONTINUED TO BE USED AT THE FACILITY AFTER THE COMPLAINT? - WHO DETERMINED THAT THERE WAS NOTHING WRONG WITH THE SCOPE (A PHYSICIAN OR A FACILITY STAFF MEMBER)? AFTER A DELAY, A RESPONSE WAS PROVIDED BY THE SITE REPRESENTATIVE ON 15-MAY-2025. THE RESPONSE INDICATED THAT THE ENDOSCOPE IS STILL IN USE AT THE FACILITY AND NO COMPLAINTS OR ISSUES HAVE BEEN REPORTED. IT WAS CONFIRMED BY A FACILITY STAFF MEMBER (ENDOSCOPY SUPERVISOR) THAT THE DEVICE FUNCTIONED NORMALLY DURING THE PROCEDURE. AFTER THE PROCEDURE WAS ABORTED DUE TO THE PATIENT'S CONDITION, THE ENDOSCOPE WAS LEAK TESTED AND REPROCESSED, AND NO ABNORMALITIES WERE FOUND. UPON RECEIVING THIS INFORMATION, THE JAPANESE INVESTIGATION REPRESENTATIVE CONFIRMED ON 16-MAY-2025 THAT THERE WERE NO PROBLEMS WITH THE ENDOSCOPE AND THAT IT COULD CONTINUE TO BE USED. SINCE THE DEVICE HAD BEEN FUNCTIONING NORMALLY AND HAD UNDERGONE STANDARD REPROCESSING WITHOUT ANY ISSUES, IT WAS DETERMINED THAT FURTHER INSPECTION OF THE DEVICE WAS UNNECESSARY. THE EVENT THAT OCCURRED WAS A DETERIORATION IN THE PATIENT'S CONDITION. BASED ON THE INVESTIGATION AND GFE, THE GASTROINTESTINAL BLEEDING HAD OCCURRED BEFORE THE ENDOSCOPE WAS INSERTED, AND IT WAS DETERMINED THAT IT WAS NOT DUE TO EQUIPMENT FAILURE. FURTHERMORE, BASED ON THE TREND ANALYSIS, THERE IS NO SIGNIFICANT INCREASE OR UPWARD TREND IN REPAIR COMPLAINTS RELATED TO THIS FAILURE MODE/HAZARD. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED BY THE MANUFACTURER. THE REVIEW BY DHR CONFIRMED THAT THE INSTRUMENT WAS MANUFACTURED UNDER NORMAL CONDITIONS ON 05-JUL-2024, PASSED ALL REQUIRED INSPECTIONS, AND WAS RELEASED ACCORDINGLY. ALSO, THERE WERE NO REWORKS OR CONCESSIONS, AND THE DATES OF APPROVAL FOR SHIPMENT AND ACTUAL DATE SHIPPED WERE CONFIRMED ON 16-JUL-2024. PENTAX MEDICAL HAS NOT RECEIVED ANY FURTHER INFORMATION FOR THIS EVENT AND THEREFORE, CONSIDERS THIS MEDWATCH REPORT CLOSED. ECTIONS, AND WAS RELEASED ACCORDINGLY. ALSO, THERE WERE NO REWORKS OR CONCESSIONS, AND THE DATES OF APPROVAL FOR SHIPMENT AND ACTUAL DATE SHIPPED WERE CONFIRMED ON 16-JUL-2024. PENTAX MEDICAL HAS NOT RECEIVED ANY FURTHER INFORMATION FOR THIS EVENT AND THEREFORE, CONSIDERS THIS MEDWATCH REPORT CLOSED.
MICRA | K132030 | IT WAS REPORTED THAT THE PATIENT IS DECEASED. DURING IMPLANT THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) EXHIBITED NON OPTIMAL TH RESHOLDS WHICH WERE ADJUDGED AS A MYOCARDIAL PROBLEM AND PLACEMENT COULD NOT BE PERFORMED. THE LEADLESS IPG WAS COLLECTED, NOT IMPLANTED AND IMPLANTATION WAS DELAYED AND PLANNED FOR A LATER DATE. DEATH WAS CONFIRMED IN THE STATE OF RESPIRATORY ARREST BEFORE THE PLANNED IMPLANT. AIRWAY OBSTRUCTION DUE TO AGING IS CONSIDERED A MAJOR FACTOR. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
MX40 1.4 GHZ SMART HOPPING | K113125 | THIS REPORT IS BASED ON INFORMATION PROVIDED BY PHILIPS FIELD SERVICE PERSONNEL AND HAS BEEN INVESTIGATED BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON AN MX40 1.4 GHZ SMART HOPPING INDICATING THAT AN ASYSTOLE ALARM WAS NOT GENERATED AT THE CENTRAL STATION FOR BED 284 AND THE PATIENT PASSED AWAY. THE DEVICE WAS IN CLINICAL USE AT THE TIME THE ISSUE WAS DISCOVERED. | A PHILIPS TECHNICAL CONSULTANT (TC) ONSITE INVESTIGATED THE REPORTED ISSUE. IT WAS NOTED THAT THE NURSE FOUND THE PATIENT IN DISTRESS, BECAUSE THE PATIENT HAD BLUISH COLOR TO THE SKIN. THE CUSTOMER PULLED UP A STRIP DISPLAY ON THE CENTRAL STATION THAT SHOWED THE INCIDENT IN QUESTION OCCURRED AT 10:29:40AM ON (B)(6) 2025 WITH A RED ASYSTOLE BANNER SHOWING ABOVE THE STRIP DISPLAY. THE TC REQUESTED MX 40 BE RELEASED TO THEM FOR SEQUESTRATION. THE TC WAS INFORMED AT THAT TIME THAT THE PATIENT HAD EXPIRED AND THAT THE MX40 WAS ON ANOTHER PATIENT. THE CUSTOMER SWAPPED OUT THE MX40 AND GAVE TO THE TC. THE TC WENT TO CENTRAL AND VERIFIED WITH STAFF THAT AUDIBLE ASYSTOLE ALARMS WERE AVAILABLE UTILIZING A SEPARATE DEPARTMENT MX40 DEVICE. THE TC CONTACTED ANOTHER TC WHO WAS ONSITE LATER IN THE DAY ON (B)(6) 2025. THIS TC WAS ABLE TO REMOTELY LOG INTO THE CENTRAL STATION AND PULL LOGS FOR THIS INCIDENT. THE LOGS SHOW THAT THE DEVICE HAD ALARMED FOR AN ASYSTOLE AND HAD BEEN ACKNOWLEDGED. (B)(6) 2025 10:30:19 (B)(6) 2-284 ASYSTOLE ENDED. (B)(6) 2025,10:30:19, (B)(6), 2-284, ASYSTOLE ENDED. (B)(6) 2025,10:30:19, (B)(6), 2-284, ASYSTOLE ENDED. (B)(6) 2025,10:30:01, (B)(6), 2-284, ACKNOWLEDGE. (B)(6) 2025,10:30:01, (B)(6), 2-284, ACKNOWLEDGE. (B)(6) 2025,10:29:57, (B)(6), 2-284, ASYSTOLE GENERATED AT 10:29:46. (B)(6) 2025,10:29:57, (B)(6), 2-284, ASYSTOLE GENERATED AT 10:29:46. (B)(6) 2025,10:29:57, (B)(6), 2-284, ASYSTOLE GENERATED AT 10:29:46. THE TC USED A SIMULATOR TO SHOW A HEART RATE OF 70/BPM AND MONITORED THE CENTRAL STATION. THE TC PLACED THE SIMULATOR INTO AN ASYSTOLE; THE HEART RATE TRANSMITTED PROPERLY, AND THE CENTRAL ALARMED FOR ASYSTOLE. THE TC THEN RETURNED THE SIMULATOR TO 70/BPM; THE CENTRAL STATION MONITORED PROPERLY, AND THE ASYSTOLE CLEARED. THIS TEST WAS REPEATED AND ALSO PASSED. THE TC WAS UNABLE TO DUPLICATE THE ORIGINAL COMPLAINT OF THE DEVICE NOT ALARMING. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
SPYSCOPE DS II | K183636 | IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT SPYSCOPE DS II ACCESS AND DELIVERY CATHETER WAS USED IN THE DUODENUM FOR THE TREATMENT OF STONES ON (B)(6) 2025. DURING THE SPYGLASS PROCEDURE, WHILE PERFORMING IRRIGATION THE ANESTHETIST WARNED OF A RAPID DETERIORATION IN THE PATIENTS VITALS. THE DEVICE WAS SAFELY REMOVED, AND DOCTORS ATTEMPTED CPR. HOWEVER, THE PATIENT PASSED AWAY SHORTLY AFTER. AIR WAS USED FOR INSUFFLATION DURING THE PROCEDURE. IN THE PHYSICIANS ASSESSMENT THE DEVICE DID NOT CONTRIBUTE TO THE PATIENTS COMPLICATIONS. THERE WERE NO REPORTED PATIENT COMPLICATIONS AS A RESULT OF THIS EVENT. | BLOCK H6: IMDRF PATIENT CODE E0602 CAPTURES THE REPORTABLE EVENT OF CARDIAC ARREST. IMDRF IMPACT CODE F2306 CAPTURES THE REPORTABLE EVENT OF CARDIOPULMONARY RESUSCITATION. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT'S DEATH.
HEMOSPRAY ENDOSCOPIC HEMOSTAT | K200972 | DURING AN ENDOSCOPIC PROCEDURE TO TREAT A BLEEDING ULCER IN THE STOMACH, THE PHYSICIAN USED A COOK HEMOSPRAY ENDOSCOPIC HEMOSTAT. IT WAS REPORTED THAT THE PATIENT HAD BEEN REPEATEDLY BLEEDING FROM A DUODENAL ULCER FOR THREE DAYS, AND PRIOR TO THIS TREATMENT, HE HAD BEEN TREATED WITH CLIPS TO STOP THE BLEEDING TWICE. DUE TO ANEMIA, LOW BLOOD PRESSURE, AND TARRY STOOLS, THE PHYSICIAN UNDERWENT ANOTHER ENDOSCOPY. THERE WAS A LARGE AMOUNT OF BLOOD CLOTS AND BLEEDING FROM THE DUODENAL ULCER. THEY USED TWO CLIPS AND SPRAYED WITH HEMOSPRAY. AT THAT POINT, THE BLEEDING HAD STOPPED, BUT ABOUT 9 HOURS LATER, BLOOD PRESSURE DROPPED DUE TO RE-BLEEDING, AND AN ENDOSCOPY WAS PERFORMED. DUE TO GUSHING BLEEDING, IT WAS NOT POSSIBLE TO STOP THE BLEEDING, AND THE SCOPE WAS REMOVED BECAUSE THE PATIENT'S VITALS COLLAPSED, AND THE PATIENT DIED WITHOUT BEING ABLE TO UNDERGO SURGERY OR AG TREATMENT. PER THE PHYSICIAN, "I BELIEVE THAT THE CAUSE OF REBLEEDING WAS THE EXCESSIVE STIMULATION OF EXPOSED BLOOD VESSELS BY APPLYING CLIPS, AND THE FACT THAT THE HEMOSTATIC AGENT WAS NOT DISTRIBUTED PROPERLY DUE TO THE PRESENCE OF CLIPS AT THE BLEEDING SITE." | ADDITIONAL INFORMATION: SECTION E PHONE: (B)(6). INVESTIGATION EVALUATION: A PRODUCT EVALUATION WAS NOT PERFORMED IN RESPONSE TO THIS REPORT BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT PROVIDED TO COOK FOR EVALUATION. THE REPORT COULD NOT BE CONFIRMED. A REVIEW OF THE DEVICE HISTORY RECORD COULD NOT BE CONDUCTED BECAUSE THE LOT NUMBER WAS NOT PROVIDED. INVESTIGATION CONCLUSION: WE COULD NOT CONDUCT A COMPLETE INVESTIGATION BECAUSE THE PRODUCT SAID TO BE INVOLVED WAS NOT RETURNED FOR EVALUATION. A DEFINITIVE CAUSE FOR THE REPORTED OBSERVATION COULD NOT BE DETERMINED. HOWEVER, A DEVICE FAILURE WAS NOT REPORTED. BASED ON THE RECEIVED INFORMATION, THE PATIENTS SUFFERED A REBLEED AFTER SUCCESSFUL HEMOSTASIS. THE IFU STATES, "ALL ENDOSCOPIC HEMOSTATIC THERAPIES, INCLUDING HEMOSPRAY, HAVE AN ASSOCIATED RISK OF REBLEEDING, PARTICULARLY IN SITUATIONS WHERE THE CAUSE OF BLEEDING IS AN UNRESOLVED UNDERLYING DISEASE. AFTER HEMOSTASIS HAS BEEN ACHIEVED, MONITOR PATIENTS FOR REBLEEDING PER THE RELEVANT CLINICAL GUIDELINES." THE IFU PRECAUTIONS, "LIKE OTHER MODALITIES, HEMOSPRAY MAY NOT BE EFFECTIVE FOR ALL TYPES OF BLEEDS. GASTROINTESTINAL BLEEDING MAY EXACERBATE EXISTING COMORBIDITIES, INCREASING THE POTENTIAL FOR ADVERSE EVENTS INCLUDING PATIENT MORTALITY." PRIOR TO DISTRIBUTION, ALL HEMOSPRAY ENDOSCOPIC HEMOSTATS ARE SUBJECTED TO A VISUAL INSPECTION TO ENSURE DEVICE INTEGRITY. CORRECTIVE ACTION: BASED ON THE QUALITY ENGINEERING RISK ASSESSMENT NO CORRECTIVE ACTION IS WARRANTED AT THIS TIME. A REVIEW OF THE COMPLAINT HISTORY WAS CONDUCTED. BASED ON THIS REVIEW, THE LIKELIHOOD OF THIS TYPE OF REPORT IS RARE. QUALITY ASSURANCE WILL CONTINUE TO MONITOR FOR COMPLAINT TRENDS.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025. NO ADDITIONAL INFORMATION WAS PROVIDED DURING THE INTAKE PROCESS. FOLLOW-UP WITH THE PATIENTS¿ PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME DURING CCPD THERAPY IN THE EARLY MORNING HOURS OF (B)(6) 2025. THE PATIENT¿S SON HAD ARISEN TO USE THE RESTROOM AT APPROXIMATELY 1:30 AM, AND NOTED HIS MOTHER WAS REPORTEDLY SLEEPING COMFORTABLY. HOWEVER, AT APPROXIMATELY 6:30 AM THE PATIENT¿S SON DISCOVERED THE PATIENT HAD EXPIRED WHEN HE WENT TO HELP THE PATIENT DISCONNECT FROM THE LIBERTY SELECT CYCLER. THE PDRN STATED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2024 THROUGH (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT DIRECTLY TO THE FUNERAL HOME (NO AUTOPSY WAS PERFORMED). THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST DUE TO HER SIGNIFICANT CARDIAC HISTORY (CORONARY ARTERY DISEASE, STENOSED AORTIC VALVE NEEDING REPLACEMENT). THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. ADDITIONALLY, THE PDRN STATED THE PATIENT¿S LIBERTY SELECT CYCLER WILL BE SCHEDULED FOR PICK-UP THIS WEEK. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENTS OCCURRED. THE PDRN REPORTED THE CAUSE OF THE SERIOUS ADVERSE EVENTS WAS THE PATIENT¿S SIGNIFICANT CARDIAC COMORBIDITIES. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. THE PDRN SERIOUS ADVERSE EVENTS WERE REPORTEDLY UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENT. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
EPIQ CVX | K132304 | IT WAS REPORTED THAT THE EPIQ CVX ULTRASOUND SYSTEM DISPLAYED AN ERROR CODE DURING A TEE (TRANSESOPHAGEAL ECHOCARDIOGRAM) EXAMINATION ON A STROKE PATIENT. THE ISSUE REMAINED AFTER THE SYSTEM WAS RESTARTED TWICE. THE PROCEDURE WAS ABLE TO BE COMPLETED AFTER AN EXCHANGE OF THE ULTRASOUND SYSTEM. THE PATIENT ULTIMATELY EXPIRED; HOWEVER, THERE IS CURRENTLY NO EVIDENCE TO SUGGEST THE PATIENT EXPIRED AS A RESULT OF THE DEVICE MALFUNCTION AND ADDITIONAL INFORMATION IS BEING OBTAINED. A PHILIPS SERVICE ENGINEER REPLACED THE SYSTEM¿S ACQUISITION MODULE TO ADDRESS THE CUSTOMER¿S IMMEDIATE CONCERN. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW UP REPORT WILL BE SUBMITTED.
APEXPRO TELEMETRY SYSTEM V6 | K080251 | IT WAS ALLEGED THAT THE PATIENT STOPPED BEING MONITORED BY THE GEHC MONITORING SYSTEM. THE PATIENT EXPIRED. | LEGAL MANUFACTURER: HCS RESEARCH PARK - 9900 INNOVATION DRIVE USA WAUWATOSA, WI 53226 A1-6: NOT PROVIDED BY THE CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING AT THIS TIME. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
ION | K182188 | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE THE PATIENT EXPIRED. THE ENDOLUMINAL SALES ASSOCIATE (ESA) REPORTED THAT UPON ARRIVAL AT THE HOSPITAL PROCEDURE ROOM, THEY SAW THAT THE ION SYSTEM WAS UNDOCKED AND BLOOD WAS COMING OUT OF THE ENDOTRACHEAL TUBE. THE PHYSICIAN HAD PLACED A BRONCHOSCOPE, WAS ADMINISTERING SALINE AND EPINEPHRINE VIA THE ENDOTRACHEAL TUBE, AND WAS ATTEMPTING TO STOP THE BLEEDING VIA SUCTIONING THE AIRWAY. THE ESA WAS LATER INFORMED THAT THE PATIENT HAD EXPIRED. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN WERE MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | AN ION SYSTEM LOG REVIEW COULD NOT BE PERFORMED AS THE LOGS ARE NOT AVAILABLE. A DEVICE HISTORY RECORD (DHR) REVIEW CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT OF UNKNOWN AGE WITH UNKNOWN MEDICAL CO-MORBIDITIES UNDERWENT AN ION BRONCHOSCOPY FOR BIOPSY. DURING THE PROCEDURE BLEEDING WAS NOTED IN THE ET TUBE PROMPTING TRANSITION TO MANUAL BRONCHOSCOPY TO INSTILL EPINEPHRINE, SALINE AND SUCTIONING. THE PATIENT DIED DURING THE PROCEDURE. THERE WAS NO ALLEGATION OF MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE REPORTED EVENT WAS LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. ATTEMPTS AT OBTAINING FURTHER INFORMATION HAS BEEN UNSUCCESSFUL. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DAVINCI SINGLE PORT | K171294 | A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE STUDY TO COMPARE THE OUTCOMES OF 96 PATIENTS UNDERGOING ROBOTIC-ASSISTED RADICAL CYSTECTOMY (RARC) WITH URINARY DIVERSION FOR BLADDER CANCER. THE STUDY AIMED TO EVALUATE THE POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF 47 SINGLE-PORT (SP) VS 49 MULTIPORT ROBOTIC-ASSISTED RADICAL CYSTECTOMY (MP RARC). THE STUDY REPORTED THAT IN THE SP COHORT, TWO MORTALITIES WERE NOTED WITH ONE PATIENT EXPIRING ON POSTOPERATIVE DAY 1 SECONDARY TO ACUTE HEMORRHAGIC SHOCK AND RESPIRATORY FAILURE FROM PULMONARY EMBOLISM AND A SECOND PATIENT EXPIRING SECONDARY TO HEMORRHAGIC SHOCK AND RESPIRATORY DISTRESS. THE STUDY CONCLUDED THAT THE SP ROBOTIC PLATFORM OFFERS A SAFE ALTERNATIVE TO THE MP ROBOTIC PLATFORM IN PERFORMING RARC WITH URINARY DIVERSION. NO CLINICALLY SIGNIFICANT DIFFERENCE WAS DEMONSTRATED IN PERIOPERATIVE OR ONCOLOGIC OUTCOMES. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS REPORTED, NOR DID THE AUTHORS ALLEGED THAT ANY INTUITIVE PRODUCTS CAUSED OR CONTRIBUTED TO THE DEATHS. THE DESIGNATED AUTHOR CONTACT HAS NOT RESPONDED TO A REQUEST FOR ADDITIONAL INFORMATION. | THIS MEDWATCH REPORT REPRESENTS THE DA VINCI SINGLE PORT (SP) DEATH REPORTS. SEE SECTION H10 FOR THE ASSOCIATED DA VINCI X MULTIPORT (MP) REPORTED DEATHS. AN INVESTIGATION COULD NOT BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATES OR DA VINCI SYSTEM IDENTIFIERS. THERE WERE NO DEVICE ISSUES DOCUMENTED IN THE ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENTS IN THIS LITERATURE ARTICLE DIED FOLLOWING A RADICAL CYSTECTOMY, ONE FROM ACUTE HEMORRHAGIC SHOCK WITH PULMONARY EMBOLISM AND THE OTHER FROM HEMORRHAGIC SHOCK AND RESPIRATORY DISTRESS. THE DETAILS OF HOW THEY DIED ARE NOT REPORTED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. CITATION: FANG, A. M., HAYEK, O., KAYLOR, J. M., PEYTON, C. C., FERGUSON, J. E., NIX, J. W., & RAIS BAHRAMI, S. (2024). POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF SINGLE-PORT VS MULTIPORT ROBOTIC-ASSISFTED RADICAL CYSTECTOMY. JOURNAL OF ENDOUROLOGY, 38(5), 438¿443. HTTPS://DOI.ORG/10.1089/END.2023.0553 SECTION A2 AGE: THE STUDY POPULATION AGE RANGE FOR DA VINCI SP WAS 67.1 +/-11.1 YEARS. MEAN 67 YEARS. SECTION A3 GENDER: THE STUDY POPULATION FOR DA VINCI SP WAS MALE - 39; FEMALE - 8. SECTION A4 WEIGHT: THE STUDY POPULATION FOR DA VINCIE SP WAS BMI OF 28.8+/-4.9 KG/M2. SECTION B3 EVENT DATE IS UNKNOWN. THE DATE REFLECTS THE DATE THE ARTICLE WAS PUBLISHED - MAY 2024.
DAVINCI XI | K131861 | A LITERATURE ARTICLE DESCRIBED A RETROSPECTIVE STUDY TO COMPARE THE OUTCOMES OF 96 PATIENTS UNDERGOING ROBOTIC-ASSISTED RADICAL CYSTECTOMY (RARC) WITH URINARY DIVERSION FOR BLADDER CANCER. THE STUDY AIMED TO EVALUATE THE POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF 47 SINGLE-PORT (SP) VS 49 MULTIPORT (MP) ROBOTIC-ASSISTED RADICAL CYSTECTOMY. THE STUDY REPORTED THAT IN THE MP COHORT, A MORTALITY WAS NOTED WITH THE PATIENT EXPIRING ON POSTOPERATIVE DAY 2 SECONDARY TO SEVERE ACIDOSIS AND HYPOTENSION. THE STUDY CONCLUDED THAT THE SP ROBOTIC PLATFORM OFFERS A SAFE ALTERNATIVE TO THE MP ROBOTIC PLATFORM IN PERFORMING RARC WITH URINARY DIVERSION. NO CLINICALLY SIGNIFICANT DIFFERENCE WAS DEMONSTRATED IN PERIOPERATIVE OR ONCOLOGIC OUTCOMES. THERE WERE NO DA VINCI DEVICE MALFUNCTIONS REPORTED, NOR DID THE AUTHORS ALLEGED THAT ANY INTUITIVE PRODUCTS CAUSED OR CONTRIBUTED TO THE DEATHS. THE DESIGNATED AUTHOR CONTACT HAS NOT RESPONDED TO A REQUEST FOR ADDITIONAL INFORMATION. | THIS MEDWATCH REPORT REPRESENTS THE DA VINCI MULTIPORT (MP) REPORTED DEATHS. SEE SECTION H10 FOR THE ASSOCIATED DA VINCI SINGLE PORT (SP) DEATH REPORTS. AN INVESTIGATION COULD NOT BE PERFORMED AS THE ARTICLE DID NOT PROVIDE ANY SPECIFIC INFORMATION REGARDING THE PROCEDURE DATES OR DA VINCI SYSTEM IDENTIFIERS. THERE WERE NO DEVICE ISSUES DOCUMENTED IN THE ARTICLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THIS ARTICLE DISCUSSED A SERIES OF PATIENTS COMPARING MULTIPORT AND SINGLE PORT ROBOTIC SURGERY WHO UNDERWENT RADICAL CYSTECTOMY. THIS REPORT IS REGARDING A MULTIPORT PATIENT WHO DIED ON POST OPERATIVE DAY 2 SECONDARY TO SEVERE ACIDOSIS AND HYPOTENSION. NO INTRAOPERATIVE DETAILS ARE PROVIDED. HOW THEY DIED AND THE CIRCUMSTANCES WHICH LED TO THEIR DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. CITATION: FANG, A. M., HAYEK, O., KAYLOR, J. M., PEYTON, C. C., FERGUSON, J. E., NIX, J. W., & RAIS BAHRAMI, S. (2024). POSTOPERATIVE OUTCOMES AND ANALGESIC REQUIREMENTS OF SINGLE-PORT VS MULTIPORT ROBOTIC-ASSISTED RADICAL CYSTECTOMY. JOURNAL OF ENDOUROLOGY, 38(5), 438¿443. HTTPS://DOI.ORG/10.1089/END.2023.0553. SECTION A2 AGE: THE STUDY POPULATION AGE RANGE FOR DA VINCI MP WAS 67.0 +/- 9.4 YEARS. MEAN AGE 67 YEARS. SECTION A3 GENDER: THE STUDY POPULATION FOR DA VINCI MP WAS MALE - 41; FEMALE - 8. SECTION A4 WEIGHT: THE STUDY POPULATION FOR DA VINCIE MP WAS BMI OF 30.3+/-6.5 KG/M2. SECTION B3 EVENT DATE IS UNKNOWN. THE DATE REFLECTS THE DATE THE ARTICLE WAS PUBLISHED - MAY 2024.
DAVINCI X | K171294 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED PULMONARY WEDGE RESECTION, THE PROCEDURE WAS CONVERTED TO OPEN AFTER APPROXIMATELY 4 HOURS OF DISSECTION FOR LYSIS OF ADHESIONS. THE SURGEON INFORMED THE CLINICAL SALES REPRESENTATIVE THAT THE CONVERSION WAS DUE TO THE SURGEON NOT BEING ABLE TO REMOVE ALL OF THE ADHESIONS, THUS WAS UNABLE TO VISUALIZE THE PATIENT ANATOMY, AND IN ORDER TO PREVENT FURTHER BLEEDING FROM THE EXTENDED DISSECTION. THE PATIENT EXPIRED TWO HOURS POST CONVERSION DUE TO CARDIAC SHOCK DURING THE OPEN PROCEDURE. PRIOR TO THE PROCEDURE, THE SURGEON ACKNOWLEDGED THAT THEY MIGHT HAVE TO CONVERT TO OPEN AS THE PATIENT WAS KNOWN TO BE UNSTABLE DUE TO A HISTORY OF CARDIAC ISSUES AND AN UNSPECIFIED INFECTION LIKELY CAUSING ADHESIONS. THERE WAS NO ORGAN DAMAGE REPORTED. THE PATIENT REQUIRED AN UNSPECIFIED AMOUNT OF BLOOD TRANSFUSIONS DURING THE OPEN SURGERY. IT WAS REPORTED THAT THE DEATH WAS NOT RELATED TO DA VINCI PRODUCTS. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, TIP-UP FENESTRATED GRASPER, MARYLAND BIPOLAR FORCEPS, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD KNOWN CARDIAC DISEASE AND PREDICTED TO HAVE CHALLENGING ANATOMY GOING INTO THE PROCEDURE. THEY CONVERTED TO OPEN AND THE PATIENT HAD AN EVENT 2 HOURS INTO THE OPEN PORTION OF THE OPERATION AND EXPIRED. NO ALLEGATION HAS BEEN MADE REGARDING ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONVERTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER AS NO LONGER NEEDED. EVALUATION RESULTS DETERMINED THERE WERE NO ERRORS OR FAULTS FOUND. THE DEVICE IS CURRENTLY AWAITING CONFIRMATION FOR RETURN TO LOAN OR TO BE SCRAPPED. IF ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. IN THE PREVIOUS FILE THE REPORT WAS SUBMITTED FOR AN INITIAL WHICH IS UPDATED TO THE FINAL IN THIS REPORT. ADDITIONALLY, EVALUATION METHOD CODE GRID, EVALUATION RESULTS CODE GRID, COMPONENT CODE GRID AND CONCLUSION CODE GRID WERE ALSO UPDATED IN THIS REPORT.
PATIENT INFORMATION CENTER IX | K211900 | IT WAS REPORTED THE PATIENT PASSED WAY WHILE CONNECTED TO THE MONITOR IN THE TIMEFRAME OF (B)(6) 2025. WHEN THE CUSTOMER WAS ATTEMPTING TO COLLECT THE LOGS ON (B)(6) 2025, THE APPLICATION SHUT DOWN; THEREFORE, THE CUSTOMER WAS REQUESTING ASSISTANCE IN DOWNLOADING THE LOGS. IN ADDITION, IT WAS INDICATED THE CUSTOMER ALLEGED NO FAILURE OR FAULT OF THE DEVICE AS IT RELATES TO THE PATIENT DEATH. A REVIEW OF THE LOGS BY PRODUCT SUPPORT ENGINEERING CONFIRMED THE DEVICE WAS FUNCTIONING AS EXPECTED DURING THE EVENT TIMEFRAME. | THE IMPELLA DEVICE WILL NOT BE RECEIVED FROM THE CUSTOMER. THEREFORE, INVESTIGATION OF THE DEVICE IS NOT POSSIBLE. SHOULD THE DEVICE OR ANY NEW INFORMATION BE RECEIVED, A SUPPLEMENTAL MDR WILL BE FILED. THE INSTRUCTIONS FOR USE FOR THE IMPELLA 5.5 WITH SMART ASSIST FOR USE DURING CARDIOGENIC SHOCK STATES THE FOLLOWING: SECTION: POTENTIAL ADVERSE EVENTS (UNITED STATES): ¿ACUTE RENAL DYSFUNCTION, AORTIC VALVE INJURY, BLEEDING, CARDIOGENIC SHOCK, CEREBRAL VASCULAR ACCIDENT/STROKE, DEATH, HEMOLYSIS, LIMB ISCHEMIA, MYOCARDIAL INFARCTION, RENAL FAILURE, THROMBOCYTOPENIA AND CARDIAC OR VASCULAR INJURY (INCLUDING VENTRICULAR PERFORATION).¿
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
NELLCOR | K120773 | IT WAS REPORTED THAT THE MONITOR'S ALARM HISTORY WAS REQUESTED FOR INVESTIGATION. THE ALARM HISTORY COULD NOT BE VIEWED AT THAT TIME, SO ONLY THE ALARM SETTINGS WAS RECORDED. THE SENSOR WAS DETACHED AND THE POWER WAS ON. THE PATIENT PASSED AWAY AT THE MEDICAL FACILITY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
RESPIRONICS | K102985 | AN AUTHORIZED SERVICE PROVIDER (ASP) REPORTED THAT A PATIENT PASSED AWAY WHILE THE V60 VENTILATOR WAS IN USE ON THE PATIENT AND THAT THE DEVICE WAS FOUND ALARMING FOR A DEPLETED BACKUP BATTERY. THE REPORTER IS UNSURE IF THE DEVICE HAD ANYTHING TO DO WITH THE PATIENT¿S OUTCOME, AS IT WAS COMMUNICATED THAT THE PATIENT WAS ALREADY IN THE PROCESS OF PASSING AWAY WHEN BEING PLACED ONTO THE VENTILATOR. ON (B)(6) 2025, THE MANAGER OF RESPIRATORY CARE OR NEURODIAGNOSTIC SERVICES AT THE HOSPITAL CONFIRMED IN A GOOD FAITH EFFORT RESPONSE THAT THE DEVICE ALARMED AS INTENDED WHEN THE DEVICE WAS CHECKED, AND THAT THE BACKUP BATTERY WAS FOUND TO BE DEPLETED. THE MANAGER ALSO STATED THAT THE PATIENT DEATH WAS NOT DIRECTLY ASSOCIATED WITH THE EVENT. THE V60 VENTILATOR WAS PLUGGED INTO AN ELECTRICAL OUTLET, AND THE V60 VENTILATOR HAD BEEN FUNCTIONING PROPERLY. THE MANAGER STATED THAT THEY JUST WANTED TO CONFIRM THAT THE BACKUP BATTERY WORKED PROPERLY AND REQUESTED AN EVALUATION OF THE DEVICE TO DETERMINE HOW LONG THE DEVICE MAY HAVE BEEN USING THE BACKUP BATTERY. THE ASP AGREED TO TRAVEL TO THE CUSTOMER¿S SITE TO PERFORM AN EVALUATION OF THE DEVICE TO CONFIRM THAT THE DEVICE AND ITS BACKUP BATTERY ARE FULLY FUNCTIONAL. A CLINICAL ASSESSMENT WAS PERFORMED BY A PHILIPS CLINICAL EXPERT ON THIS CASE. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, DEVICE CAUSE AND CONTRIBUTION TO THE EVENT CANNOT CURRENTLY BE CONFIRMED, NOR REFUTED, BASED ON THE EVIDENCE PRESENT. ACCORDING TO THE V60 USER MANUAL ABOUT THE BACKUP BATTERY, ¿TO REDUCE THE RISK OF POWER FAILURE TO THE VENTILATOR, PAY CLOSE ATTENTION TO THE BATTERY'S CHARGE LEVEL. THE BATTERY'S OPERATION TIME IS APPROXIMATE AND IS AFFECTED BY VENTILATOR SETTINGS, DISCHARGE AND RECHARGE CYCLES, BATTERY AGE, AND AMBIENT TEMPERATURE. THE BACKUP BATTERIES ARE INTENDED FOR SHORT-TERM USE ONLY. THEY ARE NOT INTENDED TO BE A PRIMARY POWER SOURCE. AS A SAFEGUARD, THE VENTILATOR PROVIDES A LOW BATTERY ALARM. IT ALSO HAS A CAPACITOR-DRIVEN BACKUP ALARM THAT SOUNDS FOR AT LEAST 2 MINUTES WHEN BATTERY POWER IS COMPLETELY LOST. THE VENTILATOR CHARGES THE BATTERY WHENEVER THE VENTILATOR IS CONNECTED TO AC, WITH OR WITHOUT THE VENTILATOR SWITCHED ON. THE BATTERY (CHARGED) LED FLASHES TO SHOW THAT THE BATTERY IS BEING CHARGED. CHECK THE BATTERY CHARGE LEVEL BEFORE PUTTING A PATIENT ON THE VENTILATOR AND BEFORE UNPLUGGING THE VENTILATOR FOR TRANSPORT OR OTHER PURPOSES. THE POWER SOURCE SYMBOL AT THE BOTTOM OF THE RIGHT-HAND CORNER OF THE SCREEN SHOWS THE POWER SOURCE IN USE AND, IF THE VENTILATOR IS RUNNING ON BATTERY, THE LEVEL OF BATTERY CHARGE. IF THE BATTERY IS NOT FULLY CHARGED, RECHARGE IT BY CONNECTING THE VENTILATOR TO AC POWER FOR A MINIMUM OF 5 HOURS. PRESSING THE HELP BUTTON SHOWS YOU THE APPROXIMATE TIME REMAINING UNTIL THE BATTERY IS FULLY CHARGED. IF THE BATTERY IS NOT FULLY CHARGED WITHIN 16 HOURS, OR THE VENTILATOR DISPLAYS A CHECK VENT: BATTERY FAILED ALARM, REPLACE THE BATTERY.¿ IN ADDITION, THE V60 USER MANUAL¿S WARNINGS, CAUTIONS, AND NOTES ON ALARMS AND MESSAGES STATES THAT ¿IF AC POWER FAILS AND THE BACKUP BATTERY IS NOT INSTALLED OR IS DEPLETED, AN AUDIBLE AND VISUAL ALARM ANNUNCIATES FOR AT LEAST 2 MINUTES. IMMEDIATELY DISCONTINUE VENTILATOR USE AND SECURE AN ALTERNATIVE MEANS OF VENTILATION. AS IN MOST VENTILATORS WITH PASSIVE EXHALATION PORTS, WHEN POWER IS LOST, SUFFICIENT AIR IS NOT PROVIDED THROUGH THE CIRCUIT, AND EXHALED AIR MAY BE REBREATHED.¿. FURTHER GOOD FAITH EFFORTS ARE BEING COMPLETED TO THE CUSTOMER TO OBTAIN ADDITIONAL INFORMATION REGARDING THE REPORTED EVENT. THE INVESTIGATION IS ONGOING. | ON 16APR2025, A 3RD PARTY AUTHORIZED SERVICE PROVIDER (ASP) WENT TO THE CUSTOMER SITE AND EVALUATED THE DEVICE. THE ASP COMPLETED A PERFORMANCE VERIFICATION TEST (PVT) WHICH THE DEVICE PASSED, CONFIRMING THAT THE DEVICE MET THE MANUFACTURER¿S SPECIFICATIONS FOR USE AND WAS NOT EXPERIENCING ANY ISSUES. THE DEVICE WAS CONFIRMED TO BE FULLY FUNCTIONAL. A COPY OF A PORTION OF THE DEVICE DRPT (DIAGNOSTIC REPORT) WAS PROVIDED AND IT CONFIRMED THAT, ACROSS APPROXIMATELY 7 AND A HALF MINUTES, THE DEVICE PRODUCED A LOW INTERNAL BATTERY ALARM 4 TIMES, CONFIRMING THAT THE DEVICE ALARMED AND NOTIFIED THE USER APPROPRIATELY, AS INTENDED. WITH THE DEVICE HAVING ALARMED AS EXPECTED AND PASSED THE PVT, IT IS CONFIRMED THAT THERE WAS NO DEVICE ISSUE AT THE TIME OF THE PATIENTS PASSING. IF THE DEVICE DID SHUT DOWN DUE TO BATTERY DEPLETION, IT IS CONCLUDED THAT IT WAS DUE TO A LACK OF CORRECTIVE ACTION BY THE USER FOLLOWING THE INITIAL LOW INTERNAL BATTERY ALARM. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
CHPV RT ANG RESERVOIR UNITIZED | K221840 | N/A | UPDATED FIELDS: G3, G6, H1, H2, H11. CORRECTED FIELD: H1. ADDITIONAL INFORMATION PROVIDED BY INTEGRA'S MEDICAL AFFAIRS DIRECTOR: "I HAD A DISCUSSION WITH DR. (B)(6) (SURGEON) CONCERNING THE CONSENSUS ON SUNKEN FLAP SYNDROME - BASED RECOMMENDATIONS FOR THE DIAGNOSIS AND SURGICAL MANAGEMENT OF CRANIOPLASTY AND POST-TRAUMATIC HYDROCEPHALUS AS OUTLINED BY THE EUROPEAN PANEL." "DR. (B)(6) HAS UNDERSTOOD THE CONSENSUS AND ITS RELEVANCE TO CRANIOPLASTY IN THE CONTEXT OF POST-TRAUMATIC HYDROCEPHALUS. SHOULD THERE BE ANY NEED FOR FURTHER CLARIFICATION FROM DR. (B)(6), I WILL PROVIDE ADDITIONAL UPDATES ACCORDINGLY."
AZURION | K200917 | IT WAS REPORTED TO PHILIPS THAT DURING AN EMERGENT PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE, INVOLVING A PATIENT WITH ST-ELEVATION MYOCARDIAL INFARCTION (STEMI), THE SYSTEM DISPLAYED AN X-RAY TUBE CURRENT OUT-OF-RANGE WARNING MESSAGE. THE PROCEDURE WAS ABORTED. THE PATIENT BEGAN TO DECOMPENSATE, AND THE PHYSICIAN INSERTED A BALLOON PUMP. THE PATIENT WAS THEN TRANSFERRED TO ANOTHER FACILITY, WHERE THEY WERE NOTED TO BE ON A VENTILATOR AND UNRESPONSIVE. ON (B)(6) 2025, THE CUSTOMER INFORMED PHILIPS THAT THE PATIENT HAD PASSED AWAY. NO OTHER INFORMATION WAS PROVIDED. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE CALLED AND WANTED TO KNOW IF SHE NEEDED TO RETURN THE REPLACEMENT DEVICE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S WIFE WAS INFORMED SHE COULD RETURN THE DEVICE IF SHE WANTED TO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAD PASSED AWAY. THE PATIENT'S WIFE CALLED AND WANTED TO KNOW IF SHE NEEDED TO RETURN THE REPLACEMENT DEVICE. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S WIFE WAS INFORMED SHE COULD RETURN THE DEVICE IF SHE WANTED TO. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H CODING IN THIS REPORT.
PLUM 360 (V 15.X) | K141789 | IT WAS REPORTED THAT A PLUM 360 SHUT DOWN UNEXPECTEDLY DURING USE. THE REPORTER STATED THAT "THE PUMP HAD A REPLACE BATTERY ALERT FROM 2 DAGO AGO BUT NO ONE PULL THE PUMP FROM SERVICE NOR SEND IN A WORK ORDER. A LEVO INFUSION WAS STARTED ON THE PATIENT AND THE PUMP WAS UNPLUGGED TO TRANSPORT THE PATIENT. THE PUMP SHUT DOWN AND THE PATIENT DIED." NO ADDITIONAL INFORMATION IS AVAILABLE AT THIS TIME. | THE DEVICE WAS RETURNED FOR EVALUATION; HOWEVER, TESTING HAS NOT YET BEEN COMPLETED.
DREAMSTATION 2 ADVANCED AUTO CPAP | K200480 | THE DAUGHTER OF THE PATIENT REPORTED THE PATIENT EXPIRED. THE REPORTED INQUIRED ABOUT THE REGISTRATION INFORMATION FOR THE DREAMSTATION 2 ADVANCED AUTO CPAP DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
FLEXOR RAABE GUIDING SHEATH | K142829 | AS REPORTED, DURING AN ANGIOGRAM WITH REVASCULARIZATION ATTEMPT, A FLEXOR RAABE GUIDING SHEATH UNRAVELED AND SEPARATED. THE PATIENT HAD SIGNIFICANT ATHEROSCLEROTIC PERIPHERAL ARTERY DISEASE, AND THE REVASCULARIZATION ATTEMPT WAS FOR IMPROVEMENT OF WOUND HEALING ON THE RIGHT FOOT. ACCESS WAS REPORTEDLY OBTAINED IN THE POPLITEAL AND LEFT COMMON FEMORAL ARTERIES (LCFA). THE COMPLAINT DEVICE WAS ADVANCED VIA A CONTRALATERAL APPROACH FROM THE LCFA TO THE MID- SUPERFICIAL FEMORAL ARTERY, WITHOUT ENCOUNTERING RESISTANCE. THE ACCESS SITE WAS NOT SCARRED; HOWEVER, THE ANATOMY WAS CALCIFIED AND MILDLY TORTUOUS. THE BIFURCATION WAS ALSO CALCIFIED. VARIOUS 0.018-INCH WIRE GUIDES WERE USED THROUGH THE SHEATH DURING THE PROCEDURE, AS WELL AS A COOK MICROCATHETER, OTHER MANUFACTURERS' BALLOON CATHETERS, AND ANOTHER MANUFACTURER¿S GUIDING CATHETER. RESISTANCE WAS NOT ENCOUNTERED DURING INSERTION OR REMOVAL OF ANY DEVICE THROUGH THE SHEATH. DURING THE PROCEDURE, WHILE ATTEMPTING TO EXCHANGE THE COMPLAINT DEVICE WITH A 6-FRENCH SHEATH OVER ANOTHER MANUFACTURER'S WIRE IN ANTICIPATION OF STENTING, APPROXIMATELY TWENTY CENTIMETERS OF THE SHEATH "SHEARED" IN THE VESSEL UPON ATTEMPTED GRADUAL REMOVAL. THE DILATOR WAS NOT REINSERTED PRIOR TO ATTEMPTED REMOVAL OF THE SHEATH. RESISTANCE WAS REPORTEDLY NOT ENCOUNTERED DURING SHEATH REMOVAL. VASCULAR SURGERY WAS CONSULTED FOR EMERGENT EVALUATION OF THE RETAINED SHEATH. THE USER CONFIRMED THAT A WIRE DID NOT SEPARATE DURING THE PROCEDURE. PER THE CUSTOMER, THE SHEATH SEPARATION DID NOT RESULT IN A LIFE-THREATENING EVENT, AND THE PATIENT DID NOT BECOME HEMODYNAMICALLY UNSTABLE AFTER THE SHEATH SEPARATED. BILATERAL SURGICAL CUTDOWNS WERE PERFORMED, AND THE RETAINED FRACTURED SHEATH WAS REMOVED. THE PATIENT WAS SENT TO RECOVERY WITH GOOD PERFUSION. TWELVE DAYS AFTER THE PROCEDURE, THE PATIENT, WHO HAD A HISTORY OF CORONARY ARTERY DISEASE, DEVELOPED CHEST PAIN, EXPERIENCED CARDIAC ARREST AND DIED. AN AUTOPSY WILL NOT BE PROVIDED TO COOK TO AID IN THE INVESTIGATION OF THE EVENT. | BLANK FIELDS ON THIS FORM INDICATE THE INFORMATION IS UNKNOWN, UNCHANGED, OR UNAVAILABLE. SUMMARY OF EVENT: AS REPORTED, DURING AN ANGIOGRAM WITH REVASCULARIZATION ATTEMPT, A FLEXOR RAABE GUIDING SHEATH UNRAVELED AND SEPARATED. THE PATIENT HAD SIGNIFICANT ATHEROSCLEROTIC PERIPHERAL ARTERY DISEASE, AND THE REVASCULARIZATION ATTEMPT WAS FOR IMPROVEMENT OF WOUND HEALING ON THE RIGHT FOOT. ACCESS WAS REPORTEDLY OBTAINED IN THE POPLITEAL AND LEFT COMMON FEMORAL ARTERIES (LCFA). THE COMPLAINT DEVICE WAS ADVANCED VIA A CONTRALATERAL APPROACH FROM THE LCFA TO THE MID- SUPERFICIAL FEMORAL ARTERY, WITHOUT ENCOUNTERING RESISTANCE. THE ACCESS SITE WAS NOT SCARRED; HOWEVER, THE ANATOMY WAS CALCIFIED AND MILDLY TORTUOUS. THE BIFURCATION WAS ALSO CALCIFIED. VARIOUS 0.018-INCH WIRE GUIDES WERE USED THROUGH THE SHEATH DURING THE PROCEDURE, AS WELL AS A COOK MICROCATHETER, OTHER MANUFACTURERS¿ BALLOON CATHETERS, AND ANOTHER MANUFACTURER¿S GUIDING CATHETER. RESISTANCE WAS NOT ENCOUNTERED DURING INSERTION OR REMOVAL OF ANY DEVICE THROUGH THE SHEATH. DURING THE PROCEDURE, WHILE ATTEMPTING TO EXCHANGE THE COMPLAINT DEVICE WITH A 6-FRENCH SHEATH OVER ANOTHER MANUFACTURER¿S WIRE IN ANTICIPATION OF STENTING, APPROXIMATELY TWENTY CENTIMETERS OF THE SHEATH ¿SHEARED¿ IN THE VESSEL UPON ATTEMPTED GRADUAL REMOVAL. THE DILATOR WAS NOT REINSERTED PRIOR TO ATTEMPTED REMOVAL OF THE SHEATH. RESISTANCE WAS REPORTEDLY NOT ENCOUNTERED DURING SHEATH REMOVAL. VASCULAR SURGERY WAS CONSULTED FOR EMERGENT EVALUATION OF THE RETAINED SHEATH. THE USER CONFIRMED THAT A WIRE DID NOT SEPARATE DURING THE PROCEDURE. PER THE CUSTOMER, THE SHEATH SEPARATION DID NOT RESULT IN A LIFE-THREATENING EVENT, AND THE PATIENT DID NOT BECOME HEMODYNAMICALLY UNSTABLE AFTER THE SHEATH SEPARATED. BILATERAL SURGICAL CUTDOWNS WERE PERFORMED, AND THE RETAINED FRACTURED SHEATH WAS REMOVED. THE PATIENT WAS SENT TO RECOVERY WITH GOOD PERFUSION. TWELVE DAYS AFTER THE PROCEDURE, THE PATIENT, WHO HAD A HISTORY OF CORONARY ARTERY DISEASE, DEVELOPED CHEST PAIN, EXPERIENCED CARDIAC ARREST AND DIED. AN AUTOPSY WILL NOT BE PROVIDED TO COOK TO AID IN THE INVESTIGATION OF THE EVENT. INVESTIGATION EVALUATION: REVIEWS OF THE COMPLAINT HISTORY, DEVICE HISTORY RECORD (DHR), INSTRUCTIONS FOR USE (IFU), MANUFACTURING INSTRUCTIONS, AND QUALITY CONTROL PROCEDURES WERE CONDUCTED DURING THE INVESTIGATION. A VISUAL INSPECTION OF THE COMPLAINT DEVICE WAS ALSO CONDUCTED. THE COMPLAINT DEVICE WAS RETURNED TO COOK FOR INVESTIGATION IN TWO SECTIONS. THE PROXIMAL SECTION WAS UNRAVELED; HOWEVER, THE SEPARATED SHEATH MATERIAL WAS CONNECTED BY THE UNRAVELED METAL COIL. THE DISTAL SECTION WAS SEPARATED, MEASURING APPROXIMATELY 48-CENTIMETERS IN LENGTH. A DOCUMENT-BASED INVESTIGATION EVALUATION WAS PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORD FOUND NO RELEVANT NON-CONFORMANCES ON THE LOT. A REVIEW OF COMPLAINT HISTORY FOUND NO ADDITIONAL COMPLAINTS FOR THIS LOT NUMBER. THE PRODUCT IFU STATES THAT REINSERTION OF THE DILATOR PRIOR TO REMOVAL OF THE SHEATH ¿INCREASES THE STRENGTH OF THE SHEATH AND LESSENS THE RISK OF DEVICE SEPARATION¿ AND SUGGESTS INSERTION OF THE DILATOR PRIOR TO REMOVAL IF RESISTANCE IS ENCOUNTERED OR ANTICIPATED DURING WITHDRAWAL OF THE DEVICE. A REVIEW OF THE DEVICE MASTER RECORD (DMR) CONCLUDED THAT SUFFICIENT INSPECTION ACTIVITIES ARE IN PLACE TO IDENTIFY THIS FAILURE MODE PRIOR TO DISTRIBUTION. THE INFORMATION PROVIDED UPON REVIEW OF THE DMR, DHR, IFU, AND INVESTIGATION OF THE RETURNED DEVICE SUGGESTS THAT THERE IS EVIDENCE THE DEVICE WAS MANUFACTURED TO SPECIFICATION. THERE IS NO EVIDENCE OF NON-CONFORMING DEVICES IN-HOUSE OR IN THE FIELD. BASED ON THE INFORMATION PROVIDED AND THE RESULTS OF THE INVESTIGATION, COOK HAS CONCLUDED THAT THE PATIENT¿S ANATOMY LIKELY CONTRIBUTED TO THIS EVENT. THE ANATOMY WAS MILDLY TORTUOUS AND THE VESSELS, INCLUDING THE BIFURCATION, WERE CALCIFIED. THE DILATOR WAS NOT REINSERTED PRIOR TO ATTEMPTED REMOVAL OF THE SHEATH. ALTHOUGH RESISTANCE WAS REPORTEDLY NOT ENCOUNTERED DURING SHEATH ADVANCEMENT OR REMOVAL, THE SHEATH MATERIAL WAS BUNCHED/TWISTED/DEFORMED, UNRAVELED, AND SEPARATED UPON RETURN TO COOK. IT IS LIKELY THAT THE SHEATH CAUGHT WITHIN THE TORTUOUS AND CALCIFIED ANATOMY UPON REMOVAL OF THE DEVICE. THE RISK ANALYSIS FOR THIS FAILURE MODE WAS REVIEWED AND NO ADDITIONAL ESCALATION WAS REQUIRED. THE APPROPRIATE PERSONNEL HAVE BEEN NOTIFIED, AND COOK WILL CONTINUE TO MONITOR FOR SIMILAR EVENTS. THIS REPORT IS REQUIRED BY THE FDA UNDER 21 CFR PART 803 AND IS BASED ON UNCONFIRMED INFORMATION SUBMITTED BY OTHERS. NEITHER THE SUBMISSION OF THIS REPORT NOR ANY STATEMENT CONTAINED HEREIN IS INTENDED TO BE AN ADMISSION THAT ANY COOK DEVICE IS DEFECTIVE OR MALFUNCTIONED, THAT A DEATH OR SERIOUS INJURY OCCURRED, NOR THAT ANY COOK DEVICE CAUSED, CONTRIBUTED TO, OR IS LIKELY TO CAUSE OR CONTRIBUTE TO A DEATH OR SERIOUS INJURY IF A MALFUNCTION OCCURRED.
TIGHTRAIL ROTATING DILATOR SHEATH | K150360 | A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) AND A RIGHT ATRIAL (RA) LEAD DUE TO CIED SYSTEM/POCKET INFECTION AND BACTEREMIA. SPECTRANETICS LEAD LOCKING DEVICES (LLDS) WERE INSERTED INTO EACH LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 16F GLIDELIGHT LASER SHEATH ON THE RV LEAD, THE LEAD WAS REMOVED SUCCESSFULLY. NEXT, TARGETING THE RA LEAD WITH THE 16F GLIDELIGHT, ADVANCEMENT WAS MADE TO THE INNOMINATE/SUPERIOR VENA CAVA (SVC) JUNCTION, WHERE PROGRESS STALLED. THEN, SWITCHING TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH, THE RA LEAD WAS REMOVED. UPON REMOVAL OF THE LEAD, THE PATIENT'S BLOOD PRESSURE DROPPED, AND RESCUE EFFORTS BEGAN, INCLUDING STERNOTOMY. A PERFORATION FROM THE SVC/INNOMINATE JUNCTION AND THE SVC WAS DISCOVERED AND REPAIRED. THE PATIENT SURVIVED THE PROCEDURE; HOWEVER, HE PASSED AWAY THE FOLLOWING DAY. THIS REPORT CAPTURES THE 13F TIGHTRAIL DEVICE IN USE WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. | D4/H4) THE LOT NUMBER WAS NOT PROVIDED, THUS THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE TIGHTRAIL WAS DISCARDED, THUS NO RETURNED PRODUCT INVESTIGATION WAS PERFORMED. H6) PERFORATION OF VESSELS, GREAT VESSEL PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI. | CORRECTION: THIS REPORT IS BEING SUBMITTED TO INCLUDE THE PRODUCT UDI.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A PATIENT'S SON STATING THE PATIENT DID NOT RECEIVE THE DEVICE. THE PATIENT PASSED AWAY. THE CAUSE OF DEATH IS NOT KNOWN. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LATERAL F3D 18 MM IMPLANT; COBB ELEVATOR | K231743 | THE PRODUCT MANAGER WAS PRESENT IN THE OPERATING ROOM FOR A CASE INVOLVING A LATERAL L1-L5 LUMBAR INTERBODY FUSION (LLIF) PROCEDURE. THE PATIENT WAS POSITIONED LATERALLY WITH THE LEFT SIDE FACING UP. THE SURGEON MADE A SMALL INCISION ON THE LATERAL FLANK FOR ACCESS TO THE L1 VERTEBRAL BODY. THE PROCEDURE COMMENCED WITH THE PLACEMENT OF A ZAVATION LATERAL RETRACTOR SYSTEM, THE DISK WAS DISSECTIONED, AND THEN THE LATERAL F3D IMPLANT WAS PLACED INTO AT THE L1 LEVEL. THE OPERATION PROGRESSED TO L2, WHERE THE SURGEON UTILIZED A COBB SPINAL ELEVATOR TO EXCISE OSTEOPHYTES FOR ADEQUATE ACCESS TO THE L2 VERTEBRA. DURING THE PROCEDURE, THE PATIENT BEGAN TO EXPERIENCE SIGNIFICANT HEMORRHAGING. A GENERAL SURGEON WAS URGENTLY CALLED TO THE OPERATING ROOM FOR LIFE-SAVING INTERVENTION. UNFORTUNATELY, THE SOURCE OF THE BLEEDING COULD NOT BE IDENTIFIED IN TIME, AND THE PATIENT ULTIMATELY SUCCUMBED TO THE HEMORRHAGE.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE TOTAL BLOWER HOURS WERE 1452.3 WITH NO ERRORS OR REPORTED FAULTS. THE BLOWER BOX ASSEMBLY WAS CONTAMINATED AND REPLACED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED AS PREVENTIVE MAINTENANCE. IN ADDITION, THE SD CARD WAS FITTED. THE DEVICE WAS CLEANED AND TESTED. THE DEVICE PASSED ALL TESTS. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS REPORTED AS BEING RETURNED TO THE DME TWO YEARS AGO, HOWEVER THE DEVICE WAS NEVER RECEIVED BY THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION CPAP PRO DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS REPORTED AS BEING RETURNED TO THE DME TWO YEARS AGO, HOWEVER THE DEVICE WAS NEVER RECEIVED BY THE MANUFACTURER. THE MANUFACTURER RECEIVED ADDITIONAL INFORMATION FROM THE PATIENT'S DAUGHTER REGARDING THE EL SETTLEMENT PAYMENT IN REFERENCE TO THE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. SHE EXPRESSED, "MY DAD SHOULD NOT HAVE PASSED IF HE RECEIVED HIS NEW MACHINE ON TIME. WE TRIED FOR MONTHS." THE DAUGHTER SHARED THAT HER FATHER HAD BEEN USING THE CPAP DEVICE FOR APPROXIMATELY 4 OR 5 YEARS BEFORE HIS DEATH. SHE EXPLAINED THAT HER FATHER REGULARLY CLEANED HIS DEVICE ACCORDING TO THE INSTRUCTIONS, USING THE RECOMMENDED CLEANING SOLUTIONS AND WATER. WHILE SHE WAS UNSURE IF HE USED A SOCLEAN OR A SIMILAR DEVICE, SHE NOTED THAT HE ALWAYS KEPT UP WITH THE LATEST EQUIPMENT, SO IT IS POSSIBLE HE DID. THE DAUGHTER FURTHER MENTIONED THAT HER FATHER NEVER APPEARED WELL-RESTED, OFTEN REQUIRING OXYGEN AT 3 OR 4 LITERS AND USED THE CPAP MACHINE AT NIGHT. SHE RECALLED THAT HE HAD USED THE CPAP THE NIGHT BEFORE HIS PASSING BUT WAS STRUGGLING TO GET IT TO WORK PROPERLY. UPON WAKING, HIS OXYGEN LEVELS WERE IN THE 70S, AND HE HAD EXPERIENCED DIFFICULTIES WITH THE CPAP DEVICE OVER THE PAST FEW NIGHTS. SHE STATED THAT HER FATHER ALWAYS SEEMED TIRED AND GROGGY, AND ALTHOUGH HE INTENDED TO CONTACT PHILIPS ABOUT HIS DEVICE, SHE BELIEVES HE NEVER MANAGED TO DO SO BEFORE HE FELL ILL AND PASSED AWAY. THE DAUGHTER DESCRIBED A PARTICULARLY DIFFICULT NIGHT BEFORE HIS PASSING, WHEN HER FATHER HAD TROUBLE USING THE MACHINE, COULD NOT SLEEP WITH IT, AND FOUND IT MALFUNCTIONING. SHE ALSO REPORTED THAT HE RELIED ON OXYGEN, THE CPAP MACHINE, AND A NEBULIZER. THE DAUGHTER STATED HER FATHER PASSED AWAY FROM COPD AND "ALL THAT STUFF" RESTRICTIVE AIRWAYS, CHRONIC COPD, THE WHOLE REASON HE WAS ON THE NEBULIZER AND OXYGEN. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX A AGE AND SEX IN THIS REPORT. THE MANUFACTURER HAS UPDATED BOX B DATE OF DEATH AND OTHER RELEVANT HISTORY IN THIS REPORT. THE MANUFACTURER HAS UPDATED PRODUCT PROBLEM OR ADVERSE EVENT TO BOTH. THE PREVIOUSLY SUBMITTED REPORT WAS FILED AS AN ADVERSE EVENT ONLY. THE MANUFACTURER HAS UPDATED BOX E INITIAL REPORTER IN THIS REPORT. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
CENTRIMAG BLOOD PUMP | K020271 | IT WAS REPORTED THAT THE PATIENT WAS ON VENO-VENOUS EXTRACORPOREAL MEMBRANE OXYGENATION (VV ECMO). ON (B)(6) 2025 AT APPROXIMATELY 23:32, IT WAS IDENTIFIED THAT THE VENOUS CANNULA CRACKED AND DISCONNECTED FROM CIRCUIT. A CIRCUIT EXCHANGE WAS PERFORMED IN ATTEMPT TO STOP EXSANGUINOUS DRAINAGE. DURING TIME OF CIRCUIT, THE PATIENT CODED AND WAS PLACED BACK ON VV ECMO BUT BUBBLES WERE IDENTIFIED IN THE CIRCUIT WITH INABILITY TO GET OUT WITHOUT CAUSING ADDITIONAL HARM TO PATIENT. THE PATIENT PASSED AWAY AT 00:05 ON (B)(6) 2025 WHEN THE CENTRIMAG BLOOD PUMP WAS "BUMPED BY STAFF" WHICH CAUSED AIR TO GET IN THE CIRCUIT. THE PATIENT WAS NON-PULSATILE ALREADY AND THE AIR CAUSED THE PATIENT TO PASS AWAY. PER THE SITE, GIVEN THE CIRCUMSTANCES, IT WAS BELIEVED THAT THE HARDWARE WAS NOT AT FAULT AND WOULD NOT BE RETURNED. THE STAFF IN THE ROOM "KNOCKED INTO THE PUMP" ALLEGEDLY CAUSING A CRACK WHICH REQUIRED SWITCHING THE CIRCUIT DUE TO THE PATIENT EXSANGUINATING. A NEW PRE-PRIMED CIRCUIT WAS INITIATED, AND AIR WAS IDENTIFIED POST INITIATION. THE PATIENT DECEASED BECAUSE THE AIR WAS UNABLE TO BE SAFELY REMOVED. THE STAFF IN THE ROOM "KNOCKED INTO THE PUMP" ALLEGEDLY CAUSING A CRACK WHICH REQUIRED SWITCHING THE CIRCUIT DUE TO THE PATIENT EXSANGUINATING. A NEW PRE-PRIMED CIRCUIT WAS INITIATED, AND AIR WAS IDENTIFIED POST INITIATION. THE PATIENT DECEASED BECAUSE THE AIR WAS UNABLE TO BE SAFELY REMOVED. THE PUMP WAS A PRE-PRIMED CIRCUIT THAT WAS READY FOR EMERGENCIES PRIOR TO THE EVENT. PER TWO SEPARATE PHONE CALLS, AIR WAS IDENTIFIED IN THE CIRCUIT BEFORE INITIATING SUPPORT AND LIKELY ENTERED THE SYSTEM DURING PRIMING. AFTER FOLLOW-UP THE ECMO COORDINATOR COMMUNICATED "THE CHANGE OUT WAS ATTEMPTED BUT WAS UNSUCCESSFUL BECAUSE THE PATIENT [HAD PASSED AWAY]¿. | D2B: CORRECTED. NO FURTHER INFORMATION WAS PROVIDED. THE MANUFACTURER IS CLOSING THE FILE ON THIS EVENT.
WATCHMAN TRUSTEER? ACCESS SYSTEM | K240018 | IT WAS REPORTED THAT AIR EMBOLISM AND CEREBRAL VASCULAR ACCIDENT (CVA) OCCURRED. THE PROCEDURE WAS CANCELLED. THE PATIENT DIED. A LEFT ATRIAL APPENDAGE CLOSURE (LAA) PROCEDURE WAS BEING PERFORMED USING INTRACARDIAC ECHOCARDIOGRAM (ICE) UNDER LOCAL AND CONSCIOUS SEDATION. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) WAS INSERTED AND ADVANCED INTO THE LEFT ATRIUM (LA) ALONG WITH THE ICE CATHETER. AIR WAS THEN VISUALIZED IN THE LA ON ICE, BUT NOT FLUOROSCOPY AS IT WAS NOT IN USE DURING TIME OF WAS INSERTION. THE PATIENT BECAME BRADYCARDIC. HIGH FLOW OXYGEN WAS APPLIED, AND THE HEAD OF THE PATIENT'S BED RAISED. THE PHYSICIAN CONTINUED THE PROCEDURE. A 24MM WATCHMAN FLX PRO CLOSURE DEVICE AND DELIVERY (WDS) SYSTEM WAS INSERTED AND ADVANCED TO THE LAA. THE PHYSICIAN ENCOUNTERED CHALLENGES WITH PATIENT ANATOMY AND THE LAA AND DECIDED IT WAS BEST TO ABORT THE PROCEDURE. AFTER THE CONCLUSION OF THE PROCEDURE, THE PATIENT WAS NOT RESPONDING TO NEUROLOGICAL STATUS CHECKS. A RAPID RESPONSE WAS CALLED, AND THE PATIENT HAD A COMPUTED TOMOGRAPHY (CT) SCAN PERFORMED WHICH FOUND A CVA. THE PATIENT WAS INTUBATED AND TAKEN TO THE INTENSIVE CARE UNIT (ICU). THE PLAN WAS TO TRANSFER THE PATIENT TO ANOTHER FACILITY TO TRY HYPERBARIC CHAMBER INTERVENTIONS, BUT THE PATIENT DIED ONE DAY POST INDEX PROCEDURE.
SMART PORT CT | K062414 | ON OR ABOUT (B)(6) 2023, (B)(6). UNDERWENT PLACEMENT OF AN ANGIODYNAMICS SMART PORT CT PRODUCT, MODEL NO. H787CT80STPDVI1; LOG NO. 4799362. THE DEVICE WAS IMPLANTED BY (B)(6), MD., AT (B)(6) HOSPITAL, TO BE USED FOR CHEMOTHERAPY. ON OR ABOUT (B)(6) 2023, (B)(6) PRESENTED HERSELF TO (B)(6) HOSPITAL WITH A CHIEF COMPLAINT OF PORT INFECTION, WHERE SHE WAS DIAGNOSED WITH SEPTIC SHOCK SECONDARY BLOODSTREAM INFECTION PORT-A-CATH WITH PULMONARY SEPTIC EMBOLI AND EXPIRED 24 HOURS AFTER ADMISSION INTO HOSPITAL. | THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF PATIENT SEPSIS INFECTION CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT (SAE). NO PORT DEVICE WAS RETURNED FOR EVALUATION SINCE THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION, DAMAGE OR PERFORMANCE ISSUE DURING USE. A DEVICE HISTORY RECORD (DHR) REVIEW COULD BE PERFORMED SINCE THERE WAS NO LEGITIMATE REPORTED LOT NUMBER AND REPORTED PACKAGED GOOD ITEM NUMBER WAS NOT CONFIRMED TO BE SHIPPED TO THE REPORTED END USER FACILITY. THE PORT WAS IMPLANTED INTO THE PATIENT 45 DAYS PRIOR TO THE PATIENT'S INFECTION. THERE WAS NO REPORT OF PORT DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING USE. IF THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE IN QUESTION WAS TO CONTRIBUTE TO AN INFECTION IN THE PATIENT IT WOULD HAVE LIKELY OCCURRED SHORTLY AFTER THE IMPLANT DATE AND NOT 45 DAYS LATER. DEVICE DIRECTIONS FOR USE CAUTIONS THAT EACH ACCESS OF THE PORT SHOULD BE PERFORMED USING ASEPTIC TECHNIQUE, FOLLOWING THE INSTITUTIONS UNIVERSAL PRECAUTIONS. THERE IS NO INDICATION FROM THE REPORTED COMPLAINT THAT THE MANUFACTURING/PACKAGING/STERILIZATION OF THE PORT DEVICE COULD HAVE CONTRIBUTED TO THE PATIENT'S INFECTION. INFECTION IS CAUTIONED IN THE DEVICE DFU AS A POTENTIAL COMPLICATION OF PORT SYSTEM USE. LABELING REVIEW: THE DIRECTIONS FOR USE (DFU) THAT IS PROVIDED IN THE SMARTPORT DEVICE CONTAINS THE FOLLOWING DIRECTIONS AND PRECAUTIONS: CONTRAINDICATIONS CATHETER INSERTION IN THE SUBCLAVIAN VEIN MEDIAL TO THE BORDER OF THE FIRST RIB, AN AREA WHICH IS ASSOCIATED WITH HIGHER RATES FOR PINCH-OFF. WARNINGS: THE DEVICE IS TO BE IMPLANTED, USED, MAINTAINED, AND REMOVED IN STRICT ACCORDANCE WITH INSTITUTIONAL AND OR CENTERS FOR DISEASE CONTROL (CDC) GUIDELINES OR POLICIES. REUSE OF SINGLE-USE DEVICES CREATES A POTENTIAL RISK OF PATIENT OR USER INFECTIONS. CONTAMINATION OF THE DEVICE MAY LEAD TO INJURY, ILLNESS OR DEATH OF THE PATIENT. DO NOT USE SYRINGES SMALLER THAN 10 ML SYRINGE WHEN ACCESSING THE PORT AS SYSTEM DAMAGE CAN OCCUR. FLUSHING OCCLUDED CATHETERS WITH SMALL SYRINGES CAN CREATE EXCESSIVE PRESSURES WITHIN THE PORT SYSTEM. ABSENCE OF A BLOOD RETURN OR A POOR BLOOD RETURN CAN BE A SIGN OF A POTENTIAL COMPLICATION SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, FIBRIN FORMATION, THROMBOSIS OR MALPOSITION. THIS SHOULD BE EVALUATED PRIOR TO DEVICE USAGE. A BLOOD RETURN SHOULD BE PRESENT PRIOR TO USAGE OF DEVICE FOR ANY THERAPY OR TESTING. DO NOT ATTEMPT TO MEASURE THE PATIENT'S BLOOD PRESSURE ON THE ARM IN WHICH A PERIPHERAL SYSTEM IS LOCATED, SINCE CATHETER OCCLUSION OR OTHER DAMAGE TO THE CATHETER COULD OCCUR. IF THE PATIENT COMPLAINS OF PAIN, OR THERE IS SWELLING WHEN THE DEVICE IS FLUSHED OR WHEN MEDICATION OR CONTRAST MEDIA IS ADMINISTERED, EVALUATE THE DEVICE FOR INFILTRATION, PROPER NEEDLE PLACEMENT, AND POTENTIAL COMPLICATIONS SUCH AS OCCLUSION, KINKING, BREAKAGE, PINCH-OFF SYNDROME, THROMBOSIS OR MALPOSITION. FAILURE TO ASSESS THESE COMPLAINTS OR OBSERVATIONS CAN LEAD TO DEVICE FAILURE. PRECAUTIONS CAREFULLY READ AND FOLLOW ALL INSTRUCTIONS PRIOR TO USE. STRICT ASEPTIC TECHNIQUE IS OF PARAMOUNT IMPORTANCE WHEN IMPLANTING ANY DEVICE. POTENTIAL COMPLICATIONS: USE OF AN ANGIODYNAMICS PORT SYSTEM INVOLVES POTENTIAL RISKS NORMALLY ASSOCIATED WITH THE INSERTION OR USE OF ANY IMPLANTED DEVICE OR INDWELLING CATHETER INCLUDING BUT NOT LIMITED TO: EROSION OF VESSEL AND SKIN, IMPLANT REJECTION, INFECTION, NECROSIS OF SCARRING OF SKIN OVER IMPLANT AREA, VESSEL TRAUMA. POST-OPERATIVE CARE: THE INCISION SITE SHOULD BE MONITORED FOR SIGNS OF INFECTION, INFLAMMATION, HEMATOMA, DEVICE ROTATION OR EROSION. ROUTINE WOUND CARE SHOULD BE GIVEN TO THESE SITES. THE SMART PORT CT IMPLANTABLE PORT MAY BE USED IMMEDIATELY AFTER VERIFICATION OF CATHETER PLACEMENT. INSTRUCT PATIENT TO AVOID ANY HEAVY EXERTION OR STRENUOUS ACTIVITY DURING THE FIRST FEW DAYS AFTER SURGERY. GENERAL GUIDELINES: EACH ACCESS OF AN ANGIODYNAMICS SMART PORT CT IMPLANTABLE PORT SHOULD BE PERFORMED USING ASEPTIC TECHNIQUE. FOLLOW INSTITUTIONAL UNIVERSAL PRECAUTIONS. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | ON (B)(6) 2025 IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY THAT MORNING WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE CAUSE OF DEATH WAS UNKNOWN. ADDITIONAL INFORMATION WAS PROVIDED THROUGH FOLLOW-UP WITH THE PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS FOUND DECEASED BY THEIR SPOUSE ON (B)(6) 2025. THE PATIENT WAS STILL CONNECTED TO THE LIBERTY SELECT CYCLER. IT WAS UNKNOWN WHAT PHASE AND CYCLE OF THE PATIENT¿S TREATMENT THEY WERE IN AT THE TIME OF DEATH. NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT. THE PATIENT¿S SPOUSE REPORTED THAT EVERYTHING WAS FINE PRIOR TO THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. IT WAS REPORTED THAT CARDIAC ISSUES ARE SUSPECTED. THE DEATH IS NOT RELATED TO THE USE OF ANY FRESENIUS DEVICE(S) OR PRODUCT(S) AND/OR THEIR DIALYSIS THERAPY. | PLANT INVESTIGATION: THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. ALTHOUGH NO CAUSE OF DEATH WAS DOCUMENTED, CARDIAC ISSUES WERE SUSPECTED. CARDIOVASCULAR DISEASE IS THE LEADING CAUSE OF DEATH IN PATIENTS WITH CHRONIC KIDNEY DISEASE (CKD) WITH AN ESTIMATED 40%-50% OF ALL DEATHS IN PATIENTS WITH CKD STAGE 4 AND 5 BEING ATTRIBUTABLE TO CARDIOVASCULAR-RELATED MORTALITY. THIS PATIENT HAD SEVERAL COMORBID CARDIAC DISEASES INCLUDING CHRONIC SYSTOLIC HEART FAILURE. BASED ON THE AVAILABLE INFORMATION AND NO EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
ENSEAL X1 CURVED 25CM SHAFT | K172580 | IT WAS REPORTED THAT DURING A THIN VESSEL WAS SEALED WITH COVE AND AFTER 24H MASSIVE BLEEDING OCCURRED. WHEN OPENING THE PATIENT, IT WAS IDENTIFIED THAT THE SEALED VESSEL WAS OPEN. THE PATIENT DIED. | (B)(4). DATE SENT: 4/22/2025. D4 BATCH #: UNKNOWN. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: WHAT WERE THE INDICATIONS FOR SURGERY? LUNG NODULES WERE THE HEMOSTASIS ISSUES IDENTIFIED POST OPERATIVELY OR INTRA OPERATIVELY? POST-OPERATIVE. WERE ANY SURGICAL CLIPS USED IN THE AREA OF THE DEVICE FIRING? AFTER OPENING THE SHEARS, IT WAS NECESSARY TO PLACE A CLIP. WHAT WAS THE TOTAL ESTIMATED BLOOD LOSS (ML)? BLEEDING AFTER OPENING THE SEAL IS ESTIMATED AT 2 LITERS. WAS A TRANSFUSION REQUIRED? YES. WHAT WAS THE PRE AND POSTOPERATIVE ANTI-COAGULANT AND PAIN MANAGEMENT PROTOCOL? NO INFORMATION. WAS THERE CALCIFICATION OF TISSUE THAT IMPEDED CLAMPING OR CUTTING? NO INFORMATION WERE THERE ANY PRE-EXITING PATIENT CONDITIONS? COAGULOPATHY PLEASE PROVIDE A TIMELINE OF EVENTS. AFTER THE CLAMPING OF THE VESSELS AND ARTERY, AFTER 1 HOUR OF THE PROCEDURE THE ARTERY CLAMPING OPENED. TWO CLIPS WERE PLACED IN THE SITE AND THERE WAS HEMOSTASIS. IN THE CASE OF THE SEAL OF THE ENSEAL, AFTER 24 HOURS BLEEDING AND OPENING OF THE SEAL WERE IDENTIFIED. IS THERE AN AUTOPSY REPORT AVAILABLE? NO. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE LOT NUMBER, AND NO NON-CONFORMANCE WERE IDENTIFIED. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
HARMONIC 1100 SHEARS 36CM | K200841 | IT WAS REPORTED THAT THE PATIENT UNDERWENT SURGERY FOR DEEP INFILTRATIVE ENDOMETRIOSIS WITH DISCOID RESECTION OF THE COLON. DURING THE SURGERY, A HARMONIC AND A CIRCULAR STAPLER WERE USED. LATER, THE PATIENT WAS SCHEDULED FOR EMERGENCY SURGERY DUE TO AN UNDETECTED LEAK IN THE SIGMOID COLON. THEY SUFFERED CARDIAC ARREST FOUR TIMES DURING THE PROCEDURE, BUT STILL MANAGED TO SURVIVE. THEY DIED POSTOPERATIVELY. THE CAUSE OF THE LEAKAGE WAS ATTRIBUTED TO A THERMAL LESION IN THE SIGMOID COLON THAT OCCURRED DURING THE MANIPULATION AND RELEASE OF STRUCTURES DURING THE FIRST SURGERY. | (B)(4). DATE SENT: 7/9/2025. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SALES REP WAS PRESENT IN THE FIRST SURGERY THAT WAS PROGRAMED AND AS PART OF AN EVENT. THIS SURGERY WAS COMPLETED SUCCESSFULLY WITHOUT ANY ISSUES OR COMPLICATIONS. THERE WERE NO GENERATOR ALARMS. THE LEAK HAPPENED 2-3 DAYS AFTER THE SURGERY. THE SURGEON FOUND THE BURN THAT HAD NOT BEEN SEEN DURING THE FIRST PROCEDURE AND HE THINKS THAT IT COULD HAVE BEEN CAUSED BY THE HARMONIC. SHE SAID THAT SURGEONS HAVE RECEIVED TRAINING IN HARMONIC AND WERE AWARE THAT THE DEVICE HEATS UP IN THE DISTAL PART DURING ACTIVATION.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT PASSED AWAY ONE YEAR AGO. A REPLACEMENT FOR THE REMSTAR PRO C-FLEX+ DEVICE WAS DECLINED, AND THE PATIENT'S SISTER HAS REQUESTED TO BE REMOVED FROM THE LIST. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
UMBILICAL CATHETER | K921352 | A PREMATURE NEWBORN 34 WEEKS + 2 DAYS, WITH A BIRTH WEIGHT OF 1480G, WAS TRANSFERRED TO NEONATAL INTENSIVE CARE AT A SECOND HOSPITAL ON 1 DAY OF LIFE FOR SURGICAL MANAGEMENT OF AN ANORECTAL MALFORMATION WITH RECTOVAGINAL FISTULA. ON ADMISSION, HER CLINICAL CONDITION WAS STABLE AND PARENTERAL NUTRITION IS ADMINISTERED THROUGH UMBILICAL VENOUS CATHETER INSERTED AT BIRTH BY THE FIRST HOSPITAL. THE DAY AFTER ADMISSION, THE PATIENT IS OPERATED ON TO DILATE THE FISTULA. THE OPERATION IS GOING WELL AND ALLOWS EVACUATION OF FAECES. ON DAY 2 OF ADMISSION, SHE SUFFERED A CARDIORESPIRATORY ARREST DUE TO TAMPONADE OF IATROGENIC ORIGIN (PERFUSO- PERICARDIUM RELATED TO MALPOSITION OF THE UMBILICAL VENOUS CATHETER). THE EVOLUTION IS UNFAVORABLE DESPITE CARDIOPULMONARY RESUSCITATION STARTED EARLY AND PERICARDIAL DRAINAGE. THE PREMATURE DIED. | THE INVOLVED UMBILICAL CATHETER IS NOT AVAILABLE. FROM THE DESCRIPTION, THIS SERIOUS ADVERSE EVENT IS CLEARLY RELATED TO THE MALPOSITION OF THE UMBILICAL CATHETER. IT IS NOT TRACED TO A DEFECT OF OUR DEVICE. THE BATCH REVIEW DOES NOT SHOW ANY NON-CONFORMITY. THE PRODUCTS ARE IN CONFORMITY WITH THE SPECIFICATIONS. ALL CONTROLS ARE COMPLIANT, INCLUDING THE 100% TESTS CARRIED OUT ON EACH CATHETER DURING THE MANUFACTURING PROCESS. THESE TESTS ARE AS FOLLOWS: THE BLUNTNESS OF THE CATHETER TIP, MARKING THE CATHETERS, TIGHTNESS AND NON-OBSTRUCTION. THE HISTORICAL DATA ANALYSIS OF COMPLAINTS THESE LAST 3 YEARS, ON THIS REFERENCE DOES NOT SHOW ANY SIMILAR INCIDENT. AT THIS TIME, NO CORRECTIVE ACTION HAS BEEN INITIATED, AS THERE IS NO EVIDENCE INDICATING A CONNECTION BETWEEN THE ISSUE AND THE DEVICE.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION FROM A DEALER ON 26 MARCH 2025 THAT THEY ARE REQUESTING ANALYSIS OF A TRILOGY EVO VENTILATOR FOR DETAILS OF RECORDED DATA BETWEEN 14:00 AND 15:00 ON (B)(6) 2025 ASSOCIATED WITH "AN INCIDENT". THEY REQUESTED TO KNOW WHAT ALARM(S) OCCURRED DURING THIS RELEVANT TIME AND THE STATUS OF ANY EVENTS. ON 27 MARCH 2025, INFORMATION WAS PROVIDED BY THE SALES REPRESENTATIVE THAT THE PATIENT HAD DIED. NO PATIENT INFORMATION WAS PROVIDED. IN REGARD TO THE PATIENT DEATH, THE SALES REPRESENTATIVE STATED THERE WAS NO CAUSAL RELATIONSHIP WITH THE MEDICAL DEVICE (VENTILATOR), AND NO INFORMATION ABOUT ANY MALFUNCTION OF THE DEVICE. THE CAUSE OF THE PATIENT'S DEATH WAS REPORTED TO HAVE BEEN DETERMINED TO BE THE DELAYED RESPONSE BY THE HEALTHCARE PROVIDERS, AND THE HOSPITAL IS CURRENTLY CONSIDERING MEASURES TO ADDRESS THE HEALTHCARE PROVIDER'S RESPONSE. AT THIS TIME, THERE IS NO ALLEGATION OF A DEVICE DEFECT, MALFUNCTION OR FAILURE. IT HAS BEEN INDICATED THAT THE VENTILATOR IS AVAILABLE FOR RETURN TO THE MANUFACTURER FOR A DEVICE EVALUATION; HOWEVER, THE VENTILATOR HAS NOT YET BEEN RECEIVED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE TYPE OF REPORTABLE EVENT FIELD HAS BEEN UPDATED FROM "PRODUCT PROBLEM" TO APPROPRIATELY REFLECT "DEATH" AS DESCRIBED IN THE REPORT NARRATIVE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. BASED ON AVAILABLE INFORMATION AT THIS TIME, THE ALLEGED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. NO FURTHER ACTION IS REQUIRED. IF ADDITIONAL INFORMATION IS OBTAINED ABOUT THIS EVENT, THE PMS CLINICAL EXPERT WILL REVIEW. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. DURING EVALUATION AN ERROR INDICATING A REBOOT OF THE CENTRAL PROCESSING UNIT (CPU). PER THE DREAMSTATION SERVICE AND TECHNICAL REFERENCE MANUAL, THIS LEVEL ERROR DOES NOT AFFECT DEVICE FUNCTIONALITY. THE SERVICE TECHNICIAN RECOMMENDED THE DEVICE BE SCRAPPED. NO FURTHER INVESTIGATION DETAILS WERE PROVIDED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
COOLSEAL¿ TRINITY | K211579 | IT WAS REPORTED THAT DURING A COOLSEAL TRINITY PROCEDURE ON (B)(6) 2024, A COOLSEAL TRINITY 30CM AND COOLSEAL GENERATOR WERE USED DURING A THORACOSCOPIC PROCEDURE ON A CHILD (6 YEARS OLD, FEMALE), DISSECTING AND SEALING VESSELS. DURING THE PROCEDURE IT WAS IDENTIFIED THAT THE PATIENT WAS BLEEDING FROM A PREVIOUSLY DISSECTED AND SEALED BRANCH OF THE PULMONARY ARTERY, AND THERE WERE CONCERNS OF A POTENTIAL SEAL FAILURE. THE PATIENT SUFFERED SIGNIFICANT BLEEDING WHICH COULD NOT BE STOPPED, RESULTING IN PATIENT DEATH. ON A FOLLOW UP EMAIL IT WAS DETERMINED THAT THE BLEEDING STARTED 3 HOURS AFTER THE SURGERY BEGAN, MAJOR BLEEDING WAS ENCOUNTERED FROM A CORNER OF A PREVIOUSLY SEALED AND DIVIDED BASAL ARTERIAL BRANCH. THE PROCEDURE WAS A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY. NO PRE-EXISTENT CONDITIONS WERE REPORTED. MULTIPLE STEPS WERE TAKEN AFTER THE BLEEDING WAS OBSERVED TO TREAT IT, COMPRESSION, TRANEXAMIC ACID AND BLOOD TRANSFUSION. NO OTHER INFORMATION WAS RECEIVED. | H6: APPROPRIATE TERM/CODE NOT AVAILABLE: SUGGESTED CODE : ELECTROSURGICAL CUTTER AND COAGULATOR. LOT NUMBER OF THE DEVICE NOT PROVIDED BY THE COMPLAINANT; THEREFORE, THE UDI, EXPIRATION AND MANUFACTURING DATES ARE NOT KNOWN. DEVICE HISTORY RECORD (DHR) REVIEW WAS UNABLE TO BE CONDUCTED FOR THE DISPOSABLE DEVICE AS THE IDENTIFICATION NUMBERS WERE NOT PROVIDED BY THE COMPLAINANT. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING INFORMATION WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT FOR A THORACIC AORTIC ANEURYSM USING NON-GORE STENT GRAFTS (VALIANT THORACIC STENT GRAFT, MEDTRONIC). THE FIRST VALIANT STENT GRAFT WAS INSERTED FROM THE RIGHT FEMORAL ARTERY, BUT IT COULD NOT BE ADVANCED THROUGH THE RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY. ONCE THE DEVICE WAS REMOVED, A 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH WAS INSERTED FROM THE RIGHT FEMORAL ARTERY AND SUCCESSFULLY ADVANCED TO THE LEVEL OF THE RENAL ARTERIES. THE DEVICE WAS THEN INSERTED THROUGH THE SHEATH, AND TWO VALIANT STENT GRAFTS WERE IMPLANTED IN THE DESCENDING AORTA AS PLANNED. ACCESS ANGIOGRAPHY REVEALED VASCULAR INJURY NEAR THE RIGHT EXTERNAL ILIAC ARTERY AND RIGHT COMMON ILIAC ARTERY. THE PATIENT WENT INTO SHOCK AND A CARDIAC ARREST. CARDIAC MASSAGE WAS INITIATED, AND A GORE® VIABAHN® ENDOPROSTHESIS WITH HEPARIN BIOACTIVE SURFACE (VIABAHN DEVICE) WAS USED AS TREATMENT. THE FIRST VIABAHN (10 MM X 10 CM) WAS IMPLANTED AT THE SITE OF VASCULAR INJURY (RIGHT EXTERNAL ILIAC ARTERY-RIGHT COMMON ILIAC ARTERY), AND THE SECOND VIABAHN (9 MM X 10 CM) WAS IMPLANTED SLIGHTLY PROXIMAL TO THE LEVEL OF THE FEMORAL HEAD. THE NATIVE VESSEL WAS REPORTEDLY ABOUT 8 MM IN DIAMETER. AFTER PLACEMENT, A TYPE III ENDOLEAK WAS OBSERVED AT THE OVERLAP SITE OF THE TWO VIABAHNS. BALLOON MOLDING WAS PERFORMED, BUT THE LEAK DID NOT RESOLVE. WHEN AN 8 FR OR 10 FR WAS INSERTED IN ORDER TO DELIVER THE THIRD VIABAHN TO THE LEAK SITE, THE OVERLAP OF THE TWO VIABAHNS BECAME UNINTENTIONALLY DISLODGED. THE VIABAHN IMPLANTED ON THE PROXIMAL SIDE WAS COMPLETELY PUSHED OUT OF THE VESSEL. THE GUIDEWIRE WAS ALSO PULLED OUT OF THE TRUE LUMEN, MAKING RE-CANNULATION IMPOSSIBLE. A BALLOON OCCLUSION WAS PERFORMED FROM THE CONTRALATERAL SIDE, AND THE NEXT COURSE OF TREATMENT WAS BEING CONSIDERED. THE PATIENT REMAINED IN CARDIAC ARREST, AND CARDIAC MASSAGE WAS CONTINUED. BLEEDING FROM THE INTUBATION TUBE WAS ALSO OBSERVED. AT THIS POINT, TWO HOURS HAD PASSED SINCE THE CARDIAC ARREST, AND THE HEARTBEAT HAD NOT RESUMED. THE PHYSICIAN DECIDED TO DISCONTINUE RESUSCITATION, AND THE PATIENT WAS DECEASED. <ADDITIONAL INFORMATION OBTAINED FROM THE GORE FSA ON APRIL 11, 2025> THE CAUSE OF DEATH IS UNKNOWN. THE PHYSICIAN'S COMMENT: THE CAUSAL RELATIONSHIP BETWEEN THE PATIENT'S DEATH AND GORE DEVICES (SHEATH/VIABAHN) IS UNKNOWN. REGARDING THE VASCULAR INJURY, THERE WAS SIGNIFICANT RESISTANCE DURING INSERTION OF THE FIRST VALIANT STENT GRAFT, AND IT WAS NOT POSSIBLE TO ADVANCE THROUGH THE AREA NEAR THE RIGHT EXTERNAL AND RIGHT COMMON ILIAC ARTERIES. THERE WAS NO RESISTANCE DURING INSERTION OF THE 24 FR GORE® DRYSEAL FLEX INTRODUCER SHEATH, SUGGESTING THAT THE VASCULAR INJURY LIKELY OCCURRED DURING THE INSERTION OF THE FIRST STENT GRAFT. AS THE HEARTBEAT DID NOT RESUME EVEN AFTER PLACEMENT OF THE VIABAHN, IT IS LIKELY THAT REGARDLESS OF WHETHER THE VIABAHN COULD HAVE CONTROLLED THE BLEEDING IT WAS ALREADY TOO LATE. | A1: NO PATIENT SPECIFIC DETAILS HAVE BEEN PROVIDED. THEREFORE, THE PATIENT INITIALS REFLECT THE W. L. GORE INTERNAL CASE NUMBER. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
COOLSEAL¿ GENERATOR | K211579 | IT WAS REPORTED THAT DURING A COOLSEAL TRINITY PROCEDURE ON (B)(6) 2024, A COOLSEAL TRINITY 30CM AND COOLSEAL GENERATOR WERE USED DURING A THORACOSCOPIC PROCEDURE ON A CHILD (6 YEARS OLD, FEMALE), DISSECTING AND SEALING VESSELS. DURING THE PROCEDURE IT WAS IDENTIFIED THAT THE PATIENT WAS BLEEDING FROM A PREVIOUSLY DISSECTED AND SEALED BRANCH OF THE PULMONARY ARTERY, AND THERE WERE CONCERNS OF A POTENTIAL SEAL FAILURE. THE PATIENT SUFFERED SIGNIFICANT BLEEDING WHICH COULD NOT BE STOPPED, RESULTING IN PATIENT DEATH. ON A FOLLOW UP EMAIL IT WAS DETERMINED THAT THE BLEEDING STARTED 3 HOURS AFTER THE SURGERY BEGAN, MAJOR BLEEDING WAS ENCOUNTERED FROM A CORNER OF A PREVIOUSLY SEALED AND DIVIDED BASAL ARTERIAL BRANCH. THE PROCEDURE WAS A VIDEO ASSISTED THORACOSCOPIC LOBECTOMY SURGERY. NO PRE-EXISTENT CONDITIONS WERE REPORTED. MULTIPLE STEPS WERE TAKEN AFTER THE BLEEDING WAS OBSERVED TO TREAT IT, COMPRESSION, TRANEXAMIC ACID AND BLOOD TRANSFUSION. NO OTHER INFORMATION WAS RECEIVED. | A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE REPORTED LOT/SERIAL NUMBER. THE DEVICE WAS RELEASED MEETING ALL QA SPECIFICATIONS. THE DEVICE WAS RETURNED FOR INSPECTION WHICH FOUND NO CONCERNS AND MET THE MANUFACTURER´S SPECIFICATIONS.
CIRCULAR MECH STD 29MM | K983536 | IT WAS REPORTED THAT THE PATIENT UNDERWENT SURGERY FOR DEEP INFILTRATIVE ENDOMETRIOSIS WITH DISCOID RESECTION OF THE COLON. DURING THE SURGERY, A HARMONIC AND A CIRCULAR STAPLER WERE USED. (B)(6), THE PATIENT WAS SCHEDULED FOR EMERGENCY SURGERY DUE TO AN UNDETECTED LEAK IN THE SIGMOID COLON. HE SUFFERED CARDIAC ARREST FOUR TIMES DURING THE PROCEDURE, BUT STILL MANAGED TO SURVIVE. HE DIED POSTOPERATIVELY. THE CAUSE OF THE LEAKAGE WAS ATTRIBUTED TO A THERMAL LESION IN THE SIGMOID COLON THAT OCCURRED DURING THE MANIPULATION AND RELEASE OF STRUCTURES DURING THE FIRST SURGERY. | (B)(4). DATE SENT: 4/23/2025. D4 BATCH #: UNKNOWN. INVESTIGATION SUMMARY: AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: ON WHAT ANATOMICAL LOCATION WAS THE THERMAL INJURY IDENTIFIED POST-OPERATIVELY? WHAT WAS THE DISTANCE BETWEEN THE ANASTOMOTIC SITE AND THERMAL LESION? DOES THE SURGEON BELIEVE THERE WAS AN ALLEGED DEFICIENCY OF THE CDH29B THAT MAY HAVE CAUSED OR CONTRIBUTED TO THE THERMAL INJURY, POST-OPERATIVE LEAK, AND/OR PATIENT¿S DEATH? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
FUSION HOLLOW FIBER OXYGENATOR | K240666 | IT WAS REPORTED THAT DURING A CARDIOPULMONARY BYPASS PROCEDURE OF LESS THAN A 6 HOUR DURATION, USING THE FUSION OXYGENATOR, STRAW-COLORED FLUID AND FOAM WERE OBSERVED COMING FROM THE GAS EXHAUST OF THE OXYGENATOR. THE CUSTOMER DESCRIBED THIS AS PLASMA LEAKAGE. CIRCULATORY ARREST OCCURRED DURING THE PROCEDURE, AND THE OXYGENATOR WAS CHANGED OUT AND REPLACED WITH ANOTHER OXYGENATOR. THE REST OF THE ORIGINAL CUSTOM TUBING PACK WAS USED. THERE WAS MINIMAL BLOOD LOSS OF 250MLS, AND NO BLOOD TRANSFUSION WAS REQUIRED. NO DAMAGE WAS SEEN TO ANY COMPONENT OR PACKAGING. THE PROCEDURE WAS AN EMERGENCY DURING THE NIGHT. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CLINICIAN DESCRIBED THE LEAK AS A PLASMA LEAK. THE PROCEDURE COMMENCED ON (B)(6) 2025 AND CONTINUED INTO (B)(6) 2025. ALL ACTIVATED COAGULATION TIMES (ACT) WERE WITHIN PROTOCOL. FOR THE SECOND CARDIOPUMONARY BYPASS (CPB) RUN, ALL ACTS WERE GREATER THAN 1000 SECONDS. THE CIRCULATORY ARREST WAS NOT PLANNED. THE CLINICIAN COOLED THE PATIENT TO 18¿ IN A CONTROLLED TIMEFRAME TO USE CIRCULATORY ARREST TO REPLACE THE OXYGENATOR. IT WAS AN ELECTIVE AORTIC VALVE REPLACEMENT (AVR). THERE WERE NO KNOWN COAGULATION CONDITIONS. THE PRIMING INCLUDED PLASMALYTE, GELOFUSINE, MANNITOL AND 10 ,000IU HEPARIN. IT WAS AN INITIAL BYPASS RUN OF 253 MINUTES. BYPASS WAS TERMINATED AND THE PUMP WAS THEN CIRCULATING FOR 2 HOURS WHEN THE SURGEON DEEMED IT NECESSARY TO RECOMMENCE CPB. AT 30 MINUTES INTO THE SECOND CPB RUN, THE OXYGENATOR STARTED TO LEAK AND AN INCREASED SWEEP SPEED WAS REQUIRED. THE SURGEON STATED THAT THE OPERATION WAS LIKELY TO CONTINUE FOR SEVERAL HOURS, THEREFORE, THE DECISION WAS MADE TO COOL DOWN AND CHANGE THE OXYGENATOR OUT. THE TOTAL SECOND CPB WAS 294 MINUTES. IT TOOK 30 MINUTES TO COOL, CHANGE OUT THE OXYGENATOR AND CONTINUE BYPASS WITH A NEW OXYGENATOR. THE PATIENT PASSED AWAY IN THE OPERATING THEATRE FROM SURGICAL COMPLICATIONS. THE CLINICIAN BELIEVED THAT IT WAS HIGHLY UNLIKELY THAT CHANGING OUT THE OXYGENATOR AFFECTED THE OUTCOME, HOWEVER, THEY CAN NOT CONFIRM IT WITH 100% CERTAINTY. THE PRESSURES AND FLOWS WITHIN THE CPB CIRCUIT INCLUDING THE OXYGENATOR WERE ALL WITHIN NORMAL PARAMETERS. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE CAUSE OF DEATH HAD NOT BEEN IDENTIFIED BUT IT WAS DESCRIBED AS COMPLICATIONS ARISING FROM SURGERY. IT WAS NOT THOUGHT TO BE DIRECTLY LINKED TO THE OXYGENATOR. MEDTRONIC RECEIVED ADDITIONAL INFORMATION THAT THE FIRST PUMP RUN WAS A STANDARD CASE. THE SECOND PUMP RUN WAS REQUIRED DUE TO COMPLICATIONS WITH THE PATIENT REQUIRING A SECOND PROCEDURE. THE SECOND PUMP RUN WAS STARTED WITH THE ORIGINAL CIRCUIT AND OXYGENATOR. AFTER 30 MINUTES OF THE SECOND PUMP RUN, THE OXYGENATOR SHOWED SIGNS OF FAILING. THE PATIENT WAS COOLED AND THE OXYGENATOR WAS CHANGED. THE OPERATION WAS COMPLETED WITH THE SECOND OXYGENATOR WHICH HAD NO ISSUES. THEY WERE NOT ON PATIENT BYPASS FOR 120 MINUTES. THEY WERE ON STANDBY RECIRCULATING THROUGH THE RECIRCULATION LINE WITH DILUTE BLOOD/ HARTMANNS MIX. THE START OF SECOND BYPASS RUN WAS 30 MINUTES. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
HLS SET | K112360 | COMPLAINT ID # (B)(4). | THE EVENT OCCURRED IN THE US. IT WAS REPORTED THAT DURING PATIENT TREATMENT ISSUES WERE PRESENT WITH ERRONEOUS PRESSURE READINGS (ARTERIAL, INTERNAL AND DELTA) FROM AN HLS-SET. NEW INFORMATION HAS BEEN RECEIVED ON 2025-05-14 BY THE SSU/CUSTOMER, THAT THE INVOLVED PATIENT PASSED AWAY. THE PATIENT¿S FAMILY DECIDED TO DISCONTINUE THE TREATMENT, RESULTING IN THE PATIENT PASSING AWAY. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. THE AFFECTED PRODUCT WAS NOT AVAILABLE FOR TECHNICAL INVESTIGATION AS IT WAS DISCARDED BY THE CUSTOMER. THUS, THE EXACT ROOT CAUSE REMAINS UNKNOWN. HOWEVER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-06-02 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES INACCURATE PRESSURE READINGS FROM A PRE-PRIMED HLS SET USED DURING ECMO THERAPY. THE CHIEF OF PERFUSION AT THE TREATING INSTITUTION, SUGGESTED DURING A CONFERENCE THAT FLUID MAY HAVE ENTERED THE WHITE CONNECTOR OF THE HLS SET PRIOR TO THERAPY, POSSIBLY CONTRIBUTING TO THE MALFUNCTION. THE TEAM REPORTEDLY ATTEMPTED TO CLEAN THE AREA BEFORE USE. ECMO THERAPY BEGAN ON (B)(6) 2025. NO ISSUES WERE NOTED DURING INITIAL PRIMING WITH PLASMALYTE, AND NO DEPOSITS OR CRYSTAL FORMATIONS WERE SEEN AT THE SENSOR PORT WHICH MAY BE IN CONFLICT WITH THE PREVIOUS SUGGESTION OF FLUID INTRUSION IN THE PORT THE ABSENCE OF VISIBLE DEPOSITS OR RESIDUE AT THE TIME OF SETUP MAY SUGGEST THAT INTRUSION MAY HAVE OCCURRED AFTER ECMO INITIATION AS INTRUSION WAS NOT IDENTIFIED DURING INITIAL DEVICE INSPECTION. WHILE THE CIRCUIT APPEARED CLEAN AT INITIATION, THE DISCOVERY OF BLOOD UPON DISASSEMBLY RAISES QUESTIONS ABOUT WHETHER CONTAMINATION MAY HAVE DEVELOPED SOMETIME DURING SUPPORT. THE REPORTED PRESSURE ISSUE WAS FIRST OBSERVED ON (B)(6), RESULTING IN THE REPLACEMENT OF THE SENSOR PROBE (CABLE) ON (B)(6). THE REPORTED BEHAVIOR SUGGESTS THE ROOT CAUSE MAY LIE IN EITHER THE CABLE ITSELF OR THE INTERNAL SENSOR PORT OF THE HLS SET. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER FOUND NEAR THE POST-OXYGENATOR MEMBRANE. LAB VALUES ON FEBRUARY 1 INDICATED STABLE PARAMETERS. THE HLS SET WAS NOT REPLACED, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES REPORTEDLY CORRELATED WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING BLOOD SAMPLING. NO REDUCTION IN BLOOD FLOW WAS OBSERVED. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER IDENTIFIED NEAR THE POST-OXYGENATOR MEMBRANE AT THE END OF THERAPY. LABORATORY VALUES OBTAINED ON FEBRUARY 1 SHOWED A HEMOGLOBIN LEVEL OF 8.6 G/DL, HEMATOCRIT OF 28%, D-DIMER CONCENTRATION OF 12.21 MICROGRAM/ML, PLATELET COUNT OF 126 × 1 BILLION/L, AND A PTT OF 26.7 SECONDS. WHILE THESE VALUES APPEARED RELATIVELY STABLE FOR AN ECMO PATIENT, THE MARKEDLY ELEVATED D-DIMER LEVEL MAY SUGGEST EARLY THROMBOTIC ACTIVITY WITHIN THE CIRCUIT, DESPITE THE ABSENCE OF VISIBLE CLOTS AT THAT TIME (1, 2). ACCORDING TO RAJSIC ET AL. (2022) (3), THE RECOMMENDED THERAPEUTIC RANGE FOR THE PREVENTION OF VENOUS THROMBOEMBOLISM IS 1.5 TO 2.5 TIMES THE PATIENT¿S BASELINE APTT (40¿80 SECONDS). A PTT OF 26.7 SECONDS MAY HAVE CONTRIBUTED TO THROMBUS FORMATION WITHIN THE ECMO CIRCUIT. THE HLS SET ITSELF WAS NOT REPLACED DURING THE ECMO RUN, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES WERE OBSERVED IN CORRELATION WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING PRE- AND POST-OXYGENATOR BLOOD SAMPLING. DESPITE THESE ANOMALIES, NO REDUCTION IN BLOOD FLOW WAS REPORTED, AND THE CIRCUIT CONTINUED TO PROVIDE EFFECTIVE SUPPORT. IN CONCLUSION, A PRODUCT MALFUNCTION INVOLVING THE INTERNAL PRESSURE SENSOR OF THE HLS SET WAS LIKELY, AS EVIDENCED BY PERSISTENT INACCURATE PRESSURE READINGS DESPITE CABLE REPLACEMENT. HOWEVER, THIS MALFUNCTION DID NOT RESULT IN A LOSS OF ECMO SUPPORT OR REDUCTION IN BLOOD FLOW. THAT SAID, REDUNDANT PRESSURE MONITORING DURING WAS, ASSUMEDLY, PRESENT VIA THE PART PRESSURE FROM THE HLS DISPOSABLE. NOTABLY, THE FIRST PRESSURE ANOMALY WAS REPORTED ON (B)(6)¿NINE DAYS AFTER ECMO INITIATION ON (B)(6)¿INDICATING THAT THE HLS SET INITIALLY PERFORMED AS EXPECTED. THE SUBSEQUENT DISCOVERY OF BLOOD RESIDUES ON THE DEVICE UPON DISASSEMBLY SUGGESTS THAT CONTAMINATION OF THE INTERNAL PRESSURE SENSOR PORT MAY HAVE OCCURRED WHILE THERAPY, POTENTIALLY CONTRIBUTING TO THE OBSERVED BEHAVIOR. THE PATIENT¿S CLINICAL COURSE ULTIMATELY ENDED WITH WITHDRAWAL OF CARE AT THE FAMILY¿S REQUEST. WHILE THROMBOTIC EVENTS MAY HAVE CONTRIBUTED TO OVERALL MORBIDITY, THE DECISION TO WITHDRAW CARE APPEARS TO BE INFLUENCED BY BROADER CLINICAL OR ETHICAL CONSIDERATIONS. BASED ON AVAILABLE INFORMATION, THE PATIENT¿S EXPIRATION WAS NOT INFLUENCED BY THE DEVICE MALFUNCTION. THE FOLLOWING IS A QUOTE FROM THE INITIATING COMPLAINT INDICATING APPARENT SYNCHRONY WITH THE ABOVE OPINION ON DEVICE INFLUENCE: "ECLS COORDINATOR AT (B)(6) HOSPITAL, MENTIONED THEY ARE HAVING ISSUES WITH ERRONEOUS PRESSURE READINGS FROM AN HLS ON A PATIENT. IT IS NOT NEGATIVELY IMPACTING THE PATIENT OR LEVEL OF CARE, NO PATIENT HARM." ACCORDING TO THE INSTRUCTION FOR USE (HLS SET, CHAPTER "PREPARATION AND INSTALLATION") THE PRESSURE SENSORS HAVE TO BE CHECKED BEFORE PRIMING. IN GENERAL, IT SHOULD BE NOTED THAT PRE-PRIMING CAN HAVE A NEGATIVE IMPACT ON THE SENSOR FUNCTION AND THE ASSOCIATED PRESSURE VALUES AND CAN LEAD TO THE OUTPUT OF UNEXPECTED / IMPLAUSIBLE PRESSURE VALUES SHORTLY BEFORE USE, AS INDICATED IN CHAPTER "PRIMING THE SYSTEM", WHERE IT SAYS "ONLY FILL THE SYSTEM SHORTLY BEFORE USE AND IN ACCORDANCE WITH THE PRIMING INSTRUCTIONS. CALIBRATION VALUES ARE LOST WHEN THE DRIVE UNIT IS SWITCHED OFF". FURTHER THE CARDIOHELP HAS A FLOW/BUBBLE SENSOR AND A VENOUS PROBE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. IN ADDITION, WITHIN THE CHAPTER ¿PRIMING THE SYSTEM¿ IT IS WRITTEN NOT TO PRE-PRIME THE CIRCUIT TOO EARLY. BASED ON THE RESULTS THE REPORTED FAILURE "PRESSURE ISSUES" COULD NOT BE CONFIRMED. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-04. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
HIGH FLOW INSUFFLATION UNIT | K243527 | IT WAS REPORTED THAT DURING A LAPAROSCOPIC SURGERY FOR DIVERTICULAR DISEASE, AFTER THE INTRODUCTION OF OPTIC TROCAR, INDUCTION OF PNEUMOPERITONEUM WITH CARBON DIOXIDE (CO2) AT 12 MMHG, PLACEMENT OF THREE ADDITIONAL TROCARS, EXPLORATION OF THE ABDOMINAL CAVITY, AND LYSIS OF ADHESIONS, THE PATIENT EXPERIENCED A DROP IN BLOOD PRESSURE AND CARDIAC ARREST, WHICH RESULTED IN DEATH. THE PATIENT'S FAMILY AGREED ONLY TO A POSTMORTEM COMPUTED TOMOGRAPHY (CT) SCAN, WHICH DOCUMENTED A PICTURE OF SYSTEMIC GAS EMBOLISM AND ABSENCE OF VASCULAR LESIONS. THE DEVICE WAS CHECKED BEFORE STARTING THE PROCEDURE WITHOUT DETECTING ANY ANOMALY. IT DID NOT GIVE ANY ERROR CODE UNTIL PATIENT¿S DEATH. NO OVERPRESSURE EVENT OCCURRED, AND NO DEVICE MALFUNCTION WAS REPORTED DURING THE PROCEDURE. IT WAS NOTED THAT CHECKS WERE PERFORMED ON THE DEVICES USED DURING THE PROCEDURE, AND THE RESULTS WERE IN NORMAL RANGE. THE CUSTOMER HYPOTHESIZED ROLE OF GRADIENT PARTIAL PRESSURE OF INSUFFLATED GAS. THE HIGH-FLOW INSUFFLATION UNIT AND THE NON-OLYMPUS ANESTHESIA VENTILATOR USED DURING THE PROCEDURE HAVE BEEN PLACED UNDER QUARANTINE. | D10: FRAME TROLLEY S/5 AISYS MONITOR RACK UPPER 3 DWR SN (B)(6) GE HEALTHCARE. ON MARCH 25, 2025, AN OLYMPUS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED TO THE USER FACILITY AND CHECKED THE DEVICE ON SITE IN THE PRESENCE OF THE HOSPITAL CLINICAL ENGINEER. NO ISSUE WAS FOUND. ON MARCH 27, 2025, AN OLYMPUS FSE WAS AT THE CUSTOMER'S SITE AGAIN FOR A DEEPER CHECK OF THE SUBJECT DEVICE WITH DEDICATED TOOLS. THE DEVICE PASSED ALL TESTS; NO FAILURES OR DEFECTS WERE FOUND. ALL CHECKS WERE PERFORMED IN THE PRESENCE OF THE HOSPITAL CLINICAL ENGINEER. OLYMPUS WAS NOT ABLE TO DETECT A FAILURE. THE DEVICE MEETS ITS SPECIFICATIONS. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
CONFIDENCE KIT, NO NEEDLES | K060300 | ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS. | DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: G1 (MANUFACTURING SITE NAME), H4. H3, H6: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. DEVICE HISTORY REVIEW (DHR): A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE. PRODUCT CODE: 283913000. LOT NUMBER: 424341. IT WAS ELECTRONICALLY REVIEWED AND NO NONCONFORMANCES/MANUFACTURING IRREGULARITIES WERE IDENTIFIED DURING THE MANUFACTURING PROCESS. THE PRODUCT WAS RELEASED ON: 13 FEB 2025. MANUFACTURING SITE: (B)(4). EXPIRY DATE: 31 JAN 2027. CEMENT NUMBER: PRODUCT CODE: 183901001/ BATCH NUMBER: 4597805. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
CONFIDENCE KIT, NO NEEDLES | K060300 | ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS. | DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: B5 AND D10. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
CONFIDENCE KIT, NO NEEDLES | K060300 | ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS. | DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: B5 AND D10. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
CONFIDENCE KIT, NO NEEDLES | K060300 | ADDITIONAL INFORMATION RECEIVED STATES THAT THE ORIGINAL KYPHOPLASTY DID INVOLVE DEPUY SYNTHES CEMENT AND DEPUY SYNTHES SCREWS. | DEPUY SYNTHES IS SUBMITTING THIS REPORT PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH DEPUY SYNTHES HAS NOT BEEN ABLE TO INVESTIGATE OR VERIFY PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY FDA, DEPUY SYNTHES OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE DEVICE, DEPUY SYNTHES, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. H11 ADDITIONAL NARRATIVE: ADDED: D4 (EXPIRY DATE), G1 (MANUFACTURING SITE NAME), H4 CORRECTED: D4 (PRIMARY UDI NUMBER). H3, H6: PRODUCT WAS NOT RETURNED. BASED ON THE INFORMATION AVAILABLE, IT HAS BEEN DETERMINED THAT NO CORRECTIVE AND PREVENTATIVE ACTION IS PROPOSED. THIS COMPLAINT WILL BE ACCOUNTED FOR AND MONITORED VIA POST MARKET SURVEILLANCE ACTIVITIES. IF ADDITIONAL INFORMATION IS MADE AVAILABLE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. DEVICE HISTORY REVIEW (DHR): A MANUFACTURING RECORD EVALUATION WAS PERFORMED FOR THE FINISHED DEVICE PRODUCT CODE: 283913000421387 LOT NUMBER: IT WAS ELECTRONICALLY REVIEWED AND NO NONCONFORMANCES/MANUFACTURING IRREGULARITIES WERE IDENTIFIED DURING THE MANUFACTURING PROCESS. THE PRODUCT WAS RELEASED ON: 23 JAN 2025 MANUFACTURING SITE: JABIL LE LOCLE EXPIRY DATE: 31 DEC 2026 CEMENT: PART # 183901001/ LOT # 4572175. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL MEDWATCH, A FOLLOW-UP MEDWATCH WILL BE FILED AS APPROPRIATE.
POWERTRIALYSIS | K133456 | IT WAS REPORTED, PATIENT REMOVED TRIALYSIS CATHETER THAT WAS PLACED IN FEMORAL VEIN AND BLED OUT. SUTUTRES WERE IN TACT ALONG WITH THE SUTURE WINGS BUT TRIALYSIS HAD BEEN PULLED THROUGH THE LOOP. PATIENT BLED OUT AND EXPIRED. NO OTHER INFORMATION WAS PROVIDED. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. THE COMPLAINT OF DETACHED SUTURE WINGS IS CONFIRMED; HOWEVER, THE EXACT CAUSE IS UNKNOWN. ONE PHOTOGRAPH OF A 20 CM POWERTRIALYSIS CATHETER WAS RETURNED FOR EVALUATION. AN INITIAL VISUAL OBSERVATION OF THE PHOTOGRAPH SHOWED THE SUTURE WINGS OF THE DEVICE WERE DISPLACED AND POSITIONED ABOUT 3 CM DISTAL TO THEIR ORIGINAL LOCATION ON THE MOLDED JOINT. NO OBVIOUS DAMAGE COULD BE SEEN ON THE SUTURE WINGS OR MOLDED JOINT IN THE RETURNED PHOTOGRAPH; HOWEVER, THE RESOLUTION OF THE PHOTOGRAPH MADE IT DIFFICULT TO DISCERN SUCH DETAILS. USE RESIDUE WAS OBSERVED ON THE CATHETER IN THE RETURNED PHOTOGRAPH. THE DISTAL END OF THE CATHETER APPEARED TO BE MISSING; HOWEVER, NO DETAILS OF THE POSSIBLE BREAK IN THE CATHETER COULD BE DISCERNED FROM THE RETURNED PHOTOGRAPH. WHILE THE SUTURE WINGS OF THE CATHETER SHOWN IN THE RETURNED PHOTOGRAPH WERE OBSERVED TO BE DISPLACED, IT WAS NOT POSSIBLE TO DETERMINE THE ROOT CAUSE OF THIS DISPLACEMENT FROM THE RETURNED PHOTOGRAPH. THE APPROPRIATE RISK DOCUMENTATION WAS REVIEWED, AND RISK MANAGEMENT TOOLS ARE IN PLACE TO ENSURE THAT POTENTIAL RISKS ASSOCIATED WITH DEVICE USE ARE IDENTIFIED AND REDUCED TO THE EXTENT POSSIBLE. NO BATCH INFORMATION WAS RECEIVED FROM THE CUSTOMER AND, CONSEQUENTLY, IT WAS NOT POSSIBLE TO PERFORM A HISTORICAL REVIEW OF SIMILAR COMPLAINTS NOR WAS IT POSSIBLE TO PERFORM A REVIEW OF MANUFACTURING RECORDS AND PROCESSES FOR THE BATCH INVOLVED. DETAILS REGARDING CATHETER PLACEMENT, INDWELL TIME, OR MAINTENANCE ARE UNKNOWN, AND NO FURTHER INFORMATION REGARDING PATIENT HISTORY OR CONDITION HAS BEEN COMMUNICATED TO BD. ADDITIONAL INFORMATION HAS BEEN REQUESTED FROM THE COMPLAINANT, AND THIS EVALUATION WILL BE UPDATED UPON RECEIPT OF A PHYSICAL SAMPLE FOR EVALUATION AND/OR ANY ADDITIONAL INFORMATION PROVIDED. THIS COMPLAINT WILL BE RECORDED FOR FUTURE TRENDING AND MONITORING PURPOSES.
B125 | K171580 | WHILE BEING USED TO MONITOR A PATIENT DURING TRANSPORT, THE BATTERY FAILED AND THE MONITOR SHUT DOWN. IT WAS ALLEGED THAT THE MONITOR DID NOT ALARM FOR LOW BATTERY. THE PATIENT EXPIRED. | LEGAL MANUFACTURER: HCS TROUT - 3114 N GRANDVIEW BLVD. USA WAUKESHA, WI 53188. A1-6: NOT PROVIDED BY CUSTOMER. GE HEALTHCARE'S INVESTIGATION IS ON-GOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
LLD EZ LEAD LOCKING DEVICE | K142116 | A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD DUE TO BACTEREMIA. A SPECTRANETICS LLD EZ LEAD LOCKING DEVICE (LLD EZ) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS GLIDELIGHT LASER SHEATH, THE LEAD WAS REMOVED; HOWEVER, THE PATIENT'S BLOOD PRESSURE DECLINED. RESCUE EFFORTS BEGAN, INCLUDING STERNOTOMY, AND A PERFORATION AT THE SUPERIOR VENA CAVA (SVC)/RA JUNCTION WAS DISCOVERED AND REPAIRED. THE PATIENT WAS TRANSFERRED TO THE ICU, BUT UNFORTUNATELY PASSED AWAY LATER THAT DAY. THIS REPORT CAPTURES THE LLD EZ PROVIDING TRACTION WITHIN THE RV LEAD WHEN THE SVC/ RA JUNCTION PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTED IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. | A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. B6) RELEVANT TEST/LABORATORY DATA - NOT AVAILABLE FROM FACILITY. D4/H4) THE LOT NUMBER WAS NOT PROVIDED; THUS, THE FOLLOWING INFORMATION IS UNKNOWN: UNIQUE ID, EXPIRATION DATE, AND MANUFACTURE DATE. H3) THE DEVICE WAS DISCARDED; THUS, NO DEVICE EVALUATION COULD BE COMPLETED. H6) GREAT VESSEL PERFORATION, CARDIAC PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE LLD. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
ION | K182188 | IT WAS REPORTED THAT A CLINICAL STUDY REGISTRY PATIENT UNDERWENT AN UNEVENTFUL 19-MINUTE ION-ASSISTED LUNG BIOPSY. THE PROCEDURE WAS COMPLETED AS PLANNED, AND THE PATIENT WAS DISCHARGED THE SAME DAY. THERE WERE NO REPORTED DEVICE ISSUES OR COMPLICATIONS ASSOCIATED WITH THE PROCEDURE. THE PATIENT¿S SPOUSE INFORMED THE PHYSICIAN THAT THE PATIENT EXPIRED 18 DAYS POST-PROCEDURE. THE REGISTRY STUDY PHYSICIAN DEEMED THE DEATH TO BE NOT RELATED TO THE PROCEDURE, NOT RELATED TO THE ION SYSTEM OR ACCESSORIES, AND PROBABLY RELATED TO THE PATIENT¿S PRE-EXISTING CO-MORBIDITIES, INCLUDING AN 80 PACK-YEAR HISTORY OF SMOKING. NO ADDITIONAL INFORMATION WAS AVAILABLE. | THE REGISTRY STUDY PHYSICIAN REPORTED THAT THE EVENT WAS PROBABLY RELATED TO THE PATIENT¿S EXISTING CONDITION(S), NOT RELATED TO THE PROCEDURE, NOT RELATED TO ION SYSTEM/ACCESSORIES. A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WAS RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT UNDERWENT AN ION ENDOLUMINAL BIOPSY AS PART OF A CLINICAL STUDY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE AND THE PATIENT WAS DISCHARGED HOME THE SAME DAY. THE BIOPSY CONFIRMED A SQUAMOUS CELL CANCER. THE PATIENT DIED 18 DAYS AFTER THE PROCEDURE. THE DEATH WAS ATTRIBUTED TO THEIR MULTIPLE UNDERLYING MEDICAL CO-MORBIDITIES INCLUDING CHRONIC DEMYELINATING POLYNEUROPATHY, HEPATITIS B, IRON DEFICIENCY ANEMIA, COPD/EMPHYSEMA AND MODERATE PULMONARY FIBROSIS. BASED ON THE AVAILABLE DATA THE DEATH WAS DUE TO UNDERLYING MEDICAL CONDITIONS AND ALTHOUGH THERE IS NO CLEAR EVIDENCE OF ASSOCIATION WITH THE INDEX PROCEDURE, THE POSSIBILITY THAT THE EVENT WAS PROCEDURE RELATED CANNOT BE EXCLUDED. THE DEATH WAS IDENTIFIED AS PART OF A CLINICAL STUDY. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
REVEAL LINQ | K132649 | IT WAS REPORTED THAT A PATIENT WITH A LINQ DEVICE EXPERIENCED A TERMINAL EPISODE ON (B)(6) 2025, RESULTING IN DEATH. THE PATIENT'S PREVIOUS DEVICE REACHED THE RECOMMENDED REPLACEMENT TIME (RRT), AND IT APPEARED THAT WHEN THE DEVICE WAS REPLACED, THE PATIENT NEVER PERFORMED AN INITIAL MANUAL TRANSMISSION WITH THE REMOTE MONITOR, WHICH WAS REQUIRED TO PAIR WITH THE NEW DEVICE, SO NO TRANSMISSIONS WERE COMING THROUGH. THE ONLY TRANSMISSIONS ON THE NETWORK WERE FROM THE LINQ MOBILE MANAGER AND CARELINK EXPRESS. THE PATIENT RECORDED A SYMPTOM ON (B)(6) 2025, BUT IT WAS NOT SEEN UNTIL THE LINQ DEVICE WAS INTERROGATED BY CARELINK EXPRESS POST-MORTEM. THERE WERE NO ACTIONS TAKEN ON THE MONITOR. IT WAS NOTED THAT THE MONITOR HAD BEEN PLUGGED IN, SUCCESSFULLY UPDATED THE FIRMWARE, AND REGULARLY CHECKED IN, INDICATING THAT THE MONITOR WAS WORKING AS EXPECTED. THE REMOTE MONITOR WAS DISCARDED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
PROXIMATE*VASC LIN STAP RELOAD | K020779 | IT WAS REPORTED THAT DURING A LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STALING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. | (B)(4). DATE SENT: 5/12/2025. CORRECTED DATA: B1, B2, H1. UPON REVIEW IT WAS IDENTIFIED THAT THIS IS A DUPLICATE REPORT. ALL REPORTED INFORMATION FOR THIS EVENT WAS REPORTED UNDER 3005075853-2025-02969. MOVING FORWARD ALL ADDITIONAL INFORMATION WILL BE FILED UNDER 3005075853-2025-0263.
PROXIMATE*VASC LIN STAP RELOAD | K020779 | IT WAS REPORTED THAT DURING A LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STALING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. | (B)(4). DATE SENT: 5/12/2025. CORRECTED DATA: B1, B2, H1. UPON REVIEW IT WAS IDENTIFIED THAT THIS IS A DUPLICATE REPORT. ALL REPORTED INFORMATION FOR THIS EVENT WAS REPORTED UNDER 3005075853-2025-02969. MOVING FORWARD ALL ADDITIONAL INFORMATION WILL BE FILED UNDER 3005075853-2025-0263.
PROXIMATE*VASC LIN STAP RELOAD | K020779 | IT WAS REPORTED THAT DURING A LIVER TRANSPLANT THE RIGHT HEPATIC VEIN WAS STAPLED USING THE TX LINEAR STAPLER. ONCE THE VEIN WAS CUT AFTER STALING A MASSIVE BLEEDING WAS ENCOUNTERED RESULTING IN MAJOR HEMORRHAGE. THE OPERATING CONSULTANT IS QUESTIONING IF THE STAPLER FIRED AS THERE WAS NO EVIDENCE OF A STAPLE LINE ON THE PATIENT VESSEL OR IN THE EXPLANTED LIVER WHEN INSPECTED. | (B)(4). DATE SENT: 5/12/2025. CORRECTED DATA: B1, B2, H1. UPON REVIEW IT WAS IDENTIFIED THAT THIS IS A DUPLICATE REPORT. ALL REPORTED INFORMATION FOR THIS EVENT WAS REPORTED UNDER 3005075853-2025-02969. MOVING FORWARD ALL ADDITIONAL INFORMATION WILL BE FILED UNDER 3005075853-2025-0263.
CONTOUR 3D | K101212 | MEDTRONIC RECEIVED INFORMATION THAT 4 YEARS AND 1 MONTH POST IMPLANT OF THIS 32MM TRICUSPID ANNULOPLASTY RING, THE PATIENT DIED. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE ANNULOPLASTY RING COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
NOVUM IQ | K211122 | IT WAS REPORTED THAT A NOVUM IQ LARGE VOLUME PUMP UNDER-INFUSED DURING PATIENT INFUSION. THERE WAS NO REPORT OF PATIENT INJURY OR MEDICAL INTERVENTION ASSOCIATED WITH THIS EVENT. NO ADDITIONAL INFORMATION IS AVAILABLE. | RELATED COMPLAINTS: DEVICE 2: 1416980-2025-02471 AND DEVICE 3: 1416980-2025-02469. ADDITIONAL INFORMATION: B1, B2, B3, B5, D10, F10, AND H1. B5: UPON FOLLOW UP INFORMATION, IT WAS FURTHER REPORTED THREE (3) PUMPS WERE IN USE ON ONE PATIENT. THE PATIENT REPORTEDLY WAS RECEIVING LEVOPHED, DOBUTAMINE, VASOPRESSIN AND IV FLUIDS ON THE NOVUM PUMPS. ON AN UNSPECIFIED DATE, THE PATIENT REPORTEDLY EXPERIENCED A ¿DECLINE IN CONDITION.¿ THIS PATIENT WAS TRANSFERRED TO A ¿HIGHER LEVEL OF CARE.¿ ON THE DATE OF THE EVENT, THE PATIENT REPORTEDLY ¿CODED¿ AND DIED. IT WAS NOT REPORTED WHICH PUMP WAS RUNNING WHICH MEDICATED SOLUTION OR WHICH PUMP UNDER-INFUSED RELATED TO THE PATIENT¿S REPORTED DEATH. ADDITIONALLY, THE CLINICAL ENGINEERING DEPARTMENT OF THE REPORTING FACILITY EVALUATED THE PUMPS ONSITE AND ¿NO ISSUES WERE IDENTIFIED WITH THE THREE PUMPS INVOLVED.¿ THIS COMPLIANT IS FOR DEVICE ONE (1). SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
TRILOGY EVO, USA | K181166 | THE MANUFACTURER RECEIVED INFORMATION "THAT WHILE BEING TRANSPORTED ON THE WAY TO THE PATIENT'S PRIMARY CARE PHYSICIAN ON (B)(6) 2025 AT APPROXIMATELY 14:00, THE PATIENT (FEMALE, AGE 9) WAS NOTED TO HAVE GONE UNRESPONSIVE. UPON ARRIVAL AT AN INSTITUTIONAL FACILITY, THE PATIENT UNDERWENT CARDIOPULMONARY RESUSCITATION (CPR) AND SUBSEQUENTLY EXPIRED. A REQUEST WAS LODGED BY THE REPORTER TO PROVIDE A SUMMARY OF THE EVENT LOGS FOR FURTHER INVESTIGATION TO DETERMINE WHAT EXACTLY HAPPENED WITH THE VENTILATOR DURING THE TIME OF THE EVENT. NO DIRECT ALLEGATION OF MALFUNCTION OR FAILURE TO PERFORM TO MANUFACTURER DECLARED SPECIFICATIONS HAS BEEN LODGED AT THIS TIME. NOR HAS ANY ALLEGATION OF DEVICE CAUSE AND/OR CONTRIBUTION HAS BEEN LODGED AT THIS TIME". CURRENTLY WITH THE INFORMATION AVAILABLE, CAUSE AND/OR CONTRIBUTION OF THE DEVICE TO THE PATIENT EXPIRATION CANNOT BE CONFIRMED NOR REFUTED PENDING ADDITIONAL GOOD FAITH EFFORTS AND DEVICE EVALUATION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". AFTER BEING INFORMED BY THE MANUFACTURER THAT THE SUMMARY OF THE VENTILATOR LOGS CANNOT BE PROVIDED WITHOUT FURTHER DEVICE EVALUATION, THE REPORTER AGREED TO SEND IN THE DEVICE TO THE MANUFACTURER FOR EVALUATION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. .
INTELLIVUE MX450 PATIENT MONITOR | K130849 | IT WAS REPORTED THE MONITOR DID NOT GENERATE A RED CRITICAL ALARM AND THE PATIENT PASSED AWAY; HOWEVER, IT WAS ALSO NOTED THE ECG ARRHYTHMIA ALARMS MAY HAVE BEEN TURNED OFF. ADDITIONAL INFORMATION RECEIVED INDICATED THERE WAS A DELAY IN CARE, BUT IT WAS NOT KNOWN HOW LONG THE DELAY IN CARE WAS FOR. IT WAS INDICATED THE LACK OF ALARMS MAY HAVE CONTRIBUTED TO THE DEATH, BUT NURSES ALSO INDICATED THE PATIENT DEATH MAY HAVE OCCURRED INDEPENDENT OF THE ALARM ISSUE. THE ALARMS WERE DEACTIVATED BUT THE NURSES STATED THE OPTION TO DEACTIVATE THE ALARMS SHOULD NOT BE READILY AVAILABLE. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE ANY MALFUNCTION OF THE DEVICE, AND IT CANNOT BE DETERMINED WHETHER THE LACK OF ALARM CONTRIBUTED TO THE PATIENT'S DEATH. | E1: REPORTING INSTITUTION PHONE #: (B)(6). E1: REPORTER PHONE #: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE WAS RECEIVED BY A THIRD-PARTY SERVICE CENTER, BUT THE EVALUATION HAS NOT YET BEGUN. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. | THIS REPORT IS BEING SUBMITTED TO PROVIDE THE DEVICE'S THIRD-PARTY EVALUATION RESULTS. IT WAS REPORTED THAT THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE RENTAL DREAMSTATION ST30 DEVICE WAS RETURNED AND EVALUATED BY A THIRD-PARTY SERVICE CENTER. THE DEVICE FOUND THAT NO FAULTS WERE REPORTED. THE FINE FILTER AND POLLEN FILTER WERE REPLACED FOR PREVENTATIVE MAINTENANCE, AND THE ACCESSORY FLIP MODULE DOOR WAS MISSING FROM THE DEVICE. THE DEVICE WAS CLEANED AND TESTED, AND ALL TESTS PASSED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
CONFIDA BRECKER CURVE GUIDEWIRE | K132623 | IT WAS REPORTED THAT DURING THE IMPLANT OF THIS TRANSCATHETER BIOPROSTHETIC VALVE, A VENTRICULAR PERFORATION OCCURRED. PER THE PHYSICIAN, THE PERFORATION WAS CAUSED BY ADJUSTMENT OF THE GUIDEWIRE POSITION. SUBSEQUENTLY, THE PATIENT PASSED AWAY. | IMAGE REVIEW: 7 FLUOROSCOPIC CINE LOOPS OF THE IMPLANT PROCEDURE WERE PROVIDED FOR REVIEW. PATIENT SUMMARY WAS ALSO PROVIDED FOR ANATOMICAL REVIEW. ACCORDING TO THE REPORT, A GUIDEWIRE GWBC30 ¿ THE MEDTRONIC CONFIDA GUIDEWIRE ¿ WAS USED. IMAGES SHOW A LAO OPEN-ARCH PROJECTION WITH THE CONFIDA FLUSH IN THE AORTA¿S OUTER CURVE, WITH ITS TIP SITTING DEEP IN THE LEFT VENTRICLE. THE LEFT VENTRICLE IS REPORTED TO BE SMALL WHICH IS SUPPORTED BY IMAGES FROM THE EXECUTIVE SUMMARY REPORT. ALTHOUGH NO VISIBLE ADJUSTMENT OF THE GUIDEWIRE WAS MADE IN THE PROVIDED CINE LOOPS, ITS DISTAL END IS NOT PROPERLY POSITIONED IN A CURVE BUT CURLED UP AGAINST THE LEFT VENTRICLE-WALL, THE APPEARING TO BE TANGLED UP IN THE MV-CORDAE OR PAPILLARY MUSCLES. THERE ARE IMAGES THAT DISPLAY THE INITIAL SITUATION TO BE A FAILED PREVIOUS TAV-IN-SAV IMPLANT ATTEMPT, WITH THE SAV SHOWING MERELY AS A RING AND THE FIRST EVOLUT VALVE BEING HELD IN PLACE BY A SNARE. THE CONFIDA GUIDEWIRE TRACKS ALONG THE GREATER CURVE OF THE AORTA, OUTSIDE OF THE FIRST EVOLUT BIOPROSTHESIS. IN THE FINAL IMPLANT IMAGE, THE FIRST EVOLUT CAN BE SEEN TO BE SEVERELY TILTED IN THE ASCENDING AORTA, POKING INTO THE VESSEL WALL WITH IT¿S OUTFLOW CROWN. IN THE EVOLUT SYSTEM¿S INSTRUCTION FOR USE (IFU), POTENTIAL ADVERSE EVENTS INCLUDE, BUT ARE NOT LIMITED TO, DEATH AND CARDIAC TAMPONADE. THE IMAGES AND INFORMATION PROVIDED IN THE REPORT DON¿T ALLOW FOR IDENTIFICATION OF THE EXACT ROOT CAUSE OF THE VENTRICULAR PERFORATION LEADING TO THE PATIENT DEATH. HOWEVER, THE DEEP INSERTION OF THE CONFIDA GUIDEWIRE IN COMBINATION WITH ITS SUB-OPTIMAL TIP-PLACEMENT AND FORWARD PUSH ON THE GUIDEWIRE EXERTED DURING THE DELIVERY SYSTEM ADVANCEMENT ARE LIKELY TO HAVE CONTRIBUTED TO THE VENTRICULAR PERFORATION. IT¿S NOT PERCEIVABLE BY THE CINE LOOPS ALONE, HOW WELL THE GUIDEWIRE WAS CONTROLLED DURING THE IMPLANT ATTEMPT. IT SHOULD ALSO BE NOTED THAT THE FINAL RESTING POSITION OF THE FIRST EVOLUT BIOPROSTHESIS AND ITS RESULTING INTERACTION WITH THE VESSEL WALL WERE NOT OPTIMAL. THERE WAS NO MENTION OF MEDICATION BEING ADMINISTERED. IT WAS REPORTED THAT DESPITE THE APPLICATION OF CARDIO SYNTHESIS, THE PATIENT PASSED AWAY AFTER THE PROCEDURE. UPDATED H.6 MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
AMPLATZER VASCULAR PLUG II | K071699 | N/A. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD (DHR) AND COMPLAINT HISTORY REVIEWS WERE NOT PERFORMED BECAUSE THIS COMPLAINT WAS BASED ON AN ARTICLE REVIEW AND NO LOT INFORMATION WAS PROVIDED. SOME OF THE COMPLICATIONS REPORTED WERE RESIDUAL SHUNT, AORTIC VALVE REGURGITATION AND DEATH; THESE COMPLICATIONS ARE ANTICIPATED FOR THE PROCEDURE. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE WAS RECEIVED FOR ANALYSIS. BASED ON AVAILABLE INFORMATION AND DUE TO THE LIMITED INFORMATION AVAILABLE FROM THE ARTICLE, THE CAUSE OF THE REPORTED RESIDUAL SHUNT AND AORTIC VALVE REGURGITATION COULD NOT BE CONCLUSIVELY DETERMINED. DEATH WAS RELATED TO COMMODITIES OF THE INFANT REQUIRING ON-GOING ECMO SUPPORT AND RESULTING IN BLEEDING. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH RESPECT TO LABELING, DESIGN, OR MANUFACTURING OF THE DEVICE. B2: DEATH DATE IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE ATTACHMENT: TRANSCATHETER OCCLUSION OF THE NEOAORTA TO AUGMENT MECHANICAL CIRCULATORY SUPPORT AFTER STAGE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT PASSED AWAY DUE TO BOWEL PERFORATION. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
TANDEM MOBI INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY | K233044 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. PER THE CONTACT, THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. IT WAS REPORTED THAT THE CAUSE OF DEATH WAS AN INTENTIONAL INSULIN OVERDOSE. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT. THANKS! | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
SAPPHIRE M.T INFUSION PUMP - GERMAN | K192860 | THIS COMPLAINT WAS REPORTED BY A CUSTOMER FROM GERMANY. A PROGRAMMING ERROR WHICH RESULTED IN PATIENT DEATH WAS REPORTED. | INVESTIGATION TYPE: EITAN MEDICAL INVESTIGATED THE PUMP'S EVENT LOG. INVESTIGATION FINDINGS: THE EVENT WAS CONFIRMED TO BE SOLELY A RESULT OF A PROGRAMMING ERROR. THE PUMP FUNCTIONED AS DESIGNED AND DELIVERED THE INFUSION IN ACCORDANCE WITH THE PROGRAMED PARAMETERS. CONCLUSION: THE EVENT WAS CONFIRMED TO BE A RESULT OF A PROGRAMMING ERROR. THE SAPPHIRE PUMP DESIGN MINIMIZES PROGRAMMING ERRORS BY REQUIRING USERS TO REVIEW AND APPROVE PARAMETERS BEFORE STARTING OR REPEATING INFUSIONS, WITH CLEAR ON-SCREEN DISPLAYS THROUGHOUT. THIS REPORTED EVENT OCCURRED OUTSIDE OF THE US ON A DEVICE THAT IS SIMILAR TO THE ONE MARKETED IN THE US AND FOR WHICH THERE IS NO DATA IN GUDID. DEVICE'S PREMARKET SUBMISSION NUMBER: K192860.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
TUBING SET | K112360 | COMPLAINT ID #(B)(4). | THE AFFECTED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION AND WAS INVESTIGATED IN THE GETINGE LABORATORY ON 2025-06-12 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES AND NO ABNORMALITIES WERE DETECTED DURING TESTING." THE INVOLVED CARDIOHELP DEVICE IS INVESTIGATED WITHIN COMPLAINT #(B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000194. A MEDICAL REVIEW WAS REQUESTED BY GETINGE MEDICAL AFFAIRS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
CARDIOHELP | K133598 | COMPLAINT ID # (B)(4). | NEW INFORMATION WAS RECEIVED ON 2025-05-27, THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT # (B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000193. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
TUBING SET | K112360 | COMPLAINT: (B)(4). | THE AFFECTED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION AND WAS INVESTIGATED IN THE GETINGE LABORATORY ON (B)(6) 2025 WITH THE FOLLOWING CONCLUSION: "THE REPORTED FAILURE COULD NOT BE REPRODUCED AND NOT BE CONFIRMED DURING THE INVESTIGATION. THE HLS SET SHOWED NO DAMAGES AND NO ABNORMALITIES WERE DETECTED DURING TESTING." THE INVOLVED CARDIOHELP DEVICE IS INVESTIGATED WITHIN COMPLAINT #(B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000194. A MEDICAL REVIEW WAS REQUESTED BY GETINGE MEDICAL AFFAIRS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE GERMAN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
CARDIOHELP | K133598 | COMPLAINT ID # (B)(4). | NEW INFORMATION WAS RECEIVED ON 2025-05-27, THAT THE AFFECTED CARDIOHELP DEVICE WAS INVESTIGATED BY A GETINGE FIELD SERVICE TECHNICIAN AND RECEIVED PREVENTIVE MAINTENANCE. NO ABNORMALITIES WERE FOUND WITH THE CARDIOHELP DEVICE. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE INVOLVED HLS SET WAS REQUESTED FOR FURTHER INVESTIGATION WITHIN THE COMPLAINT #(B)(4) AND IS REPORTED UNDER MFG REPORT NUMBER #8010762-2025-0000193. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE.
ENDO CLIP III | K100242 | IT WAS REPORTED THAT ON A POST VETERINARY LAPAROSCOPIC GALLBLADDER SURGERY; WHILE CLOSING THE BILE DUCT, THE CLIP APPLIER DID NOT LOAD THE CLIPS PROPERLY. THE PRODUCT WAS EXPLANTED. IT WAS ALSO NOTED THAT THE HOSPITALIZATION AND SURGICAL TIME WAS EXTENDED FOR MORE THAN 30 MINUTES. ALSO, THE INCISION EXTENDED DUE TO THE PRODUCT PROBLEM. TO RESOLVE THE ISSUE, EMERGENCY REINTERVENTION WAS DONE. IT WAS REPORTED THAT THE PATIENT EXPIRED. | ADDITIONAL INFORMATION: D9, G3, H3, H6 H3 EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. THE DEVICE WAS AVAILABLE FOR EVALUATION. THE EVALUATION FOUND NO POTENTIALLY CONTRIBUTING FACTORS, AND THE SAMPLE MET ALL RELATED SPECIFICATIONS. IT WAS REPORTED THAT THE CLIPS DID NOT LOAD PROPERLY. THE REPORTED ISSUE COULD NOT BE CONFIRMED. THE MOST LIKELY CAUSE COULD NOT BE IDENTIFIED BECAUSE NO RELATED PROBLEM WAS DETECTED WITH THE DEVICE. THE MANUFACTURING RECORDS FOR EACH DEVICE ARE THOROUGHLY REVIEWED PRIOR TO RELEASE TO ENSURE THAT IT MEETS ALL MEDTRONIC QUALITY SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
AMPLATZER TORQVUE DELIVERY SYSTEM | K072313 | N/A | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. AN EVENT OF USE OF DEVICE PROBLEM WAS REPORTED. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED, AND THE PRODUCT MET ALL SPECIFICATIONS. BASED ON THE AVAILABLE INFORMATION, THE REPORTED USE OF DEVICE RELATED TISSUE DAMAGED COULD NOT BE DETERMINED. THE CAUSE OF DEATH WAS VENTRICULAR RUPTURE DURING A PERCUTANEOUS ATTEMPT TO CLOSE A POST INFARCT VENTRICULAR SEPTAL RUPTURE (VSD). THE PATIENT EFFECTS OF PERICARDIAL EFFUSION AND CARDIAC TAMPONADE COULD NOT BE CONCLUSIVELY DETERMINED. BLEEDING APPEARS TO RELATE TO PERICARDIAL EFFUSION/TAMPONADE, UNABLE TO TREAT, AND SUBSEQUENT CARDIOGENIC SHOCK (CARDIAC ARREST). THE REPORTED UNEXPECTED MEDICAL INTERVENTIONS WERE A RESULT OF CASE-SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT QUALITY ISSUE WITH REGARDS TO MANUFACTURE, DESIGN, OR LABELING.
ULRICH MEDICAL | K171392 | ACCORDING TO REPORTING PERSON: "WE HAVE BEEN INFORMED TODAY OF SERIOUS ALLERGIC INCIDENT RESULTING IN THE DEATH OF 1 PATIENT AT THE (B)(6) HOSPITAL. THIS INCIDENT HAS OCCURRED IN FEBRUARY, A PATIENT WAS INJECTED WITH IOMERON USING THE ULRICH INJECTOR. THE PATIENT TUBING XD 2040 ANNOUNCED TO BE INVOLVED COMES FROM BATCH E000095." | THE MENTIONED ARTICLE XD 2040 IS NOT SOLD IN THE USA. HOWEVER, XD 2042 IS FUNCTIONALLY IDENTICAL AND IS SOLD IN THE USA. BASED ON THE INFORMATION AVAILABLE, THE DEVICE CAN BE EXCLUDED AS A SOURCE OF ERROR. IT IS APPROVED BY THE AUTHORITIES FOR USE WITH THE SAID CONTRAST MEDIUM. ALL DEVICES ON SITE HAVE SUCCESSFULLY PASSED A TECHNICAL SAFETY TEST, THERE IS NO INDICATION OF A ENDOGENOUS CAUSE OF THE INJECTION SYSTEM.
UL_INJECT | K171392 | ACCORDING TO USER INFORMATION: "ON (B)(6) 2025 AT 11:45 A.M., MS. (B)(6) REPORTED THE DEATH OF A PATIENT DUE TO AIR IN THE HEART IN RELATION TO (B)(6). A CT THORAX ABDOMEN EXAMINATION WITH ULTRAVIST 300 IN 500 ML GLASS BOTTLES AT A FLOW RATE OF 3ML/S WITH 21ML NACL / 100ML KM / 30ML NACL. THE CONTRAST AGENT WAS NOT PREHEATED. A 320 CM PATIENT TUBE WAS USED. ACCESS WAS VIA A PORT SYSTEM THAT IS ALSO USED FOR CHEMOTHERAPY. MS. (B)(6) WAS UNABLE TO STATE THE AMOUNT OF AIR. THE INJECTOR DID NOT SHOW ANY ERROR MESSAGES."
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT FAILED TO PRESS THE START BUTTON BEFORE GOING TO SLEEP. THE PATIENT WAS FOUND IN THE MORNING IN CARDIOPULMONARY ARREST AND WAS ADMITTED TO THE HOSPITAL. THE DEVICE WAS RETRIEVED FROM THE HOSPITAL. NO EVALUATION OF THE DEVICE HAS BEEN PERFORMED YET. THE CLINICAL ASSESSMENT WAS PERFORMED, AND THE FOLLOWING RATIONALE WAS PROVIDED. THE FAMILY REPORTED THAT PATIENT TESTS HAVE BEEN COMPLETED, AND IT WAS DETERMINED THERE IS NO BRAIN ACTIVITY, SO A DECISION WAS MADE NOT TO PURSUE LIFE-PROLONGING TREATMENT. THEREFORE, DEATH IS PRESUMED BASED ON "4/2025-NO BRAIN ACTIVITY AND THEY DECIDED NOT TO PURSUE LIFESAVING INTERVENTIONS". BASED ON INFORMATION PROVIDED AND AVAILABLE, THIS EVENT IS LIKELY A FORESEEABLE UNINTENTIONAL USE ERROR. THEREFORE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH CAN BE CONFIRMED BASED ON THE EVIDENCE PRESENT. AT THIS TIME, WE ARE UNABLE TO CONFIRM IF ANY MALFUNCTION OF THE DEVICE OCCURRED. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED A PATIENT FAILED TO PRESS THE START BUTTON BEFORE GOING TO SLEEP. THE PATIENT WAS FOUND IN THE MORNING IN CARDIOPULMONARY ARREST AND WAS ADMITTED TO THE HOSPITAL. THE DEVICE WAS RETRIEVED FROM THE HOSPITAL. THE DEVICE EVALUATION FOUND NO EVIDENCE OF PRODUCT MALFUNCTION IN WHICH THE ABILITY OF THE DEVICE TO DELIVER THE PRESCRIBED THERAPY TO THE PUBLISHED SPECIFICATIONS WOULD BE IMPACTED. IN ADDITION, THERE IS NO EVIDENCE THAT THE DEVICE MALFUNCTIONED IN A MANNER THAT WOULD BE LIKELY TO CAUSE OR CONTRIBUTE TO DEATH OR SERIOUS INJURY IF IT WERE TO RECUR. THE DEVICE WAS TESTED AND PASSED ALL TESTING.
GALAXY SYSTEM | K223144 | IT WAS REPORTED THAT A PATIENT DEVELOPED BRADYCARDIA AND OXYGEN DESATURATION DURING A GALAXY-ASSISTED BIOPSY PROCEDURE FOR A RIGHT UPPER LOBE LESION. THE INJURY OCCURRED DURING THE CASE AND THE ROBOT WAS SAFELY UNDOCKED TO ADDRESS THE PATIENT. MEDICATIONS WERE ADMINISTERED TO RAISE THE PATIENT'S VITALS (NOT PROVIDED BY THE HOSPITAL STAFF) AND THE PATIENT WAS ADMITTED FOR OBSERVATION. TWO WEEKS LATER, NOAH MEDICAL WAS INFORMED THAT THE PATIENT PASSED AWAY 2 DAYS AFTER THE PROCEDURE. THERE WERE NO MALFUNCTIONS REPORTED. THE PHYSICIAN DOES NOT BELIEVE THE GALAXY SYSTEM CONTRIBUTED TO THE BRADYCARDIA AND OXYGEN DESATURATION, ATTRIBUTING IT INSTEAD TO THE BREATH HOLD DURING THE TOOL INSERTION IN THE LESION TILT. IT IS IMPORTANT TO NOTE THAT THE PATIENT HAD PRE-EXISTING VENTRICULAR ISSUES. NOAH MEDICAL WAS INFORMED THAT THE PATIENT EXPIRED 2 DAYS AFTER THE BRONCHOSCOPY PROCEDURE. THE PHYSICIAN ATTRIBUTED THE DEATH TO MULTIORGAN FAILURE. ADDITIONALLY, THE PATIENT HAD RIGHT VENTRICULAR DYSFUNCTION, ELEVATED POTASSIUM LEVELS AND MULTIPLE MYOCARDIAL INFARCTIONS THE NIGHT OF PASSING. ADDITIONAL INFORMATION WAS REQUESTED FROM THE HOSPITAL, A SUPPLEMENTAL WILL BE SUBMITTED IF THE INFORMATION IS ACQUIRED BY NOAH MEDICAL. | SYSTEM AND BRONCHOSCOPE MANUFACTURING RECORDS WERE REVIEWED AND FOUND NO ISSUES ASSOCIATED WITH THIS CASE. FAILURE ANALYSIS OF THE SCOPE COULD NOT BE PERFORMED AS THE SCOPE WAS DISCARDED. VIDEO LOGS WERE REVIEWED AND NO USE ERRORS WERE IDENTIFIED. AT TIMESTAMP 22:21, A SUDDEN SCOPE DISCONNECTION AND DISENGAGEMENT OF THE CART BRAKES WERE NOTED, WHICH IS LIKELY DUE TO THE NEED TO ADDRESS THE PATIENT'S CONDITION. THE PHYSICIAN DOES NOT BELIEVE THE GALAXY SYSTEM CONTRIBUTED TO THE BRADYCARDIA AND OXYGEN DESATURATION, ATTRIBUTING IT INSTEAD TO THE BREATH HOLD DURING THE TOOL INSERTION IN THE LESION TILT. IT IS IMPORTANT TO NOTE THAT THE PATIENT HAD PRE-EXISTING VENTRICULAR ISSUES. WHILE PERFORMING BREATH HOLDS DURING BRONCHOSCOPY IS NOT INHERENTLY HIGH-RISK, IT MAY PRESENT POTENTIAL RISKS, SUCH AS DECREASED OXYGEN LEVELS, PARTICULARLY IN PATIENTS WITH UNDERLYING CARDIAC CONDITIONS, AS SEEN IN THIS CASE. 10 DAYS AFTER THE EVENT, NOAH MEDICAL WAS INFORMED THAT THE PATIENT EXPIRED 2 DAYS AFTER THE BRONCHOSCOPY PROCEDURE. THE PHYSICIAN ATTRIBUTED THE DEATH TO MULTIORGAN FAILURE. ADDITIONALLY, THE PATIENT HAD RIGHT VENTRICULAR DYSFUNCTION, ELEVATED POTASSIUM LEVELS AND MULTIPLE MYOCARDIAL INFARCTIONS THE NIGHT OF PASSING. THIS CASE IS BEING REPORTED DUE TO THE FOLLOWING CONCLUSIONS: THE PATIENT DEVELOPED BRADYCARDIA AND OXYGEN DESATURATION DURING A GALAXY-ASSISTED BIOPSY PROCEDURE THAT REQUIRED MEDICAL INTERVENTION TO STABILIZE THE VITAL SIGNS AND THE PATIENT WAS ADMITTED. 2 DAYS LATER, THE PATIENT PASSED AWAY. THE PHYSICIAN DID NOT ATTRIBUTE THE BRADYCARDIA AND OXYGEN DESATURATION TO THE GALAXY SYSTEM AND STATES IT IS DUE TO THE BREATH HOLD AND THE PHYSICIAN ATTRIBUTED THE DEATH TO MULTIORGAN FAILURE.
CORTRAK 2 ENTERAL ACCESS SYSTEM | K240965 | IT WAS REPORTED VIA FDA USER FACILITY REPORT MDR NUMBER MW5168383 THE FOLLOWING INFORMATION: PT ADMITTED FOR A FALL AT HOME RESULTING IN A FRACTURE OF THE DISTAL CLAVICLE W MINIMAL DISPLACEMENT. NGT (NASOGASTRIC TUBE)/ CORTRACK WAS ORDERED THAT RESULTED IN 3-MINUTE CARDIAC ARREST. INJURY TO THE RIGHT MAINSTEM AND THROUGH THE LUNG. ABDOMINAL X-RAY SHOWED FEEDING TUBE WAS LOOPED OVER RIGHT HEMITHORAX CONCERNING FOR TRAUMATIC BRONCHOPLEURAL FISTULA AND POSSIBLE EXTENDING INTO THE PLEURAL SPACE. TYPE OF EVENT: DEATH, HEALTH EFFECT: CARDIAC ARREST, FISTULA, PLEURAL EFFUSION. | THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. THE LOT NUMBER WAS NOT PROVIDED BY THE CUSTOMER. ALL INFORMATION REASONABLY KNOWN AS OF 01-MAY-2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL, INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL, INC. AVANOS MEDICAL, INC. HAS NO INDEPENDENT KNOWLEDGE OF THE EVENT REPORTED BUT IS RELAYING THE INFORMATION THAT WAS PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL, INC. COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL, INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION A USER OF A DREAMSTATION CPAP PRO PASSED AWAY THREE YEARS AGO. THE PATIENT'S SISTER STATES THEY ARE RECEIVING LETTERS REGARDING THE DEVICE AND THE PATIENT WANTED THE LETTERS TO STOP BEFORE PASSING AWAY. THE PATIENT'S SISTER COULD NOT RECALL WHAT HAPPENED WITH THE RECALLED DEVICE OR THE REPLACEMENT DEVICE. IT WAS NOT RECEIVED BY THE MANUFACTURER. THE CAUSE OF DEATH WAS NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION VIA EMAIL TO THE NHS (NATIONAL HEALTH SERVICE) (SBS.APINVOICING@NHS.NET) ON 06/30/2025, 07/03/2025, 07/09/2025. NO RESPONSE WAS RECEIVED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
PHILIPS CPAP DEVICE | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING USER OF UNKNOWN DEVICE PASSED AWAY. THE PATIENT'S WIFE REPORTED THEY NEED TO RETURN THE PATIENT'S REPLACEMENT DEVICE SINCE THE PATIENT HAD PASSED AWAY BEFORE HE RECEIVED THE REPLACEMENT. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THIS REPORT WILL CAPTURE THE DEATH OF THE PATIENT ON THE DEVICE THAT WAS RECALLED SINCE THE PATIENT DID NOT USE THE REPLACEMENT DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING USER OF UNKNOWN DEVICE PASSED AWAY. THE PATIENT'S WIFE REPORTED THEY NEED TO RETURN THE PATIENT'S REPLACEMENT DEVICE SINCE THE PATIENT HAD PASSED AWAY BEFORE HE RECEIVED THE REPLACEMENT. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THIS REPORT WILL CAPTURE THE DEATH OF THE PATIENT ON THE DEVICE THAT WAS RECALLED SINCE THE PATIENT DID NOT USE THE REPLACEMENT DEVICE. THE DEVICE WAS RETURNED TO THE MANUFACTURER FOR EVALUATION. THE DEVICE WAS IN GOOD CONDITION AND PASSED EVALUATION. THE CUSTOMER COMPLAINT WAS NOT ADDRESSED. THE DEVICE WAS SCRAPPED.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH WAS NOT PROVIDED. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION VIA EMAIL TO THE NHS (NATIONAL HEALTH SERVICE) (SBS.APINVOICING@NHS.NET) ON 06/30/2025, 07/03/2025, 07/09/2025. NO RESPONSE WAS RECEIVED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
INDIGO SYSTEM FLASH ASPIRATION CATHETER | K240030 | THE PATIENT WAS UNDERGOING A THROMBECTOMY PROCEDURE IN THE LEFT ILIAC AND FEMORAL VEIN FOR A DEEP VEIN THROMBOSIS (DVT) USING AN INDIGO SYSTEM LIGHTNING FLASH ASPIRATION TUBING (FLASH TUBING), AN INDIGO SYSTEM FLASH ASPIRATION CATHETER (FLASH CATHETER), A BALLOON (10X40), GUIDEWIRES, A MICROSHEATH, AND SHEATHS (6FR AND 16FR) IN A PRONE POSITION VIA THE LEFT LOWER EXTREMITY. DURING THE PROCEDURE, AN INITIAL VENOGRAM WAS TAKEN WHICH SHOWED EXTENSIVE CLOT IN THE FEMORAL, COMMON FEMORAL, AND ILIAC VENOUS SYSTEM. THE 6FR SHEATH WAS UPSIZED TO A 16FR SHEATH. SEVERAL PASSES WERE MADE IN THE TARGET LOCATION USING THE FLASH CATHETER. A VENOGRAM SHOWED SOME RESIDUAL NARROWING OF THE COMMON FEMORAL AND FEMORAL VEIN SO AN ANGIOPLASTY WAS PERFORMED WITH A BALLOON CATHETER. A POST-INTERVENTION VENOGRAM SHOWED NO RESIDUAL STENOSIS. THE VEIN WAS WIDELY PATENT. THE FLASH CATHETER AND SHEATH WERE THEN REMOVED FROM THE PATIENT. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS PLACED IN A SUPINE POSITION ON THE STRETCHER. IT WAS REPORTED THAT THE PATIENT WAS GETTING READY TO BE EXTUBATED AND BEGAN TO EXPERIENCE OXYGEN DESATURATION AND TACHYCARDIA, WHICH BECAME PROGRESSIVELY WORSE. THE PATIENT THEN BECAME HYPOTENSIVE AND BRADYCARDIC. A CODE WAS CALLED, MULTIPLE ROUNDS OF MEDICATIONS WERE GIVEN, AND 3 ROUNDS OF CARDIOPULMONARY RESUSCITATION (CPR) WERE PERFORMED. TWO UNITS OF PACKED RED BLOOD CELLS (PRBCS) WERE TRANSFUSED. THE PATIENT WAS ABLE TO REGAIN HER BLOOD PRESSURE. THE PHYSICIAN SUSPECTED THAT THE PATIENT HAD A PULMONARY EMBOLISM, SO AN INTERVENTIONAL RADIOLOGIST WAS CONSULTED. ONCE THE PATIENT WAS STABLE, SHE WAS MOVED BACK ONTO THE TABLE. THE PATIENT UNDERWENT A THROMBECTOMY PROCEDURE IN BILATERAL PULMONARY ARTERIES FOR PULMONARY EMBOLISM USING NON-PENUMBRA CATHETERS, A NON-PENUMBRA THROMBECTOMY SYSTEM, AND A SHEATH (6FR) VIA THE RIGHT GROIN. POST-INTERVENTION ANGIOGRAM DEMONSTRATED SCATTERED BILATERAL SEGMENTAL AND SUBSEGMENTAL EMBOLI. THE CATHETER AND SHEATH WERE REMOVED, AND HEMOSTASIS WAS ACHIEVED. THE PROCEDURE WAS COMPLETED AND THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU). IT WAS REPORTED THAT THE PATIENT'S CONDITION CONTINUED TO DETERIORATE WITH MAXIMUM PRESSOR SUPPORT, AND THE PATIENT WAS SHOCKED TWICE WITH NO IMPROVEMENT IN STATUS. THE FAMILY DECIDED TO WITHDRAW CARE, AND THE PATIENT EXPIRED ON (B)(6) 2025. THE CEC CHAIR ADJUDICATED SHOCK/HEMODYNAMIC INSTABILITY DUE TO PULMONARY EMBOLISM TO BE A SERIOUS ADVERSE EVENT WITH A PROBABLE RELATIONSHIP TO THE INDEX PROCEDURE AND A POSSIBLE RELATIONSHIP TO THE INDIGO ASPIRATION SYSTEM. | THE PRODUCT WAS NOT RETURNED FOR EVALUATION. THEREFORE, A PHYSICAL DEVICE INVESTIGATION WAS UNABLE TO BE PERFORMED. FROM THE INFORMATION PROVIDED, THERE IS NO INDICATION THAT THERE WAS ANY DEVICE MALFUNCTION, NONCONFORMANCE, OR MISUSE THAT CONTRIBUTED TO THE REPORTED EVENT. THE MANUFACTURING RECORDS FOR THIS LOT WERE REVIEWED AND DID NOT REVEAL ANY OUTSTANDING DISCREPANCIES, DESIGN, OR QUALITY CONCERNS.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE USER OF A REMSTAR PRO C-FLEX+ PASSED AWAY. THE CAUSE OF THE DEATH WAS NOT PROVIDED. THE DEVICE WAS RETURNED TO THE MANUFACTURER IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. DURING THE EVALUATION OF THE DEVICE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. IF ADDITIONAL INFORMATION IS AVAILABLE A FOLLOW-UP REPORT WILL BE SUBMITTED.
REMSTAR AUTO A-FLEX W/HUMID, SD CARD | K091319 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REMSTAR AUTO DEVICE DEVELOPED CANCER AND HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION DEVICE DEVELOPED MELANOMA AND HAS PASSED AWAY. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED VENTRAL HERNIA ¿ COMPLEX TRANSVERSUS ABDOMINUS RELEASE (TAR) REPAIR, THE PATIENT WAS KEPT OVERNIGHT FOR OBSERVATION AND LABORATORY RESULTS. THE SURGEON INFORMED THE INTUITIVE SURGICAL CLINICAL SALES REPRESENTATIVE (CSR) THAT THE PATIENT WAS DOING WELL POST-OPERATIVELY; HOWEVER, EXPIRED OVER NIGHT DUE TO UNKNOWN CAUSES. THE SURGEON REPORTED THAT THE CAUSE OF DEATH WAS LIKELY A HEART ATTACK OR A PULMONARY EMBOLISM. THE SURGEON SPECIFIED THAT THE PATIENT OUTCOME WAS NOT CAUSED BY THE DA VINCI-ASSISTED PROCEDURE AND THERE WERE NO DA VINCI ISSUES DURING THE PROCEDURE. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE REQUESTED; NO RESPONSE WAS RECEIVED. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, MARYLAND BIPOLAR FORCEPS, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING OR SUBSEQUENT TO THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED UNEXPECTEDLY AFTER AN UNEVENTFUL HERNIA REPAIR. NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT HAS BEEN MADE AND THE CASE HAD NO ISSUES. THE SURGEON STATED THE PATIENT MAY HAVE DIED OF A MYOCARDIAL INFARCTION OR PULMONARY EMBOLISM BUT NO SPECIFIC INFORMATION ALLUDING TO THESE CAUSES IS PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED - SURGEON REPORTED THE PATIENT EXPIRED DURING THE NIGHT, TIME NOT PROVIDED.
DREAMSTATION BIPAP AUTOSV, DS | K090539 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION BIPAP AUTOSV DEVICE HAS PASSED AWAY AND THE WIFE WOULD LIKE TO RETURN IT. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REP DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY AND THE SON WOULD LIKE TO HAVE THE FATHER'S INFORMATION CLEARED AND HIS ENTERED SO HE COULD USE THE DEVICE. THE RESPIRATORY THERAPIST DIRECTED THE SON TO THE DME FOR THOSE CHANGES. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WILL NOT BE RETURNED FOR INVESTIGATION AS THE SON WOULD LIKE TO USE THE DEVICE. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION ST30. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DS2ADV AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT OF THE DREAMSTATION AUTO BIPAP DEVICE PASSED AWAY IN 2021 WHILE ON HOSPICE CARE. THE PATIENT WAS NOT USING THE DEVICE TOWARD THE END OF HER LIFE, AND THE DAUGHTER CONTACTED PHILIPS TO RECYCLE THE DEVICE. A SHIPPING LABEL WAS CREATED, AND THE DEVICE WAS RETURNED TO PHILIPS. THE DEVICE WAS EVALUATED, AND NO VISIBLE FOAM DEGRADATION OR PROBLEMS WERE FOUND. THERE IS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT PASSED AWAY ONE MONTH AGO AND THE REP RECRT DREAMSTATION AUTO CPAP DEVICE WAS THROWN AWAY AND THE PATIENT "WOULD NOT HAVE BEEN USING THIS DEVICE". A REPLACEMENT FOR THE DEVICE WAS DECLINED. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT THE PATIENT OF THE REMSTAR AUTO A-FLEX DEVICE PASSED AWAY FROM RESPIRATORY FAILURE. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FOLLOW UP REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED.
COMBISET SMARTECH BLOODLINE (NO HEPARIN) | K201207 | IT WAS REPORTED TO FRESENIUS THAT A BLOOD LEAK OCCURRED DURING A HEMODIALYSIS (HD) PATIENT¿S TREATMENT, WHICH SUBSEQUENTLY RESULTED IN THEIR DEATH. ADDITIONAL DETAILS WERE PROVIDED UPON FOLLOW-UP AND REVIEW OF THE PATIENT¿S TREATMENT SHEET. THE PATIENT¿S PRE-TREATMENT VITAL SIGNS AT 10:33AM WERE CONFIRMED AS BLOOD PRESSURE (BP) 115/80, PULSE (P) 73, RESPIRATIONS (R) 16, TEMPERATURE (T) 96.6°. THE TREATMENT WAS INITIATED UTILIZING A CENTRAL VENOUS CATHETER (CVC) IN THE CHEST, ALONG WITH FRESENIUS COMBISET SMARTECH BLOODLINES, 2008T HD MACHINE, AND PATIENT CONNECTOR CLIP (NON-STERILE). THE CVC WAS REPORTEDLY WIPED WITH ALCOHOL PRIOR TO CONNECTING THE BLOODLINES. THE TREATMENT SHEET INDICATED THAT THE CVC CATHETER DRESSING WAS GAUZE WITH A SOLUTION OF 2% CHLORHEXIDINE GLUCONATE AND 70% ISOPROPYL ALCOHOL. THE ARTERIAL PRESSURE SETTING ON THE 2008T MACHINE WAS 160MMHG AND THE VENOUS PRESSURE WAS 100 ASYMMETRICAL. THE PATIENT DID NOT REPORT ANY CONCERNS PRIOR TO THE START OF TREATMENT. THE TREATMENT SHEET DOCUMENTED THE ULTRAFILTRATION RATE (UFR) 760, FLUID REMOVED 38, FLUID GIVEN 300, THE ACCESS AND CONNECTOR CLIP WERE CHECKED, THE PATIENT WAS RESTING COMFORTABLY, THE PATIENT WAS ALERT, THE ACCESS WAS VISIBLE, AND TREATMENT WAS INITIATED WITHOUT PROBLEMS. THE GREEN AMP LIGHT WAS ON. THE FOLLOWING VITAL SIGNS AND NOTES WERE RECORDED DURING THE TREATMENT: 11:00AM BP 144/54, P 64, ACCESS AND CONNECTOR CLIP CHECKED; 11:02AM BP 114/64, P 64, ACCESS AND CONNECTOR CLIP CHECKED, PATIENT WAS ALERT WITH NO COMPLAINTS, THE GREEN AMP LIGHT WAS ON; 11:31AM BP 154/71, P 66, ARTERIAL PRESSURE 182, VENOUS PRESSURE 125, GREEN AMP LIGHT ON, ACCESS VISIBLE, PATIENT DENIES COMPLAINTS, ACCESS AND CONNECTOR CLIP CHECKED, UFR 760, FLUID REMOVED 723. AT 11:45AM THE MACHINE ALARMED WITH AN UNSPECIFIED MACHINE ALARM (INITIALLY REPORTED AS AN ARTERIAL PRESSURE ALARM). THE PATIENT¿S BP WAS DOCUMENTED AS BP 87/31, P 125. THE TREATMENT SHEET DOCUMENTED THAT THE ACCESS WAS CHECKED, AND THE GREEN AMP LIGHT WAS ON. AT THIS TIME, IT WAS NOTED THERE WAS A BLOOD LEAK. THE LEAK WAS REPORTEDLY COMING FROM THE CONNECTION OF THE CATHETER LUMEN TO THE COMBISET VENOUS BLOOD LINE. THE CONNECTION WAS LOOSE, BUT NOT COMPLETELY DISCONNECTED. THE PATIENT CONNECTOR CLIP WAS REPORTEDLY STILL CONNECTED TO THE BLOOD LINE AND CATHETER. WHEN THE BLOOD LEAK WAS DETECTED, THE PATIENT WENT UNRESPONSIVE. A PULSE WAS NOT PALPABLE. TREATMENT WAS DISCONTINUED AND THE PATIENT¿S BLOOD FROM THE EXTRACORPOREAL CIRCUIT WAS IMMEDIATELY RETURNED WITH NORMAL SALINE (NS). CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED. ANOTHER STAFF MEMBER ACTIVATED 911. THE PATIENT WAS ADMINISTERED OXYGEN VIA AMBU MASK. 3L OF NS WERE ADMINISTERED WHILE WAITING FOR EMERGENCY MEDICAL SERVICES (EMS). NO PATIENT VITALS WERE DOCUMENTED AFTER BECOMING UNRESPONSIVE. EMS ARRIVED AND IMMEDIATELY TRANSPORTED THE PATIENT TO THE HOSPITAL. AT THE HOSPITAL, THE PATIENT RECEIVED CPR, MULTIPLE MEDICATIONS (UNKNOWN), AND WAS INTUBATED. THE PATIENT EXPIRED AT 4:00PM. THE CAUSE OF DEATH DOCUMENTED IN THE MEDICAL RECORD DISCHARGE IS CARDIAC ARREST, ANEMIA, AND END STAGE RENAL DISEASE (ESRD). NO ESRD DEATH NOTIFICATION FORM CMS-2746-U2 WAS AVAILABLE. THE MEDICAL RECORD DOCUMENTED THE PATIENT¿S ESTIMATED BLOOD LOSS (EBL) AT 2L. IT WAS DOCUMENTED THAT PRIOR TO THE START OF THE PATIENT¿S TREATMENT, THE CONNECTION AT THE COMBISET VENOUS BLOODLINE AND CATHETER WAS SECURE. THE PATIENT CONNECTOR CLIP WAS REPORTEDLY ON AT THE START OF TREATMENT AND AT THE TIME OF THE BLOOD LEAK, HOWEVER; IT COULD NOT BE CONFIRMED IF THE CLIP REMAINED ON THE BLOODLINE POST EVENT. THE PATIENT DID NOT COMPLAIN OF ANY SYMPTOMS AT ANY TIME PRIOR TO THE ADVERSE EVENT. THERE WERE NO ADDITIONAL ALARMS OR TREATMENT ISSUES DOCUMENTED PRIOR TO THE UNKNOWN MACHINE ALARM. THE PATIENT¿S ACCESS SITE WAS REPORTED TO BE VISIBLE DURING THE TREATMENT. NO FURTHER INFORMATION WAS PROVIDED. THE SAMPLE WAS REPORTED TO BE AVAILABLE FOR MANUFACTURER EVALUATION. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN HD THERAPY UTILIZING THE COMBISET SMARTECH BLOODLINES AND THE PATIENT EVENT OF BLOOD LOSS WITH SUBSEQUENT DEATH. ALTHOUGH IT WAS REPORTED THAT A BLOOD LEAK OCCURRED FROM THE CONNECTION OF THE CATHETER LUMEN AND THE COMBISET VENOUS BLOOD LINE, THERE IS NO DOCUMENTATION OF ANY DEFECTS OR DAMAGE NOTED AT THE CONNECTION SITE PRIOR TO THE START OF TREATMENT. ADDITIONALLY, THERE WERE NO OTHER REPORTED ISSUES LEADING UP TO THE EVENT THAT WOULD CAUSE THE REPORTED LOOSE CONNECTION. IT WAS REPORTED THAT THE CONNECTION WAS SECURE PRIOR TO THE START OF TREATMENT. IN ADDITION, THERE ARE SEVERAL DOCUMENTED CHECKS THROUGHOUT THE TREATMENT THAT INDICATE THE PATIENT ACCESS, AND THE PATIENT CONNECTOR CLIP WERE CHECKED. THE COMBISET SMARTECH BLOODLINE IFU STATES TO ¿ENSURE ALL CONNECTIONS ARE SECURE BEFORE USE AND MONITOR FOR LEAKS REGULARLY DURING PATIENT USE. BLOOD LEAKS CAN RESULT IF CONNECTIONS ARE NOT SECURE.¿ ALTHOUGH IT WAS REPORTED THAT THE PATIENT CONNECTOR CLIP WAS PLACED PRIOR TO THE INITIATION OF TREATMENT AND DOCUMENTED AS STILL BEING CONNECTED AT THE TIME OF THE EVENT, IT COULD NOT BE CONFIRMED IF THE CLIP WAS STILL ATTACHED POST EVENT. THE COMBISET SMARTECH BLOODLINES IFU ADDITIONALLY STATES ¿USE OF CATHETERS, AVF NEEDLE SETS OR OTHER DEVICES NOT COMPATIBLE WITH THIS SET MAY RESULT IN BLOOD/AIR LEAKAGE OR ACCIDENTAL DISCONNECTION, AIR EMBOLUS, SIGNIFICANT BLOOD LOSS, PATIENT INJURY OR DEATH.¿ IT WAS ALSO DOCUMENTED SEVERAL TIMES THAT THE PATIENT¿S ACCESS SITE WAS UNCOVERED AND VISIBLE DURING TREATMENT.
ALARIS SYSTEM | K133532 | IT WAS INITIALLY REPORTED THE DEVICE HAD AN EVENT WHERE THE PUMP DELIVERED AN INCORRECT DOSE. THE DEVICE WAS INVESTIGATED BY THE FACILITY'S BIOMEDICAL ENGINEERING AND PER REPORT, IT "PASSED PREVENTATIVE MAINTENANCE AND RETURNED BACK TO SERVICE". THERE WAS PATIENT INVOLVEMENT HOWEVER OUTCOME IS UNKNOWN. BD RECEIVED ADDITIONAL INFORMATION THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT THE EVENT INVOLVED A "COMFORT CARE PATIENT" ON A CONTINUOUS MORPHINE DRIP INFUSING AT 5MG/HR. FROM A 100ML (CONCENTRATION: 1MG/1ML) BAG. THE INFUSION WAS INITIATED AT 12:14. AT 14:09 THE NURSE RESPONDED TO THE ROOM AND FOUND THAT THE BAG WAS EMPTY. THE BAG AND TUBING WERE INSPECTED BY CUSTOMER; NO LEAKS WERE NOTED PER REPORT. IT WAS REPORTED THE PATIENT LATER PASSED AWAY. THE PATIENT'S FAMILY REQUESTED NO FURTHER INTERVENTIONS. THE EVENT OCCURRED ON 11 JANUARY 2025. PER CUSTOMER, THE EVENT WAS NOT IMMEDIATELY REPORTED TO BD AS THEY "DID NOT HAVE REASON TO BELIEVE THERE WAS AN EQUIPMENT MALFUNCTION." THE CUSTOMER, AT THE TIME, BELIEVED THAT THE ISSUE WAS DUE TO "DRUG DIVERSION". HOWEVER, THE SAME DEVICE HAS BEEN INVOLVED IN ANOTHER (CAPTURED UNDER SEPARATE RECORD MFR REPORT # 2016493-2025-77596) AND THIS PROMPTED THE CUSTOMER TO REPORT THE EVENT TO BD. | THE ACTUAL DATE OF DEATH IS UNKNOWN. BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
FLEXCATH CONTOUR¿ STEERABLE SHEATH | K232321 | IT WAS LATER REPORTED THAT THE PATIENT WAS IN PULSELESS ELECTRICAL ACTIVITY WHEN THE BLOOD PRESSURE DROPPED. PRIOR TO THE PROCEDURE, INTRACARDIAC ECHOCARDIOGRAPHY (ICE) DID NOT REVEAL ANY PERICARDIAL EFFUSION. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
INTELLISPACE EVENT MGMT PLATFORM | K102974 | IT WAS REPORTED THE DEVICE DID NOT GENERATE AN AUDIBLE ALARM ON THE EMERGIN PAGER AFTER THE PATIENT WENT INTO CARDIAC ARREST. IT WAS INDICATED THE VISUAL ALARM MESSAGE WAS GENERATED BUT NO AUDIBLE BEEP ON THE PAGING DEVICE. BIOMED TESTED THE ALARM CHAIN, REVEALING ALARMS OCCURRED VISUALLY AND AUDIBLY AT THE CENTRAL STATION AND ON THE EMERGIN PAGING DEVICE. IT WAS REPORTED THE NURSES BEGAN CARDIAC MASSAGE. IMMEDIATELY FOLLOWING THE ALARM AT THE CENTRAL STATION, THE USER DISCOVERED THE PATIENT IN CARDIAC ARREST, FROM WHICH THE PATIENT PASSED AWAY. THE LOGS WERE REVIEWED, REVEALING THE ALARM MESSAGE WAS DISTRIBUTED TO THE THREE PAGERS, WITH THE ALARM MESSAGE BEING FOUND ON THE PAGER SCREENS. THE CUSTOMER WAS NOT SURE WHETHER THE DEVICES PROVIDED AN AUDIBLE ALERT. | THE LOGS HAVE BEEN REVIEWED BY A PRODUCT SUPPORT ENGINEER (PSE). IT HAS BEEN ESTABLISHED THAT MULTIPLE ALERTS WERE SENT OUT OF THE IEM SERVER DURING THE DATE AND TIME OF THE EVENT. THE IEM SERVER PROCESSED THE ALERTS AND SENT THEM TO ASCOM FOR DELIVERY TO THE BEEPERS. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
HOMECHOICE APD SYSTEM | K102936 | IT WAS REPORTED THAT A PATIENT PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. IT WAS NOT REPORTED IF THE PATIENT WAS HOSPITALIZED PRIOR TO DEATH. IT WAS REPORTED THAT THE PATIENT WAS CONNECTED DURING THE TIME OF THE EVENT. IT WAS NOT REPORTED IF AN AUTOPSY WAS PERFORMED. IT WAS NOT REPORTED IF PD THERAPY WAS ONGOING PRIOR TO DEATH OR AT THE TIME OF DEATH. NO ADDITIONAL INFORMATION WAS AVAILABLE. | ADDITIONAL INFORMATION: H6 AND H11. H11: THE DEVICE WAS NOT RECEIVED FOR EVALUATION; THEREFORE A DEVICE ANALYSIS COULD NOT BE COMPLETED. A SERVICE HISTORY REVIEW REVEALED NO INDICATION THAT THE PARTS REPLACED DURING SERVICING CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE REPORTED EVENT WAS NOT VERIFIED. THE CAUSE OF THE CONDITION COULD NOT BE DETERMINED. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
FLEXCATH CONTOUR¿ STEERABLE SHEATH | K232321 | IT WAS REPORTED THAT THE PATIENT PASSED AWAY. IT WAS REPORTED THAT POST A PULSE FIELD ABLATION PROCEDURE, THE PATIENT EXPERIENCED APPARENT STROKE-LIKE SYMPTOMS AFTER WAKING FROM ANESTHESIA IN THE HOLDING AREA. THE PATIENT WAS SUBSEQUENTLY TAKEN TO INTERVENTIONAL RADIOLOGY (IR) FOR REVASCULARIZATION AND ADMITTED TO THE INTENSIVE CARE UNIT. DESPITE THE INTERVENTION, THE PATIENT CONTINUED TO SUFFER FROM GLOBAL DYSPHASIA AND STROKE-LIKE SYMPTOMS. THE PATIENT THEN UNDERWENT HEMORRHAGIC CONVERSION AND EXPERIENCED A SEIZURE. THREE DAYS LATER DO NOT RESUSCITATE (DNR) ORDERS WERE PUT IN PLACE FOR THE PATIENT AND THE PATIENT THEN PASSED AWAY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
HLS SET | K112360 | THE EVENT OCCURRED IN SPAIN. MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON (B)(6) 2025: THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON (B)(6) 2024, IN HOSPITAL (B)(6), SPAIN. THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." THE GETINGE SALES AND SERVICE UNIT (SSU) CONFIRMED ON (B)(6) 2025, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. COMPLAINT ID #(B)(4). | MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON (B)(6) 2025: "THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON (B)(6) 2024, IN HOSPITAL (B)(6), SPAIN. THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. HOWEVER, THE SSU CONFIRMED ON (B)(6) 2025, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. AFTER CONSULTATION WITH THE CUSTOMER, THE SSU CONFIRMED ON (B)(6) 2025, THAT THE INVOLVED HLS SET WAS FUNCTIONING AS INTENDED AND HAD NO MALFUNCTION. THE PATIENT PASSED AWAY DUE TO HIS DISORDERS. THE AFFECTED HLS SET WAS DISCARDED BY THE CUSTOMER. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "DEATH" COULD BE CONFIRMED BUT WAS NOT RELATED TO A DEVICE MALFUNCTION. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS AS A SINGLE EVENT (TIMEFRAME FROM (B)(6) 2024 TILL (B)(6) 2025). THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE SPANISH MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER (B)(4).
SAFARI2¿ | K151244 | EVENT DESCRIPTION: ON (B)(6) 2025, DURING TAVI CASE AFTER THE FINAL RELEASE OF THE ACURATE PRIME VALVE 27MM, THE VALVE WAS IMPLANTED AT HIGH POSITION (PROBABLY BECAUSE THERE WAS A RAPH BETWEEN RIGHT AND NON-CORONARY CUSP. HOWEVER, THE CT SCAN ANALYSIS WAS SUBOPTIMAL AND THIS IS NOT 100% CONFIRMED. BEFORE CASE, THE PHYSICIANS WERE ALSO AWARE OF THE RAPHE). THE PHYSICIANS ESTIMATED THAT THE HEMODYNAMIC RESULT NEEDED TO BE IMPROVED, SO THEY DECIDED TO DO POST-DILATATION (IGNORING MY WARNINGS OF THE RISK TO DO POST-DILATATION IN A HIGH POSITION) AND THE VALVE MOVED AORTIC. THEN, AFTER CHECKING THAT THE CORONARY FLOW WAS GOOD, THEY DECIDED TO DO VALVE IN VALVE. AFTER THE VALVE IN VALVE, THE HEMODYNAMICS WAS GOOD. ON (B)(6), THE PATIENT HAD A STROKE AND DISSECTION OF AORTA. ON (B)(6), THE PHYSICIANS ANNOUNCED DEATH. PATIENT PRESENT AT TIME OF EVENT? YES. PATIENT COMPLICATIONS: DEATH. IN THE PHYSICIAN'S OPINION, DID THE DEVICE OR PROCEDURE CAUSE OR CONTRIBUTE TO THE PATIENT COMPLICATION? NO. MEDICAL OR SURGICAL INTERVENTIONS: VALVE IN VALVE. PATIENT OUTCOME: DEATH. 23APR2025: THE FOLLOWING INFORMATION RECEIVED VIA A RELATED ACURATE PRIME COMPLAINT INSTIGATED THE NEED FOR A SAFARI2 COMPLAINT: DID THE PHYSICIAN ALLEGE ANY DIFFICULTIES WERE EXPERIENCED WITH THE GUIDEWIRE DURING THE PROCEDURE? DURING THE SECOND VALVE (VIV), THE SAFARI WIRE'S LOOP WAS NOT INTACT. ADDITIONAL INFORMATION: EVENT SUMMARY: IT WAS FURTHER REPORTED THE SEVERELY CALCIFIED NATIVE AORTIC ANNULUS MEASURED 25.2MM AND A SAFARI2 GUIDEWIRE AND 14F ISLEEVE INTRODUCER SHEATH WERE USED DURING THE TAVR PROCEDURE. PRE-DILATATION BALLOON AORTIC VALVULOPLASTY (BAV) WAS PERFORMED ON THE FIRST ACURATE PRIME VALVE WITH A 23MM NON-BOSTON SCIENTIFIC BALLOON CATHETER. AFTER IMPLANTATION OF THE FIRST ACURATE PRIME VALVE, MODERATE AORTIC REGURGITATION WAS IDENTIFIED WHICH NECESSITATED POST-DILATATION BAV. WHEN THE FIRST ACURATE PRIME VALVE MOVED AORTICALLY IT LOST CONTACT WITH THE NATIVE AORTIC ANNULUS. DURING THE VALVE-IN-VALVE PROCEDURE THE LOOP OF THE SAFARI2 GUIDEWIRE WAS NOT INTACT. THE AORTIC DISSECTION WAS IDENTIFIED AS THE SUSPECTED SOURCE OF THE STROKE. 24APR2025: PLEASE REQUEST INFORMATION TO HELP OUR TEAM UNDERSTAND WHAT IS MEANT BY "DURING THE SECOND VALVE (VIV), THE SAFARI WIRE'S LOOP WAS NOT INTACT". WAS THERE DAMAGE TO THE FORMED CURVE DISTAL REGION OF THE WIRE AND IF SO, PLEASE SPECIFY THE TYPE OF DAMAGE AND ASK FOR IMAGES SHOWING ANY DAMAGE. FINAL RELEASE OF THE ACURATE PRIME VALVE WAS PERFORMED AND THE VALVE FRAME WAS SHAPED LIKE A CONE, COMPRESSED AT LOWER CROWNS AND IN A HIGH POSITION. THE SAFARI2 GUIDEWIRE WAS INVERTED AT FINAL VALVE RELEASE. AFTER CONFIRMING CORONARIES WERE PATENT A VALVE-IN-VALVE WAS PREPARED. THE SECOND PRIME VALVE WAS PREPARED AND THE SAFARI2 NOT IN AN OPTIMAL POSITION AT RELEASE. 25APR2025: QUESTION: LISTING AND MODEL OF OTHER DEVICES USED DURING THE PROCEDURE, INCLUDE THE MODEL, BRAND NAME, SIZES, ETC. ANSWER: ACURATE PRIME VALVE, ISLEEVE INTRODUCER SHEATH, ACURATE PRIME DELIVERY SYSTEM, 23MM OSYPKA VACS III. QUESTION: DID THE END USER MONITOR THE WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF THE CURVE AND DISTAL TIP? ANSWER: YES. QUESTION: WAS THE CURVE OF THE SAFARI2 GUIDEWIRE CONSTRAINED WITHIN A CATHETER DURING INSERTION INTO AND WITHDRAWAL FROM THE BODY OR TREATMENT SITE? ANSWER: NO. QUESTION: DID THE END USER ADVANCE THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE? ANSWER: UNKNOWN. QUESTION: WAS THE WIRE POSITION MAINTAINED WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE? ANSWER: YES. QUESTION: IS THERE ANY ADDITIONAL IMAGING (E.G., INTRAOPERATIVE FLUOROSCOPY, ECG) THAT WAS PERFORMED DURING THE PROCEDURE THAT MIGHT PROVIDE MORE INSIGHTS INTO THE GUIDEWIRE POSITIONING AND VALVE DEPLOYMENT PROCESS? ANSWER: ALL AVAILABLE MEDIA HAS BEEN SUBMITTED. QUESTION: IS THIS A NEW USER OR EXPERIENCED USER? ANSWER: EXPERIENCED. QUESTION: WAS THE AORTIC DISSECTION IDENTIFIED BEFORE OR AFTER THE STROKE OCCURRED? ANSWER: AFTER, STROKE WAS SUSPECTED AND MRI CONFIRMED DISSECTION. QUESTION: WERE THERE ANY KNOWN ANATOMICAL OR PATHOLOGICAL FINDINGS PRE-PROCEDURE THAT MIGHT HAVE PREDISPOSED THE PATIENT TO DISSECTION? ANSWER: FUNCTIONALLY BICUSPID ANATOMY. QUESTION: WAS THE PRESENCE OF A RAPHE BETWEEN THE RIGHT AND NON-CORONARY CUSP CONFIRMED POST EVENT OR IS IT STILL PRESUMED? ANSWER: CONFIRMED VIA MEDIA. QUESTION: IS THE WIRE AVAILABLE TO BE RETURNED AND IF NOT, PLEASE PROVIDE A REASON FOR NO RETURN? ANSWER: DEVICE WAS DISCARDED AT HEALTHCARE FACILITY. ALTHOUGH THERE IS NO KNOWN RELATIONSHIP BETWEEN THE SAFARI2 DEVICE AND ADVERSE EVENT, PERFORATION AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE SAFARI2 DEVICE AND ARE LISTED IN THE INSTRUCTIONS FOR USE (IFU). THEREFORE, THIS INCIDENT IS BEING FILED AS A DEATH AS A GOOD FAITH MEASURE. | PRODUCT WAS DISCARDED AT THE HEALTHCARE FACILITY; THEREFORE, NO VISUAL OR PHYSICAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. AS DESCRIBED IN THE DEVICE INSTRUCTIONS FOR USE (IFU) WARNINGS SECTION: MONITOR THE WIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF CURVE AND DISTAL TIP. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH-RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. NEVER ADVANCE THE GUIDEWIRE AGAINST THE RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING THE GUIDEWIRE AFTER DEVICE DEPLOYMENT. THE WIRE SHOULD ONLY BE INTRODUCED INTO OR WITHDRAWN FROM THE VENTRICLE THROUGH A CATHETER ALREADY POSITIONED INTO THE VENTRICLE. THE CURVE OF THE SAFARI2TM GUIDEWIRE SHOULD BE CONSTRAINED WITHIN A CATHETER DURING INSERTION INTO OR WITHDRAWAL FROM THE BODY OR TREATMENT SITE. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE: 1. ENSURE A CATHETER IS APPROPRIATELY PLACED IN THE VENTRICLE OR OTHER INTENDED TREATMENT AREA. 2. CAREFULLY REMOVE THE SAFARI2 GUIDEWIRE, FROM THE HOOP BY GRASPING THE PROXIMAL END OF THE J-STRAIGHTENER SEPARATING THE STRAIGHTENER FROM THE HOOP. 3. AFTER INSPECTION OF THE SAFARI2 GUIDEWIRE, ADVANCE THE J-STRAIGHTENER OVER THE DISTAL SECTION OF THE SAFARI2 WIRE TO STRAIGHTEN THE CURVE. 4. INSERT THE SAFARI2 GUIDEWIRE INTO THE HUB OF THE CATHETER WITH THE AID OF THE J- STRAIGHTENER. 5. CAREFULLY ADVANCE THE DISTAL TIP OF THE SAFARI2 GUIDEWIRE INTO THE LUMEN OF THE CATHETER. 6. REMOVE THE SAFARI2 GUIDEWIRE J- STRAIGHTENER BY WITHDRAWING IT OVER THE LENGTH OF THE SAFARI2 GUIDEWIRE. 7. ADVANCED THE SAFARI2 GUIDEWIRE THROUGH THE CATHETER UNDER FLUOROSCOPIC GUIDANCE. 8. CONFIRM SAFARI2 GUIDEWIRE CURVE POSITION TO ASSURE SAFARI2 GUIDEWIRE PLACEMENT AND STABILITY. 9. MAINTAIN WIRE POSITION WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE. 10. TO REMOVE THE GUIDEWIRE FROM THE TREATMENT AREA: A. A CATHETER MUST BE ADVANCED INTO THE TREATMENT AREA OVER THE SAFARI2 GUIDEWIRE PRIOR TO WITHDRAWING THE WIRE. B. THE SAFARI2 GUIDEWIRE IS PULLED BACK COMPLETELY INTO THE CATHETER. C. THE SAFARI2 GUIDEWIRE MAY THEN BE WITHDRAWN FROM THE CATHETER. AT THIS TIME, IT IS NOT POSSIBLE TO ASSIGN A DEFINITIVE ROOT CAUSE FOR THE EVENT AS REPORTED. BASED ON THE INFORMATION PROVIDED, IT APPEARS THAT PROCEDURAL AND/OR CLINICAL FACTORS IMPACTED ON THE EVENT AS REPORTED. SHOULD ADDITIONAL INFORMATION BE RECEIVED, A FOLLOW UP MEDWATCH REPORT WILL BE SUBMITTED. ALTHOUGH ANNEX G INSTRUCTIONS STATE THAT WITH STAND-ALONE DEVICES ANNEX G TERM SHOULD NOT BE USED, CODE 4755 WAS SELECTED DUE TO COMPONENT CODE (G) SHOWING UP ON THE MISSING DATA REPORT.
HLS SET | K112360 | THE EVENT OCCURRED IN SPAIN. MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON 2025-04-01: THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON 2024-11-16, IN HOSPITAL (B)(6). THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." THE GETINGE SALES AND SERVICE UNIT (SSU) CONFIRMED ON 2025-04-20, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. COMPLAINT ID#: (B)(4). | MAQUET CARDIOPULMONARY GMBH HAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICE ON 2025-04-01: "THE VIGILANCE UNIT OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT, INVOLVING ¿PLS SET¿, UDI-DI (B)(4) MANUFACTURED BY GETINGE WHICH OCCURRED ON 2024-11-16, IN HOSPITAL (B)(6). THE FOLLOWING INFORMATION WAS PROVIDED: INCIDENT DESCRIPTION: DEATH." BASED ON THE INFORMATION AVAILABLE THAT THE PATIENT EXPIRED, A REPORT IS REQUIRED. HOWEVER, THE SSU CONFIRMED ON 2025-04-20, THAT THE AFFECTED PRODUCT IN QUESTION WAS AN HLS SET NOT A PLS SET. AFTER CONSULTATION WITH THE CUSTOMER, THE SSU CONFIRMED ON 2025-04-24, THAT THE INVOLVED HLS SET WAS FUNCTIONING AS INTENDED AND HAD NO MALFUNCTION. THE PATIENT PASSED AWAY DUE TO HIS DISORDERS. THE AFFECTED HLS SET WAS DISCARDED BY THE CUSTOMER. BASED ON THE AVAILABLE INFORMATION THE REPORTED EVENT "DEATH" COULD BE CONFIRMED BUT WAS NOT RELATED TO A DEVICE MALFUNCTION. THIS COMPLAINT WAS FOUND IN THE DATABASE OF CUSTOMER COMPLAINTS AS A SINGLE EVENT (TIMEFRAME FROM 2024-03-31 TILL 2025-03-31). THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: THIS EVENT OCCURRED ON THE SPANISH MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER: 701069073.
CHORDX UNKNOWN | K174060 | ACCORDING TO PUBLICATION "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE. INSIGHTS FROM A 6 YEAR SINGLE CENTER STUDY¿, THIS PUBLICATION REPORTS ON A RETROSPECTIVE COHORT STUDY THAT WAS CONDUCTED ON CONSECUTIVE PATIENTS WHO UNDERWENT MIMVR USING ARTIFICIAL CHORDAE AT A SINGLE CENTER IN VIETNAM BETWEEN 2016 AND 2022. 90 PATIENTS WERE TREATED, 41 SIMPLE AND 49 INTERMEDIATE TO COMPLEX REPAIRS. NO CASES REQUIRED MR REPLACEMENT AND ALL OUTCOMES WERE FAVORABLE. THE MEDIAN FOLLOW UP TIME WAS 30.3 MONTHS AND THERE WERE TWO CARDIAC DEATHS, ONE IN EACH GROUP. THE KAPLAN¿MEIER SURVIVAL ESTIMATES FOR GROUPS 1 AND 2 AT 12 AND 24 MONTHS WERE 97% VS. 100% AND 97% VS. 96%, RESPECTIVELY, AND ESTIMATES FOR FREEDOM FROM RECURRENT MR WERE 97% VS. 92% AND 97% VS. 88%, RESPECTIVELY. THERE WAS A NEGATIVE ASSOCIATION BETWEEN THE VOLUME OF OPERATIONS AND THE DURATION OF OPERATION AND AORTIC CROSS CLAMP TIME, LEADING TO SHORTER DURATIONS. CONCLUSION. BASED ON OUR SINGLE-CENTER EXPERIENCE, MIMVR USING ARTIFICIAL CHORDAE VIA RIGHT MINI-THORACOTOMY CA. BE SAFELY AND ELECTIVELY PERFORMED IN RESOURCE-LIMITED COUNTRIES FOR PATIENTS WITH MR. THIS APPROACH HAS BEEN SHOWN TO BE APPLICABLE FOR A RANGE OF MR COMPLEXITIES, FROM SIMPLE TO INTERMEDIATE-TO-COMPLEX MV REPAIRS, AND HAS DEMONSTRATED PROMISING RESULTS IN TERMS OF MIDTERM FREEDOM FROM MR RECURRENCE. THIS INVESTIGATION IS RELEGATED TO CHORD-X UNKNOWN CONFIGURATION FOR DEATH AND MITRAL VALVE REGURGITATION.. NO PRODUCT IDENTIFYING INFORMATION PROVIDED. | ACCORDING TO PUBLICATION "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE. INSIGHTS FROM A 6 YEAR SINGLE CENTER STUDY¿, THIS PUBLICATION REPORTS ON A RETROSPECTIVE COHORT STUDY THAT WAS CONDUCTED ON CONSECUTIVE PATIENTS WHO UNDERWENT MIMVR USING ARTIFICIAL CHORDAE AT A SINGLE CENTER IN VIETNAM BETWEEN 2016 AND 2022. 90 PATIENTS WERE TREATED, 41 SIMPLE AND 49 INTERMEDIATE TO COMPLEX REPAIRS. NO CASES REQUIRED MR REPLACEMENT AND ALL OUTCOMES WERE FAVORABLE. THE MEDIAN FOLLOW UP TIME WAS 30.3 MONTHS AND THERE WERE TWO CARDIAC DEATHS, ONE IN EACH GROUP. THE KAPLAN¿MEIER SURVIVAL ESTIMATES FOR GROUPS 1 AND 2 AT 12 AND 24 MONTHS WERE 97% VS. 100% AND 97% VS. 96%, RESPECTIVELY, AND ESTIMATES FOR FREEDOM FROM RECURRENT MR WERE 97% VS. 92% AND 97% VS. 88%, RESPECTIVELY. THERE WAS A NEGATIVE ASSOCIATION BETWEEN THE VOLUME OF OPERATIONS AND THE DURATION OF OPERATION AND AORTIC CROSS CLAMP TIME, LEADING TO SHORTER DURATIONS. CONCLUSION. BASED ON OUR SINGLE-CENTER EXPERIENCE, MIMVR USING ARTI5CIAL CHORDAE VIA RIGHT MINI-THORACOTOMY CAN BE SAFELY AND EFFECTIVELY PERFORMED IN RESOURCE-LIMITED COUNTRIES FOR PATIENTS WITH MR. THIS APPROACH HAS BEEN SHOWN TO BE APPLICABLE FOR A RANGE OF MR COMPLEXITIES, FROM SIMPLE TO INTERMEDIATE-TO-COMPLEX MV REPAIRS AND HAS DEMONSTRATED PROMISING RESULTS IN TERMS OF MIDTERM FREEDOM FROM MR RECURRENCE. THIS INVESTIGATION IS RELEGATED TO CHORD-X UNKNOWN CONFIGURATION FOR THE FOLLOWING: DEATH (2). MITRAL VALVE REGURGITATION (33). MULTIPLE ATTEMPTS FOR ADDITIONAL INFORMATION HAVE GONE UNMET. IF INFORMATION IS PROVIDED IN THE FUTURE, A SUPPLEMENTAL REPORT WILL BE ISSUED. A REVIEW OF MANUFACTURING RECORDS COULD NOT BE PERFORMED AS A DEFINITIVE LOT/SERIAL NUMBER AND DATE OF SURGERY WERE NOT PROVIDED BY THE COMPLAINANT. A REVIEW OF THE AVAILABLE INFORMATION WAS PERFORMED. THIS INVESTIGATION REVIEWS CLINICAL OUTCOMES REPORTED IN THE 2023 PUBLICATION TITLED "MINIMALLY INVASIVE MITRAL VALVE REPAIR WITH ARTIFICIAL CHORDAE: INSIGHTS FROM A 6-YEAR SINGLE-CENTER STUDY" BY BUI ET AL. THE STUDY REPORTS A RETROSPECTIVE COHORT ANALYSIS OF 90 PATIENTS UNDERGOING MINIMALLY INVASIVE MITRAL VALVE REPAIR (MIMVR) AT A SINGLE CENTER IN VIETNAM BETWEEN APRIL 2016 AND APRIL 2022. ALL PROCEDURES UTILIZED ARTIFICIAL CHORDAE (CHORD-X LOOPS OR GORE CV-5 SUTURES) THROUGH A RIGHT MINITHORACOTOMY APPROACH. PATIENTS WERE DIVIDED INTO TWO GROUPS BASED ON THE ANYANWU COMPLEXITY SCORE: GROUP 1 INCLUDED SIMPLE REPAIRS (N=41), AND GROUP 2 INCLUDED INTERMEDIATE-TO-COMPLEX REPAIRS (N=49). THE MEAN AGE OF THE COHORT WAS 50.5 ± 12.9 YEARS, AND 77.8% WERE MALE. ALL CASES INVOLVED CARPENTIER TYPE II MITRAL REGURGITATION, PRIMARILY DUE TO FIBROELASTIC DEFICIENCY. BOTH CHORD-X AND GORE CV-5 SUTURES WERE USED FOR NEOCHORDAE IMPLANTATION. CHORD-X WAS USED IN 56% OF CASES, SINGLE GORE CV-5 IN 33%, AND BOTH TECHNIQUES COMBINED IN 11%. ALL PATIENTS ALSO UNDERWENT RING ANNULOPLASTY USING THE CARPENTIER PHYSIO II RING. THE STUDY REPORTS NO CONVERSIONS TO MITRAL VALVE REPLACEMENT AND NO REOPERATIONS RELATED TO THE MITRAL VALVE WERE OBSERVED DURING THE FOLLOW UP PERIOD. POSTOPERATIVE OUTCOMES INCLUDED TWO CARDIAC DEATHS (2.2%) DURING THE MEDIAN 30.3-MONTH FOLLOW-UP PERIOD, WITH ONE PATIENT READMITTED AFTER FOUR MONTHS DUE TO ACUTE HEART FAILURE AND THE OTHER PRESENTING WITH INFECTIVE ENDOCARDITIS TWO MONTHS POST-IMPLANT. ONE PATIENT (1.1%) EXPERIENCED A STROKE WITH SEQUELAE. REOPERATION FOR BLEEDING WAS REQUIRED IN THREE CASES (3.3%). MODERATE OR SEVERE RECURRENT MITRAL REGURGITATION OCCURRED IN 11.1% OF PATIENTS OVERALL, WITH A HIGHER INCIDENCE IN GROUP 2 (18.4%) COMPARED TO GROUP 1 (2.4%). SURVIVAL ESTIMATES AT 48 MONTHS WERE 97% FOR GROUP 1 AND 96% FOR GROUP 2 (P=0.850). KAPLAN-MEIER ESTIMATES FOR FREEDOM FROM RECURRENT MR AT 48 MONTHS WERE 97% FOR GROUP 1 AND 88% FOR GROUP 2 (P=0.260). THE AUTHORS CONCLUDED THAT MIMVR USING ARTIFICIAL CHORDAE (INCLUDING BOTH GORE CV-5 SUTURES AND CHORD-X LOOPS) IS SAFE AND EFFECTIVE ACROSS A RANGE OF MITRAL REGURGITATION COMPLEXITIES, EVEN IN RESOURCE-LIMITED SETTINGS. THIS CLINICAL OUTCOME ANALYSIS IS BASED ON A SINGLE-SURGEON EXPERIENCE USING ARTIFICIAL CHORDAE (GORE CV-5 AND CHORD-X) FOR MINIMALLY INVASIVE MITRAL VALVE REPAIR. THE STUDY DEMONSTRATED FAVORABLE MIDTERM SURVIVAL, LOW RATES OF REOPERATION, AND ACCEPTABLE RECURRENCE RATES OF MITRAL REGURGITATION. WHILE BOTH GORE CV-5 SUTURES AND CHORD-X LOOPS WERE EMPLOYED, THE PUBLICATION DOES NOT DISAGGREGATE OUTCOMES BY DEVICE TYPE. THEREFORE, IT IS NOT POSSIBLE TO ASSOCIATE SPECIFIC COMPLICATIONS OR EVENTS WITH EITHER PRODUCT INDIVIDUALLY. ALL REPORTED COMPLICATIONS, INCLUDING RECURRENT MR, PERIOPERATIVE STROKE, BLEEDING, AND MORTALITY, ARE CONSISTENT WITH THE KNOWN RISKS FOR MITRAL VALVE REPAIR PROCEDURES USING ARTIFICIAL CHORDAE AND ARE IDENTIFIED IN THE IFU. A COMPLAINT AND RECALL QUERY COULD NOT BE PERFORMED AS PRODUCT IDENTIFYING INFORMATION IS UNKNOWN. BASED ON THE AVAILABLE INFORMATION, A DEFINITIVE ROOT CAUSE FOR THIS EVENT CANNOT BE DETERMINED. THIS EVENT DOES NOT IDENTIFY ADDITIONAL HAZARDS OR MODIFY THE PROBABILITY AND SEVERITY OF EXISTING HAZARDS. THERE IS NO INDICATION THAT AN ERROR OR DEFICIENCY OCCURRED AT ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC AND THE IFU ADEQUATELY COMMUNICATES RISK. THIS COMPLAINT WAS REVIEWED FOR A CAPA EVALUATION AND A CAPA IS NOT WARRANTED AT THIS TIME. ARTIVION WILL CONTINUE TO MONITOR SIMILAR COMPLAINTS TO DETERMINE IF ADDITIONAL ACTIONS ARE WARRANTED; HOWEVER, AT THIS TIME NO FURTHER ACTIONS ARE NECESSARY. THIS REPORT IS BEING SUBMITTED AS REQUIRED BY FEDERAL REGULATIONS AND DOES NOT CONSTITUTE AN ADMISSION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. FURTHERMORE, THIS REPORT REFLECTS THE EVENT AS ALLEGED BY THE COMPLAINANT AND DOES NOT IMPLY THAT THE INFORMATION REPORTED TO ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC IS ACCURATE OR HAS BEEN CONFIRMED BY ARTIVION ¿ FORMERLY CRYOLIFE/JOTEC.
EPIQ CVXI | K132304 | IT WAS REPORTED THAT THE EPIQ CVXI ULTRASOUND SYSTEM BECAME UNRESPONSIVE AND LOCKED UP DURING AN EXAMINATION WHERE A PATIENT WAS CODING. THE CLINICIAN HAD TO BRING IN ANOTHER SYSTEM IN AN ATTEMPT TO COMPLETE THE EXAMINATION. UNFORTUNATELY, THE PATIENT DID NOT RECOVER FROM THE CODE AND ULTIMATELY EXPIRED; HOWEVER, THE CUSTOMER HAS CLEARLY STATED THE ULTRASOUND DID NOT CAUSE THE PATIENT INCIDENT. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW UP REPORT WILL BE SUBMITTED.
EPIQ CVXI | K132304 | IT WAS REPORTED THAT THE EPIQ CVXI ULTRASOUND SYSTEM BECAME UNRESPONSIVE AND LOCKED UP DURING AN EXAMINATION WHERE A PATIENT WAS CODING. THE CLINICIAN HAD TO BRING IN ANOTHER SYSTEM IN AN ATTEMPT TO COMPLETE THE EXAMINATION. UNFORTUNATELY, THE PATIENT DID NOT RECOVER FROM THE CODE AND ULTIMATELY EXPIRED; HOWEVER, THE CUSTOMER HAS CLEARLY STATED THE ULTRASOUND DID NOT CAUSE THE PATIENT INCIDENT. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW UP REPORT WILL BE SUBMITTED.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEVICE WAS COLLECTED BECAUSE THE PATIENT PASSED AWAY. THE DEVICE EXTERIOR HAD TURNED BLACK AND TOBACCO SMELLS WERE OBSERVED. A REQUEST HAS BEEN MADE FOR REPAIR OF THE DEVICE. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. THE REQUEST IS FOR THE REPAIR OF THE DEVICE NOT RELATED TO THE PATIENT DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING A DEVICE WAS COLLECTED BECAUSE THE PATIENT PASSED AWAY. THE DEVICE EXTERIOR HAD TURNED BLACK AND TOBACCO SMELLS WERE OBSERVED. A REQUEST HAS BEEN MADE FOR REPAIR OF THE DEVICE. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED. THE REQUEST IS FOR THE REPAIR OF THE DEVICE NOT RELATED TO THE PATIENT DEATH. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND A LIFE-RELATED SMELL (TOBACCO SMELL) WAS CONFIRMED TO BE FOUND ON THE MUFFLER ASSEMBLY. THE MUFFLER ASSEMBLY WAS REPLACED AND THE DEVICE WAS CLEANED. DEVICE PASSED TESTING AND CONFIRMED TO OPERATE PROPERLY.
EXTRACTOR PRO RX | K102082 | IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT AN EXTRACTOR PRO RX RETRIEVAL BALLOON WAS USED IN THE BILE DUCT DURING AN ENDOSCOPIC RETROGRADE CHOLANGIOPANCREATOGRAPHY (ERCP) PROCEDURE PERFORMED IN (B)(6) 2025. THE EXACT PROCEDURE DATE WAS NOT REPORTED. DURING THE PROCEDURE, A PERFORATION OCCURRED NEAR THE DUODENUM. IT WAS REPORTED THAT THE PHYSICIAN SUSPECTS THAT THE EXTRACTOR PRO RX RETRIEVAL BALLOON MAY BE A CONTRIBUTING FACTOR IN CAUSING THE PERFORATION. THE PROCEDURE WAS COMPLETED USING THE SAME ORIGINAL DEVICE. THE PATIENT PASSED AWAY IN (B)(6) 2025 DUE TO PERITONITIS. THE EXACT PATIENT DEATH DATE WAS NOT REPORTED. | BLOCK B2: APPROXIMATED BASED ON THE REPORTED PATIENT DEATH DATE OF (B)(6) 2025. BLOCK B3: APPROXIMATED BASED ON THE REPORTED EVENT DATE OF JANUARY 2025. BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. BLOCK H6: IMDRF PATIENT CODE E2114 IS BEING USED TO CAPTURE THE REPORTABLE EVENT OF PERFORATION. IMDRF PATIENT CODE E1024 IS BEING USED TO CAPTURE THE REPORTABLE EVENT OF PERITONITIS. IMDRF IMPACT CODE F02 IS BEING USED TO CAPTURE THE REPORTABLE EVENT OF PATIENT DEATH.
PHENOM CATHETER | K180959 | ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT A BALLOON (NON-MEDTRONIC) USED WITH THE PIPELINE WAS DAMAGED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
BIOFLO VORTEX | K190559 | ON (B)(6) 2021, THE DECEDENT, (B)(6), UNDERWENT PLACEMENT OF THE DEFENDANTS' SMARTPORT POWER INJECTABLE PORT PRODUCT, I.E., SMARTPORT PLASTIC PORT: MODEL/CAT NUMBER: H787CT80LPPDVI0; LOT #: 5708017, FOR HIS CHEMOTHERAPY TREATMENT. UPON INFORMATION AND BELIEF, THE AFOREMENTIONED DEVICE'S DESCRIPTION IS AS FOLLOWS: DETACHED 8F POLY CATH, VI INTRODUCER, & FILLED SH OR SMART PORT PLASTIC, 8F SINGLE PLASTIC PORT FILLED NON-VALVED PG, OR 8FRENCH ANGIODYNAMICS SMARTPORT POWER INJECTABLE PORT. UPON INFORMATION AND BELIEF, ON (B)(6) 2021, THE DEVICE AT ISSUE WAS IMPLANTED BY DR. (B)(6), MD, AT (B)(6) HOSPITAL IN (B)(6), WISCONSIN. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, THE DECEDENT PRESENTED TO THE EMERGENCY DEPARTMENT AT (B)(6) HOSPITAL IN (B)(6), WISCONSIN, IN AN AMBULANCE, WITH COMPLAINTS OF WEAKNESS, AND WAS THEN ADMITTED WITH PROGRESSIVE WEAKNESS AND OTHER ACCOMPANYING SYMPTOMS. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, THE CULTURE OF THE DECEDENT'S PORT-A-CATH WAS ORDERED. UPON INFORMATION AND BELIEF, THE DECEDENT'S BLOOD CULTURE AND THE BLOOD CULTURE AEROBIC PANEL CAME BACK AS "ABNORMAL". UPON INFORMATION AND BELIEF, THE DECEDENT'S AFOREMENTIONED "ABNORMAL" BLOOD CULTURES' RESULTS SHOWED THE PRESENCE OF STAPHYLOCOCCUS AND STAPHYLOCOCCUS AUREUS. UPON INFORMATION AND BELIEF, ON (B)(6) 2022, THE CT CHEST WAS ORDERED BY THE INFECTIOUS DISEASE DEPARTMENT. UPON INFORMATION AND BELIEF, ON (B)(6) 2022, THE AFOREMENTIONED CT CHEST PERFORMED ON THE DECEDENT SHOWED THE FOLLOWING RELEVANT FINDINGS: "INFECTED PORT/ENDOCARDITIS"; "MSSA BACTEREMIA PRELIMINARY". "VEGETATION ON PORT", "PERICARDIAL EFFUSION [OR ITS POTENTIAL COMPLICATION PERICARDIAL] TAMPONADE", ETC. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, DR. (B)(6), MD, AN INFECTIOUS DISEASES PHYSICIAN, NOTED THAT THE DECEDENT'S BLOOD CULTURE CAME BACK AS STAPHYLOCOCCUS AUREUS AND THAT THE DECEDENT HAD VEGETATION ON [HIS] PORT. DR. (B)(6) ALSO NOTED THAT THE DECEDENT'S BLOOD CULTURE CAME BACK AS MSSA [MSSA IS A COMMON MEDICAL ABBREVIATION FOR METHICILLIN-SUSCEPTIBLE STAPHYLOCOCCUS AUREUS]. DR. (B)(6) DID ALSO UNDERLINE THAT THERE WERE CLEARLY SEPTIC EMBOLI [FROM STAPHYLOCOCCUS]. UPON INFORMATION AND BELIEF, ON (B)(6) 2022, THE DECEDENT'S EC ECHO COMPLETE W/DOPPLER (I.E., DOPPLER ECHOCARDIOGRAPHY), ORDERED BY (B)(6), PA-C, FOUND LIKELY VEGETATION ON THE DECEDENT'S RIGHT ATRIAL CATHETER. ON (B)(6) 2022, DR. (B)(6), MD, WHO, UPON INFORMATION AND BELIEF, REVIEWED & ANALYZED THE RESULTS OF THE AFOREMENTIONED DOPPLER ECHOCARDIOGRAPHY COULD NOT "RULE OUT SMALL VEGETATION ON THE AORTIC VALVE" OF THE DECEDENT. UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2022, DR. (B)(6), M.D., FCCM UNDERLINED THAT THE DECEDENT'S ILLNESS AT ISSUE WAS DUE TO "MSSA ENDOCARDITIS ASSOCIATED WITH THE PORT INFECTION, PULMONARY SEPTIC EMBOLI, POSSIBLY PLEURAL AND PERICARDIAL INFECTION." UPON INFORMATION AND BELIEF, ON OR ABOUT (B)(6) 2025, THE DECEDENT'S PHYSICIAN, DR. (B)(6), M.D., FCCM, WHILE DOCUMENTING THE DECEDENT'S PORT-RELATED DIAGNOSES, NOTED THAT THE DECEDENT WOULD REQUIRE CARDIAC SURGERY AND PROLONGED ANTIBIOTIC SUPPORT, HOWEVER DR. (B)(6) ALSO OPINED THAT THE DECEDENT WOULD NOT HAVE SURVIVED THE REQUIRED CARDIAC SURGERY. ON (B)(6) 2022, (B)(6) DIED AT (B)(6) HOSPITAL, IN (B)(6), WISCONSIN. | THE DEVICE IS NOT AVAILABLE TO BE RETURNED TO THE MANUFACTURER FOR EVALUATION. AN INVESTIGATION INTO THE ROOT CAUSE FOR THE EVENT IS CURRENTLY IN PROGRESS. THE RESULTS OF THE INVESTIGATION WILL BE SENT VIA A FOLLOW UP MEDWATCH. REFERENCE (B)(4).
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION ST30 DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. AT THIS TIME, THERE IS NO PATIENT INFORMATION AVAILABLE; THEREFORE, NO GOOD FAITH EFFORT ATTEMPTS CAN BE MADE, AND PHILIPS IS NOT ABLE TO FULLY INVESTIGATE THIS COMPLAINT. IF ANY ADDITIONAL INFORMATION IS RECEIVED OR THE DEVICE IS RETURNED, A SUPPLEMENTAL REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THREE ATTEMPTS TO HAVE THE DEVICE RETURNED FOR EVALUATION AND INVESTIGATION WERE UNSUCCESSFUL. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
BIPAP AUTOSV ADVANCED | K090539 | THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A BIPAP AUTOSV ADV DEVICE HAD PASSED AWAY SEVERAL YEARS AGO. THE PATIENT¿S WIFE REQUESTED TO STOP RECEIVING CORRESPONDENCE AND INDICATED THAT A REPLACEMENT DEVICE WAS NOT NEEDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT¿S DEATH, AND NO FURTHER DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A DREAMSTATION BIPAP PRO DEVICE HAS PASSED AWAY. THE PATIENT¿S HUSBAND CALLED IN REQUESTING A RETURN LABEL TO RETURN THE RECALLED DEVICE. NO FURTHER DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RECEIVED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A DREAMSTATION BIPAP PRO DEVICE HAD PASSED AWAY. THE PATIENT¿S HUSBAND CALLED IN REQUESTING A RETURN LABEL TO RETURN THE RECALLED DEVICE. NO FURTHER DETAILS WERE PROVIDED. THE DEVICE HAS BEEN RETURNED FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND NO ERRORS OR DEVICE PROBLEMS. THERE WAS NO EVIDENCE OF FOAM PARTICLES OR DEGRADATION. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW-UP REPORT WILL BE FILED. THE MANUFACTURER HAS UPDATED BOX H IN THIS REPORT.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION THAT PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | IT WAS REPORTED THAT THE PATIENT OF THE REMSTAR AUTO A-FLEX DEVICE PASSED AWAY. ACCORDING TO THE PATIENT'S NIECE, THE PATIENT PASSED AWAY IN (B)(6) 2024 AND SHE WOULD LIKE TO STOP GETTING LETTERS. THERE IS NO ALLEGATION OF MALFUNCTION, OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
EDISON SYSTEM | K233466 | A 81-YEAR-OLD MALE PATIENT WITH ADVANCED METASTATIC CHOLANGIOCARCINOMA RECEIVED REPEAT HISTOTRIPSY TREATMENT (INITIAL TREATMENT (B)(6) 2025, TOTAL PTV 49.0 CC) ON (B)(6) 2025, TO A SINGLE LESION IN SEGMENT V, WITH TWO PARTIALLY OVERLAPPING PLANNED TREATMENT VOLUMES (PTVS) REPORTED AS PTV 1: TOTAL PTV = 31 CC, PTV 2: TOTAL PTV = 29.5CC, FOR A TOTAL PTV OF 60.5 CC. THE PATIENT WAS DISCHARGED POST-PROCEDURE WITHOUT INCIDENT AND ULTIMATELY RETURNED HOME. ON POST-OPERATIVE DAY 3 ((B)(6) 2025), HE DEVELOPED ANURIA, AND HIS CREATININE ROSE TO 7. THE TREATING PHYSICIAN REPORTED THAT THE PATIENT UNDERWENT HEMODIALYSIS AND HAD SOME URINE OUTPUT, WITH A DIAGNOSIS OF ACUTE TUBULAR NECROSIS (ATN) ATTRIBUTED TO A COMBINATION OF ANESTHESIA, HISTOTRIPSY, AND NSAID USE, WITH POSSIBLE CONTRIBUTION FROM ACUTE INTERSTITIAL NEPHRITIS (AIN). THE PATIENT REMAINED DIALYSIS-DEPENDENT, BUT ON-GOING HYPOTENSION COMPLICATED MANAGEMENT. THE FAMILY ULTIMATELY ELECTED TO WITHDRAW CARE AND THE PATIENT DIED ON (B)(6) 2025. NO DEVICE MALFUNCTIONS OR OTHER NOTEWORTHY EVENTS OCCURRED DURING THE CASE. | NO DEVICE MALFUNCTIONS OR OTHER NOTEWORTHY EVENTS OCCURRED DURING THE CASE.
DREAMSTATION AUTO CPAP | K131982 | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP. THE SPOUSE OF THE PATIENT REPORTED THE PATIENT HAS SADLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
MICRA | K132030 | IT WAS REPORTED THAT THE PATIENT DIED. IT WAS NOTED THAT DURING THE IMPLANTATION OF THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) THE PATIENT EXPERIENCED FEMORAL COMPLICATIONS RESULTING IN DISSECTION AND BLOOD LOSS. IT WAS FURTHER REPORTED THAT THE MICRA INTRODUCER APPEARED TO "JUMP FORWARD". ONCE THE INTRODUCER WAS REMOVED ARTERIAL BLOOD WAS PUMPING OUT OF THE GROIN. MANUAL PRESSURE WAS ATTEMPTED, A FEMORAL ARTERY COMPRESSION DEVICE WAS APPLIED, BUT HEMOSTASIS WAS NOT MAINTAINED. THE VASCULAR SURGEON WAS CALLED IN AND THE PATIENT WAS TRANSPORTED TO THE SURGICAL SUITE. HOWEVER, THE PATIENT PASSED AWAY FROM FEMORAL COMPLICATIONS. | CORRECTED: H6. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION ST30 | K102465 | THIS NOTIFICATION IS BEING REVIEWED DUE TO A USER OF A DREAMSTATION AUTO CPAP. IT HAS BEEN REPORTED THE PATIENT HAS SADLY PASSED AWAY. THE CAUSE OF DEATH WAS NOT REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
HLS SET | K112360 | COMPLAINT ID # (B)(4). | THE EVENT OCCURRED IN THE US. IT WAS REPORTED THAT DURING PATIENT TREATMENT ISSUES WERE PRESENT WITH ERRONEOUS PRESSURE READINGS (ARTERIAL, INTERNAL AND DELTA) FROM AN HLS-SET. NEW INFORMATION HAS BEEN RECEIVED ON 2025-05-14 BY THE SSU/CUSTOMER, THAT THE INVOLVED PATIENT PASSED AWAY. THE PATIENT¿S FAMILY DECIDED TO DISCONTINUE THE TREATMENT, RESULTING IN THE PATIENT PASSING AWAY. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. THE AFFECTED PRODUCT WAS NOT AVAILABLE FOR TECHNICAL INVESTIGATION AS IT WAS DISCARDED BY THE CUSTOMER. THUS, THE EXACT ROOT CAUSE REMAINS UNKNOWN. HOWEVER, A MEDICAL REVIEW WAS PERFORMED BY GETINGE MEDICAL AFFAIRS ON 2025-06-02 WITH FOLLOWING CONCLUSION: "THE REPORTED INCIDENT INVOLVES INACCURATE PRESSURE READINGS FROM A PRE-PRIMED HLS SET USED DURING ECMO THERAPY. THE CHIEF OF PERFUSION AT THE TREATING INSTITUTION, SUGGESTED DURING A CONFERENCE THAT FLUID MAY HAVE ENTERED THE WHITE CONNECTOR OF THE HLS SET PRIOR TO THERAPY, POSSIBLY CONTRIBUTING TO THE MALFUNCTION. THE TEAM REPORTEDLY ATTEMPTED TO CLEAN THE AREA BEFORE USE. ECMO THERAPY BEGAN ON (B)(6) 2025. NO ISSUES WERE NOTED DURING INITIAL PRIMING WITH PLASMALYTE, AND NO DEPOSITS OR CRYSTAL FORMATIONS WERE SEEN AT THE SENSOR PORT WHICH MAY BE IN CONFLICT WITH THE PREVIOUS SUGGESTION OF FLUID INTRUSION IN THE PORT THE ABSENCE OF VISIBLE DEPOSITS OR RESIDUE AT THE TIME OF SETUP MAY SUGGEST THAT INTRUSION MAY HAVE OCCURRED AFTER ECMO INITIATION AS INTRUSION WAS NOT IDENTIFIED DURING INITIAL DEVICE INSPECTION. WHILE THE CIRCUIT APPEARED CLEAN AT INITIATION, THE DISCOVERY OF BLOOD UPON DISASSEMBLY RAISES QUESTIONS ABOUT WHETHER CONTAMINATION MAY HAVE DEVELOPED SOMETIME DURING SUPPORT. THE REPORTED PRESSURE ISSUE WAS FIRST OBSERVED ON (B)(6), RESULTING IN THE REPLACEMENT OF THE SENSOR PROBE (CABLE) ON (B)(6). THE REPORTED BEHAVIOR SUGGESTS THE ROOT CAUSE MAY LIE IN EITHER THE CABLE ITSELF OR THE INTERNAL SENSOR PORT OF THE HLS SET. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER FOUND NEAR THE POST-OXYGENATOR MEMBRANE. LAB VALUES ON FEBRUARY 1 INDICATED STABLE PARAMETERS. THE HLS SET WAS NOT REPLACED, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES REPORTEDLY CORRELATED WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING BLOOD SAMPLING. NO REDUCTION IN BLOOD FLOW WAS OBSERVED. CLOTS WERE NOT INITIALLY OBSERVED BUT WERE LATER IDENTIFIED NEAR THE POST-OXYGENATOR MEMBRANE AT THE END OF THERAPY. LABORATORY VALUES OBTAINED ON FEBRUARY 1 SHOWED A HEMOGLOBIN LEVEL OF 8.6 G/DL, HEMATOCRIT OF 28%, D-DIMER CONCENTRATION OF 12.21 MICROGRAM/ML, PLATELET COUNT OF 126 × 1 BILLION/L, AND A PTT OF 26.7 SECONDS. WHILE THESE VALUES APPEARED RELATIVELY STABLE FOR AN ECMO PATIENT, THE MARKEDLY ELEVATED D-DIMER LEVEL MAY SUGGEST EARLY THROMBOTIC ACTIVITY WITHIN THE CIRCUIT, DESPITE THE ABSENCE OF VISIBLE CLOTS AT THAT TIME (1, 2). ACCORDING TO RAJSIC ET AL. (2022) (3), THE RECOMMENDED THERAPEUTIC RANGE FOR THE PREVENTION OF VENOUS THROMBOEMBOLISM IS 1.5 TO 2.5 TIMES THE PATIENT¿S BASELINE APTT (40¿80 SECONDS). A PTT OF 26.7 SECONDS MAY HAVE CONTRIBUTED TO THROMBUS FORMATION WITHIN THE ECMO CIRCUIT. THE HLS SET ITSELF WAS NOT REPLACED DURING THE ECMO RUN, BUT THE SENSOR CABLE WAS EXCHANGED TWICE. ELEVATED PINT PRESSURES WERE OBSERVED IN CORRELATION WITH PHYSICAL DISTURBANCES OF THE CABLE, SUCH AS DURING PRE- AND POST-OXYGENATOR BLOOD SAMPLING. DESPITE THESE ANOMALIES, NO REDUCTION IN BLOOD FLOW WAS REPORTED, AND THE CIRCUIT CONTINUED TO PROVIDE EFFECTIVE SUPPORT. IN CONCLUSION, A PRODUCT MALFUNCTION INVOLVING THE INTERNAL PRESSURE SENSOR OF THE HLS SET WAS LIKELY, AS EVIDENCED BY PERSISTENT INACCURATE PRESSURE READINGS DESPITE CABLE REPLACEMENT. HOWEVER, THIS MALFUNCTION DID NOT RESULT IN A LOSS OF ECMO SUPPORT OR REDUCTION IN BLOOD FLOW. THAT SAID, REDUNDANT PRESSURE MONITORING DURING WAS, ASSUMEDLY, PRESENT VIA THE PART PRESSURE FROM THE HLS DISPOSABLE. NOTABLY, THE FIRST PRESSURE ANOMALY WAS REPORTED ON FEBRUARY 1¿NINE DAYS AFTER ECMO INITIATION ON (B)(6)¿INDICATING THAT THE HLS SET INITIALLY PERFORMED AS EXPECTED. THE SUBSEQUENT DISCOVERY OF BLOOD RESIDUES ON THE DEVICE UPON DISASSEMBLY SUGGESTS THAT CONTAMINATION OF THE INTERNAL PRESSURE SENSOR PORT MAY HAVE OCCURRED WHILE THERAPY, POTENTIALLY CONTRIBUTING TO THE OBSERVED BEHAVIOR. THE PATIENT¿S CLINICAL COURSE ULTIMATELY ENDED WITH WITHDRAWAL OF CARE AT THE FAMILY¿S REQUEST. WHILE THROMBOTIC EVENTS MAY HAVE CONTRIBUTED TO OVERALL MORBIDITY, THE DECISION TO WITHDRAW CARE APPEARS TO BE INFLUENCED BY BROADER CLINICAL OR ETHICAL CONSIDERATIONS. BASED ON AVAILABLE INFORMATION, THE PATIENT¿S EXPIRATION WAS NOT INFLUENCED BY THE DEVICE MALFUNCTION. THE FOLLOWING IS A QUOTE FROM THE INITIATING COMPLAINT INDICATING APPARENT SYNCHRONY WITH THE ABOVE OPINION ON DEVICE INFLUENCE: "ECLS COORDINATOR AT (B)(6) HOSPITAL, MENTIONED THEY ARE HAVING ISSUES WITH ERRONEOUS PRESSURE READINGS FROM AN HLS ON A PATIENT. IT IS NOT NEGATIVELY IMPACTING THE PATIENT OR LEVEL OF CARE, NO PATIENT HARM." ACCORDING TO THE INSTRUCTION FOR USE (HLS SET, CHAPTER "PREPARATION AND INSTALLATION") THE PRESSURE SENSORS HAVE TO BE CHECKED BEFORE PRIMING. IN GENERAL, IT SHOULD BE NOTED THAT PRE-PRIMING CAN HAVE A NEGATIVE IMPACT ON THE SENSOR FUNCTION AND THE ASSOCIATED PRESSURE VALUES AND CAN LEAD TO THE OUTPUT OF UNEXPECTED / IMPLAUSIBLE PRESSURE VALUES SHORTLY BEFORE USE, AS INDICATED IN CHAPTER "PRIMING THE SYSTEM", WHERE IT SAYS "ONLY FILL THE SYSTEM SHORTLY BEFORE USE AND IN ACCORDANCE WITH THE PRIMING INSTRUCTIONS. CALIBRATION VALUES ARE LOST WHEN THE DRIVE UNIT IS SWITCHED OFF". FURTHER THE CARDIOHELP HAS A FLOW/BUBBLE SENSOR AND A VENOUS PROBE TO MEASURE AND CONTROL THE BLOOD FLOW AND PARAMETERS. IF THE MEASURED VALUES ARE ABOVE HIGH LIMIT OR BELOW LOW LIMIT OF THE SET LIMITS THE SYSTEM GENERATES A VISUAL AND ACOUSTICAL ALARM. IN ADDITION, WITHIN THE CHAPTER ¿PRIMING THE SYSTEM¿ IT IS WRITTEN NOT TO PRE-PRIME THE CIRCUIT TOO EARLY. BASED ON THE RESULTS THE REPORTED FAILURE "PRESSURE ISSUES" COULD NOT BE CONFIRMED. THE PRODUCTION RECORDS OF THE AFFECTED PRODUCT WERE REVIEWED ON 2025-06-04. ACCORDING TO THE FINAL TEST RESULTS, THE PRODUCT PASSED THE TESTS AS PER SPECIFICATIONS. PRODUCTION RELATED INFLUENCES ARE UNLIKELY. THE REVIEW OF SCRAP, REWORK, ENHANCEMENTS AND DESIGN CHANGES WERE REVIEWED AN NO ABNORMALITIES IN REGARD TO THE REPORTED FAILURE WERE FOUND. IN ADDITION, A REVIEW FOR POTENTIAL FIELD ACTIONS AND CAPAS RELATED TO THE FAILURE AND PRODUCT IN THIS COMPLAINT WAS PERFORMED. THIS COMPLAINT IS NOT IN SCOPE OF ANY ONGOING FIELD ACTIONS AND/OR CAPAS. THE REVIEW OF THE NON-CONFORMITIES HAS BEEN PERFORMED FOR THE REVIEWED TIME PERIOD. IT DOES NOT SHOW ANY NON-CONFORMITY IN REGARD TO THE REPORTED FAILURE. THE CUSTOMER WILL BE INFORMED ABOUT THE RESULTS BY THE GETINGE SALES AND SERVICE UNIT. THE OCCURRENCE RATE WAS CALCULATED FOR THE REPORTED ISSUE AND IT WAS DETERMINED THAT THIS IS NOT A SYSTEMIC ISSUE. THEREFORE, NO REMEDIAL ACTION IS REQUIRED. THE OCCURRENCE RATE RELATED TO THE REPORTED ISSUE IS CURRENTLY BEING MONITORED AS PART OF MAQUET CARDIOPULMONARY¿ S TRENDING PROGRAM AND ADDITIONAL INVESTIGATIONS OR CORRECTIONS WILL BE IMPLEMENTED IN CASE OF ADVERSE TRENDING. NOTE: NO UDI NUMBER HAS BEEN INCLUDED IN THE SECTION D4 UNIQUE DEVICE IDENTIFIER (UDI) SINCE THE LOT NUMBER OF THE AFFECTED SET WAS NOT PROVIDED. THEREFORE, IT IS NOT POSSIBLE TO IDENTIFY THE SET LOT NUMBER OF THE DEVICE IN QUESTION AND THEREFORE ITS UDI NUMBER.
REMSTAR PRO C-FLEX+ SYSTEM ONE 60 SERIES | K113068 | THE REPORTER CALLED IN STATING THE DEVICE IS NO LONGER AVAILABLE FOR RETURN. THE REPORTER STATED THE PATIENT EXPIRED 3 YEARS AGO (EXACT DATE OF DEATH NOT REPORTED). THE CAUSE OF DEATH WAS NOT REPORTED. FOLLOW-UP EFFORTS HAVE BEEN MADE TO ACQUIRE INFORMATION SURROUNDING THE PATIENT'S DEATH. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
MRIDIUM | K143369 | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A POTENTIAL UNDER-INFUSION EVENT. ACCORDING TO BIOMED, THE MRI TECHNICIAN OBSERVED A BUBBLE OF AIR IN THE IV LINE. HOWEVER, NO ADDITIONAL INFORMATION WAS PROVIDED REGARDING WHETHER ANY ALARMS OCCURRED DURING THE EVENT. PER BIOMED THE PATIENT CODED DURING THE EVENT. | THE DEVICE HISTORY LOG SUGGESTED THAT THE USER LIKELY HAD ISSUES PRIMING THE IV SET LINE. A "BUBBLE DETECTOR ACTIVITY" WAS LOGGED, INDICATING THAT THE BUBBLE DETECTOR WAS ACTIVE AND DETECTED A BUBBLE BELOW THE ALARM THRESHOLD IMMEDIATELY AFTER THE INFUSION STARTED. SUBSEQUENTLY, A "BUBBLE DETECTED" ALARM OCCURRED SHORTLY AFTER THE INFUSION BEGAN. THIS WAS CONSISTENT WITH THE EVENT DESCRIPTION PROVIDED BY THE BIOMED, WHERE THE USER OBSERVED AIR IN THE IV SET WHEN THE "BUBBLE DETECTED" ALARM WAS TRIGGERED. THE PRESENCE OF THE "BUBBLE DETECTED" ALARM, OCCURRING AFTER THE INFUSION PARAMETERS WERE SET AND IMMEDIATELY AFTER THE INFUSION BEGAN, STRONGLY INDICATED THAT AIR WAS PRESENT IN THE IV LINE AT THE START OF THE INFUSION. PROPER PRIMING OF THE IV SET IS CRUCIAL TO REMOVE ALL AIR FROM THE TUBING BEFORE CONNECTING IT TO THE PATIENT AND INITIATING THE INFUSION. THE OPERATOR'S MANUAL AND THE IV SET POUCH LABEL CONTAINED INSTRUCTIONS ON HOW TO PROPERLY PRIME THE IV SET. THE DEVICE HISTORY LOG SHOWED THAT THE USER OPENED THE DOOR AFTER THE BUBBLE ALARM, POTENTIALLY ATTEMPTING TO TROUBLESHOOT THE ISSUE. ALL ALARMS WERE SILENCED BY THE USER BEFORE THE INFUSION WAS CANCELED. A BUBBLE ALARM IN THIS CONTEXT IS A COMMON SIGN THAT THE PRIMING STEP MAY HAVE BEEN INSUFFICIENT, MISSED, OR THAT THE USER EXPERIENCED DIFFICULTY PRIMING AND SETTING UP THE LINE. THIS COULD HAVE LED TO AIR BEING DETECTED BY THE PUMP AND MAY HAVE RESULTED IN A DELAY OF THERAPY. THE CUSTOMER CONFIRMED THAT THE PATIENT CODED DURING THE EVENT AND THAT MEDICAL INTERVENTION WAS REQUIRED. THEREFORE, IRADIMED IS REPORTING THIS EVENT.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI ASSISTED NEPHRECTOMY, THERE WAS AN UNSPECIFIED SURGICAL COMPLICATION RESULTING IN MAJOR HEMORRHAGING. THE SURGEON REPORTED THAT THE BLEEDING REQUIRED A CONVERSION TO OPEN LAPAROTOMY AND AN OPEN THORACOTOMY. DESPITE ALL RESUSCITATIVE EFFORTS THE PATIENT EXPIRED IN THE OPERATING ROOM. THE SURGEON REPORTED THAT THE SURGICAL COMPLICATION WAS NOT RELATED TO THE ROBOTIC SYSTEM AND THERE WAS NO MALFUNCTION OF THE SYSTEM OR INSTRUMENTS. THE STATED CAUSE OF DEATH WAS MAJOR HEMORRHAGE. ADDITIONAL EVENT INFORMATION WAS REQUESTED BUT WAS NOT PROVIDED BY THE SURGEON. | THE FOLLOWING CONCOMITANT DEVICES WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, VESSEL SEALER EXTEND. A SITE REVIEW FOUND THAT NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THE PRODUCTS USED. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICES USED DURING THE EVENT FOUND NO RELATED NON-CONFORMANCES. DIGITAL SIGNAL PROCESSOR (DSP) LOGS SHOW THAT THE VESSEL SEALER EXTEND (VSE) INSTRUMENT WAS INSTALLED 3 TIMES AND PASSED HOMING 3 TIMES. ALL CUT COMPLETE AND SEAL EVENTS WERE NOTED WITH NO ERRORS. THE MASTER SUPERVISORY CONTROLLER (MSC) LOGS DID NOT SHOW ANY VSE RELATED ERRORS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT, WITH A SIGNIFICANT COMORBID DIAGNOSIS OF XANTHOGRANULOMATOUS PYELONEPHRITIS, WAS UNDERGOING A LEFT NEPHRECTOMY WHEN SIGNIFICANT BLEEDING OCCURRED. WHAT BLED AND HOW THEY GOT INTO THAT BLEEDING WAS NOT PROVIDED. THE PROCEDURE WAS CONVERTED TO AN OPEN LAPAROTOMY AND ALSO REQUIRED A THORACOTOMY WHICH SUGGESTS THEY HAD TO EMERGENTLY GET CONTROL OF THE AORTA. HOW THESE EVENTS UNFOLDED ARE NOT PROVIDED, HOWEVER, THE PATIENT EVENTUALLY DIED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
ION | K182188 | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, THE PATIENT'S BLOOD PRESSURE DROPPED SIGNIFICANTLY, CAUSING HEMODYNAMIC INSTABILITY LEADING TO DEATH. THE PROCEDURE INVOLVED ION NAVIGATION AND RADIAL ENDOBRONCHIAL ULTRASOUND (REBUS), IT WAS INTERRUPTED BEFORE THE BIOPSY COULD BE PERFORMED. APPROXIMATELY 10-15 MINUTES INTO NAVIGATION, THE ANESTHESIOLOGIST INFORMED THE PROCEDURE TEAM OF THE PATIENT¿S ¿DANGEROUSLY LOW¿ BLOOD PRESSURE. THE ION SYSTEM WAS UNDOCKED, AND CARDIOPULMONARY RESUSCITATION WAS INITIATED. THE PATIENT WAS PRONOUNCED DEAD APPROXIMATELY 45 MINUTES LATER. THERE WAS NO REPORTED DEVICE MALFUNCTION ASSOCIATED WITH THE EVENT. MULTIPLE ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION FROM THE PHYSICIAN WERE MADE; NO RESPONSE HAS BEEN RECEIVED. | A REVIEW OF THE ION SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND NO RELATED SYSTEM ERRORS OCCURRED DURING THE PROCEDURE THAT WERE RELEVANT TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION BRONCHOSCOPY, AND DURING NAVIGATION AND REBUS IMAGING, PRIOR TO BIOPSY, THE PATIENT DEVELOPED HYPOTENSION. EFFORTS AT RESUSCITATION WERE UNSUCCESSFUL AND THE PATIENT DIED. THERE WERE NO REPORTED MALFUNCTIONS OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. NO FURTHER DATA IS AVAILABLE DESPITE MULTIPLE EFFORTS. BASED ON THE AVAILABLE DATA THE EVENTS WERE PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. --FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. --FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. --KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. --BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. .
ION | K182188 | IT WAS REPORTED THAT DURING AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE, A PATIENT EXPERIENCED SIGNIFICANT BLEEDING, PULMONARY EMBOLISM, CARDIAC ARREST AND DEATH. THE INTUITIVE ENDOLUMINAL SALES MANAGER (ESM) WAS INFORMED BY HOSPITAL PERSONNEL THAT DURING THE BIOPSY PROCEDURE, A THIRD PARTY CRYOPROBE WAS USED TO COLLECT THE BIOPSY SAMPLES; THE BIOPSY SPECIMEN APPEARED SUGGESTIVE OF BLOOD VESSEL INVOLVEMENT. FOLLOWING THE BIOPSY, THE PATIENT HAD SIGNIFICANT BLEEDING, PROMPTING A BRONCHOSCOPIC AIRWAY INSPECTION AND SUCTIONING OF THE BLEEDING. CONE BEAM CT SCAN SHOWED A PULMONARY EMBOLISM. SHORTLY THEREAFTER, THE PATIENT SUFFERED CARDIAC ARREST AND REQUIRED RESUSCITATION, INCLUDING TRANSFUSION OF BLOOD PRODUCTS. THE PATIENT WAS STABILIZED; HOWEVER, EXPIRED AT AN UNSPECIFIED LATER TIME. THE ESM WAS INFORMED THAT THE ADVERSE EVENTS WERE NOT ATTRIBUTED TO THE ION SYSTEM, AND NO DEVICE MALFUNCTION WAS IDENTIFIED. THE COMPLICATIONS WERE BELIEVED TO BE RELATED TO THE CRYOPROBE BIOPSY THAT INDICATED VESSEL INVOLVEMENT. ADDITIONAL INFORMATION WAS REQUESTED FROM THE RESPIRATORY MEDICINE PHYSICIAN WHO STATED THE FACTS PROVIDED WERE INCORRECT, AND THAT THE MATTER IS CONFIDENTIAL AND INTERNAL REVIEW IS IN PROGRESS. NO CLARIFICATION OF THE INCORRECT FACTS WAS PROVIDED. | SYSTEM LOGS ARE NOT AVAILABLE FOR REVIEW. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY WHERE A CRYO-PROBE WAS USED TO OBTAIN BIOPSY SAMPLES. DURING THE PROCEDURE, AFTER BIOPSIES WERE OBTAINED, SIGNIFICANT BLEEDING WAS NOTED AND A PULMONARY EMBOLISM WAS DIAGNOSED ON CONE BEAM CT. THE PATIENT FURTHER DECOMPENSATED WITH A CARDIAC ARREST. DESPITE AGGRESSIVE INTERVENTIONS AND EFFORTS AT RESUSCITATION, THE PATIENT ULTIMATELY DIED. BASED ON THE AVAILABLE DATA THE BLEEDING, PULMONARY EMBOLISM, AND DEATH POST BIOPSY WERE LIKELY PROCEDURE RELATED. THERE IS NO COMPELLING EVIDENCE THAT THE EVENTS WERE DEVICE RELATED. OF NOTE, THE PHYSICIAN REPORTED DISCREPANCIES IN THE DETAILS BUT DECLINED TO PROVIDE FURTHER INFORMATION. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW-RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE-CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. THE RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). THE PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). THE RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). THERE WERE NO REPORTED EVENTS OF VENOUS THROMBOEMBOLISM (VTE) INCLUDING DEEP VENOUS THROMBOSIS OR PULMONARY EMBOLISM. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS AND NO VTE. HOWEVER, IT IS POSSIBLE THAT IMMOBILITY DURING GENERAL ANESTHESIA AND BLEEDING WITH SUBSEQUENT CLOTTING ASSOCIATED WITH LUNG BIOPSIES MAY CONFER SOME DEGREE OF INCREASED RISK FOR VTE EVENTS. LIKELY OF MORE RELEVANCE, IT IS ESTIMATED THAT CANCER CONFERS A 4-7X HIGHER RISK FOR VTE. THE HIGHER RATE OF VTE HAS BEEN REPORTED TO RANGE BETWEEN 3-13.9% IN PATIENTS WITH LUNG CANCER. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. BLOM JW, DOGGEN CJ, OSANTO S, ROSENDAAL FR. MALIGNANCIES, PROTHROMBOTIC MUTATIONS, AND THE RISK OF VENOUS THROMBOSIS. JAMA. 2005. BAGCHI A, KHAN MS, SARASWAT A, ANSARI A, NAI Q, IYER V, HAMOUDA D, KHUDER S, VERGHESE C. INCREASED INCIDENCE OF THROMBOTIC COMPLICATIONS WITH NON-SMALL CELL LUNG CANCER NECESSITATES CONSIDERATION OF PROPHYLACTIC ANTICOAGULATION IN YOUNG INDIVIDUALS. CUREUS. 2021. CHEW HK, DAVIES AM, WUN T, HARVEY D, ZHOU H, WHITE RH. THE INCIDENCE OF VENOUS THROMBOEMBOLISM AMONG PATIENTS WITH PRIMARY LUNG CANCER. J THROMB HAEMOST. 2008. CONNOLLY GC, DALAL M, LIN J, KHORANA AA. INCIDENCE AND PREDICTORS OF VENOUS THROMBOEMBOLISM (VTE) AMONG AMBULATORY PATIENTS WITH LUNG CANCER. LUNG CANCER. 2012.
CS300 INTRA-AORTIC BALLOON PUMP, SPANISH, 110V | K063525 | IT WAS REPORTED THAT CS300 INTRA-AORTIC BALLOON PUMP (IABP) SHUT DOWN DURING THE PATIENT'S TRANSFER FROM THE CORONARY ICU TO THE CT SCAN. EVENT OCCURRED ON (B)(6) 2025, 4:30 PM. THE PATIENT DIED, ALTHOUGH IT IS UNKNOWN IF IT WAS DUE TO THE EQUIPMENT. | DUE TO CHARACTER LIMITATION IN E1, FULL EVENT SITE NAME IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
REMSTAR AUTOA-FLEX | K113068 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A REMSTAR AUTOA-FLEX DEVICE HAS PASSED AWAY AND THE DAUGHTER WANTED TO RETURN THE DEVICE. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PORTEX ENDOTRACHEAL TUBE, SILICONIZED | K790312 | IT WAS REPORTED THAT AN PRETERM INFANT REQUIRED INTUBATION FOR ONGOING RESPIRATORY MANAGEMENT. DURING A PLANNED UPSIZING OF THE ENDOTRACHEAL (ET) TUBE, THE INFANT HAD A 3RD EPISODE OF A TYPICAL INTUBATION FINDING AND REQUIRED USE OF A RECALLED SIZE 2.0 ET TUBE. THE TUBE WAS USED IN THE KNOWLEDGE THAT IT HAD BEEN RECALLED BUT WAS USED AS A LIFESAVING INTERVENTION. THE INFANT WAS SUBSEQUENTLY TRANSFERRED TO ANOTHER NEONATAL UNIT FOR EAR, NOSE, AND THROAT (ENT) REVIEW, AND THE ET TUBE WAS STRAIGHTFORWARD, AND A SUCTION CATHETER COULD BE PASSED THROUGH THE ET TUBE TO PROVIDE ET SUCTION. THE INFANT HAD A SIZE 2.0 ENT TUBE INSERTED IN THE EARLY HOURS ON (B)(6) 2025, (B)(6) 2025 AS AN EMERGENCY PROCEDURE PRIOR TO THE ABOVE INTUBATION. IT WAS REPORTED TO BE UNKNOWN WHAT BRAND OF ENT TUBE THAT WAS, AND IT WAS NOT POSSIBLE TO A SUCTION CATHETER VIA THAT ET TUBE. THE INFANT HAD BILATERAL GRADE 4 INTRAVENTRICULAR HEMORRHAGES (IVHS) WITH EARLY HYDROCEPHALUS EVIDENT ON SCAN. CARE WAS RE-ORIENTED ON (B)(6) 2025, AND THE INFANT DIED. | H3, H6: NO PRODUCT OR PICTURES WERE RETURNED. THE REPORTED COMPLAINT COULD NOT BE CONFIRMED, AND A PROBABLE CAUSE COULD NOT BE DETERMINED. IF THE PRODUCT IS RETURNED THIS COMPLAINT WILL BE REOPENED FOR FURTHER INVESTIGATION.
VERSACROSS ACCESS SOLUTION | K150709 | IT WAS REPORTED THAT THE PATIENT EXPERIENCED A STROKE. THE PATIENT PASSED AWAY. PER LITERATURE REVIEW, IT WAS REPORTED THAT: PER LITERATURE REVIEW, THIS IS A SINGLE-CENTER RETROSPECTIVE STUDY OF CONSECUTIVE PATIENTS WHO UNDERWENT MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR (MTEER) WITH THE PASCAL (NON-BOSTON SCIENTIFIC DEVICE). TRANSSEPTAL PUNCTURE WAS UNDERTAKEN WITH EITHER A MECHANICAL NEEDLE OR THE VERSACROSS WIRE. THE PRIMARY ENDPOINTS WERE SUCCESS OF TRANSSEPTAL PUNCTURE (TSP) AND SUCCESSFUL DELIVERY OF THE EDWARDS SHEATH ON THE CHOSEN DELIVERY WIRE. SECONDARY ENDPOINTS INCLUDED NUMBER OF WIRES USED, TAMPONADE RATE, INTERVAL FROM FEMORAL VENOUS ACCESS TO TSP AND FIRST PASCAL DEVICE DEPLOYMENT, PROCEDURAL DEATH, AND STROKE. BETWEEN MARCH 2021 AND OCTOBER 2022, 33 CONSECUTIVE M-TEER CASES WERE IDENTIFIED. THIRTY-THREE CONSECUTIVE PATIENTS (10 WITH MECHANICAL NEEDLE, 23 WITH VERSACROSS) WHO UNDERWENT MTEER WITH THE EDWARDS PASCAL DEVICE WERE IDENTIFIED. ALL PATIENTS HAD SUCCESSFUL TSP. IN THE VERSACROSS ARM, THE RADIOFREQUENCY WIRE WAS USED SUCCESSFULLY FOR DELIVERY OF THE SHEATH IN ALL CASES. THERE WERE NO CASES OF PERICARDIAL EFFUSION/TAMPONADE IN EITHER ARM. INTERVAL FROM FEMORAL VENOUS ACCESS TO TSP AND TO DEPLOYMENT OF THE FIRST PASCAL DEVICE WAS SHORTER WITH THE VERSACROSS SYSTEM. SIGNIFICANTLY FEWER WIRES WERE USED WITH VERSACROSS. IN THE MECHANICAL NEEDLE GROUP, SIGNIFICANTLY MORE GUIDEWIRES WERE REQUIRED IN THE PROCEDURE. THERE WERE NO PROCEDURAL STROKES OR DEATHS. ONE LATE STROKE OCCURRED IN THE VERSACROSS GROUP AT 36 HOURS POST MTEER. THE SAME PATIENT SUBSEQUENTLY DIED BEFORE DISCHARGE. THE EMBOLIC STROKE WAS NOT DEEMED TO BE RELATED TO THE VERSACROSS DEVICE AND WAS ATTRIBUTED TO ATRIAL FIBRILLATION, A DIFFICULT PROCEDURE WITH MULTIPLE REPOSITIONINGS, AND THE PASCAL DEVICE ITSELF. THERE WERE NO PROCEDURAL DEATHS OR STROKES IN EITHER GROUP. IN CONCLUSION, VERSACROSS APPEARS A SAFE AND EFFECTIVE METHOD OF TSP AND FOR DELIVERY OF THE 22FR SHEATH FOR M-TEER WITH PASCAL. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS. ARTICLE CITATION: DOSHI, S. N., SAVVOULIDIS, P., MECHERY, A., LAWTON, E., AND NADIR, M. A. (2023). VERSACROSS TRANSSEPTAL SYSTEM FOR MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR WITH THE PASCAL REPAIR PLATFORM. STRUCTURAL HEART, 7(6), 100203. HTTPS://DOI.ORG/10.1016/J.SHJ.2023.100203. | THE FIELD B3: DATE OF EVENT WAS UPDATED WITH ESTIMATED DATE AS 22JUN2023, SINCE AOI HAS NO SPECIFIC PROCEDURE DATE, DATE OF WHEN AOI WAS PUBLISHED WAS USED AS EVENT DATE. IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. THIS WAS NOT AVAILABLE BECAUSE THIS WAS RELATED TO A LITERATURE STUDY. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. ARTICLE CITATION: DOSHI, S. N., SAVVOULIDIS, P., MECHERY, A., LAWTON, E., AND NADIR, M. A. (2023). VERSACROSS TRANSSEPTAL SYSTEM FOR MITRAL TRANSCATHETER EDGE-TO-EDGE REPAIR WITH THE PASCAL REPAIR PLATFORM. STRUCTURAL HEART, 7(6), 100203. HTTPS://DOI.ORG/10.1016/J.SHJ.2023.100203.
EUPHORA RX | K143480 | IT WAS REPORTED THAT DURING A PROCEDURE INVOLVING A EUPHORA BALLOON RX, THE DEVICE BECAME STUCK IN THE INFLATED POSITION WITHIN THE LEFT MAIN TO CIRCUMFLEX ARTERY. THE COMPLICATION AROSE DURING FINAL POST-DILATATION OF STENTS IN THE LEFT MAIN CORONARY ARTERY EXTENDING INTO THE LEFT CIRCUMFLEX ARTERY, WHEN A 3.5X15 MM NON-MEDTRONIC NC BALLOON FRACTURED. THE DISTAL (MONORAIL) OPAQUE TIP SEPARATED FROM THE HYPOTUBE AND REMAINED WITHIN THE CORONARY ARTERY, STUCK IN THE PREVIOUSLY IMPLANTED STENT IN THE LEFT MAIN TO THE CIRCUMFLEX ARTERY. MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO RETRIEVE THE FRACTURED NON-MEDTRONIC BALLOON. DURING EFFORTS TO RETRIEVE THE NON-MEDTRONIC DEVICE, A EUPHORA BALLOON RX WAS INTRODUCED AND SUBSEQUENTLY BECAME STUCK IN THE SAME SEGMENT, ALSO IN THE INFLATED POSITION. AN ATTEMPT WAS MADE TO DEFLATE THE EUPHORA BALLOON BY CUTTING THE SHAFT, BUT THIS WAS UNSUCCESSFUL. SEVERAL ADDITIONAL RETRIEVAL EFFORTS FAILED. THE PATIENT WAS ULTIMATELY REFERRED FOR SURGICAL INTERVENTION FOR POTENTIAL DEVICE RETRIEVAL AND CORONARY ARTERY BYPASS GRAFTING (CABG). | ADDITIONAL INFORMATION: THE LESION BEING TREATED WAS VERY CALCIFIED, AT A 90 DEGREE ANGLE, BIFURCATED AND HAD 90% STENOSIS. IT WAS NOT DIFFICULT TO REMOVE THE PROTECTIVE SHEATH OR THE PACKAGING STYLETTE. THE EUPHORA DEVICE WAS INSPECTED BEFORE USE WITH NO ISSUES NOTED. RESISTANCE WAS NOT NOTED WHILE ADVANCING THE DEVICE TO THE LESION. A CONCENTRATION OF 30/50 CONTRAST/SALINE WAS USED. THE PREVIOUSLY IMPLANTED STENTS WERE NON-MEDTRONIC DEVICES. THERE WERE NO DIFFICULTIES NOTED DURING INFLATION OF THE DEVICE. FOUR INFLATIONS WERE PERFORMED PRIOR TO THE DEFLATION DIFFICULTIES. INFLATION PRESSURES OF 15, 19, 18, AND 12 ATM WERE APPLIED, RESPECTIVELY, FOR EACH INFLATION. THE DEVICE WAS MOVED OR REPOSITIONED WHILE INFLATED. THE BALLOON FAILED TO DEFLATE. THE DETACHED PORTION OF THE EUPHORA DEVICE WAS NOT SUCCESSFULLY REMOVED FROM THE PATIENT. THE CABG WAS SUCCESSFULLY PERFORMED BUT WITH THE BALLOON STILL IN THE PATIENT ANATOMY. THE PATIENT LATER PASSED AWAY. DEVICE LOT NUMBER PROVIDED. PATIENT WEIGHT PROVIDED. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CLOTTRIEVER BOLD CATHETER | K230494 | A 72-YEAR-OLD FEMALE UNDERWENT A THROMBECTOMY WITH INARI DEVICES TO ADDRESS HER DEEP VEIN THROMBOSIS. THE PATIENT HAD CLOT RANGING FROM THE FEMORAL VEIN INTO THE INFERIOR VENA CAVA (IVC). THE PLAN WAS TO PERFORM THE PROCEDURE WITHOUT EMBOLIC PROTECTION. WIRE POSITIONING WAS OBTAINED WITHOUT ISSUE AND A CLOTTRIEVER BOLD CATHETER WAS ADVANCED INTO PLACE AFTER A 13 FR CLOTTRIEVER SHEATH WAS INSERTED. 3 PASSES WERE PERFORMED THEN THE VITALS MACHINE STOPPED PICKING UP THE PATIENT'S OXYGEN SATURATION. THE PATIENT THEN BECAME TACHYCARDIC AND BEGAN HAVING PREMATURE VENTRICULAR CONTRACTIONS. THE PROCEDURE WAS ABORTED IN ATTEMPTS TO STABILIZE THE PATIENT. A RAPID RESPONSE WAS INITIATED, AND A CODE WAS CALLED SOON AFTER. CARDIOPULMONARY RESUSCITATION WAS CONDUCTED FOR 45 MINUTES BEFORE THE PATIENT'S FAMILY DECIDED TO DISCONTINUE IT. THE PATIENT EXPIRED. FOLLOW-UP INFORMATION REVEALED THAT THE PATIENT'S INITIAL COMPUTED TOMOGRAPHY SCAN SHOWED A SMALL DISTAL PULMONARY EMBOLISM, THE PHYSICIAN DECIDED NOT TO INTERVENE. THE PHYSICIAN ALSO INDICATED THAT THE PATIENT'S DETERIORATION COULD POTENTIALLY BE THE RESULT OF SEDATION AND/OR CLOT BEING SENT TO THE PULMONARY ARTERIES. | THE INARI DEVICES WERE DISCARDED BY THE USER FACILITY AND ARE NOT AVAILABLE FOR EVALUATION. THE DEVICE IDENTIFIERS WERE NOT PROVIDED; THE LOT HISTORY RECORDS OF ALL POTENTIAL LOTS SHIPPED TO THE FACILITY WERE REVIEWED. THERE WERE NO DISCREPANCIES OR UNUSUAL FINDINGS. THERE WAS NO REPORT OF ANY DEVICE MALFUNCTION. BASED ON THE INFORMATION REVIEWED, THERE IS NO EVIDENCE TO INDICATE THE PRESENCE OF A POTENTIAL QUALITY ISSUE WITH RESPECT TO MANUFACTURING, DESIGN, OR LABELING. THE CASE WAS REVIEWED BY INARI'S CHIEF MEDICAL OFFICER, WHO CONCLUDED THAT THE PATIENT'S DEATH MAY HAVE BEEN THE RESULT OF INADVERTENT CLOT EMBOLISM. DISTAL EMBOLIZATION OF BLOOD CLOTS IS IDENTIFIED IN THE DEVICE LABELING AS POSSIBLE ADVERSE EVENTS/COMPLICATIONS. THE MANUFACTURER'S REFERENCE #: (B)(4).
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A NOTIFICATION RECEIVED REQUESTING THE RETURN OF A DEVICE. THE MANUFACTURER WAS INFORMED THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT PASSED AWAY AND WIFE REPORTED THAT SHE RECEIVED A LETTER STATING TO RETURN THE RECALLED DEVICE. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND VISIBLE FOAM PARTICLES WERE NOT OBSERVED. ADDITIONALLY, THE PATIENT¿S COMPLAINT WAS NOT CONFIRMED, AND THE DEVICE WAS SCRAPPED. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THERE WAS AN OVER INFUSION OF HEPARIN TO A PATIENT ADMITTED WITH UPPER BACK PAIN, WHO "CODED" FOLLOWING ADMISSION AND WAS ADMITTED TO ICU WITH POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION. THE HEPARIN (HEPARIN 25,000 UNITS/500ML) WAS INTENDED TO INFUSE AT A RATE OF 16ML/HOUR (800 UNITS/HOUR), HOWEVER THE MEDICATION BAG WAS FOUND TO BE EMPTY UNEXPECTEDLY QUICKLY. THE HEPARIN INFUSION WAS STARTED AT 0401, AND THE ENTIRE 500ML BAG WAS NOTED TO BE EMPTY AT 0520. THE PATIENT'S PARTIAL THROMBOPLASTIN TIME WAS 25.3 SECONDS ON THE INITIAL BLOOD DRAW AT 0120, THEN JUMPED TO >169 SECONDS FOLLOWING THE OVER INFUSION. THE PATIENT DID NOT RECEIVE THE ANTIDOTE BECAUSE THEIR BLOOD PRESSURE WAS TOO LOW, ACCORDING TO THE INTENSIVIST. THE PATIENT EXPIRED LATER THAT MORNING. THE CLINICIAN INDICATED THE CAUSE OF DEATH WAS MULTIFACTORIAL. THE CLINICIAN INDICATED THE PATIENT HAD A LARGE POSTERIOR SEGMENT ELEVATION MYOCARDIAL INFARCTION, "CODED", ACHIEVED RETURN OF SPONTANEOUS CIRCULATION, AND WAS ON MULTIPLE VASOPRESSORS. A REPORT WAS RECEIVED FROM HEALTH CANADA'S CANADA VIGILANCE PROGRAM WHICH STATES, "THE PUMP WAS PROGRAMED FOR 800UTS/HR. AND SHOULD HAVE INFUSED AT 16ML/HR. BUT THE PATIENT RECEIVED AN ENTIRE BAG OF HEPARIN IN APPROXIMATELY 1 HOUR. THE CONSEQUENCES OF THAT WERE AN EXTREMELY HIGH PTT. THE PATIENT WAS NOT ABLE TO RECEIVE THE ANTIDOTE DUE TO CLINICAL CONDITION (LOW BP) AS DECIDED BY ICU. THE PATIENT PASSED BUT THE CAUSE OF THIS WAS MULTIFACTORIAL." | ADDITIONAL INFORMATION: OTHER RELEVANT HISTORY, DEVICE EVAL BY MANUFACTURER?, IMDRF ANNEX B CODE, REMEDIAL ACTION REQUIRED, REMEDIAL ACTION # AND MANUFACTURER NARRATIVE. A DEVICE HISTORY RECORD, COMPLAINT HISTORY REVIEW, AND RISK REVIEW ON FAILURE MODES ARE PERFORMED ON EACH DEVICE WITH REPORTABLE MALFUNCTION(S) ALONG WITH OTHER METHODS OF INVESTIGATION AS CODED IN SECTION H6 OF THIS MDR REPORT. INVESTIGATION SUMMARY: THE REPORT OF AN OVER-INFUSION OF HEPARIN WAS CONFIRMED THROUGH LOG ANALYSIS AND ATTRIBUTED TO A USER PROGRAMMING ERROR, WHICH REVEALED A DISCREPANCY BETWEEN THE REPORTED HEPARIN CONCENTRATION OF 25000U/500 ML AND THE ENTERED CONCENTRATION OF 800 U/500 ML. THE CONCENTRATION ENTRY ERROR RESULTED IN A PUMP CALCULATED RATE OF 500 ML/H WHEN A DOSE OF 800 U/H WAS ENTERED. THE REPORT INDICATED THE DESIRED RATE WAS 16 ML/HOUR. THE EVENT LOG REVIEW COULD NOT CONFIRM THE REPORT THAT ON THE DATE (B)(6) 2024 AT THE TIME OF 0401 AN INFUSION HEPARIN WAS STARTED AND AT THE TIME OF 0520 THE HEPARIN BAG WAS FOUND EMPTY. THE INSPECTION AND TESTING PROCESS DID NOT FIND ANY EXISTING MALFUNCTIONS. THE SUSPECT PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATIONS WITH NO OBSERVANCES OF UNREGULATED FLOW OCCURRING. THE SEAR WAS ALSO FOUND TO BE PROPERLY CLOSING THE ADMINISTRATION SET SAFETY CLAMP WHEN THE DOOR WAS OPENED. AFTER UNREGULATED FLOW TESTING, THE SUSPECT PUMP MODULE WAS PROGRAMMED TO PERFORM A ONE-HOUR SYSTEM LEVEL RATE ACCURACY TEST AT THE INCIDENT INFUSION RATE OF 16ML/HR AND TEST RESULTS MEASURED THE ACTUAL RATE AT 15.43ML/HR (-3.55% ERROR) INDICATING THE DEVICE IS WITHIN SPECIFICATION AFTER RATE CALIBRATION HAVING NO INCIDENTS OF UNREGULATED FLOW OCCURRING DURING THE TESTING. NO ERRORS WERE OBSERVED ON BOTH PCU AND SUSPECT PUMP MODULE ON THE REPORTED DATE OF THE EVENT. REVIEW OF THE SUSPECT PUMP MODULE ERROR LOG SHOWED NO MALFUNCTIONS RELEVANT TO THE FACILITY¿S COMPLAINT. ON THE DATE (B)(6) 2024 AT THE TIME OF 2:25 AM, THE CLINICIAN WAS RECORDED TO SELECT HEPARIN (ACS) (DRUGID=251) FROM GUARDRAIL DRUGS. THE HEPARIN SELECTION WAS A CUSTOM CONCENTRATION WITH THE CLINICIAN ENTERING A DRUG AMOUNT OF 800 UNITS AND A DILUENT VOLUME 500 ML. THE INTENDED CONCENTRATION HAD BEEN 25000U/500 ML. THE ENTRY ERROR RESULTED IN A SOFT GUARDRAIL CONCENTRATION ALERT WHICH THE USER ACCEPTED AND CONTINUED PROGRAMMING THE INFUSION. THE DOSE WAS ENTERED CORRECTLY AT 800 U/HR BUT BECAUSE OF THE CONCENTRATION ERROR THE INFUSION RATE CALCULATED TO 500 ML/H INSTEAD OF THE INTENDED 16 ML/H. THE INFUSION STARTED AT 2:27 AM BUT WAS PAUSED A SHORT TIME LATER BY THE CLINICIAN WITH A PRIMARY VOLUME INFUSED OF 4.171 ML. AT 2:31 AM THE SUSPECT WAS PROGRAMED TO CONTINUE THE INFUSION WITH A VOLUME TO BE INFUSED OF 495.004ML. THE INFUSION FINISHED AT 3:33 AM WITH A REPORTED PRIMARY VOLUME INFUSED OF 511.403ML ML AFTER INFUSING FOR 61 MINUTES. THE ADMINISTRATION SET WAS REQUESTED BUT WAS NOT RETURNED; THEREFORE, IT COULD NOT BE INSPECTED OR TESTED. ROOT CAUSE: THE ROOT CAUSE FOR THE REPORT OF OVER INFUSION OF HEPARIN WAS ABLE TO BE IDENTIFIED FROM THE LOG ANALYSIS AND IS BEING ATTRIBUTED TO USER PROGRAMMING ERROR. THE SUSPECT PUMP MODULE WAS FOUND TO BE INFUSING WITHIN SPECIFICATION, AND THE SEAR WAS FOUND TO BE PROPERLY CLOSING THE SAFETY CLAMP WHEN THE ADMINISTRATION SET IS REMOVED.
POWERPORT SLIM IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 6F | K072549 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH PULMONARY EMBOLISM, BLOOD CLOT AND OTHER SEVERE PAIN WITH ONGOING MEDICAL TREATMENT. LATER THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H10: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS OF (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
SIMUFORM | K201449 | MEDTRONIC RECEIVED INFORMATION THAT 28 DAYS POST IMPLANT OF THIS SIMUFORM RING, THE PATIENT DIED. THE CAUSE OF DEATH WAS REPORTED AS A BRAIN ANEURYSM. RELATEDNESS OF THE DEATH TO THE SIMUFORM RING COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REP DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION IN REFERENCE TO THE "REP DREAMSTATION AUTO CPAP" DEVICE ALLEGING DEATH OF THE PATIENT, EYE IRRITATION, NOSE IRRITATION, HYPERSENSITIVITY AND CANCER. THERE WAS A REPORT OF SERIOUS OR PERMANENT HARM OR INJURY. NO MEDICAL INTERVENTION WAS SPECIFIED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
EDWARDS ESHEATH+ INTRODUCER SET | K200258 | AS REPORTED BY AN EDWARDS LIFESCIENCES JAPAN ASSOCIATE AND SUPPLEMENTED BY INFORMATION FROM THE JAPANESE TAVI REGISTRY, THE PATIENT UNDERWENT A LEFT TRANSFEMORAL TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE USING A 23 MM SAPIEN 3 ULTRA RESILIA VALVE. THE INITIAL APPROACH INVOLVED INSERTION OF A 14 FR ESHEATH+ WITHOUT RESISTANCE. HOWEVER, THE 23 MM DELIVERY SYSTEM WAS UNABLE TO ADVANCE THROUGH THE SHEATH DESPITE SEVERAL ATTEMPTS, PROMPTING THE PHYSICIAN TO SWITCH TO A LEFT SUBCLAVIAN APPROACH. A SECOND DEVICE KIT WAS OPENED, AND THE VALVE WAS SUCCESSFULLY IMPLANTED VIA THE SUBCLAVIAN ARTERY WITHOUT FURTHER TECHNICAL DIFFICULTIES. DURING THE ATTEMPTED TRANSFEMORAL ACCESS, A VASCULAR RUPTURE OCCURRED IN THE LEFT COMMON FEMORAL ARTERY (CFA), WHICH REQUIRED SURGICAL REPAIR. POST-DEPLOYMENT ANGIOGRAPHY REVEALED THROMBOEMBOLISM IN THE VESSEL DISTAL TO THE KNEE, FOR WHICH ENDOVASCULAR TREATMENT (EVT), INCLUDING THROMBECTOMY, WAS PERFORMED WITH SUCCESSFUL RECANALIZATION. ON POSTOPERATIVE DAY (POD) 2, THE PATIENT DEVELOPED ACUTE OCCLUSION OF THE LEFT LOWER LIMB ARTERY AND UNDERWENT EMERGENCY THROMBECTOMY AND EVT. HOWEVER, LIMB ISCHEMIA PERSISTED. ON POD 3, CONTRAST-ENHANCED COMPUTED TOMOGRAPHY SHOWED STENOSIS OF THE LEFT SUPERFICIAL FEMORAL ARTERY (SFA) AND CFA. EMERGENCY EVT WAS AGAIN PERFORMED WHILE THE PATIENT WAS ON CONTINUOUS HEMODIAFILTRATION. ALTHOUGH PARTIAL IMPROVEMENT IN BLOOD FLOW WAS OBSERVED, THE PATIENT DEVELOPED WORSENING ACIDOSIS AND REFRACTORY HYPOTENSION, ULTIMATELY RESULTING IN DEATH ON POD 3. THE CAUSE OF DEATH WAS REPORTED AS ACUTE LOWER LIMB ARTERIAL OCCLUSION. PER THE ATTENDING MEDICAL TEAM, THE PATIENT'S LONG-TERM DIALYSIS LIKELY CONTRIBUTED TO VASCULAR FRAGILITY, WHICH MAY HAVE BEEN A FACTOR IN BOTH THE DIFFICULTY ADVANCING THE DELIVERY SYSTEM AND THE SUBSEQUENT VASCULAR COMPLICATIONS. | A SUPPLEMENTAL MDR IS BEING SUBMITTED, FOLLOWING COMPLETION OF THE ENGINEERING EVALUATION. B4, G3, G6, AND H2 HAVE BEEN UPDATED. H6: HEALTH EFFECT - CLINICAL, TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, INVESTIGATION CONCLUSIONS, AND H11 HAVE BEEN CORRECTED. THE EVENTS REPORTED ARE ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT AND ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE DEVICE WAS NOT RETURNED FOR EVALUATION. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY, THE INSTRUCTIONS FOR USE (IFU), TRAINING MANUALS AND CURRENT RISK MITIGATIONS HAVE BEEN REVIEWED. NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. IMAGERY OF THE PATIENT'S ANATOMY WAS PROVIDED AND REVEALED THE PRESENCE OF CALCIFICATION AND TORTUOSITY IN THE PATIENT'S LEFT ACCESS VESSEL. PER THE INSTRUCTIONS FOR USE (IFU), POTENTIAL ADVERSE EVENTS ASSOCIATED WITH THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE INCLUDE THROMBUS FORMATION, PLAQUE DISLODGMENT, AND EMBOLIZATION THAT MAY RESULT IN MYOCARDIAL INFARCTION, STROKE, DISTAL PERIPHERAL OCCLUSION, AND/OR DEATH. THROMBOEMBOLISM IS THE OBSTRUCTION OF A BLOOD VESSEL BY A BLOOD CLOT THAT HAS BECOME DISLODGED FROM ANOTHER SITE IN THE CIRCULATION. IT IS THE NATURAL TENDENCY OF THE BODY TO FORM A CLOT ON FOREIGN OBJECTS IN THE VASCULAR SPACE. THESE PATIENTS ARE ANTICOAGULATED FOR THE PROCEDURE AND INTERVENTIONAL BEST PRACTICES MANDATE METICULOUS WIPING AND FLUSHING OF THE DEVICES TO PREVENT AND/OR REMOVE THE CLOT. DISPLACEMENT OF CALCIUM DEPOSITS OR ATHEROMA WITH EMBOLIZATION OF DEBRIS CAN ALSO TRIGGER THE FORMATION OF A THROMBUS WHICH CAN THEN EMBOLIZE AND CAUSE BLOCKAGE IN A SMALLER VESSEL. CALCIFICATION AND ATHEROMA (CELLULAR DEBRIS, INFLAMMATORY CELLS, AND CHOLESTEROL) CAN ACCUMULATE ALONG THE WALLS OF THE VASCULATURE AND WITHIN CARDIAC STRUCTURES AND CAN VARY IN SIZE. THERE MAY BE CASES WHERE A PATIENT HAS A PROTRUDING ATHEROMA OR CALCIFIED PLAQUE PRIOR TO THE PROCEDURE. IT IS ALSO POSSIBLE FOR ONE OF THE INTERVENTIONAL DEVICES USED DURING THE THV PROCEDURE TO DISRUPT THE MATERIAL, RESULTING IN MOBILE DEBRIS. ANNULAR STRUCTURE TISSUE OR CALCIFICATION CAN ALSO BE DISRUPTED DURING SHEATH INSERTION, BALLOON VALVULOPLASTY, AND DEPLOYMENT OF THE PROSTHETIC VALVE. THE DEVICE TRAINING MANUALS AND IFU INSTRUCT THE OPERATOR TO ADMINISTER HEPARIN AND MAINTAIN THE ACT AT = 250 SEC. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN THV (ALL MODELS). TRAINING INCLUDES PATIENT SCREENING, DEVICE PREPARATION, APPROACH, DEPLOYMENT, IMAGING, PROCEDURE-SPECIFIC TRAINING MANUALS, AND PROCTORED PROCEDURES. THE TRAINING MANUAL INSTRUCTS THE OPERATOR ON PROPER SHEATH INSERTION AND WITHDRAWAL TECHNIQUES, INCLUDING PRE-DILATING THE VESSEL WITH THE EDWARDS DILATORS, AS NEEDED. IN THIS CASE, THERE WAS NO ALLEGATION OR INDICATION A PRODUCT MALFUNCTION CONTRIBUTED TO THE REPORTED THROMBOEMBOLISM. THE EXACT CAUSE OF THE REPORTED EVENT CANNOT BE DETERMINED. HOWEVER, IN ADDITION TO THE MECHANISMS DESCRIBED ABOVE, ANOTHER VASCULAR COMPLICATION (RUPTURE IN THE LEFT COMMON FEMORAL ARTERY) MAY HAVE CONTRIBUTED TO THE EVENT. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. IN THIS CASE, THE INABILITY TO ADVANCE THROUGH THE SHEATH WAS UNABLE TO BE CONFIRMED DUE TO UNAVAILABILITY OF RETURNED DEVICE/APPLICABLE IMAGERY. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT, AS CONFIRMED VIA IMAGERY OF THE PATIENT'S ANATOMY. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. SCRATCHES OBSERVED ON THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF CALCIFIED NODULES WITHIN THE ACCESS VESSEL. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AN INVESTIGATION HAS BEEN OPENED TO DETERMINE THE ROOT CAUSE AND IMPLEMENT ANY NECESSARY CORRECTIVE ACTIONS.
INVACARE® PLATINUM® OXYGEN CONCENTRATOR | K160630 | THE DEVICE'S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THE FOLLOWING EVENT: "CONCENTRATOR STOPPED WORKING, WAS ALARMING, TECH ARRIVED WITHIN 2HRS AND PATIENT HAD DIED. FAMILY MEMBER WOULD NOT LET THE TECH TAKE THE CONCENTRATOR. TECH WAS ABLE TO CHECK THE CONCENTRATOR, HE MENTIONED WHEN HE TOOK THE TUBING OFF THE CONCENTRATOR WOULD RUN BUT WHEN HE PUT THE TUBING ON IT WOULD SHUT DOWN AND START ALARMING. THERE WAS NO HUMIDIFIER BOTTLE ATTACHED TO THE CONCENTRATOR." THE REPORTED EVENT INVOLVED AN INVACARE® PLATINUM® OXYGEN CONCENTRATOR. NO FURTHER INFORMATION ABOUT THE PATIENT WAS PROVIDED. | H6: VENTEC HAS DETERMINED THAT THIS DEVICE WAS PLACED INTO COMMERCIAL DISTRIBUTION BACK IN AUGUST OF 2014, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ VENTEC SPOKE TO THE INITIAL REPORTER (DISTRIBUTOR) WHO ADVISED THAT THEY HAVE THE TUBING AND THE O2 CONCENTRATOR IN THEIR POSSESSION. THE DISTRIBUTOR FURTHER ADVISED THAT THEY WOULD NOT BE SENDING IT TO VENTEC (REACT HEALTH) FOR AN EVALUATION. THE DISTRIBUTOR STATED THAT IF THAT CHANGES IN THE FUTURE, THEY WOULD CONTACT VENTEC AND REQUEST A RETURN AUTHORIZATION (RMA). THE DEVICE WAS NOT RETURNED TO VENTEC FOR EVALUATION. VENTEC HAS DETERMINED THAT THERE IS INSUFFICIENT INFORMATION TO DETERMINE IF THE DEVICE USE CONTRIBUTED TO THE PATIENT'S OUTCOME. THE CAUSE OF THE REPORTED ISSUE COULD NOT BE DETERMINED. IF VENTEC RECEIVES ANY NEW INFORMATION REGARDING THIS EVENT, A FOLLOW-UP REPORT SHALL BE SUBMITTED AS DEFINED BY 21 CFR 803.56.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A PERITONEAL DIALYSIS REGISTERED NURSE [(PD)RN] REPORTED TO FRESENIUS CUSTOMER SERVICE THIS PD PATIENT ON CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER EXPIRED WHILE CONNECTED TO THE CYCLER. THERE WAS NO SPECIFIC ALLEGATION THAT THIS EVENT WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) IN THE INITIAL REPORTING. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REPORTED THIS PATIENT EXPIRED AT HOME ON (B)(6) 2025 DUE TO AN UNREPORTED CAUSE. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO THEIR LIBERTY SELECT CYCLER AT THE TIME OF DEATH. THE PATIENT WAS FOUND BY FAMILY TO BE UNRESPONSIVE AND PULSELESS ON THE MORNING OF (B)(6) 2025. EMERGENCY SERVICES WERE ACTIVATED BUT THE PATIENT WAS NOT TRANSPORTED AND PRONOUNCED DECEASED IN HOME. THE CAUSE OF DEATH WAS NOT REPORTED TO THE OUTPATIENT CLINIC; HOWEVER, THE PATIENT HAD A RECENT DECLINE IN HEALTH INVOLVING ADVANCED DEMENTIA THAT PREEXISTED HIS START ON PD THERAPY IN (B)(6) 2025. IT WAS CONFIRMED, THOUGH THE CAUSE OF DEATH WAS NOT REPORTED TO THE CLINIC, THE PATIENT¿S DEATH WAS DETERMINED TO BE UNRELATED TO PD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH. THE CAUSE OF THIS PATIENT¿S DEATH CANNOT BE DETERMINED; HOWEVER, THERE WAS NO INDICATION THE PATIENT¿S DEATH WAS DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S) AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
REMSTAR PRO C-FLEX+ | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE PATIENT ALLEGED DEVICE TURNS OFF RANDOMLY. THE PATIENT WAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER WAS PREVIOUSLY CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP/ BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE CUSTOMER ALLEGED DEVICE TURNS OFF RANDOMLY. THE PATIENT PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. AFTER MULTIPLE ATTEMPTS, THE CUSTOMER CONFIRMED THAT THE DEVICE DID NOT CONTRIBUTE TO THE PATIENT'S DEATH. THERE WAS NO PATIENT HARM OR INJURY ASSOCIATED WITH THIS REPORT AND NO INCREASED RISK TO THE INTENDED USER. BASED ON THE INFORMATION AVAILABLE, IT IS A NON-REPORTABLE COMPLAINT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE DEVICE WAS RETURNED TO THE MANUFACTURER AND INSPECTED. NO ERRORS WERE FOUND; THEREFORE, NO COMPONENTS WERE REPLACED TO CORRECT A MALFUNCTION. THE DEVICE PASSED ALL TESTS AND WAS RETURNED TO STOCK.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION B5 ADDITIONAL INFORMATION: PRIOR TO THE PATIENT'S CARDIAC ARREST, BLOOD WAS NOTED IN THE SURGICAL DRAIN; THERE WAS NO BLOOD IN THE DRAIN EARLIER IN THE DAY. AN EMERGENCY EXPLORATORY LAPAROTOMY WAS PERFORMED, REVEALING A HEMOPERITONEUM. THE BLEEDING SOURCE WAS IDENTIFIED AT THE JUNCTION BETWEEN THE PANCREATIC HEAD AND BODY, SPECIFICALLY AT THE SECTIONED RAW TRANSITION AREA OF THE "RESIDUAL PANCREATIC PARENCHYMA". THIS HEMORRHAGE WAS ATTRIBUTED TO THROMBOCYTOPENIA. A REVIEW OF THE ADDITIONAL EVENT INFORMATION PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT DIED FOLLOWING A LATE BLEEDING EVENT SEVERAL DAYS AFTER A PANCREATICODUODENECTOMY. THE LOCATION OF THE BLEED WAS NOTED TO BE THE RAW SURFACE OF THE PANCREATIC TRANSECTION. THE CAUSE OF THE BLEED IS UNCLEAR ALTHOUGH IT WAS NOTED THE PATIENT HAD BECOME THROMBOCYTOPENIC AND MAY HAVE DEVELOPED COAGULOPATHY. THE REASON FOR THE COAGULOPATHY IS UNCLEAR. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAM STATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. .THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO BIPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS ALLEGED DEATH. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. .
INNOSIGHT | K192254 | IT WAS REPORTED THAT A THROMBUS WAS NOT VISUALIZED ON THE INNOSIGHT ULTRASOUND SYSTEM DURING A PORTABLE FOCUSED ECHO SCAN AND PATHOLOGY. A REPEAT ECHO WAS PERFORMED AT A LATER DATE USING A DIFFERENT SYSTEM. THE THROMBUS WAS VISUALIZED. THE PATIENT RECEIVED A BONE MARROW TRANSPLANT AND SUBSEQUENTLY DIED. AN INTERNAL SITE INVESTIGATION IS CURRENTLY UNDERWAY TO DETERMINE WHETHER THE LACK OF VISUALIZATION OF THE SUBSTRATE CONTRIBUTED TO CAUSE OF DEATH. PHILIPS HAS STARTED AN INVESTIGATION, AND UPON COMPLETION, A FOLLOW-UP REPORT WILL BE SUBMITTED.
REP DREAMSTATION AUTO BIPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A REP DREAMSTATION AUTO BIPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
PATIENT INFORMATION CENTER IX | K211900 | IT WAS REPORTED THE ALARM DID NOT SOUND AND THE INTENSIVE CARE UNIT (ICU) PATIENT PASSED AWAY. | A PHILIPS REMOTE SERVICE ENGINEER (RSE) SPOKE WITH THE CUSTOMER. THE EVENT TOOK PLACE ON (B)(6) 2025 AT BED D137. WHILE A SPECIFIC TIME FOR THE FAILURE WAS NOT PROVIDED IN THE ALLEGATION, THE REMOTE INVESTIGATION FOCUSED ON THE TIMELINE OF 13:22 THROUGH THE PATIENT DISCONNECTION FROM THE MONITOR AT 13:58. THE PATIENT WAS BEING MONITORED FOR ECG AND FOR SPO2. PER THE LOG REVIEW, 'DURING THE OBSERVATION PERIOD, A TOTAL OF EIGHT EXTREME BRADYCARDIA RED ALARMS WERE RECORDED, ACCOMPANIED BY SEVERAL YELLOW ALARMS, IN A CONTEXT MARKED BY REPEATED TECHNICAL ALERTS CONCERNING THE SPO2 SENSOR (SIGNAL DIFFICULT TO CAPTURE, AGITATED PATIENT, POORLY POSITIONED SENSOR). THE FIRST RED ALARMS OCCURRED FROM 13:22:23, WITH A SUCCESSION OF CRITICAL EVENTS UNTIL 13:26:02. NO MANUAL ACKNOWLEDGMENT WAS PERFORMED DURING THIS INITIAL SEQUENCE. THE FIRST MANUAL ACKNOWLEDGMENT OF A RED ALARM AT THE CENTRAL UNIT WAS RECORDED AT 13:27:41, FOR AN ALARM GENERATED AROUND 13:26. THE ASYSTOLE RED ALARMS BEGAN AROUND 13:29, WITH A MANUAL ACKNOWLEDGMENT AT THE CENTRAL UNIT 30 SECONDS LATER. OTHER REPETITIONS OF THE ASYSTOLE ALARM WERE OBSERVED, NOTABLY AT 1:32 P.M., 1:35 P.M., 1:39 P.M., 1:42 P.M. AND 1:46 P.M., WITH SUCCESSIVE ACKNOWLEDGMENTS AT THE CENTRAL UNIT BUT ALSO AT THE MONITOR IN THE PATIENT'S ROOM, THE FIRST BEING NOTED AT 1:40:54 AND THE SECOND AT 1:46 P.M. THE SPO2 SENSOR, OFTEN DISCONNECTED OR POORLY POSITIONED, COMPLICATED THE QUALITY OF MONITORING. THE EPISODE ENDED WITH THE PATIENT BEING DISCONNECTED AT 1:58 P.M., PUTTING AN END TO THE ASYSTOLE ALARMS.' THE SPO2 SENSOR, WHICH WAS OFTEN DISCONNECTED OR POORLY POSITIONED, COMPLICATED THE QUALITY OF THE MONITORING. THE EPISODE ENDED WITH THE PATIENT BEING DISCONNECTED AT 13:58, ENDING THE ASYSTOLE ALARMS. A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) AND A PHILIPS CLINICAL PRODUCT SPECIALIST (CPS) REVIEWED THE PROVIDED INFORMATION AND CONFIRMED THE RSE FINDINGS. THE PSE ALSO REVIEWED PIC IX TECHNICAL LOG DATA FROM THE "SURVEILLANCE" HOST. THE PSE NOTED SEVERAL THREE STAR (***) XBRADY ALARMS AND AN ASYSTOLIE SENT TO THE RECORDER. THE TIMESTAMPS CORRELATE WITH THE SOUND MANAGER ENTRIES FROM THE PIC, INDICATING THE RED ALARM SOUND BEING PLAYED AT THE CENTRAL STATION. THE REPORTED ISSUE WAS NOT CONFIRMED. THE CAUSE OF THE DEATH AND THE RELATION TO THE DEVICE REMAINS UNKNOWN DUE TO LIMITED INFORMATION FROM THE CUSTOMER. THE ORIGINAL EVENT DESCRIPTION INDICATES THE PATIENT WAS NOT TREATED IN TIME. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE A DEVICE MALFUNCTION WHICH CAUSED OR CONTRIBUTED TO THE REPORTED EVENT; HOWEVER, ALARM MANAGEMENT, CLINICAL WORKFLOW, OR OTHER USE ISSUE MAY HAVE BEEN FACTORS.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE HAD FOUND THE DEVICE AFTER THE PATIENT PASSED AWAY A FEW YEARS AGO AND WOULD LIKE IT REPLACED. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIFEPAK 15 DEFIBRILLATOR/MONITOR | K142430 | THE CUSTOMER CONTACTED STRYKER TO REPORT THAT THEIR DEVICE WAS CONNECTED TO A PATIENT IN CARDIAC ARREST AND FAILED TO FLAG THE PATIENT'S ECG RHYTHM AS VENTRICULAR FIBRILLATION. AS A RESULT, DEFIBRILLATION THERAPY WAS NOT DELIVERED TO THE PATIENT. THE PATIENT IS DECEASED AND THE CUSTOMER IS UNSURE IF THE DEVICE CONTRIBUTED TO THE PATIENT'S OUTCOME. | STRYKER CONTACTED THE CUSTOMER TO REQUEST ADDITIONAL INFORMATION ON THE PATIENT. STRYKER REQUESTS PATIENT INFORMATION NECESSARY FOR REGULATORY COMPLIANCE, STRICTLY ADHERING TO HIPAA'S PRIVACY RULE. THE CUSTOMER PROVIDED STRYKER WITH ALL THE AVAILABLE PATIENT INFORMATION. PATIENT FIELDS IN WHICH INFORMATION WAS NOT PROVIDED WERE INTENTIONALLY LEFT BLANK. SECTION A1 PATIENT IDENTIFIER CHARACTER LIMIT REACHED. PATIENT IDENTIFIER IS (B)(6) STRYKER CONTINUES TO INVESTIGATE THE REPORTED FAILURE AND WILL SUBMIT A SUPPLEMENTAL REPORT ON THIS EVENT TO THE FDA AS PROVIDED BY 21 CFR 803.56.
EVERFLO INTL OPI | K061261 | THE MANUFACTURER RECEIVED INFORMATION THAT PATIENT HAS PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO CPAP DEVICE'S SOUND ABATEMENT FOAM. THE PATIENT HAS PASSED AWAY. THE DEVICE HAS BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. DURING THE EVALUATION IT IS OBSERVED THAT THERE ARE 0 ERRORS FOUND IN LOGS, NO FAULTS FOUND, AND ALL TESTING PASSED.
MEDICALYTE | K012547 | THE PATIENT STARTED DIALYSIS TREATMENT AND THE NURSE NOTICED CONTAMINATION IN THE BICARBONATE DIALYSATE. THE TREATMENT WAS STOPPED AND BLOOD CULTURES PROVIDED ELEVATED RESULTS. THE PATIENT WAS GIVEN ANTIBIOTICS, AND LATER PASSED AWAY. THE RN STATED THE DEATH DID NOT APPEAR TO BE RELATED TO THE CONTAMINATED PRODUCT.
OPTICROSS? HD | K173284 | IT WAS REPORTED THAT A PATIENT DEATH OCCURRED. THE STENOSED TARGET LESION WAS LOCATED IN THE DISTAL LEFT MAIN CORONARY ARTERY. THE OPTICROSS HD IMAGING CATHETER WAS ADVANCED FOR INTRAVASCULAR ULTRASOUND (IVUS) OF THE TARGET LESION. AS THE PHYSICIAN ATTEMPTED TO PASS THE IVUS CATHETER ACROSS THE LESION, IMAGING WAS ACTIVATED AND AN ARTIFACT WAS OBSERVED ON THE SCREEN. THE TECHNICIAN WAS ASKED WHETHER FLUSHING WAS OCCURRING, AND THEY CONFIRMED THAT IT WAS. THE TECHNICIAN WAS THEN INSTRUCTED TO STOP FLUSHING, AND THE PHYSICIAN AND TECHNICIAN DECIDED TO REMOVE AND RE-PREP THE CATHETER. THE PHYSICIAN NOTED ELEVATED ST-SEGMENTS, AND CONTRAST INJECTION REVEALED A SLOW-FLOW PHENOMENON. A DECISION WAS MADE TO ASPIRATE THE LEFT ANTERIOR DESCENDING ARTERY (LAD). ASPIRATION WAS PERFORMED, WHICH RESTORED SOME FLOW TO THE PROXIMAL LAD. HOWEVER, THE PATIENTS CONDITION WORSENED, AND THEY WENT INTO CARDIAC ARREST WITH PULSELESS ELECTRICAL ACTIVITY (PEA). CARDIOPULMONARY RESUSCITATION (CPR) WAS INITIATED, AND A PERCUTANEOUS LEFT VENTRICULAR ASSIST DEVICE (LVAD) WAS INSERTED. CORONARY ANGIOGRAPHY PERFORMED AFTER IMPELLA PLACEMENT SHOWED NORMAL CORONARY FLOW. DESPITE EFFORTS, THE PATIENT DIED FOLLOWING A PROLONGED PEA ARREST. POST-PROCEDURE INSPECTION OF THE CATHETER REVEALED AIR IN THE CLEAR LUMEN. IT WAS NOTED THAT THE SUPPLIED EXTENSION TUBING HAD NOT BEEN USED; INSTEAD, TUBING FROM A DIFFERENT MANUFACTURER HAD BEEN CONNECTED TO THE BOSTON SCIENTIFIC STOPCOCK AND SYRINGES. THE PHYSICIAN BELIEVED THAT IMPROPER CATHETER PREPARATION LED TO AN AIR EMBOLISM. THE PROCEDURE WAS NOT COMPLETED DUE TO THIS EVENT. | G4: PREMARKET / 510(K) # K173284, K213593.
GALAXY SYSTEM | K223144 | IT WAS REPORTED A PATIENT'S DEATH OCCURRED FOLLOWING A GALAXY SYSTEM BIOPSY PROCEDURE FOR A LESION IN THE LEFT LOWER LOBE. THE PATIENT WAS INITIALLY REPORTED TO HAVE BLED IN ACCORDANCE WITH TYPICAL BIOPSY PROTOCOL, WITH BLEEDING APPEARING TO BE CONTROLLED PRIOR TO THE GALAXY SYSTEM BEING UNDOCKED. SUBSEQUENTLY, A HEMOTHORAX WAS DISCOVERED, FORMING A CONDUIT THROUGH NECROTIC TISSUE INTO THE PLEURAL SPACE. CODE PROTOCOL WAS INITIATED, AND CONSULTATIONS WERE CONDUCTED WITH AN INTERVENTIONAL RADIOLOGIST AND THORACIC SURGEON. IT IS IMPORTANT TO NOTE THAT THE PATIENT HAD A PRE-EXISTING CHEST TUBE FROM A HEALED PNEUMOTHORAX DUE TO A NECROTIC LESION IN THE LEFT LUNG PRIOR TO THE PROCEDURE AND WAS ALSO UNDERGOING CHEMOTHERAPY. NO MALFUNCTIONS WERE REPORTED AND THE SCOPE WAS DISCARDED. EFFORTS WERE MADE TO DETERMINE THE DEFINITIVE CAUSE OF DEATH; HOWEVER, THE PHYSICIAN DID NOT PROVIDE ADDITIONAL COMMENTS AND THE FAMILY HAS CHOSEN NOT TO PURSUE AN AUTOPSY REPORT. THE PATIENT WAS ALSO DESIGNATED AS DO NOT RESUSCITATE (DNR). | SYSTEM AND BRONCHOSCOPE MANUFACTURING RECORDS WERE REVIEWED AND FOUND NO ISSUES ASSOCIATED WITH THIS CASE. REVIEW OF THE VIDEO LOGS REVEALED BLOOD OBSERVED DURING REGISTRATION AT 0:44. AT 00:56, BLEEDING AND TISSUE TRAUMA WERE NOTED IN THE LEFT LOWER LOBE AIRWAYS, INDICATING THAT THESE ISSUES WERE PRESENT IN THE PATIENT PRIOR TO THE GALAXY PROCEDURE. THE USER CONTINUED NAVIGATION, WITH BLEEDING OBSERVED IN THE SAME AREA AS THE TARGET LESION. THE USER WAS ABLE TO COMPLETE ALL BIOPSIES AND CONCLUDE THE PROCEDURE. IT WAS CONFIRMED BY A NOAH REPRESENTATIVE ON-SITE THAT THE EXCESSIVE BLEEDING WAS FIRST IDENTIFIED BY THE USER DURING STANDARD BRONCHOALVEOLAR LAVAGE (BAL) FOLLOWING THE GALAXY PROCEDURE. NO USE ERRORS WERE IDENTIFIED. THE SCOPE PASSED ALL THE MANUFACTURING PROCESS AND FINAL INSPECTIONS. THE PRODUCT WAS MANUFACTURED AS DESIGNED. FAILURE ANALYSIS OF THE SCOPE COULD NOT BE PERFORMED AS THE SCOPE WAS DISCARDED. THE PHYSICIAN DOES NOT ATTRIBUTE THE DEATH TO THE GALAXY SYSTEM, INSTEAD STATING THAT IT WAS CAUSED BY BLEEDING FROM THE BIOPSY. THE PATIENT'S PRE-EXISTING CONDITION CONTRIBUTED TO AN INCREASED RISK. THIS CASE IS BEING REPORTED BASED ON THE FOLLOWING CONCLUSIONS: A PATIENT'S DEATH OCCURRED FOLLOWING A GALAXY-ASSISTED BIOPSY PROCEDURE, AND STANDARD CODE PROTOCOLS WERE INITIATED. THE PHYSICIAN DOES NOT ATTRIBUTE THE DEATH TO THE GALAXY SYSTEM, INSTEAD STATING THAT IT WAS CAUSED BY BLEEDING FROM THE BIOPSY. ADDITIONALLY, THE PATIENT HAD A PRE-EXISTING CHEST TUBE FROM A HEALED PNEUMOTHORAX DUE TO A NECROTIC LESION IN THE LEFT LUNG PRIOR TO THE PROCEDURE AND WAS ALSO UNDERGOING CHEMOTHERAPY. NO MALFUNCTIONS WERE REPORTED.
DAVINCI XI | K131861 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION B5 ADDITIONAL INFORMATION: THE RESECTIONS AND ANASTOMOSES WERE PERFORMED ROBOTICALLY USING A THIRD-PARTY STAPLER, LIGASURE, AND A DA VINCI ROBOTIC MARYLAND BIPOLAR FORCEPS. THE PANCREATICOJEJUNOSTOMY AND HEPATICOJEJUNOSTOMY WERE SUTURED, WHILE THE GASTROJEJUNOSTOMY WAS STAPLED. THE PROCEDURE WAS THEN CONVERTED FROM ROBOTIC TO LAPAROSCOPY PRIOR TO PERFORMING LYMPHADENECTOMY AND RESECTING THE SECOND MESENTERIC TUMOR. THIS DECISION WAS MADE DUE TO THE TUMOR'S LOCATION, WHICH WOULD HAVE NECESSITATED REPOSITIONING ALL ROBOTIC PORT SITES. DURING THE LYMPHADENECTOMY, APPROXIMATELY 700 ML OF RETROPERITONEAL BLEEDING OCCURRED, RESULTING IN THE PATIENT EXPERIENCING TACHYCARDIA AND HYPOTENSION. HEMOSTASIS WAS ACHIEVED, AND THE LYMPHADENECTOMY WAS SUCCESSFULLY COMPLETED. THE MESENTERIC TUMOR, INITIALLY IDENTIFIED AS A MESENTERIC LYMPH NODE, WAS RESECTED. PATHOLOGY CONFIRMED IT TO BE A NEUROENDOCRINE TUMOR CONCURRENT WITH DUODENAL CARCINOMA. THE RESECTION RESULTED IN APPROXIMATELY 1 LITER OF BLOOD LOSS DUE TO THE TUMOR'S LOCATION AND ADHERENCE. HEMOSTASIS WAS ACHIEVED USING A LIGASURE DEVICE AND SUTURING. THIRTY-SIX HOURS POST-OPERATIVELY, THE PATIENT DEVELOPED MESENTERIC ISCHEMIA, SECONDARY TO THE RESECTION OF THE NEUROENDOCRINE MESENTERIC TUMOR. A LAPAROSCOPIC PROCEDURE WAS PERFORMED FOR AN ENTERECTOMY AND AN ENTEROANASTOMOSIS USING A THIRD-PARTY STAPLER INSTRUMENT. THE PATIENT LATER DEVELOPED A PANCREATIC FISTULA WITH A SUBHEPATIC COLLECTION, WHICH REQUIRED TWO ADDITIONAL LAPAROSCOPIC PROCEDURES FOR REVISION AND DRAINAGE. ON POST-OPERATIVE DAY 19, THE PATIENT DEVELOPED SEPTIC SHOCK AND WAS EXPERIENCING A LEAKAGE OF SECRETIONS THROUGH ONE OF THE ROBOTIC PORT SITES. AN EXPLORATORY LAPAROTOMY FOUND A FISTULA HAD FORMED AT THE ENTERO-ETEROANASTOMOSIS. LATER THAT DAY, THE PATIENT¿S CONDITION FURTHER DETERIORATED, RESULTING IN BLOOD DYSCRASIA AND INSTABILITY. A SECOND EXPLORATORY LAPAROTOMY WAS PERFORMED REVEALING A HEMOPERITONEUM, THOUGH THE SOURCE OF THE BLEEDING COULD NOT BE IDENTIFIED. DESPITE THESE EFFORTS, THE SEQUENCE OF EVENTS LEAD TO THE PATIENT DYING THE NEXT MORNING, ON POST-OPERATIVE DAY 20. A REVIEW OF THE ADDITIONAL EVENT INFORMATION PERFORMED BY AN INTUITIVE MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT HAD A COMPLEX AND CHALLENGING TUMOR IN WHICH THEY UNDERWENT AN INITIAL RESECTION WHICH INCORPORATED ROBOTIC INSTRUMENTATION AS WELL AS SEVERAL 3RD PARTY DEVICES. A MESENTERIC EXTENSION OF THE TUMOR IN THE FORM OF A LARGE LYMPH NODE NECESSITATED A BOWEL RESECTION WHICH WAS PERFORMED USING A 3RD PARTY DEVICE AND THAT BOWEL WENT ON TO BECOME ISCHEMIC REQUIRING ADDITIONAL EXPLORATION AND RESECTION. A LEAK ALSO DEVELOPED BETWEEN AN ANASTOMOSIS FOLLOWING THE SECOND OPERATION AND FURTHER COMPLICATIONS INCLUDING A FISTULA OCCURRED. THE ULTIMATE CAUSE OF DEATH WAS SEPTIC AND HYPOVOLEMIC SHOCK. HOW THESE COMPLICATIONS MAY OR MAY NOT HAVE BEEN RELATED TO THE INITIAL SURGERY AND ANY ROBOTIC DEVICE IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED CHOLECYSTECTOMY PROCEDURE, THE PATIENT EXPIRED. THE SURGEON STATED THAT THE ROBOTIC PROCEDURE WAS COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS; THERE WERE NO DA VINCI-RELATED ERRORS OR ISSUES. POSTOPERATIVELY, WHILE THE PATIENT WAS RECOVERING IN THE POST ANESTHESIA CARE UNIT (PACU), NARCOTICS WERE ADMINISTERED FOR PAIN. DUE TO A MISADMINISTRATION OF NARCOTICS IN THE PACU, THE PATIENT EXPIRED. THE PARTNER OF THE SURGEON CLARIFIED THAT THE PATIENT WAS RECOVERING IN THE OUTPATIENT SURGICAL CENTER¿S PACU AND WAS ADMINISTERED FENTANYL. THE SURGICAL CENTER'S PACU DID NOT HAVE NARCAN AVAILABLE AT THE TIME OF THE CODE. PARAMEDICS WERE CALLED AND ADMINISTERED NARCAN UPON ARRIVAL. ADVANCED RESUSCITATION ACTIVITIES AND CHEST COMPRESSIONS WERE CONTINUED FOR TWO HOURS. AN INITIAL POST-PROCEDURE ULTRASOUND SHOWED NO SIGNS OF BLOOD WITHIN THE ABDOMEN. AT AN UNSPECIFIED TIME DURING RESUSCITATION, A REPEAT ULTRASOUND DID SHOW SIGNS OF BLOOD IN THE ABDOMEN. IT WAS PLANNED TO TRANSFER THE PATIENT TO THE HOSPITAL'S MAIN OPERATING ROOM FOR A REOPERATION DUE TO THE FINDINGS; HOWEVER, THE PATIENT EXPIRED IN THE PACU AND THE SUBSEQUENT PROCEDURE DID NOT OCCUR. IT WAS STATED THAT THE NEW BLEEDING WITHIN THE ABDOMEN WAS BELIEVED TO BE DUE TO THE CHEST COMPRESSIONS, POSSIBLY CRACKED RIBS. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES FOUND THERE WERE NO INTERRUPTIONS OR ERROR EVENTS IMPACTING ANY OF THE PROCEDURES. THE SYSTEM WAS FUNCTIONING TO MANUFACTURING SPECIFICATIONS. THE DAY AFTER THE EVENT, THE CUSTOMER REQUESTED A FIELD SERVICE ENGINEER (FSE) TO CHECK OUT THE SYSTEM EVEN THOUGH ¿THEY DID NOT BELIEVE THE ROBOT HAD ANYTHING TO DO WIGHT THE ADVERSE EVENT¿ TO VERIFY THE EQUIPMENT. THE FSE TESTED THE SYSTEM, INCLUDING THE E-100 AND ERBE GENERATORS WITH NO ISSUES FOUND. ELECTRICAL SAFETY TESTING OF ALL THE CARTS AND GENERATORS WAS ALSO PERFORMED, FOLLOWED BY A SECOND SYSTEM TEST DRIVE. NO ISSUES WERE FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, PERMANENT CAUTERY HOOK, MARYLAND BIPOLAR FORCEPS, CADIERE FORCEPS, PROGRASP FORCEPS, MEDIUM-LARGE CLIP APPLIERS X2. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING OR SUBSEQUENT TO THIS PROCEDURE. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICES USED DURING THE REPORTED EVENT FOUND NO RELATED NON-CONFORMANCES. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT A CHOLECYSTECTOMY, WHERE NO INTRAOPERATIVE COMPLICATIONS OR ISSUES WERE REPORTED. POSTOPERATIVELY, THE PATIENT RECEIVED FENTANYL AND CODED. THE TEMPORAL RELATIONSHIP OF THE CARDIOVASCULAR COLLAPSE SUGGESTS THE EVENT WAS PRECIPITATED BY THE NARCOTIC AS AN INITIAL ULTRASOUND SHOWED NO EVIDENCE OF FLUID IN THE ABDOMEN. ALTHOUGH A SUBSEQUENT ULTRASOUND EVENTUALLY SHOWED FLUID IN THE ABDOMEN, THIS FINDING WAS SPECULATED TO BE A RESULT OF THE RESUSCITATION EFFORTS. CONTINUED RESUSCITATION EFFORTS WERE UNSUCCESSFUL, AND THE PATIENT EXPIRED. THE CAUSE OF DEATH WAS SPECULATED TO HAVE BEEN PRECIPITATED BY THE ADMINISTRATION OF THE FENTANYL NARCOTIC. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
DAVINCI X | K171294 | IT WAS REPORTED THAT THE PATIENT UNDERWENT DA VINCI-ASSISTED RECTOSIGMOIDECTOMY AND PLANNED CHOLECYSTECTOMY PROCEDURE. THERE WERE NO INTRAOPERATIVE COMPLICATIONS DURING THE PROCEDURE WHICH WAS SUCCESSFULLY COMPLETED. THE PATIENT SUBSEQUENTLY EXPIRED ON POST-OPERATIVE DAY (POD) 6 FROM COMPLICATIONS. IT WAS REPORTED THAT BY POD 4, THE PATIENT EXHIBITED SIGNS OF SEPSIS. A CT SCAN SHOWED FREE FLUID IN THE PELVIS, AND ALTHOUGH THERE WAS NO FISTULA DETECTED, THE FINDINGS WERE SUSPECTED TO BE CAUSED BY FOURNIER'S GANGRENE. THE PATIENT UNDERWENT AN EXPLORATORY LAPAROTOMY "BY VIDEO" FOR PELVIC DRAINAGE. TWO DAYS LATER, AS A RESULT OF THE FOURNIER'S GANGRENE, THE PATIENT WAS RETURNED TO THE SURGICAL CENTER ON MECHANICAL VENTILATION AND WITH VASOACTIVE INFUSIONS (NOREPINEPHRINE AND VASOPRESSIN) FOR EMERGENCY DEBRIDEMENT OF THE SCROTAL SAC. THE PATIENT EXPIRED DURING THE PROCEDURE. AN AUTOPSY WAS NOT PERFORMED. THE CAUSE OF DEATH RECORDED ON THE DEATH CERTIFICATE WAS FOURNIER'S GANGRENE. ADDITIONAL DETAILS REGARDING THE SEQUENCE OF EVENTS AND TREATMENT WERE REQUESTED; NO RESPONSE HAS BEEN RECEIVED. | SEVERAL PROCEDURES WERE PERFORMED ON THE PROVIDED INITIAL ROBOTIC PROCEDURE DATE. INSUFFICIENT INFORMATION IS AVAILABLE FOR PROCEDURE VERIFICATION; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. CONFIRMATION OF THE SPECIFIC PROCEDURE TIME HAS BEEN REQUESTED. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT UNDERWENT 2 CONCURRENT ROBOTIC PROCEDURES, A CHOLECYSTECTOMY AND A RECTOSIGMOIDECTOMY. THE PATIENT WAS ON INOTROPIC SUPPORT AND MECHANICAL VENTILATION FOLLOWING THE PROCEDURE, BUT THE SPECIFIC DETAILS OF WHY THIS LEVEL OF SUPPORT WAS REQUIRED IS NOT PROVIDED. THE PATIENT THEN DEVELOPED SEPSIS AND FOURNIER¿S GANGRENE WHICH IS A CONDITION WITH HIGH MORTALITY. HOW THEY DEVELOPED THIS CONDITION IS NOT PROVIDED. THE PATIENT EVENTUALLY DIED AS A RESULT OF THESE COMPLICATIONS. HOW THESE EVENTS WERE RELATED TO THE INITIAL PROCEDURE IS NOT PROVIDED. BASED ON THE INFORMATION IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
OMNIPOD 5 AUTOMATED INSULIN DELIVERY SYSTEM | K203768 | IT HAS BEEN REPORTED BY THE PATIENT'S BROTHER, THAT THE PATIENT PASSED AWAY ON (B)(6) 2025. THE PATIENT'S BROTHER STATES THE PATIENT WENT INTO DIABETIC KETOACIDOSIS (DKA) AND PASSED AWAY. THE AUTOPSY REPORT STATES DKA AS THE CAUSE OF DEATH. | PER THE REPORTER, THE PATIENT EXPIRED (B)(6) 2025. AFTER A REVIEW OF CLOUD DATA, THE LAST POD ACTIVATED WAS ON (B)(6) 2025 AND WAS DEACTIVATED (B)(6) 2025. BASED ON THE AVAILABLE INFORMATION, THE DEVICE WAS NOT IN USE THE DAYS PRECEDING OR AT THE TIME OF DEATH. IF ADDITIONAL INFORMATION BECOMES AVAILABLE, THE CASE WILL BE REOPENED AND REASSESSED. CLOUD DATA WAS REVIEWED FOR THIS USER FOR THE PERIOD OF (B)(6) 2025. DATA WAS AVAILABLE UP TO (B)(6) 2025. THE LAST POD USED WAS ACTIVATED AT 22:22 ON (B)(6) 2025. THE POD WAS PAIRED WITH A DEXCOM G6 SENSOR AND THE SYSTEM WAS USED IN AUTOMATED MODE. IN THE DAYS PRIOR, THE SYSTEM WAS USED IN MANUAL MODE WITH NO SENSOR PAIRED. THE AVAILABLE DATA SHOWED THAT THE AUTOMATED ALGORITHM WAS ABLE TO APPROPRIATELY ADAPT THE MICROBOLUS AMOUNTS BASED ON THE ESTIMATED GLUCOSE VALUES AND USER INPUTS. THE POD STARTED RECEIVING ABERRANT SENSOR VALUES FROM THE SENSOR ON (B)(6) 2025. THE SYSTEM RESPONDED APPROPRIATELY TO THESE ABERRANT VALUES, PUTTING THE SYSTEM INTO AUTOMATED MODE: LIMITED AFTER FOUR CONTINUOUS CYCLES AND GENERATING MISSING SENSOR VALUES MESSAGES AFTER ONE HOUR OF CONTINUOUSLY ABERRANT VALUES. THE POD WAS DEACTIVATED AT 22:12 ON (B)(6) 2025 AND DATA ENTRIES FOR THE REMAINING TWO DAYS CONSIST OF NO ACTIVE POD MESSAGES AND TOTAL DAILY INSULIN SUMMARIES. NO EVIDENCE OF ISSUES WERE OBSERVED WITH THE POD OR CONTROLLER ON THE LAST DAYS OF USE BASED ON THE AVAILABLE CLOUD DATA.
PHILIPS CPAP DEVICE | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING HUSBAND PASSED AWAY FROM CANCER. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
WECK VISTA | K152663 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION H11 ADDITIONAL INFORMATION: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
ALPHAVAC | K240397 | AN ANGIODYNAMICS TERRITORY MANAGER (TM) REPORTED THAT A PATIENT, WITH SEVERAL COMORBIDITIES, EXPIRED DURING A PULMONARY EMBOLISM (PE) REMOVAL PROCEDURE, AFTER USE OF AN ALPHAVAC MULTIPURPOSE MECHANICAL ASPIRATOR F18 C85. THE PROCEDURE WAS PERFORMD IN ACCORDANCE WITH THE DEVICE'S IFU AND ALL APPEARED NORMAL ACCORDING TO THE PHYSICIANS; HOWEVER, THE ALPHAVAC WAS REMOVED TO INSPECT IF THE CANNULA OR SHEATH WERE OBSTRUCTED; NEITHER WERE, SO THE GORE DRYSEAL WAS REMOVED. AFTER REMOVAL OF THE DRYSEAL, THE PATIENT'S PRESSURES BEGAN TO DROP AND THEY DECOMPENSATED [CARDIAC ISSUE]. AT THIS TIME, AN ACLS PROTOCOL WAS ACTIVATED. THE PHYSICIANS CHECKED THE PATIENT FOR A PLEURAL EFFUSION, PERICARDIAL EFFUSION AND HEMATOMA; HOWEVER, NONE WERE FOUND. AFTER SEVERAL RESUSCITATION ATTEMPTS, THE PATIENT ULTIMATELY EXPIRED. THIS PROCEDURE WAS THE FIRST ALPHAVAC PROCEDURE FOR BOTH ATTENDING PHYSICIANS; HOWEVER, ONE OF THE ATTENDINGS IS A REGULAR ANGIOVAC USER. THE PHYSICIANS DO NOT BELIEVE THAT ANY ANGIODYNAMICS DEVICE CONTRIBUTED, IN ANY WAY, TO THE PATIENT'S EXPIRATION. | THE CUSTOMER'S REPORTED COMPLAINT DESCRIPTION OF THE PATIENT DECOMPENSATED [CARDIAC ISSUES] AND EXPIRED AFTER REMOVAL OF THE ALPHAVAC CANNULA/SHEATH AND GORE DRYSEAL SHEATH CANNOT BE CONFIRMED GIVEN THE PATIENT CENTRIC NATURE OF THIS SERIOUS ADVERSE EVENT. NO ALPHAVAC PRODUCT WAS RETURNED TO ANGIODYNAMICS FOR EVALUATION SINCE THERE WAS NO REPORT OF DEVICE MALFUNCTION OR PERFORMANCE ISSUE DURING USE IN THE PROCEDURE. THE DEVICE HISTORY RECORD (DHR) REVIEW OF THE PACKAGING/CANNULA/SHEATH LOTS REVEALED NO QUALITY RELATED ISSUES OR MANUFACTURING DEFICIENCIES AT THE TIME OF MANUFACTURE. THE PHYSICIANS STATED THEY DO NOT BELIEVE THAT ANY ANGIODYNAMICS' DEVICE CONTRIBUTED, IN ANY WAY, TO THE PATIENT'S EXPIRATION. THERE WAS NO REPORT OF ALPHAVAC DEVICE MALFUNCTION OR USE ERROR DURING THE PROCEDURE AND PATIENT EXPIRATION IS POTENTIALLY DUE TO THEIR CO-MORBIDITIES. PULMONARY EMBOLISM, CARDIAC ARREST AND DEATH ARE LISTED IN THE DEVICE DIRECTIONS FOR USE (DFU) AS POSSIBLE COMPLICATIONS/ADVERSE EVENTS DUE TO USE OF EMBOLECTOMY SYSTEM. LABELING REVIEW: DIRECTIONS FOR USE (10903689) IS PROVIDED WITH THIS DEVICE AND CONTAINS THE FOLLOWING STATEMENTS: INDICATIONS FOR USE THE CANNULA IS INDICATED FOR: THE NON-SURGICAL REMOVAL OF THROMBI OR EMBOLI FROM THE VENOUS VASCULATURE. ASPIRATION OF CONTRAST MEDIA AND OTHER FLUIDS FROM THE VENOUS VASCULATURE. THE CANNULA IS INTENDED FOR USE IN THE VENOUS SYSTEM AND FOR THE TREATMENT OF PULMONARY EMBOLISM. INTENDED USE: THE ALPHAVAC SYSTEM IS INTENDED TO BE USED WITH COMMONLY AVAILABLE VASCULAR ACCESS TOOLS (E.G., GUIDEWIRE, VASCULAR INTRODUCER, ETC.) TO FACILITATE THE REMOVAL OF THROMBOEMBOLI, BUT NOT LIMITED TO, THROMBUS, EMBOLUS, OR CLOT DURING MINIMALLY INVASIVE PERCUTANEOUS PROCEDURES. WARNINGS: SELECTION OF THE PATIENT AS A CANDIDATE FOR USE WITH THIS DEVICE AND FOR SUCH PROCEDURES AS IT IS INTENDED IS THE PHYSICIANS' SOLE RESPONSIBILITY. THE OUTCOME IS DEPENDENT ON MANY VARIABLES INCLUDING, PATIENT PATHOLOGY, SURGICAL PROCEDURE, AND PROCEDURE/TECHNIQUE. THE BENEFITS OF USE OF THIS DEVICE MUST BE WEIGHED AGAINST THE RISKS INCLUDING RISKS OF SYSTEMIC ANTICOAGULATION AND MUST BE ASSESSED BY THE PRESCRIBING PHYSICIAN. ADVERSE EVENTS: THIS DEVICE, AS DO ALL EMBOLECTOMY SYSTEMS AND DEVICES, HAS POSSIBLE SIDE EFFECTS, WHICH INCLUDE, BUT ARE NOT LIMITED TO INFECTIONS, BLOOD LOSS, THROMBUS FORMATION, EMBOLIC EVENTS, VESSEL, VENTRICULAR OR VALVULAR DAMAGE AND COMPLICATIONS OF PERCUTANEOUS OR SURGICAL INSERTION TECHNIQUES. THESE MAY OCCUR IF THE DIRECTIONS FOR USE ARE NOT FOLLOWED. POSSIBLE COMPLICATIONS INCLUDE THOSE NORMALLY ASSOCIATED WITH LARGE BORE SURGICAL AND/OR PERCUTANEOUS VESSEL CATHETERIZATION/CANNULATION, ANTICOAGULATION, AND APPLICATION OF INTRAVASCULAR INTRODUCER SYSTEMS WHICH INCLUDE BUT ARE NOT LIMITED TO: ARRHYTHMIAS, CARDIAC ARREST, DEATH, DISTAL EMBOLIZATION OF THROMBUS, PULMONARY EMBOLISM, TACHYCARDIA. A REVIEW OF THE ANGIODYNAMICS COMPLAINT SYSTEM NOTED NO ADVERSE TRENDS FOR THIS COMPLAINT TYPE AND PRODUCT FAMILY. THIS TYPE OF COMPLAINT CONTINUES TO BE MONITORED FOR TRENDS. REFERENCE (B)(4).
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) WAS HOSPITALIZED AND EXPIRED ON (B)(6) 2025 APPROXIMATELY 26 MINUTES AFTER CCPD THERAPY BEGAN. FOLLOW-UP WITH THE PATIENT¿S INPATIENT PROGRAM MANAGER (IPM) REVEALED THE PATIENT PRESENTED TO THE EMERGENCY ROOM ON (B)(6) 2025 WITH COMPLAINTS OF DYSPNEA, DIZZINESS, AND A RASH. ALTHOUGH THE TIMELINE/SPECIFICS OF THE SERIOUS ADVERSE EVENTS ARE LARGELY UNKNOWN (PATIENT¿S ELECTRONIC MEDICAL RECORD CLOSED), THE IPM WAS ABLE TO RECALL THE PATIENT WAS DIAGNOSED WITH AN ALMOST COMPLETE ATRIOVENTRICULAR (AV) BLOCK (WHICH WOULD HAVE REQUIRED A PACEMAKER), AND A SEVERE CASE OF ADULT-ONSET HAND/FOOT/MOUTH DISEASE. THE IPM STATED THE PATIENT HAD A DO NOT RESUSCITATE (DNR) ORDER IN PLACE AND PASSED AWAY ON (B)(6) 2025. THE PATIENT WAS UNDERGOING CCPD AT THE TIME OF HER PASSING (26 MINUTES INTO TREATMENT), HOWEVER THE IPM WAS CONFIDENT THE CARDIAC ARREST OCCURRED DUE TO THE PATIENT¿S CARDIAC COMORBID CONDITIONS WHILE THE BODY WAS FIGHTING THE HAND/FOOT/MOUTH INFECTION. PER THE IPM, THE EVENTS WERE UNRELATED TO ANY FRESENIUS PRODUCT(S) AND/OR DEVICE(S) DEFICIENCY OR MALFUNCTION. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINIC REVIEW: BASED ON THE AVAILABLE INFORMATION, THE PATIENT¿S LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM HAVING A CAUSAL OR CONTRIBUTORY ROLE IN THE ADVERSE EVENTS EXPERIENCED BY THE PATIENT. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A SERIOUS INJURY, PATIENT DEATH, OR OTHER SERIOUS ADVERSE EVENT(S) RELATED TO A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) OCCURRED. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS.
STEERABLE GUIDE CATHETER | K172394 | IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT FUNCTINOAL MITRAL REGURGITATION (MR) WITH A GRADE OF 4+. ONE CLIP WAS PREPARED AND ADVANCED INTO THE ANATOMY. THE CLIP WAS DEPLOYED, REDUCING MR TO <1. IT WAS LATER FOUND THAT THE PATIENT HAD A SEPTAL INJURY AND PERICARDIAL EFFUSION. ULTIMATELY, THE PATIENT PASSED AWAY THE SAME DAY. | INVESTIGATION IS NOT YET COMPLETE. A FOLLOW-UP REPORT WILL BE SUBMITTED WITH ALL ADDITIONAL RELEVANT INFORMATION.
MICRA | K132030 | IT WAS REPORTED THAT THE PATIENT EXPERIENCED CARDIAC PERFORATION IN THE RIGHT VENTRICLE DURING A PROCEDURE TO IMPLANT A LEADLESS IMPLANTABLE PULSE GENERATOR (IPG). ECHOCARDIOGRAM RIGHT VENTRICLE GRAM, PERIOCARDIOCENTESIS AND RESUSCITATIVE MEASURES WERE PERFORMED. IT WAS REPORTED THAT THE PATIENT DIED. CAUSE OF DEATH WAS PROVIDED AS PERFORATION FROM THE LEADLESS IPG. | CONTINUATION OF D10: MC2AVR1 LEADLESS IPG; IMPLANT DATE (B)(6) 2025. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION REGARDING A DREAMSTATION AUTO CPAP. THE PATIENT HAS PASSED AWAY. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER. THE TECHNICIAN COULD NOT CONFIRM FOAM PARTICLES IN THE DEVICE EVALUATION. THE MANUFACTURER IS SUBMITTING A FINAL REPORT AT THIS TIME. IF ANY ADDITIONAL INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
OPTICROSS HD | K173284 | IT WAS REPORTED THAT PATIENT DEATH OCCURRED. THE 90% STENOSED TARGET LESION WAS LOCATED IN THE MODERATELY TORTUOUS AND SEVERELY CALCIFIED LEFT CIRCUMFLEX (LCX) ARTERY. THE OPTICROSS HD IMAGING CATHETER WAS SELECTED FOR ULTRASOUND EXAMINATION OF THE TARGET LESION. AFTER COMPLETING THE ULTRASOUND IMAGING, THE PHYSICIAN ENCOUNTERED RESISTANCE WHILE ATTEMPTING TO WITHDRAW THE CATHETER. UPON APPLYING FORCE, THE CATHETER BROKE, LEAVING APPROXIMATELY 10 CM OF THE DISTAL PORTION INSIDE THE PATIENT. THIS COMPLICATION LED TO RAPID DETERIORATION OF THE PATIENT CONDITION, PREVENTING THE CONTINUATION OF THE PROCEDURE. DUE TO THE COMBINED CIRCUMSTANCES, THE PATIENT DID NOT RECOVER AND PASSED AWAY IN THE INTENSIVE CARE UNIT (ICU). | THE DEVICE WAS RETURNED FOR ANALYSIS. VISUAL AND MICROSCOPIC INSPECTION REVEALED THE TELESCOPE, SHEATH, AND IMAGING WINDOW WERE KINKED AND THE DISTAL SECTION OF THE IMAGING WINDOW WAS STRETCHED AND DETACHED. THE DETACHED PORTION, INCLUDING THE TIP, WAS NOT RETURNED FOR ANALYSIS.
DREAMSTATION BIPAP S/T 30 W/HEATED HUMIDIFIER, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT HAS SADLY PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
MICRA | K132030 | DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF LEADLESS IMPLANTABLE PULSE GENERATORS (IPGS). THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED ACUTE, THIRTY-DAY COMPLICATIONS WHICH INCLUDED VASCULAR ANEURYSMS. THERE WERE PATIENTS WHO EXPERIENCED DEEP VEIN THROMBOSIS, PULMONARY EMBOLISMS, CARDIAC TAMPONADE, POST IMPLANT HEMATOMAS AND HEMORRHAGES, PERICARDITIS, AND UNKNOWN VASCULAR COMPLICATIONS. THERE WERE DISLODGMENTS OF LEADLESS IPGS AND OTHER UNKNOWN MALFUNCTIONS IN WHICH THERE WERE REVISIONS, REPLACEMENTS, AND OTHER INTERVENTIONS. THE STATUS OF THE DEVICES, DELIVERY SYSTEMS, AND INTRODUCERS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
NA | K092797 | ACCORDING TO THE REPORTER, ON THE DAY OF THE EVENT, THE PATIENT WAS IN THE ER (EMERGENCY ROOM) FOR EMESIS, ELEVATED WBC (WHITE BLOOD CELLS), AND HGB (HEMOGLOBIN) 6.4. BASED ON LAB WORK, DIALYSIS WAS NEEDED, AND THE PATIENT WAS TAKEN TO THE DIALYSIS UNIT. THE BLOOD FLOW FOR DIALYSIS WAS 400, WHICH WAS CONSISTENT WITH NORMAL TREATMENT PRESCRIPTION. ACCORDING TO THEIR RECORDS, THE BLOODLINE CONNECTION WAS SECURE AND HAD A CONNECTOR ON IT THAT SECURED THE BLOODLINES TO THE CATHETER (WAS REFERRED TO AS A SECURE LOCK). SHORTLY AFTER THE START OF LONG-TERM DIALYSIS, THE DIALYSIS MACHINE ALARMED. THE NURSE RESPONDED TO THE MACHINE ALARM, AND IT WAS FOUND THAT THE VENOUS PORT/HUB HAD DISCONNECTED, WHICH LED TO SIGNIFICANT BLOOD LOSS. THE RAPID RESPONSE TEAM WAS ENGAGED. THE MACHINE WAS MAINTAINED BY THE 3RD PARTIES BIOMED. FOR THE HOSPITAL, IT STATED THE MACHINE WAS NOT A CONCERN; IT DID WHAT IT WAS SUPPOSED TO AND ALARMED ACCORDING TO THE 3RD PARTY. THE MACHINES WERE MAINTAINED BY A 3RD PARTY, AS THEY WERE THE 3RD PARTY¿S EQUIPMENT. THE HOSPITAL DID NOT HAVE ACCESS TO VERIFY THE MACHINE ALARMED. IT WAS CONFIRMED THAT THE CATHETER HAD PREVIOUSLY UNDERGONE A REPAIR ON (B)(6) 2024. THE LOCUM PROVIDER¿S NOTE DID NOT PROVIDE SPECIFIC INFORMATION ABOUT THE REPAIR PROCEDURE, BUT THERE WAS NO NEED FOR BOTH L IMBS TO BE REPAIRED SINCE IT WAS THE CLAMP PART FROM THE CATHETER THAT WAS REPAIRED. THE DISCONNECTION OF THE LUER ADAPTER OCCURRED ON THE VENOUS LIMB OF THE CATHETER, ALIGNING WITH THE SIDE WHERE THE CLAMP WAS REPAIRED IN NOVEMBER. THE HOSPITAL NOTED THAT THE ONLY AVAILABLE REPAIR KIT AT THE TIME OF THE NOVEMBER REPAIR WAS THE ONE NOW IDENTIFIED AS THE KIT FROM WHICH THE LUER ADAPTER THAT DISCONNECTED IS BELIEVED TO HAVE ORIGINATED. THERE WAS NO DOCUMENTATION TO INDICATE THE CATHETER HAD BEEN REPAIRED OR ALTERED AGAIN AFTER THAT DATE. THE PATIENT WAS NON-COMPLIANT, AND THERE WAS NO INFORMATION ON HOW THE CATHETER CARED FOR/MAINTAINED OUTSIDE OF DIALYSIS SESSIONS. THE PATIENT HAD A PRE-EXISTING CONDITION OF CHRONIC PAIN, AND STAFF REPORTED THAT THE PATIENT CONSTANTLY MOVED AROUND IN THE CHAIR/BED DURING DIALYSIS TREATMENT, WHICH MIGHT HAVE CONTRIBUTED TO SOME TENSION ON THE CATHETER. NO CRACKS OR OBVIOUS DAMAGE TO THE CATHETER WERE OBSERVED UPON INSPECTION BY STAFF. THEY INVESTIGATED IF THE CATHETER SHOULD HAVE BEEN REMOVED AND REPLACED EARLIER BUT COULD NOT FIND ANYTHING THAT INDICATED A USAGE/TIMING LIMIT. THE PATIENT BECAME PULSELESS, AND CPR (CARDIOPULMONARY RESUSCITATION) STARTED. THEY NOTED THE RHYTHM WAS VFIB (VENTRICULAR FIBRILLATION) AND THEN ASYSTOLE. THE PATIENT UNDERWENT A MASSIVE BLOOD TRANSFUSION PROTOCOL. ROSC (RETURN OF SPONTANEOUS CIRCULATION) WAS UNABLE TO BE ACHIEVED. THE PATIENT PASSED AWAY ON THE SAME DAY AS THE EVENT. | ADDITIONAL INFORMATION: G3, H3, H6 H3. EVALUATION SUMMARY: MEDTRONIC CONDUCTED AN INVESTIGATION BASED UPON ALL INFORMATION RECEIVED. THE DEVICE WAS NOT RETURNED, BUT A PHOTO WAS AVAILABLE FOR EVALUATION. VISUAL INSPECTION NOTED A DETACHED VENOUS PORT AND CLIP, WITH VISIBLE SIGNS OF USE SUCH AS BLOOD. THERE APPEARED TO BE NO OTHER ABNORMALITIES WITH THE DEVICE. IT WAS REPORTED THAT LUER ADAPTER DETACHED FROM THE EXTENSION TUBE. THE REPORTED ISSUE WAS CONFIRMED. THE MOST LIKELY CAUSE COULD NOT BE ESTABLISHED FROM THE INFORMATION AVAILABLE. THE MANUFACTURING RECORDS FOR EACH DEVICE ARE THOROUGHLY REVIEWED PRIOR TO RELEASE TO ENSURE THAT IT MEETS ALL MEDTRONIC QUALITY SPECIFICATIONS. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
MICRA | K132030 | DATA WAS REVIEWED REGARDING IMPLANTS AND OUTCOMES OF LEADLESS IMPLANTABLE PULSE GENERATORS (IPGS). THEY DESCRIBED PATIENT DEATHS; HOWEVER, THE CAUSES OF DEATH WERE UNKNOWN AND DEFINED AS ALL-CAUSE, THIRTY-DAY MORTALITY. THERE WERE PATIENTS WHO EXPERIENCED ACUTE, THIRTY-DAY COMPLICATIONS WHICH INCLUDED PUNCTURE SITE EVENTS SUCH AS ARTERIOVENOUS FISTULA AND VASCULAR ANEURYSMS. THERE WERE PATIENTS WHO EXPERIENCED DEEP VEIN THROMBOSIS, PULMONARY EMBOLISMS, CARDIAC PERFORATION, CARDIAC TAMPONADE, STENOSIS DUE TO DEVICE OR IMPLANT, POST IMPLANT HEMATOMAS OR HEMORRHAGES, INTRAOPERATIVE CARDIAC ARREST, PERICARDITIS, AND UNKNOWN VASCULAR COMPLICATIONS. THERE WERE DISLODGMENTS OF LEADLESS IPGS AND OTHER UNKNOWN MALFUNCTIONS IN WHICH THERE WERE REVISIONS, REPLACEMENTS, AND OTHER INTERVENTIONS. THE STATUS OF THE DEVICES, DELIVERY SYSTEMS, AND INTRODUCERS IS UNKNOWN. NO ADDITIONAL ADVERSE PATIENT EFFECTS OR PRODUCT PERFORMANCE ISSUES WERE REPORTED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
HEMOSIL HIT- AB(PF4-H) | K153137 | A COMPLAINT HAS BEEN FILED FOR A DISCREPANT HIT TEST RESULT. THE PATIENT SAMPLE ID (B)(6), TESTED ON THE ACL TOP USING REAGENT/CALIBRATOR PN 00020014600, YIELDED A NEGATIVE RESULT. THE SAME SAMPLE, WHEN SENT TO MAYO CLINIC, RETURNED A POSITIVE SRA RESULT (CUT-OFF: >0.400 POSITIVE). THE SAMPLE WAS INITIALLY SENT TO MAYO AFTER PHARMACY QUESTIONED THE NEGATIVE RESULT, GIVEN THE PATIENT'S CLINICAL HISTORY OF A LEFT APICAL THROMBUS, STROKE, AND SUBSEQUENT DEATH UNDER PALLIATIVE CARE. THE SITE HAS NOW DISCONTINUED HIT TESTING ON THE ACL TOP AND IS SENDING SAMPLES TO QUEST. | QC AND MANUFACTURING DATA FOR HEMOSIL HIT-AB(PF4-H) LOT B37250 WERE REVIEWED AND NO ISSUES OR TRENDS WERE IDENTIFIED WITH THE IMPACTED LOT. REVIEW OF SUBMITTED CALIBRATION REPORT CONFIRMS THAT HIT-AB LOT B37250 WAS CALIBRATED ON 4/17/2025 APRIL 17, 2025, USING THE CORRECT TARGET VALUE OF 5.5 U/ML, AND THE RESULTS WERE PROPRELY VALIDATED. QC RESULTS WERE PERFORMED AFTER THE CALIBRATION WAS VALIDATED AND WERE CONFIRMED WITHIN ACCEPTABLE RANGE FOR HIT_AB CONTROLS LOT B37286. REVIEW OF THE SUBMITTED QC LEVEY JENNINGS CHARTS FOR 05/01/2025 TO 05/19/2025 SHOWS THE HIT ANTIBODIES FOR ACL TOP 350, HIT LOW AND HIGH LOT # B37286 REPORTED ACCEPTABLE RESULTS ON 05/08/2025 PRIOR TO SAMPLE TESTING. REVIEW OF THE SERVICE HISTORY SHOWS CALL 2005283457 ON 5/27/2025 - THE FIELD SERVICE ENGINEER VERIFIED THE INSTRUMENT FOR THE HIT DISCREPANT RESULTS. PREVENTATIVE MAINTENANCE FOR THE INSTRUMENT WAS LAST COMPLETED ON 4/25/25. NO INSTRUMENT ISSUES FOUND AT THAT TIME. REVIEWED GENERAL LOG LIST, CLEANED AND INSPECTED INSTRUMENT; NO ISSUES FOUND. FLUIDIC PRECISION PASSES BOTH ARMS. FLOW RATE PERFORMED AND PASSED. NO LEAKS FOUND. OPTICAL READING UNIT (ORU) REFERENCE READINGS IN RANGE AND STABLE. LIQUID LEVEL DETECTION (LLD) COUNTS FOR BOTH PROBES WERE NORMAL AND STABLE. COORDINATE DELTAS WERE LOW AND STABLE. NO INSTRUMENT RELATED ISSUES WERE IDENTIFIED. SINCE THE AFFECTED SAMPLE IS NOT AVAILABLE TO PERFORM TESTING, NO FURTHER INVESTIGATIONAL TESTING CAN BE DONE TO DETERMINE THE ROOT CAUSE OF THE REPORTED ISSUE. THE INSTRUCTIONS FOR USE (IFU) OF THE PRODUCT REPORTS A TOTAL PERCENTAGE AGREEMENT OF 82% IN METHOD COMPARISON AND 90% IN EXTERNAL VALIDATION, VS ELISA ASSAY, 87% VS SRA. THESE FIGURES SUGGEST THAT DISCREPANCIES MAY ARISE IN SPECIFIC SAMPLES WHEN COMPARING HEMOSIL HIT-AB(PF4-H) WITH OTHER METHODOLOGIES. FOR THIS REASON, THE IFU STATES THAT CONFIRMATION OF RESULTS WITH A FUNCTIONAL TEST IS RECOMMENDED AND THAT THE ASSAY RESULTS SHOULD BE USED WITH OTHER INFORMATION, INCLUDING THE CLINICAL CONTEXT IN FORMING A DIAGNOSIS. BASED ON THIS ASSESSMENT, NO REMEDIAL ACTION IS REQUIRED. THIS INCIDENT WILL BE MONITORED THROUGH TRENDING ANALYSIS.
DREAMSTATION | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING BLACK FILTER, BREATHING ISSUE, AND DEATH. MEDICAL INTERVENTION WAS NOT SPECIFIED. THERE ARE NO REPORTABLE MALFUNCTION. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR PRO SERIES C-FLEX + ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH AND LIVER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP PRO SERIES ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING EYE IRRITATION, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA / VOMITING, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, CANCER, DEATH AND LUNG. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K091319 | (B)(4) - ALLEGATION OF PATIENT HARM. | IN THE PREVIOUS REPORT, THE MANUFACTURER CAPTURED INCORRECT INFORMATION IN GENERAL INFORMATION, BOX B, BOX G AND BOX H. THE FOLLOWING CORRECTION HAS BEEN CORRECTED IN THIS REPORT. THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. GENERAL INFORMATION: 510K HAS BEEN CORRECTED. IN BOX B: DESCRIBE EVENT OR PROBLEM HAS BEEN CORRECTED. IN BOX G: 510K HAS BEEN CORRECTED. IN BOX H: TYPE OF REPORTED COMPLAINT HAS BEEN CORRECTED.
AZURION | K163715 | IT WAS REPORTED TO PHILIPS THAT AN AZURION TABLE WAS NOT WORKING PROPERLY. THE REPORT INDICATED THAT THERE WAS A PATIENT DEATH; HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, RESPIRATORY ISSUES AND SARCOMA CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS SERIES ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS SERIES ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING ESOPHAGEAL CANCER AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | IN BOX B: ADVERSE EVENT/PRODUCT PROBLEM CHANGED TO BOTH AND DEATH. DATE OF DEATH WAS UPDATED IN THIS REPORT FROM NOT APPLICABLE TO NO INFORMATION.
T:SLIM X2 INSULIN PUMP WITH INTEROPERABLE TECHNOLOGY (WITH T:CONNECT MOBILE APP) | K203234 | VOLUNTARY MEDWATCH RECEIVED ON 05/19/2025 REPORTED: IN ACCORDANCE WITH 21 CFR 803.22 (B)(2), WE ARE NOTIFYING YOU THAT ON (B)(6) 2025 THE PATIENTS WIFE REPORTED THE PATIENT DIED DUE TO A FAULTY INSULIN PUMP. NO FURTHER INFORMATION WAS PROVIDED. THIS REPORT REFLECTS INFORMATION RECEIVED BY FDA IN THE FORM OF A NOTIFICATION PER 803.22 (B)(2). REF REPORT: MW5170109. IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY "DUE TO A FAULTY INSULIN PUMP." DUE TO THE ANONYMOUS NATURE OF THE REPORT, TANDEM CUSTOMER TECHNICAL SERVICE UNABLE TO OBTAIN ADDITIONAL INFORMATION REGARDING THE ISSUE OR VERIFY AN EVENT DATE. NO ADDITIONAL PATIENT OR EVENT INFORMATION WAS AVAILABLE. | NO PRODUCT WAS RETURNED FOR EVALUATION. SHOULD NEW RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. RELATED MEDWATCH REPORT # MW5170108/MW5170109.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT REPORTED THE PATIENT PASSED AWAY. IT COULD NOT BE CONFIRMED IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. ADDITIONAL DETAILS WERE OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT¿S SON REPORTED THAT THE PATIENT WAS FOUND UNRESPONSIVE ON (B)(6) 2025. THE SON WAS NOT AT THE PATIENT¿S HOME AT THE TIME OF THE EVENT. 911 WAS CALLED BY SOMEONE IN THE HOME. EMERGENCY MEDICAL SERVICES (EMS) RESPONDED TO THE HOME AND RESUSCITATIVE EFFORTS WERE ADMINISTERED, BUT NOT SUCCESSFUL. THE PATIENT¿S SON REPORTED THAT HE DID NOT ENTER THE ROOM WHERE THE PATIENT WAS FOUND. HE COULD NOT CONFIRM IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE EVENT. THE CAUSE OF DEATH IS UNKNOWN. THE DEATH WAS REPORTEDLY UNEXPECTED (NO HEALTH ISSUES REPORTED). NO ISSUES WERE REPORTED WITH THE LIBERTY SELECT CYCLER OR PD THERAPY PRIOR TO THE DEATH. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A POSSIBLE TEMPORAL RELATIONSHIP BETWEEN THE PERITONEAL DIALYSIS (PD) AND USE OF THE LIBERTY SELECT CYCLER AND THE PATIENT¿S DEATH PRECEDED BY UNRESPONSIVENESS. HOWEVER, IT IS NOT CONFIRMED IF THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF THE ADVERSE EVENT. THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE PATIENT DEATH AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. ALTHOUGH NO CAUSE OF DEATH WAS CONFIRMED, PATIENTS ON DIALYSIS HAVE A SUBSTANTIALLY HIGHER MULTIVARIABLE-ADJUSTED MORTALITY THAN PATIENTS NOT ON DIALYSIS WHO HAVE CANCER, DIABETES, OR CARDIOVASCULAR DISEASE. A LARGE NUMBER OF RISK FACTORS THAT ARE UNRELATED TO THE DIALYSIS PROCEDURE HAVE BEEN ASSOCIATED WITH DECREASED SURVIVAL AMONG PATIENTS ON DIALYSIS. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, LUNG DISEASE, DEATH AND LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX + | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, INFLAMMATORY RESPONSE, DEATH, SARCOIDOSIS, STROKE, PNEUMONIA, DIFFICULTY BREATHING AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX + | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING INFLAMMATORY RESPONSE, CANCER, DEATH, LUNG CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION STATING THE PATIENT HAS PASSED AWAY DUE TO LUNG CANCER. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND NON-HODGKIN'S LYMPHOMA (FOLLICULAR). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, UNSPECIFIED KIDNEY PROBLEM AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING LUNG CANCER AND DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), KIDNEY DISEASE/TOXICITY, LUNG DISEASE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CANCER NON-HODGKINS LYMPHOMA, BACTERIAL LUNG INFECTION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, ASTHMA (NEW OR WORSENING), DEATH, HEART ATTACK, STROKE AND OTHER. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PLUS | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LIVER DISEASE/TOXICITY, CANCER, DEATH, OTHER, LIVER AND LUNG CANCER AND HEART ATTACK. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, CANCER, LIVER (FATAL). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SERVICE THAT A PERITONEAL DIALYSIS (PD) PATIENT EXPIRED WHILE CONNECTED TO THEIR LIBERTY SELECT CYCLER. UPON FOLLOW-UP WITH THE PATIENT¿S PDRN, IT WAS REVEALED THAT THE PATIENT EXPIRED IN THE EARLY MORNING ON (B)(6) 2025 DUE TO A CARDIAC ARREST ATTRIBUTED TO A SIGNIFICANT CARDIAC DISEASE HISTORY. IT WAS AFFIRMED THAT THE PATIENT WAS CONNECTED TO THEIR CYCLER AT THE TIME OF THE ARREST AND SUBSEQUENT DEATH. IT WAS EXPLAINED THAT THE PATIENT RECENTLY UNDERWENT AN EMERGENCY AORTIC VALVE REPLACEMENT (EXACT DATE NOT REPORTED) AND WAS DISCHARGED HOME WITH A HOLTER HEART MONITOR. THE OVERSEEING FACILITY NOTIFIED THE PATIENT¿S SPOUSE THAT THE PATIENT WAS PULSELESS, AND EMERGENCY MEDICAL SERVICES WERE ACTIVATED. THE PATIENT WAS NOT TRANSPORTED AND THEY WERE PRONOUNCED DECEASED IN THEIR HOME. IT WAS CONFIRMED THAT THE PATIENT¿S CARDIAC ARREST LEADING TO THEIR DEATH WAS UNRELATED TO PD THERAPY AND NOT DUE TO A DEFICIENCY OR MALFUNCTION OF ANY FRESENIUS PRODUCT(S) OR DEVICE(S). THE PDRN STATED THE PATIENT'S CYCLER WOULD BE RETURNED TO FRESENIUS. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CONTINUOUS CYCLIC PD (CCPD) THERAPY UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT¿S CARDIAC ARREST, LEADING TO THEIR DEATH. THE CAUSE OF THIS PATIENT¿S CARDIAC ARREST AND DEATH CAN BE ATTRIBUTED TO A SIGNIFICANT HISTORY OF CARDIOVASCULAR DISEASE AS REPORTED BY A MEDICAL PROFESSIONAL. IT IS WELL KNOWN THE END-STAGE RENAL DISEASE (ESRD) POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION, PARTICULARLY IN THE ENVIRONMENT OF CARDIOVASCULAR DISEASE AND RECENT INVASIVE CARDIAC SURGERY. THEREFORE, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS A ROOT CAUSE OR CONTRIBUTOR TO THIS PATIENT¿S ADVERSE EVENT. BASED ON THE AVAILABLE INFORMATION, THERE IS NO SPECIFIC ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) OR PRODUCT(S) DEFICIENCY OR MALFUNCTION, CAUSED OR CONTRIBUTED TO THE PATIENT¿S ADVERSE EVENTS.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING THROAT CANCER, RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY AND PATIENT HAS PASSED AWAY (DEATH). THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTOSV ADV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH, CONGESTIVE HEART FAILURE, FLUID RETENTION AND OTHER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP PRO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY DISEASE/TOXICITY, CANCER, DEATH AND GLIOBLASTOMA. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO SERIES AFLEX ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE, HEART DISEASE, OTHER (UNSPECIFIED EVENT) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING NOSE IRRITATION, RESPIRATORY TRACT IRRITATION, LIVER DISEASE/TOXICITY, LIVER CANCER (FATAL) AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
BIPAP AUTO SERIES ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH AND TONSIL CANCER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO SERIES AFLEX ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER RESULTING IN DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE/TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR PRO C-FLEX+ | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING CANCER, DEATH, OTHER (UNSPECIFIED EVENT), PROSTATE CANCER AND PULMONARY EMBOLISM. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING RESPIRATORY TRACT IRRITATION, KIDNEY DISEASE / TOXICITY, LUNG DISEASE, REDUCED CARDIOPULMONARY RESERVE AND DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, CUSTOMER COMPLAINT WAS NOT CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
CS300 INTRA-AORTIC BALLOON PUMP, PORTUGUESE, 220V | K063525 | IT WAS REPORTED THAT CS300 INTRA-AORTIC BALLOON PUMP (IABP) EXPERIENCED FAILURE IN FIBER OPTICAL SENSOR MODULE. BALLOON BURSTED INSIDE THE PATIENT. THERE ARE NO REPORTS OF COMPLICATIONS DIRECTLY RELATED TO THE BALLOON, BUT THE PATIENT DIED. | DUE TO CHARACTER LIMITATION IN E1, FULL INITIAL REPORTER NAME (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
OMNIPOD 5 POD | K203768 | IT WAS REPORTED ON (B)(6) 2025 BY THE PATIENT'S SISTER THAT THE PATIENT HAD PASSED AWAY. THE DATE OF THE PATIENT'S PASSING WAS NOT PROVIDED. THE PATIENT'S SISTER REPORTED; WHILE WEARING A POD, THE PATIENT EXPERIENCED A HYPOGLYCEMIC EPISODE. BLOOD GLUCOSE LEVELS DECLINED TO 43 MG/DL. LENGTH OF TIME THE POD WAS WORN WAS NOT PROVIDED. THE PATIENT'S SISTER REPORTED WHILE SHE WAS ASSISTING THE PATIENT WITH TREATING HER HYPOGLYCEMIA (SPECIFIC TREATMENT WAS NOT PROVIDED), THE PATIENT EXPERIENCED A HEART ATTACK. THE PATIENT WAS ALSO REPORTED TO HAVE AN IMMUNE DISORDER, CVID (COMMON VARIABLE IMMUNE DISORDER). THE PATIENT WAS TRANSPORTED BY AMBULANCE TO THE HOSPITAL, THE BLOOD GLUCOSE LEVEL HAD INCREASED TO 90 MG/DL, HOWEVER THE PATIENT REPORTEDLY DID NOT HAVE A PULSE WHILE IN TRANSPORT. THE PATIENT WAS PRONOUNCED DEAD UPON ARRIVAL TO THE HOSPITAL. NO ALLEGATIONS OF A DEVICE MALFUNCTION WERE REPORTED. | PLEASE DISREGARD PREVIOUSLY SUBMITTED (B)(4)/ REFERENCE REPORT NUMBER 3004464228-2025-23658-00. THIS WAS NOT A REPORTABLE MEDICAL EVENT RELATED TO AN INSULET DEVICE, AS THERE WAS NO ALLEGATION OR INDICATION THAT OUR DEVICE CAUSED OR CONTRIBUTED TO THE MEDICAL EVENT. THE EMDR REPORT WAS SUBMITTED IN ERROR. EMS (EMERGENCY MEDICAL SERVICES) WERE CONTACTED INITIALLY DUE TO THE PATIENT HAVING HYPOGLYCEMIA WITH A BLOOD GLUCOSE OF 43 MG/DL; HOWEVER, WHEN EMS ARRIVED THE PATIENT'S BLOOD GLUCOSE WAS 90 MG/DL, AND NO TREATMENT WAS PROVIDED. THE PATIENT LATER EXPERIENCED A HEART ATTACK AND WAS TRANSPORTED TO THE HOSPITAL. WHILE IN ROUTE TO THE HOSPITAL, THE PATIENT PASSED AWAY FROM THE HEART ATTACK. REVIEW OF CLOUD DATA FOUND OUR SYSTEM WAS CORRECTLY REDUCING AND STOPPED AUTOMATED INSULIN DELIVERY APPROPRIATELY AND THE AUTOMATED ALGORITHM WAS DETERMINED TO FUNCTION AS INTENDED. OUR PRODUCT DID NOT CAUSE OR CONTRIBUTE TO THE MEDICAL EVENT.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED USING A GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) SYSTEM. A GORE® EXCLUDER® AAA ENDOPROSTHESIS CONTRALATERAL LEG WAS ALSO USED IN THE AORTA.. PROXIMAL ACCESS WAS GAINED THROUGH A GORE® DRYSEAL FLEX INTRODUCER SHEATH IN THE RIGHT SUBCLAVIAN ARTERY. THE PROCEDURE WAS COMPLETED SUCCESSFULLY. EARLY ON (B)(6) 2025, WITHIN 24 HOURS OF THE INITIAL PROCEDURE, THE PATIENT HAD SOME WATERSHED INFARCTS. A CT OF THE CHEST SHOWED NOTHING UNUSUAL. ON (B)(6) 2025, THE PATIENT STARTED HAVING SEIZURES AND WAS INTUBATED. ACCORDING TO THE FSA, THERE WAS NO NOTABLE THROMBUS. ON (B)(6) 2025, THE PHYSICIAN REPORTED THAT THE PATIENT HAD DIED, SAYING THAT THE PATIENT NEVER RECOVERED FROM THE STROKE. | B2. OUTCOMES UPDATED TO DEATH. B4. EVENT DESCRIPTION UPDATED. G3/G4. DATE MANUFACTURER RECEIVED BY MANUFACTURER IS THE DATE WE LEARNED OF THE DEATH. H1. TYPE OF REPORTABLE EVENT UPDATED TO DEATH. H.6. HEALTH EFFECT - IMPACT CODE - ADDED F02.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING PANCREATIC CANCER, DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REMSTAR AUTO A-FLEX | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
VERSACROSS ACCESS SOLUTION | K150709 | IT WAS REPORTED A PATIENT DEATH OCCURRED. DURING A PULMONARY VEIN ISOLATION ABLATION PROCEDURE, A VERSACROSS ACCESS SOLUTION KIT WAS SELECTED FOR USE. THE PHYSICIAN OBTAINED ACCESS IN BOTH RIGHT AND LEFT FEMORAL VEIN. INTRACARDIAC ECHOCARDIOGRAPHY (ICE) IMAGINING WAS INSERTED IN LEFT FEMORAL VEIN. THE GUIDEWIRE WAS ADVANCED TO SUPERIOR VENA CAVA (SVC) AND THEN THE VERSACROSS TRANSSEPTAL SHEATH AS WELL AS THE RF WIRE. THE TRANSSEPTAL PUNCTURE (TSP) AND THE RF WIRE WERE ADVANCED IN THE LEFT ATRIUM. THERE WAS NO DIFFICULTY NOTED WITH THE TSP WORKFLOW. THE PATIENT HAD AN INCREASED HEART RATE OF 215 BPM THAT SUSTAINED FOR A MINUTE. THEN PERICARDIAL EFFUSION (PE) WAS NOTED ON ICE IMAGING AND A PERICARDIOCENTESIS BEGAN. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED AND CODE BLUE ALARM WAS CALLED. THE PATIENT WAS ASYSTOLE. THE PHYSICIAN OPENED PATIENT CHEST, AND BLOOD TRANSFUSION BEGAN. CARDIAC MESSAGE WAS PERFORMED. SODIUM BICARB WAS GIVEN THROUGH THE SHEATH. VASOPRESSORS GIVEN BY IC PHYSICIAN. AORTIC CANNULATION WAS ADMINISTERED AND VENOUS PERFUSION BEGAN. THE BYPASS SURGERY WAS INITIATED. A 90 VOLT SHOCK ADMINISTERED TO HEART, HOWEVER LATER THE PATIENT WAS DECLARED DEAD. THE AUTOPSY HAS NOT BEEN PERFORMED YET HOWEVER IT IS EXPECTED THAT DEATH WAS CAUSED BY COMPLICATIONS DUE TO PERFORATION OF CARDIAC TISSUE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). THE OFFICIAL CAUSE OF DEATH WAS PERICARDIAL EFFUSION. THE PATIENT HAD A SMALL LEFT ATRIUM. THERE ARE NO PRE-EXISTING PATIENT CONDITIONS THAT MAY HAVE CAUSED A COMPLICATION DURING THE PROCEDURE. MEDWATCH REFERENCE MW5170596. IT WAS ADDITIONALLY REPORTED THAT IT WAS PULLED OUT OVER A LITER OF FLUID. THE PATIENT CONTINUED TO BLEED. WHEN THE PATIENT'S CHEST WAS OPENED UP AND IT WAS NOTICED ONE PERFORATION IN THE RIGHT VENTRICLE (RV) THAT THEY BELIEVED MAY HAVE BEEN CAUSED BY THE TAP FOR THE PERICARDIOCENTESIS. THE LEFT APPENDAGE HAD A COUPLE OF PERFORATIONS AS WELL THAT THEY BELIEVE WAS FROM THE DILATOR. THE CALLER STATED THEY ASSUMED THE FARADRIVE CATHETER AND THE VERSACROSS WIRE WERE THE CAUSE OF THE INJURY. FARADRIVE SHEATH AND A BOSTON CORONARY SINUS (CS) WERE IN THE BODY WHEN THE EVENT OCCURRED. THE CALLER STATED THEY BELIEVE THE FARADRIVE DILATOR PERFORATED THE LEFT APPENDAGE WHILE GOING TRANSSEPTAL. | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
FREEDOM PERIPHERAL NERVE STIMULATOR | K171366 | THE COMMERCIAL TEAM MEMBER WAS NOTIFIED OF THE PATIENT'S DEATH WHICH OCCURRED ON (B)(6) 2025. | THE OTHER ADVERSE EVENTS ISSUES QUESTIONNAIRE WAS REVIEWED. THE CAUSE OF THE PATIENT'S DEATH IS DUE TO A BRAIN HEMORRHAGE WHICH WAS UNRELATED TO THE CURONIX DEVICE. THERE ARE NO ALLEGED DEFICIENCIES ASSOCIATED WITH THE CURONIX DEVICE. A CURONIX REPRESENTATIVE CONDUCTED A REVIEW OF STERILIZATION AND PACKAGING RECORDS FOR THE RESPECTIVE PRODUCT LOT; CURONIX HAS CONFIRMED THAT THE PRODUCT WAS DELIVERED STERILE, VALIDATED STERILIZATION PARAMETERS WERE USED, AND STERILE BARRIERS WERE VERIFIED TO BE INTACT FOLLOWING PACKAGING. THE STIMULATOR IS USED TO TREAT PAIN. THE CAUSE OF THE PATIENT'S DEATH IS UNRELATED TO THE CURONIX DEVICE. THE PROVIDED INFORMATION DOES NOT EVIDENCE THAT THE CURONIX DEVICE CONTRIBUTED TO THE ISSUE (NO FAULT FOUND). RATES REVIEWED AT THE MOST RECENT COMPLAINT TRENDING MEETING DO NOT INDICATE A SIGNIFICANT INCREASE IN OTHER ADVERSE EVENTS ISSUES. OTHER ADVERSE EVENTS ISSUE RATES REMAIN ACCEPTABLY LOW; THUS, A CAPA IS NOT REQUIRED. OTHER ADVERSE EVENTS ISSUE RATES WILL CONTINUE TO BE TRACKED AND TRENDED.
BIPAP AUTO SERIES ASSY | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH, EYE IRRITATION, NOSE IRRITATION, SKIN IRRITATION, RESPIRATORY TRACT IRRITATION, DIZZINESS, HEADACHE, NAUSEA, VOMITING, ASTHMA, KIDNEY DISEASE/TOXICITY, LIVER DISEASE/TOXICITY, LUNG DISEASE AND REDUCED CARDIOPULMONARY RESERVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
REP DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER WAS INFORMED THAT A PATIENT USING A REP DREAMSTATION AUTO BIPAP DEVICE PASSED AWAY ON (B)(6) 2024. ACCORDING TO THE PATIENT¿S WIFE, HE WAS NOT USING THE DEVICE AT THE TIME OF HIS DEATH, AS HE WAS UNDER HOSPICE CARE. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR INVESTIGATION. DURING THE EVALUATION OF THE DEVICE, THE THIRD-PARTY SERVICE CENTER VISUALLY INSPECTED THE DEVICE AND FOUND THE DEVICE WAS WORKING PROPERLY. NO DEVICE ERRORS OR MALFUNCTIONS WERE OBSERVED. THE DEVICE WAS SCRAPPED. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
SUREFORM | K173721 | IT WAS REPORTED THAT THE PATIENT UNDERWENT A DA VINCI-ASSISTED ROUX-EN-Y GASTRIC BYPASS AND HIATAL HERNIA REPAIR. THE PROCEDURE WAS DESCRIBED AS CHALLENGING DUE TO THE PATIENT¿S WEIGHT; HOWEVER, THERE WERE NO SPECIFIC COMPLICATIONS REPORTED AND THE PROCEDURE WAS COMPLETED WITHOUT ANY DA VINCI SYSTEM OR INSTRUMENT ISSUES. THE INTUITIVE CLINICAL SALES REPRESENTATIVE WAS LATER INFORMED BY THE SURGEON THAT ONE-WEEK POSTOPERATIVE, THE PATIENT WAS HOSPITALIZED FOR UNSPECIFIED SYMPTOMS. AN OPEN LAPAROTOMY WAS PERFORMED WHERE A LEAK WAS IDENTIFIED AT THE JEJUNOJEJUNOSTOMY (JJ) ANASTOMOSIS. AT AN UNSPECIFIED TIME, THE PATIENT REPORTEDLY CODED DUE TO POSTOPERATIVE COMPLICATIONS AND EXPIRED APPROXIMATELY 2 WEEKS AFTER THE INITIAL ROBOTIC SURGERY. NO FURTHER DETAILS WERE PROVIDED. ADDITIONAL INFORMATION WAS REQUESTED; A RESPONSE HAS NOT BEEN RECEIVED. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS OCCURRED DURING THE PROCEDURE THAT WOULD BE RELATED TO THE REPORTED EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER, MONOPOLAR CURVED SCISSORS, TIP-UP FENESTRATED GRASPER, VESSEL SEALER EXTEND, SUCTION IRRIGATOR, SUREFORM 60 STAPLER WITH WHITE RELOADS. A SITE HISTORY REVIEW SHOWS NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE SHOWED NO RELATED NON-CONFORMANCES WERE IDENTIFIED. A REVIEW OF THE SUREFORM 60 STAPLER LOG FOUND THAT THE INSTRUMENT WAS INSTALLED FIVE TIMES AND FIRED 5 WHITE RELOADS. THERE WERE NO INCOMPLETE CLAMPS OR UNCLAMPS BY THIS INSTRUMENT AND ALL FIRINGS WERE COMPLETED. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD MORBID OBESITY WITH A BMI OF 74 AND UNDERWENT A GASTRIC BYPASS WITH HIATAL HERNIA REPAIR. DETAILS ON HOW THE PROCEDURE WAS PERFORMED, AND WHICH INSTRUMENTS WERE USED AT THE ANASTOMOSIS WERE NOT PROVIDED. NO INTRAOPERATIVE ISSUES WERE REPORTED. THE PATIENT RETURNED ONE WEEK LATER AND WAS FOUND TO HAVE A LEAK. THE LOCATION OF THE LEAK WAS IDENTIFIED AT THE JEJUNOJEJUNOSTOMY (JJ) SITE. THE PATIENT SUBSEQUENTLY CODED AND DIED. WHEN THEY DIED AND THE EVENTS THAT LED TO THE DEATH ARE NOT PROVIDED. THE CAUSE OF DEATH IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED MALIGNANT HYSTERECTOMY SURGICAL PROCEDURE FOR ENDOMETRIAL CANCER, THE PATIENT CODED AND EXPIRED. THE SURGEON REPORTED THAT THE HYSTERECTOMY PORTION OF THE CASE WAS COMPLETED. DURING THE COLPOTOMY VAGINAL CUFF CLOSURE THE PATIENT CODED; THE OR TEAM EMERGENTLY REMOVED THE INSTRUMENTS, UNDOCKED THE SYSTEM AND PERFORMED CARDIOPULMONARY RESUSCITATION WITHOUT SUCCESS, AND DURING WHICH THERE WAS SOME UNSPECIFIED CHEST BLEEDING. THE SURGEON STATED HE BELIEVES THAT THE DA VINCI SYSTEM AND INSTRUMENTS HAD NOTHING TO DO WITH THE EVENT AS THERE WERE NO ISSUES REPORTED DURING THE HYSTERECTOMY PORTION; HE BELIEVES THAT THE PATIENT HAD AN UNDIAGNOSED CLOT, SUFFERED A PULMONARY EMBOLISM AND EXPIRED. NO ADDITIONAL DETAILS WERE PROVIDED. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 0 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, FENESTRATED FORCE BIPOLAR. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED FOR THE DEVICE(S) USED DURING THE PROCEDURE AND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS EVENT. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT SUDDENLY CODED DURING A MALIGNANT HYSTERECTOMY PROCEDURE. NO ISSUES OR COMPLAINTS WERE REPORTED WITH ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS UP TO THIS POINT. THEY EMERGENTLY UNDOCKED AND ATTEMPTED CPR BUT THE PATIENT DIED DESPITE THESE EFFORTS. THE CAUSE OF DEATH WAS ATTRIBUTED TO A PULMONARY EMBOLISM. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING INFORMATION WAS REPORTED TO GORE: THIS PATIENT WITH ABDOMINAL AORTIC ANEURYSM (AAA) HAD EXTENSIVE THROMBUS WITHIN THE AORTA FROM THE DESCENDING AORTA TO THE RENAL ARTERY BIFURCATION, MAKING IT AN EXTREMELY HIGH-RISK CASE. ALTHOUGH TREATMENT HAD BEEN DECLINED AT ANOTHER HOSPITAL, ENDOVASCULAR AORTIC REPAIR (EVAR) WAS TO BE PERFORMED AT THE REPORTING HOSPITAL CONSIDERING THE RISK OF ANEURYSM RUPTURE. ON (B)(6) 2025, THE PATIENT UNDERWENT AN EVAR FOR AAA USING GORE® EXCLUDER® AAA ENDOPROSTHESIS (EXCLUDER DEVICES). PRIOR TO THE PLACEMENT OF EXCLUDER DEVICES, AN ATTEMPT WAS MADE TO ACCESS THE RIGHT INTERNAL ILIAC ARTERY FOR COIL EMBOLIZATION, BUT IT WAS NOT SUCCEEDED AND TOOK APPROXIMATELY TWO HOURS. GIVEN THE HIGH-RISK NATURE OF THE CASE, EXTREME CAUTION WAS REQUIRED IN MANIPULATING GUIDEWIRES AND CATHETERS, AND ANGIOGRAPHY HAD TO BE MINIMIZED. AS A RESULT, THE PROCEDURE TOOK A TOTAL OF FOUR HOURS TO COMPLETE. THE PROCEDURE WAS SUCCESSFULLY DONE WITHOUT ANY ENDOLEAKS. ON (B)(6) 2025, THE PATIENT DEVELOPED PARAPLEGIA IN THE MORNING. ON (B)(6) 2025, THE PATIENT EXPERIENCED BLOODY STOOLS. ON (B)(6) 2025, IN THE EARLY MORNING, THE PATIENT WENT INTO SHOCK STATE AND SUBSEQUENTLY PASSED AWAY. THE REPORTING PHYSICIAN COMMENTED AS FOLLOWS: EXTREME CARE WAS TAKEN DURING GUIDEWIRE MANIPULATION AND BALLOON INFLATION. ANGIOGRAPHY WAS KEPT TO A MINIMUM, AND ALL POSSIBLE MEASURES WERE TAKEN. HOWEVER, THE THING THAT HAD BEEN FEARED OCCURRED. IT IS BELIEVED THAT THROMBUS DISPERSED DURING THE INDEX PROCEDURE, LEADING TO INTESTINAL ISCHEMIA AND ULTIMATELY RESULTING IN THE PATIENT¿S DEATH. ADDITIONAL PHYSICIAN'S COMMENT: THE ASSOCIATION BETWEEN THE GORE DEVICES AND INTESTINAL ISCHEMIA IS UNCLEAR. FUNDAMENTALLY, THE PATIENT HAD A HIGH-RISK BACKGROUND, AND THE PROCEDURE ITSELF WAS CHALLENGING TO PERFORM. REGARDING THE PARAPLEGIA, THERE WERE NUMEROUS THROMBI FROM THE DESCENDING AORTA TO THE RENAL ARTERY BIFURCATION. IT IS PRESUMED THAT WHEN CONTRAST MEDIUM WAS INJECTED, THE THROMBI WERE DISLODGED AND EMBOLIZED SMALLER ARTERIES SUCH AS THE LUMBAR ARTERIES. | AS GORE WAS UNABLE TO DETERMINE WHICH DEVICE/DEVICE COMPONENT IS INVOLVED IN THIS REPORTABLE ADVERSE EVENT, THE FOLLOWING ADDITIONAL DEVICE WILL BE IDENTIFIED IN THIS REPORT: CATALOG #DSF1233 WITH UNKNOWN SERIAL # OR UDI. H6: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE PERFORMED AS A VALID LOT NUMBER WAS NOT PROVIDED. H6: CODE D12: ACCORDING TO THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE (IFU), ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION OR ADDITIONAL INTRAOPERATIVE PROCEDURE TIME INCLUDE, BUT ARE NOT LIMITED TO: EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
INTELLIVUE MX450 PATIENT MONITOR | K130849 | IT WAS REPORTED DUE TO A DESATURATION EVENT, THE CUSTOMER REQUESTED ASSISTANCE TO MODIFY THE CRITICALITY OF THE TECHNICAL INOP ALARM FOR 'SPO2 SENSOR DISCONNECTED'. IT WAS INDICATED A RED OR YELLOW ALARM FOR THIS INOP WAS DESIRED AS CAREGIVERS DO NOT PAY ATTENTION TO THE BLUE INOP AND THE PATIENT SEVERELY DESATURATED AND BECAME BRADYCARDIC. A RESUSCITATION WAS UNSUCCESSFUL AND THE VENTILATED PATIENT ULTIMATELY PASSED AWAY. THERE WAS NO ALLEGATION OF MALFUNCTION; HOWEVER, AS THE STAFF DOES NOT PAY ATTENTION TO INOP ALARMS, THE REQUEST FOR A CONFIGURATION CHANGE FOR SPO2 SENSOR DISCONNECTED ALARMS WAS MADE. THE DEVICE WAS IN USE ON A PATIENT. THERE WAS A REPORT OF PATIENT OR USER HARM. | D4: UDI IS NOT AVAILABLE DUE TO INCOMPLETE SERIAL NUMBER. SERIAL NUMBER WAS REQUESTED, BUT IS UNKNOWN. E1: REPORTING INSTITUTION PHONE # (B)(6). THE FOLLOWING FUNCTIONAL TESTS WERE PERFORMED: THE ISSUE WAS EVALUATED BY A CLINICAL APPLICATION SPECIALIST. IT TURNED OUT THAT THERE WAS NO MALFUNCTION OF THE DEVICE. UPON FURTHER INSPECTION THE CUSTOMER WAS INFORMED THAT A CONFIGURATION FOR "SPO2 SENSOR DISCONNECTED" IN YELLOW OR RED ALARM, INSTEAD OF A BLUE ALARM IS NOT POSSIBLE. THIS CASE WILL BE HANDLED AS AN ENHANCEMENT REQUEST. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS A USER ERROR. THE ENGINEER PROVIDED THEIR ANALYSIS FINDINGS AND CONFIRMED THAT DEVICE TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
CARDIOHELP | K133598 | THE EVENT OCCURRED IN DENMARK DURING TREATMENT. IT WAS REPORTED THAT THE PATIENT ARRIVED IN THE HOSPITAL WITH A LUCAS PUMP PUMPING FOR ONE HOUR AND THE PATIENT HAS A CLOSED LEFT ANTERIOR DESCENDING (LAD). IT WAS HARD TO CANNULATE THE PATIENT AS THE PATIENT WAS BLEEDING ACCORDING TO THE CUSTOMER FROM ¿EVERYWHERE¿. THIS WAS CAUSED DUE TO THE LONG-TERM USE OF THE LUCAS PUMP. THE PATIENT WAS PUT ON ECMO. A CORONARY ARTERIOGRAPHY (KAG) WAS DONE AND THE LAD WAS OPENED. THE CUSTOMER WANTED TO DO A SMALL BYPASS WITH THE CANNULATION TO SECURE THE BLOOD FLOW TO THE LEG AND THEN THE STAFF CLAMPED THE HLS SET. A DEVICE DEFECTIVE ERROR MESSAGE OCCURRED, AS THE CUSTOMER HAS REMOVED THE SET WHILE THE PUMP WAS RUNNING. THE PATIENT WAS HAND CRANKED AND THE CARDIOHELP DEVICE WAS REPLACED. NEW INFORMATION WAS RECEIVED ON (B)(6) 2025 THAT THE PATIENT REMAINED FOR 3.5 HOURS ON THE REPLACED CARDIOHELP DEVICE BEFORE THE PATIENT PASSED AWAY. THE GETINGE SERVICE TECHNICIAN CONFIRMED THAT THERE WAS NO FAULT ON THE SECOND HARDWARE DEVICE NOR ON THE DISPOSABLE USED. THE PATIENT PASSED HAD MASSIVE BLEEDING, WHICH WAS BEFORE AND AFTER COMING ON ECMO. FURTHER IT WAS DIFFICULT TO CANNULATE THE PATIENT. THE PATIENT PASSED AWAY DUE TO BLEEDING AND LOW BLOOD PRESSURE. AS THE PATIENT PASSED AWAY AND THE CARDIOHELP DEVICE WAS REPLACED DURING TREATMENT, A REPORT IS REQUIRED. NEW INFORMATION RECEIVED ON (B)(6) 2025 THAT THE PATIENT DID NOT DIE DUE TO RELATED WRONG USER HANDLING ON THE DEVICE, BUT BECAUSE THE PATIENT WAS VERY SICK WHEN HE ARRIVED IN THE HOSPITAL. FURTHER THE LUCAS DEVICE WAS ON THE PATIENT FOR TWO HOURS AND LEAD TO A LOT IF INTERNAL BLEEDING. COMPLAINT ID#: (B)(4). | A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-05-27. NO PART WAS REPLACED. THE FAILURE OCCURRED DUE TO AN USER ERROR. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED AND THE DEVICE DEFECTIVE COULD BE CONFIRMED ON THE DATE OF EVENT. FURTHER THE GETINGE FST CONFIRMED THAT THE USER ERROR WAS VISIBLE IN THE LOGFILES. THE GETINGE FST CONFIRMED THAT THE HOSPITAL STAFF WILL RECEIVE AN TRAINING ON HOW TO HANDLE THE CARDIOHELP IN THIS KIND OF SITUATIONS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE DENMARK MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012.
CARDIOHELP | K133598 | THE EVENT OCCURRED IN DENMARK DURING TREATMENT. IT WAS REPORTED THAT THE PATIENT ARRIVED IN THE HOSPITAL WITH A LUCAS PUMP PUMPING FOR ONE HOUR AND THE PATIENT HAS A CLOSED LEFT ANTERIOR DESCENDING (LAD). IT WAS HARD TO CANNULATE THE PATIENT AS THE PATIENT WAS BLEEDING ACCORDING TO THE CUSTOMER FROM ¿EVERYWHERE¿. THIS WAS CAUSED DUE TO THE LONG-TERM USE OF THE LUCAS PUMP. THE PATIENT WAS PUT ON ECMO. A CORONARY ARTERIOGRAPHY (KAG) WAS DONE AND THE LAD WAS OPENED. THE CUSTOMER WANTED TO DO A SMALL BYPASS WITH THE CANNULATION TO SECURE THE BLOOD FLOW TO THE LEG AND THEN THE STAFF CLAMPED THE HLS SET. A DEVICE DEFECTIVE ERROR MESSAGE OCCURRED, AS THE CUSTOMER HAS REMOVED THE SET WHILE THE PUMP WAS RUNNING. THE PATIENT WAS HAND CRANKED AND THE CARDIOHELP DEVICE WAS REPLACED. NEW INFORMATION WAS RECEIVED ON (B)(6) 2025 THAT THE PATIENT REMAINED FOR 3,5 HOURS ON THE REPLACED CARDIOHELP DEVICE BEFORE THE PATIENT PASSED AWAY. THE GETINGE SERVICE TECHNICIAN CONFIRMED THAT THERE WAS NO FAULT ON THE SECOND HARDWARE DEVICE NOR ON THE DISPOSABLE USED. THE PATIENT PASSED HAD MASSIVE BLEEDING, WHICH WAS BEFORE AND AFTER COMING ON ECMO. FURTHER IT WAS DIFFICULT TO CANNULATE THE PATIENT. THE PATIENT PASSED AWAY DUE TO BLEEDING AND LOW BLOOD PRESSURE. AS THE PATIENT PASSED AWAY AND THE CARDIOHELP DEVICE WAS REPLACED DURING TREATMENT, A REPORT IS REQUIRED. NEW INFORMATION RECEIVED ON (B)(6) 2025 THAT THE PATIENT DID NOT DIE DUE TO RELATED WRONG USER HANDLING ON THE DEVICE, BUT BECAUSE THE PATIENT WAS VERY SICK WHEN HE ARRIVED IN THE HOSPITAL. FURTHER THE LUCAS DEVICE WAS ON THE PATIENT FOR TWO HOURS AND LEAD TO A LOT IF INTERNAL BLEEDING. COMPLAINT ID# (B)(4). | A GETINGE FIELD SERVICE TECHNICIAN (FST) WAS SENT FOR INVESTIGATION ON 2025-05-27. NO PART WAS REPLACED. THE FAILURE OCCURRED DUE TO AN USER ERROR. THE FST PERFORMED SAFETY, CALIBRATION, AND FUNCTIONALITY CHECKS TO FACTORY SPECIFICATIONS. ALL FUNCTION TESTS ARE PASSED. THE LOG FILES OF THE REPORTED CARDIOHELP DEVICE WERE REVIEWED AND THE DEVICE DEFECTIVE COULD BE CONFIRMED ON THE DATE OF EVENT. FURTHER THE GETINGE FST CONFIRMED THAT THE USER ERROR WAS VISIBLE IN THE LOGFILES. THE GETINGE FST CONFIRMED THAT THE HOSPITAL STAFF WILL RECEIVE AN TRAINING ON HOW TO HANDLE THE CARDIOHELP IN THIS KIND OF SITUATIONS. A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE DENMARK MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012.
INTELLIVUE MX800 PATIENT MONITOR | K150310 | IT WAS REPORTED THE PATIENT PASSED AWAY WHILE CONNECTED TO THE MONITOR. IT WAS INDICATED THE MONITOR WAS DEFECTIVE OR IMPAIRED; HOWEVER, THE SPECIFIC DEVICE ISSUE WAS NOT PROVIDED. THE DEVICE WAS TESTED AND FOUND TO BE FULLY FUNCTIONAL. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | THE AVAILABLE INFORMATION WAS REVIEWED BY THE PHILIPS PMS CLINICAL EXPERT. THE INFORMATION INDICATED AN ECG OR SPO2 RELATED ALARM WAS LIKELY ACKNOWLEDGED AND THERE WAS A DELAY IN TREATMENT. THE PATIENT WAS FOUND WITH ECG AND SPO2 UNPLUGGED. A LOG REVIEW REVEALED YELLOW ALARMS FOLLOWED BY TECHNICAL INOP ALARMS FOR 'LEADS OFF' WERE OCCURRING PRIOR TO THE PATIENT'S DEATH. THE ALARMS WERE LATER ACKNOWLEDGED AT THE CENTRAL STATION. THE PATIENT MONITOR AND CENTRAL STATION GENERATED ALL ALARMS ACCORDING TO THE CONFIGURATION. BASED ON THIS INFORMATION, THERE DOES NOT APPEAR TO BE A DEVICE MALFUNCTION THAT CAUSED OR CONTRIBUTED TO THE REPORTED DEATH; HOWEVER, USER ERRORS RELATED TO ALARM MANAGEMENT AND CLINICAL WORKFLOW MAY HAVE BEEN FACTORS. ADDITIONAL INFORMATION HAS BEEN REQUESTED, A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A DS2ADV AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH.
EDWARDS SAPIEN 3 ULTRA RESILIA TRANSCATHETER VALVE | K200258 | AS REPORTED BY EDWARDS JAPANESE AFFILIATE, DURING A TRANSFEMORAL TAVR PROCEDURE USING A 23 MM SAPIEN 3 ULTRA RESILIA (S3UR) VALVE, A STENT STRUT PEELED BACK, PREVENTING DEVICE REMOVAL. DESPITE CALCIFICATION IN THE LEFT COMMON ILIAC ARTERY (CIA) WITH A MINIMUM LUMINAL DIAMETER OF 5.9 MM, POBA WAS PERFORMED WITH AN 8 MM BALLOON BEFORE INSERTING THE 14 FR ESHEATH+. THE ESHEATH+ WAS INSERTED WITH SOME RESISTANCE BUT PASSED THROUGH. INTRALIPOS WAS USED AS A LUBRICANT, AND THE DELIVERY SYSTEM WAS INSERTED. THE S3UR VALVE PASSED THROUGH THE CALCIFIED CIA AREA WITH USUAL RESISTANCE, BUT A STENT STRUT PEELED BACK AFTER EXITING THE SHEATH, LEADING TO THE PROCEDURE'S ABORTION. ATTEMPTS TO WITHDRAW THE SHEATH AND DELIVERY SYSTEM AS ONE UNIT FAILED AS THE VALVE GOT STUCK AT THE CIA CALCIFICATION. BLOOD PRESSURE DROPPED TO 50/30 MMHG, PROMPTING ECMO INITIATION AND CARDIAC MASSAGE. ECMO WAS STARTED 10 MINUTES LATER, BUT ADEQUATE FLOW WAS NOT ACHIEVED, AND BLOOD PRESSURE REMAINED AROUND 60/40 MMHG. ANGIOGRAPHY REVEALED CIA BLEEDING, BUT HEMOSTASIS ATTEMPTS FAILED. SURGERY WAS DEEMED NECESSARY BUT NOT FEASIBLE DUE TO THE PATIENT'S AGE AND HYPOTENSION. THE FAMILY CONSENTED TO NO FURTHER INVASIVE TREATMENT, AND THE PATIENT WAS RETURNED TO THE ICU WITH THE DEVICES IN PLACE. THE PATIENT PASSED AWAY THE SAME DAY, AND POST-MORTEM EXAMINATION REVEALED A HOLE IN THE SHEATH WHERE THE STENT STRUT HAD PEELED BACK. | THIS IS 1 OF 3 REPORTS. THE INVESTIGATION IS ONGOING.
DS2ADV AUTO CPAP W/HUMID CELL/BT | K200480 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A USER OF A DREAMSTATION 2 ADVANCED AUTO CPAP DEVICE PASSED AWAY LAST YEAR. THE PATIENT'S SPOUSE IS REQUESTING INFORMATION ON WHAT TO DO WITH THE PATIENT'S UNUSED DEVICES. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ION | K182188 | IT WAS REPORTED THAT THE PATIENT UNDERWENT AN UNEVENTFUL ION ENDOLUMINAL LUNG BIOPSY PROCEDURE. THE PHYSICIAN REPORTED THAT THE PATIENT WAS STABLE THE ENTIRE TIME AND WAS EXTUBATED. THE BIOPSIED LESION WAS 5 CM IN SIZE, LOCATED IN THE LEFT UPPER LOBE FAR FROM THE PLEURA, AND WAS DETERMINED TO BE NECROTIC TISSUE. APPROXIMATELY 2.5 HOURS POST-PROCEDURE THE PHYSICIAN AND ANESTHESIOLOGIST WERE NOTIFIED THAT THE PATIENT WAS STILL DROWSY AND WAS NOT FOLLOWING COMMANDS; THE PATIENT WAS OTHERWISE HEMODYNAMICALLY STABLE. THE PATIENT WAS TRANSFERRED TO THE EMERGENCY DEPARTMENT FOR A STROKE WORK UP. THE PATIENT WAS INTUBATED AT SOME POINT OVERNIGHT WHILE IN THE ICU FOR DECLINING MENTAL STATUS AND EXPIRED AT AN UNSPECIFIED TIME LATER. THE CAUSE OF DEATH WAS UNKNOWN. AN AUTOPSY IS PENDING. PER THE PHYSICIAN, THE ION SYSTEM DID NOT EXHIBIT ANY ISSUES OR UNEXPECTED BEHAVIORS THAT CONTRIBUTED TO THE DEATH. NO ADDITIONAL INFORMATION WAS PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE SHOWED THERE WERE NO RELATED SYSTEM ERRORS TO HAVE OCCURRED DURING THE PROCEDURE THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT A PATIENT WITH A HISTORY OF ENDOMETRIAL CANCER UNDERWENT AN ION LUNG BIOPSY. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. POST EXTUBATION IN RECOVERY THE PATIENT¿S SENSORIUM REMAINED ALTERED 2.5 HOURS POST PROCEDURE PROMPTING WORK UP FOR A STROKE AND ADMISSION TO THE ICU. THE PATIENT WAS INTUBATED OVERNIGHT FOR WORSENING MENTAL STATUS BUT REMAINED OTHERWISE HEMODYNAMICALLY STABLE. THE PATIENT EVENTUALLY DIED AND AN AUTOPSY IS PENDING AND EFFORTS AT OBTAINING THE RESULTS ARE ONGOING. AT THIS JUNCTURE THE CAUSE OF DEATH IN THE SETTING OF AN ALTERED SENSORIUM POST PROCEDURE IS UNCLEAR BUT THERE IS NO EVIDENCE OF MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS, OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENTS WERE LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 ASSOCIATED DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 ASSOCIATED DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023. SECTION B2: PATIENT DATE OF DEATH: INTUBATION WAS THE DAY AFTER THE PROCEDURE (05/01/2025); HOWEVER, THE DATE OF DEATH WAS NOT PROVIDED.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT¿S DAUGHTER REPORTED THAT AS THEY WERE ATTEMPTED TO DISCONNECT THE PATIENT FROM THE CYCLER AFTER THE EVENT OCCURRED, FLUID DID COME OUT AS SHE WAS REMOVING THE CASSETTE FROM CASSETTE DOOR. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS IN TREATMENT ON (B)(6) 2025 CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT HAD TO GET UP AND USE THE RESTROOM. THE PATIENT WAS IN THE RESTROOM WHEN SHE EXPRESSED TO HER SPOUSE THAT SHE COULD NOT BREATHE. THE PATIENT¿S SPOUSE CALLED 911. EMERGENCY MEDICAL SERVICES (EMS) RESPONDED. THE PATIENT HAD BECOME UNRESPONSIVE PRIOR TO EMS ARRIVAL. EMS ATTEMPTED RESUSCITATION, BUT IT WAS UNSUCCESSFUL. THE PATIENT WAS PRONOUNCED DECEASED IN THE HOME. THE CAUSE OF THE PATIENT¿S DEATH IS UNKNOWN. THERE WERE NO ISSUES OR REPORTED ALARMS DURING THE TREATMENT PRIOR TO THE PATIENT¿S DEATH. THE FLUID LEAK REPORTED DURING THE DISCONNECTION OF THE PATIENT AND THE SUPPLIES OCCURRED BECAUSE OF THE PATIENT¿S DAUGHTER NOT KNOWING WHAT SHE WAS DOING. THE PATIENT¿S DAUGHTER WAS ATTEMPTING TO DISCONNECT THE CASSETTE, AND SHE CAUSED THE FLUID TO LEAK. THERE WAS NO LEAK INDICATED DURING THE PATIENT¿S TREATMENT. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS AND UTILIZATION OF THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF NOT BEING ABLE TO BREATH, BECOMING UNRESPONSIVE, AND DEATH. HOWEVER, THERE IS NO DOCUMENTATION OF A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE REPORTED FLUID LEAK WAS A RESULT OF AN ERROR THAT THE PATIENT¿S DAUGHTER MADE WHEN ATTEMPTING TO DISCONNECT THE SUPPLIES FROM THE CYCLER. THERE WERE NO REPORTED ALARMS OR ISSUES WITH THE TREATMENT PRIOR TO THE PATIENT EVENT. ALTHOUGH THE CAUSE OF DEATH WAS NOT CONFIRMED, SUDDEN CARDIAC DEATH (SCD) IS RESPONSIBLE FOR THE MAJORITY OF CARDIOVASCULAR-RELATED DEATHS IN PATIENTS WITH ESRD WITH STUDIES REPORTING THAT UP TO 25 % OF ALL DEATHS IN THIS HIGH-RISK POPULATION IS ATTRIBUTABLE TO SCD. BASED ON THE AVAILABLE INFORMATION AND NO EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH.
PULSE | K170121 | IT WAS REPORTED TO BOSTON SCIENTIFIC VIA AN ARTICLE PUBLISHED IN THE WORLD JOURNAL OF UROLOGY THAT A COMPARATIVE RETROSPECTIVE ANALYSIS WAS CONDUCTED TO EVALUATE THE SAFETY AND EFFICACY OF HOLEP IN PATIENTS AGED OVER 85 YEARS WITH INDWELLING CATHETER (IDC). A OF A TOTAL OF 144 PATIENTS, WITH 71 IN THE GROUP OF PATIENTS AGED 85 OR GREATER AND 73 IN THE CONTROL GROUP WITH PATIENTS UNDER 70 YEARS OF AGE, WERE EVALUATED FOR RATE OF POSTOPERATIVE SPONTANEOUS MICTURITION, WITH SUCCESS DEFINED AS THE ABILITY TO RESUME SATISFACTORY SPONTANEOUS VOIDING WITHOUT SIGNIFICANT POST-VOID RESIDUAL VOLUME. MEDICAL FILES OF PATIENTS OPERATED BETWEEN JUNE 2012 AND APRIL 2020, AT TWO ACADEMIC CENTERS IN FRANCE, WERE REVIEWED TO RETRIEVE THOSE WITH HISTORY OF ACUTE OR CHRONIC URINARY RETENTION AND TRIAL WITHOUT CATHETER (TWOC) FAILURE REQUIRING MAINTENANCE OF IDC AT THE TIME OF SURGERY. FOLLOW-UP ASSESSMENTS WERE CONDUCTED AT BASELINE OF 3 MONTHS, 6 MONTHS, AND 12 MONTHS. FOLLOWING THE PROCEDURES, THE FOLLOWING COMPLICATIONS WERE REPORTED FOR THE OLDER PATIENTS GROUP: 8 PATIENTS REQUIRED BLOOD TRANSFUSIONS, 17 PATIENTS HAD TWOC FAILURE DURING HOSPITALIZATION, AT 3-MONTHS EVALUATION 5 PATIENTS COULD NOT VOID SPONTANEOUSLY, AND AT 12-MONTHS EVALUATION 2 PATIENTS REMAINED WITH AN IDC, 8 PATIENTS HAD URINARY INCONTINENCE, AND 10 PATIENTS HAD DIED. FOLLOWING THE PROCEDURES, THE FOLLOWING COMPLICATIONS WERE REPORTED FOR THE YOUNGER PATIENTS GROUP: 7 PATIENTS REQUIRED BLOOD TRANSFUSIONS, 7 PATIENTS HAD TWOC FAILURE DURING HOSPITALIZATION, AND AT 12-MONTHS EVALUATION 4 PATIENTS HAD URINARY INCONTINENCE AND 2 PATIENTS HAD DIED. IT WAS NOTED THAT ONLY ONE DEATH WAS DIRECTLY RELATED TO SURGERY (SEVERE SEPSIS DUE TO PNEUMONIA, COMPLICATED BY MULTIORGAN FAILURE, LEADING TO DEATH ON POSTOPERATIVE DAY 7). THIS REPORT IS FOR THE DEATH DIRECTLY RELATED TO SURGERY REPORTED IN THE ARTICLE. REFERENCE LITERATURE ARTICLE "COMPARATIVE STUDY OF HOLEP IN ELDERLY PATIENTS WITH INDWELLING CATHETERS: A RETROSPECTIVE DUAL-CENTER STUDY" INCLUDED WITH THIS REPORT FOR A FULL LISTING OF PHYSICIANS AND HEALTH CARE FACILITIES. | DATE OF DEATH (B2) WAS APPROXIMATED BASED ON AN ESTIMATED EARLIEST PROCEDURE DATE NOTED IN THE STUDY AND OTHER INFORMATION PROVIDED REGARDING THE DEATH EVENT. KLEIN, C., ANRACT, J., PINAR, U., LACROIX, X., MANSOUR, R., ROBERT, G., DELONGCHAMPS, N.B. (2025). "COMPARATIVE STUDY OF HOLEP IN ELDERLY PATIENTS WITH INDWELLING CATHETERS: A RETROSPECTIVE DUAL-CENTER STUDY". WORLD JOURNAL OF UROLOGY 43, 75. HTTPS://DOI.ORG/10.1007/S00345-024-05437-9 DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
ALARIS SYSTEM | K211218 | INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST. | THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K211218 | INITIALLY, THE CUSTOMER REQUESTED LOG REVIEW FOR AN UNSPECIFIED ISSUE. AFTER FOLLOW UP BY CLINICAL SPECIALIST, IT WAS REPORTED THAT THERE WAS A PROGRAMMING ERROR WITH A STAT INFUSION OF EPINEPHRINE TO A PATIENT ADMITTED FOR CARDIAC ARREST. THE EPINEPHRINE (0.1MCG/KG/MIN THEN TITRATABLE) WAS INITIALLY PROGRAMMED USING INTEROPERABILITY. THE TOTAL BAG VOLUME AND CONCENTRATION WAS 4MG IN 250ML. AFTER THE INITIAL PROGRAMMING, THE PUMP WAS STOPPED. LATER, THE PUMP WAS RESTARTED WITH MANUAL PROGRAMMING. AFTER THE EPINEPHRINE INFUSION WAS RESTARTED, THE CLINICIAN NOTICED THAT THE PUMP WAS SET TO NON-WEIGHT-BASED DOSING (MCG/MINUTE) HOWEVER THE MEDICATION ORDER AND TITRATIONS WERE INTENDED TO BE ADMINISTERED USING WEIGHT-BASED DOSING. IT WAS REPORTED THAT THE "PATIENT ENDED UP BEING STARTED ON ADDITIONAL MEDICATIONS BECAUSE IT WAS THOUGHT THAT THEY WERE MAXED ON THE CURRENT MEDICATION." FOLLOWING THE EVENT, THE PATIENT EXPERIENCED HYPOTENSION AND WAS STARTED ON DOBUTAMINE. THE PATIENT EXPIRED. IT WAS NOT CLEAR TO WHAT DEGREE THE INFUSION ERROR IMPACTED THE PATIENT ALREADY ADMITTED FOR CARDIAC ARREST. | THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
OMNIPOD 5 CONTROLLER | K203768 | ON 22MAY25 INSULET RECEIVED AN EMAIL FROM A HEALTHCARE PROFESSIONAL (HCP) NOTIFYING INSULET OF THEIR PATIENT'S DEATH DUE TO DIABETIC KETOACIDOSIS (DKA) AND INQUIRING ABOUT THE PERSONAL DIABETES MANAGER (PDM). INSULET FOLLOWED UP WITH THE HCP AND OBTAINED THE ADDITIONAL BACKGROUND. THE PATIENT HAD COMORBIDITIES THAT INCLUDED RENAL DISEASE AND A PREVIOUS CEREBROVASCULAR ACCIDENT. REPORTEDLY, THE PATIENT WAS AWAITING A RENAL TRANSPLANT. THE PATIENT HAD STRUGGLED WITH MANAGEMENT OF DIABETES PRIOR TO USE OF OMNIPOD 5. APPROXIMATELY A WEEK BEFORE THE PATIENT EXPIRED, SHE WAS EXPERIENCING INCREASED BLOOD GLUCOSE LEVELS. THE PATIENT PASSED AWAY ON (B)(6) 2024. THE PATIENT EXPIRED AT HOME AND WAS WEARING THE DEVICE AT TIME OF DEATH. THE CLINIC HAS POSSESSION OF THE PDM AND WILL MAINTAIN THE DEVICE. INSTRUCTIONS FOR UPLOADING LOGS WERE PROVIDED. THE REQUESTED GLOOKO INFORMATION IS NOT AVAILABLE AS THE PATIENT DID NOT CONSENT TO GLOOKO/OP5 DATA SHARING. THE OP5 WAS BEING USED WITH A DEXCOM G6. DEXCOM HAS BEEN NOTIFIED. | THE PHYSICAL DEVICE WAS NOT RETURNED. ANALYSIS OF THE AVAILABLE CLOUD DATA FOUND NO EVIDENCE OF ISSUES WITH THE SYSTEM. IT WAS NOTED THAT THE USER DELIVERED THE LARGEST AMOUNTS OF TOTAL INSULIN AND TOTAL BOLUS INSULIN ON THE DAY BEFORE THE DATE OF OCCURRENCE, HOWEVER THERE IS NO EVIDENCE OF A SYSTEM ISSUE CONTRIBUTING TO THIS INSULIN DELIVERY. THE DATA SHOWED THAT THE AUTOMATED ALGORITHM APPROPRIATELY ADAPTED THE AUTOMATED INSULIN DELIVERY BASED ON THE ESTIMATED GLUCOSE VALUES AND BOLUS DELIVERIES. THE DATA SHOWED THAT THE SYSTEM WAS USED PRIMARILY IN AUTOMATED MODE IN THE DAYS LEADING UP TO THE DATE OF OCCURRENCE ((B)(6) 2024). THE DATA SHOWED THAT THE AUTOMATED ALGORITHM WAS ABLE TO APPROPRIATELY ADAPT THE MICROBOLUS AMOUNTS BASED ON THE ESTIMATED GLUCOSE VALUES RECEIVED BY THE POD FROM THE SENSOR. THE DATA SHOWED THAT GLUCOSE VALUES REACHED URGENT HIGH VALUES AT 03:25:39 ON (B)(6) 2024. AT THAT TIME, THE SYSTEM STARTED DELIVERING AT THE MAX AUTOMATED RATE AND CONTINUED TO DELIVER AT THE MAX RATE DUE TO URGENT HIGH BLOOD GLUCOSE LEVELS UNTIL AN AUTOMATED DELIVERY RESTRICTION ALERT WAS GENERATED AT 05:30:40 ON (B)(6) 2024. THE ALERT WAS ACKNOWLEDGED BY THE PATIENT AND THE SYSTEM WAS PUT INTO MANUAL MODE. THE SYSTEM STARTED DELIVERING THE USER'S MANUAL BASAL PROGRAM AS EXPECTED. AN 8.35 U BOLUS WAS DELIVERED SHORTLY AFTER THE SYSTEM MODE CHANGE, RESULTING IN BLOOD GLUCOSE LEVELS DECREASING, REACHING URGENT LOW VALUES AT 07:40:40 ON (B)(6) 2024. BLOOD GLUCOSE LEVELS STARTED INCREASING AGAIN, REACHING URGENT HIGH AT 12:20:41 ON (B)(6) 2024. SENSOR VALUES RECEIVED BY THE POD STAYED URGENT HIGH UNTIL THE SENSOR BECAME INACTIVE AT 22:15:43 ON (B)(6) 2024. NO SENSOR VALUES WERE RECEIVED BY THE POD THROUGH THE LAST DATA POINT AT 11:30:47 ON (B)(6) 2024. THE SYSTEM CONTINUED TO DELIVER THE USER'S MANUAL BASAL PROGRAM DURING THIS PERIOD. MULTIPLE BOLUSES WERE DELIVERED ON (B)(6) 2024, HOWEVER NO BOLUSES WERE DELIVERED ON (B)(6) 2024. NO EVIDENCE OF ISSUES WITH THE SYSTEM WERE OBSERVED IN THE AVAILABLE CLOUD DATA. LOT RELEASE RECORDS WERE REVIEWED AND THE PRODUCT LOT MET ALL ACCEPTANCE CRITERIA. SPECIFICALLY, A POD IS PAIRED TO A PDM AND PUT THROUGH SIMULATED USE TESTING INCLUDING COMMUNICATING WITH THE PDM, DEPLOYMENT, DELIVERING FLUID, OCCLUSION DETECTION, AND FREEDOM FROM HAZARD ALARMS. IF ADDITIONAL INFORMATION IS PROVIDED A SUPPLEMENTAL REPORT WILL BE SUBMITTED.
KANGAROO | K221603 | THE VP OF NURSING REACHED OUT TO A CARDINAL HEALTH SALES REP TO INFORM THAT A PATIENT WHO HAD BEEN ON A KANGAROO OMNI PUMP DID NOT APPEAR TO BE RECEIVING THE CORRECT FLUSHING VOLUMES. THE PATIENT WAS STARTED ON (B)(6) 2025 FOR GLUCERNA ENTERAL FEEDINGS. THE ORDERS FOR PUMP DELIVERY WERE TO FEED 40ML/HR AND FLUSH 200ML EVERY 6H. THE DAUGHTER SAID THE PUMP WAS NOT FLUSHING AND THE PATIENT BECAME DEHYDRATED. THE DAUGHTER SAID WHEN SHE WOULD CHANGE THE BAG EVERY DAY SHE SAW THAT THE WATER WAS STILL IN THE BAG BUT DIDN'T CALL TO ASK WHY. THE PATIENT WAS ADMITTED TO ICU ON (B)(6) 2025 FOR DEHYDRATION. AND PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION RECEIVED: THE PATIENT WAS ADMITTED (B)(6) 2025, FOR ENTERAL THERAPY WITH DYSPHAGIA RELATED TO CVA AND IV ANTIBIOTIC THERAPY FOR POSITIVE BLOOD CULTURES. MEDICAL HISTORY DOCUMENTED AT THAT TIME INCLUDED HYPERLIPIDEMIA, HYPERCHOLESTEROLEMIA, TYPE II DIABETES MELLITUS, OSTEOPOROSIS, DEGENERATIVE LUMBAR SPINE DISEASE, AND SUSPECTED DEMENTIA. | THE DEVICE WAS RECEIVED MAY 27TH, 2025. THE DEVICE WAS RECEIVED ON SOFTWARE VERSION 1.5.4612. UPON RECEIPT, THE DEVICE WAS DISINFECTED USING ISOPROPANOL ALCOHOL. THE DEVICE WAS SUBJECTED TO A VISUAL INSPECTION TO DETERMINE IF ANY OBVIOUS DAMAGE OR ABNORMALITIES COULD BE DETECTED. THE DEVICE APPEARED TO BE NORMAL WITH NO DAMAGE OR MISSING PARTS DETECTED. DEVICE INTAKE, VISUAL INSPECTION, AND CERTIFICATION WAS COMPLETED. THE DEVICE HAD A STICKER ON THE BACK INDICATING THAT IT HAD UNDERGONE PREVENTATIVE MAINTENANCE AND PASSED TESTING BY (B)(4) ON 3/20/2025. THIS IS SUPPORTED BY THE DEVICE STATING THAT IT HAD LAST PASSED CERTIFICATION 68 DAYS AGO (AS OF 5/27/2025). AS PART OF END OF LINE TESTING, A KANGAROO OMNI PUMP MUST PASS FACTORY CERTIFICATION. THE FACTORY CERTIFICATION FOR THIS SERIAL NUMBER PASSED ON THE FIRST ATTEMPT ON NOVEMBER 12TH, 2024. AFTER THE DEVICE WAS RECEIVED ON MAY 27TH, A SERVICE CERTIFICATION WAS RUN AND PASSED. THE DEVICE WAS POWERED ON AND THE HISTORY FUNCTION WAS ENGAGED. THE HISTORY FEATURE RECORDS AND STORES IN MEMORY 30 DAYS OF DELIVERY VOLUMES. THE HISTORY FEATURE ALLOWS REVIEW IN 1-HOUR INCREMENTS FOR THE LAST 72 HOURS AND REVIEW OF 30 DAYS OF HISTORY IN 1-DAY INCREMENTS. REVIEW OF THIS DEVICE INDICATED NO HISTORY RECORDED UNTIL DAY 24. HISTORY DATA WAS PRESENT FOR DAYS 24 THROUGH 30. THE FEED TOTALIZER HAS A TOTAL FED OF 5601ML AND A TOTAL FLUSHED OF 9ML. THIS INDICATES THAT THE PUMP FED AT LEAST 2,534ML THAT WAS CAPTURED OUTSIDE OF THE HISTORY PERIOD. WHEN THE DEVICE WAS RECEIVED, THE KEEP SETTINGS BUTTON WAS PRESSED TO RECORD HOW THE USER PROGRAMMED THE DEVICE. THE DEVICE WAS PROGRAMMED IN CONTINUOUS MODE WITH A RATE OF 40ML/HR USING A STANDARD FEED/FLUSH SET. THE FLUSH SETTINGS WERE INCOMPLETE WITH ONLY THE FLUSH VOLUME OF 40ML PROGRAMMED AND NO FLUSH INTERVAL PROGRAMMED. FAILURE TO PROGRAM BOTH THE FLUSH VOLUME AND FLUSH INTERVAL WILL RESULT IN THE DEVICE NOT PERFORMING THE FLUSH OPERATION AS THE INPUT WAS NOT FULLY PROGRAMMED. NAVIGATING TO THE SHOW FLUSH PAGES ON EITHER THE READY TO FEED OR FEEDING SCREENS ALSO SHOW THAT THE FLUSH WAS NOT PROPERLY PROGRAMMED. THE KANGAROO OMNI ENTERAL FEEDING PUMP INVOLVED IN COMPLAINT PASSED ALL TESTING AND WAS DETERMINED TO BE FUNCTIONING AS EXPECTED. ANALYSIS OF THE PROGRAM SETTINGS, ALONG WITH REVIEW OF THE HISTORY INDICATE INCOMPLETE PROGRAMMING OF THE FLUSH SETTINGS BY THE USER RESULTED IN THE PUMP NOT FLUSHING.
INTELLIVUE MULTI-MEASUREMENT MODULE X3 | K171801 | IT WAS REPORTED THE PATIENT PASSED AWAY AFTER LOSS OF MONITORING WHILE TRANSPORTING THE PATIENT. THE CUSTOMER INDICATED WHEN THE PATIENT WAS TRANSPORTED TO OTHER AREAS OF THE HOSPITAL FOR TESTING, THE MONITOR LOSES NETWORK CONNECTIVITY; THEREFORE, THE PATIENT WAS NOT MONITORED AT THE CENTRAL STATION. THE ALARMS WERE PRESENT ON THE X3 DURING TRANSPORT. IT WAS INDICATED THE CUSTOMER WAS NOT ALLEGING THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE PATIENT WAS BEING MONITORED FOR HEART RATE, SPO2, AND RESPIRATION. THEY WERE DNR AND IN COMFORT CARE AT THE TIME OF THE EVENT, AND THEY WERE IN POOR CONDITION. WHILE BEING TRANSPORTED TO CT, THE PATIENT'S STATUS DECLINED, AND THE NURSES DID NOT SEE THE CHANGE IN CONDITION AS THE X3 MONITOR DROPPED FROM THE CENTRAL STATION. THE RADIOLOGY TECHNICIAN HEARD THE X3 MONITORING BUT DID NOT KNOW WHAT THE MONITOR ALARMS MEANT. NURSES DO NOT TYPICALLY STAY WITH PATIENTS DURING TRANSPORT. NO RESUSCITATION ATTEMPTS WERE MADE DUE TO THE PATIENT'S DNR STATUS. IN ADDITION, IT WAS NOTED THERE ARE NETWORK ACCESS POINTS LOCATED OUTSIDE OF THE EMERGENCY DEPARTMENT TO THE CT LOCATION. | THE PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE AND ASSESSED THE WIRELESS SYSTEM. IT WAS DETERMINED THAT TWO ACCESS POINT (AP) CABLES WERE UNPLUGGED IN THE ER. A COVERAGE AREA VERIFICATION (CAV) FROM (B)(6) 2024 REFLECTS THAT THE WIRELESS SYSTEM WAS HOOKED UP CORRECTLY AND RUNNING AT THE TIME. IT IS UNKNOWN WHY THE ACCESS POINTS WERE LATER DISCONNECTED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE ISSUE WAS THE DISCONNECTION OF THE ACCESS POINT CABLES. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE FSE PERFORMED A WIRELESS REMEDIATION IN WHICH THE ACCESS POINTS WERE RETURNED TO FACTORY DEFAULT AND RECONFIGURED. WHEN THE FSE WAS ONSITE ON MAY 13, 2025, HE VERIFIED THAT ALL AP¿S WERE ONLINE, AND NO ERRORS WERE DETECTED BY THE UPGRADER TOOL. TO ENSURE PROPER WIRELESS COVERAGE FROM THE INTELLIVUE X3 MONITOR IN THE ER, THE FSE ALSO PERFORMED A CAV CHECK WITH ANOTHER WIRELESS DEVICE (AN MX40).
DURAMATRIX-ONLAY COLLAGEN DURA SUBSTITUTE MEMBRANE | K061487 | LIMITED INFORMATION PROVIDED. CLINICAL DATA COLLECTION SITE REPORTED A "DM/DMO" PRODUCT WAS USED ON A PATIENT FOR A MASS RESECTION ON (B)(6) 2025. NO PRODUCT ITEM NUMBER, PRODUCT LOT NUMBER, OR SPECIFIC PRODUCT NAME WAS PROVIDED. MEDICAL INTERVENTION REQUIRED AND A DECOMPRESSIVE HEMICRANIECTOMY WAS PERFORMED ON (B)(6) 2025. VARIOUS HOSPITAL COMPLICATIONS WERE REPORTED (SEE ATTACHED DOCUMENT FOR ALL REPORTED PATIENT INFORMATION). PATIENT EXPERIENCED ACUTE SEVERE NEUROLOGIC DECLINE ON (B)(6) 2025 AND BECAME HYPOTENSIVE, LIKELY FROM SEPTIC SHOCK. PATIENT WAS TERMINALLY EXTUBATED AND PASSED SHORTLY THEREAFTER ON (B)(6) 2025. IT WAS NOTED THAT "RELATIONSHIP OF THE EVENT" (EVENT OF SEPTIC SHOCK) TO THE INTERVENTION IS UNRELATED. CLINICIAN DOES NOT BELIEVE THERE IS CORRELATION BETWEEN THE "DM/DMO" PRODUCT AND THE SEPTIC SHOCK. NO ADDITIONAL EVENT, PRODUCT, OR PATIENT INFORMATION PROVIDED.
UNKNOWN_NEUROVASCULAR_PRODUCT | K223305 | IN DISCOUNT (EVALUATION OF MECHANICAL THROMBECTOMY IN ACUTE ISCHEMIC STROKE RELATED TO A DISTAL ARTERIAL OCCLUSION) STUDY, THERE WAS HEMORRHAGIC TRANSFORMATION WITH WORSENING OF CLINICAL CONDITION IN A CONTEXT OF FAILED RECANALIZATION ON (B)(6) 2023. PATIENT DIED ON (B)(6) 2024. NO ADDITIONAL INFORMATION IS AVAILABLE. THE CAUSE OF DEATH IS UNKNOWN AS THE PATIENT WAS LOST OF VIEW.
X-PORT ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K022983 | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE, THE PATIENT ALLEGEDLY DEVELOPED WITH UNSPECIFIED INFECTION. REPORTEDLY, THE PATIENT EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H10: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE/WAS PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. H11: SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
WALLFLEX ESOPHAGEAL | K233939 | IT WAS REPORTED TO BOSTON SCIENTIFIC CORPORATION THAT A WALLFLEX ESOPHAGEAL STENT WAS IMPLANTED IN THE ESOPHAGUS TO TREAT A MALIGNANT OBSTRUCTION DURING A GASTROSCOPY WITH STENT PLACEMENT PROCEDURE PERFORMED ON (B)(6) 2025. THE WALLFLEX ESOPHAGEAL STENT WAS SUCCESSFULLY DEPLOYED. HOWEVER, UNIDENTIFIED LIQUID IN THE PATIENT'S STOMACH WAS OBSERVED. WHEN THE STENT WAS ABLE TO OPEN UP THE OBSTRUCTION IN THE ESOPHAGUS, THE LIQUID CAME UP CAUSING THE PATIENT TO ASPIRATE. EMERGENCY RESUSCITATION WAS ATTEMPTED, BUT THE PATIENT PASSED AWAY. NO FURTHER INFORMATION HAS BEEN OBTAINED DESPITE GOOD FAITH EFFORTS. | BLOCK D4, H4: THE COMPLAINANT WAS UNABLE TO PROVIDE THE SUSPECT DEVICE UPN AND LOT NUMBER. THEREFORE, THE MANUFACTURE AND EXPIRATION DATES ARE UNKNOWN. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H6: IMDRF PATIENT CODE E0704 CAPTURES THE REPORTABLE EVENT OF ASPIRATION. IMDRF IMPACT CODE F2306 CAPTURES THE REPORTABLE EVENT OF EMERGENCY RESUSCITATION WAS ATTEMPTED. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT PASSED AWAY.
ANGIOJET? SOLENT? OMNI | K111182 | IT WAS REPORTED THAT THE PATIENT DIED. THE TARGET LESION WAS LOCATED IN THE POPLITEAL ARTERY. THE PROCEDURE WENT WELL, HOWEVER, THE DOCTOR NOTED THAT THE PATIENT MIGHT DEVELOP COMPARTMENT SYNDROME. POST-PROCEDURE, THE PATIENT EXHIBITED SOME UNKNOWN COMORBIDITIES. NO COMPLICATIONS RELATED TO THE DEVICE, OR THE PROCEDURE WERE NOTED; HOWEVER, THE PATIENT SUBSEQUENTLY DECEASED. | H6 - PATIENT CODES: UPDATED.
OP TELESCOPE, 7°, 11 FR. X 220 MM, 7.5 FR. CHANNEL, ANGLED OCULAR | K853486 | ADDITIONAL INFORMATION RECEIVED THE PATIENT WAS RE-ADMITTED TO THE FACILITY, RECEIVED BLOOD TRANSFUSIONS, WAS TREATED FOR SEPTIC SHOCK, AND WAS ADMINISTERED THE ANTIBIOTICS ZOSYN AND LINEZOLID. THE PATIENT WAS SUBSEQUENTLY TRANSFERRED FOR SLEDD (SUSTAINED LOW-EFFICIENCY DAILY DIALYSIS) AND HAS SINCE EXPIRED. | THIS SUPPLEMENTAL REPORT IS BEING SUBMITTED TO UPDATE FIELDS (A2, A3, A4, A5, A6, B1, B2, B5, B7, H1, H6) AND ATTACH THE RELATED MEDWATCH REPORT. ADDITIONALLY, THIS SUPPLEMENTAL REPORT IS PROVIDING FURTHER INFORMATION SHARED BY THE CUSTOMER. IT WAS DETERMINED BY THE OLYMPUS UROLOGY MEDICAL DIRECTOR THAT IT IS UNLIKELY THE OLYMPUS DEVICE CAUSED OR CONTRIBUTED TO THE HEMORRHAGE/BLOOD LOSS/BLEEDING IN THIS CASE. BLEEDING IS A KNOWN PROCEDURAL COMPLICATION AND IS MORE LIKELY RELATED TO THE PATIENT¿S COMPLEX ANATOMY AND COMPLEX UNDERLYING MEDICAL HISTORY. IT WAS ALSO DETERMINED THAT IT IS UNLIKELY THAT THE OLYMPUS DEVICE CAUSED OR CONTRIBUTED TO SEPSIS AS THE PATIENT PRESENTED WITH AN INFECTED OR ERODED ARTIFICIAL URINARY SPHINCTER THAT REQUIRED REMOVAL. THE SEPSIS IN THIS CASE IS MORE LIKELY RELATED TO THE PATIENT¿S PRESENTING DIAGNOSIS AS NOTED ABOVE. BASED ON THE INFORMATION PROVIDED IT IS UNKNOWN IF THE DEATH IN THIS CASE IS A RESULT OF THE BLEEDING OR SEPSIS REPORTED. HOWEVER, BASED UPON THE CURRENT KNOWN INFORMATION IT WAS DETERMINED THE DEATH AS UNLIKELY RELATED TO THE OLYMPUS DEVICE.
AZURION | K163715 | IT WAS REPORTED TO PHILIPS THAT DURING A STEMI (ST-ELEVATION MYOCARDIAL INFARCTION) PROCEDURE, IMAGING BECAME UNAVAILABLE. THE PATIENT WAS SUBSEQUENTLY MOVED TO ANOTHER ROOM. THE REPORT INDICATED THAT THE PATIENT PASSED AWAY; HOWEVER, DETAILED INFORMATION REGARDING THE DEATH HAS NOT BEEN PROVIDED TO DATE. AN INVESTIGATION INTO THIS COMPLAINT HAS BEEN INITIATED BY PHILIPS.
75MM SELECTABLE NEW TLC | K092577 | IT WAS REPORTED THAT POST OPERATION TO A COLOSTOMY CLOSING PROCEDURE THE PATIENT DIED 5 DAYS AFTER USING THE STAPLER AND TWO CHARGES. IN CONVERSATION WITH THE SURGEON, HE INFORMED ME THAT HE WAS A VERY SERIOUS PATIENT, WITH SEVERAL COMORBIDITIES AND CHRONIC RENAL. HOWEVER, THE SAME WAS REPROACHED APPROXIMATELY 5 DAYS AFTER SURGERY, AND IT WAS NOTICED THAT IN THE STAPLING LINE THERE WAS A DESCENT. | (B)(4), DATE SENT 6/9/2025, D4 BATCH # UNK. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AS PART OF OUR COMPANY QUALITY SYSTEM PROCESS, ALL DEVICES ARE MANUFACTURED, INSPECTED, AND DISTRIBUTED TO APPROVED SPECIFICATIONS. HOWEVER, IF THE PRODUCT IS RECEIVED AT A LATER DATE, THE INVESTIGATION WILL BE UPDATED AS APPLICABLE. THE MANUFACTURING RECORDS WERE REVIEWED AND THE MANUFACTURING/PACKAGING CRITERIA WERE MET PRIOR TO THE RELEASE OF THIS LOT. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT: HOW MANY NTLC¿S WERE USED? HOW MANY RELOADS WERE USED (WHAT WAS THE PRODUCT CODE)? WHAT WERE THE INDICATIONS FOR SURGERY? WHAT SURGICAL PROCEDURE WAS PERFORMED? WHAT COLOR SETTING WAS SELECTED ON THE DEVICE AND WHAT WAS THE SEQUENCE OF THE FIRINGS? WHAT ANATOMICAL STRUCTURES WAS THE DEVICE FIRED ON? WHAT DEVICE WAS USED TO CREATE THE COMMON CHANNEL? WHAT WAS USED TO CLOSE THE COMMON OPENING? WERE THERE ANY ISSUES EXPERIENCED WITH THE DEVICE IN THE INITIAL SURGICAL PROCEDURE? WAS THE DEVICE DIFFICULT TO FIRE? HOW WERE THE POST-OPERATIVE ISSUES IDENTIFIED? WERE THERE ANY ISSUES NOTED WITH STAPLE FORMATION AT ANY POINT THROUGHOUT THE CARE OF THE PATIENT? IF SO, PLEASE DESCRIBE. DOES THE SURGEON BELIEVE THAT THE ALLEGED DIFFICULTIES WITH THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENTS DEATH? IS THERE AN AUTOPSY REPORT AVAILABLE? IS THE SURGEON INTERESTED IN SPEAKING WITH ETHICON MEDICAL AND ENGINEERING STAFF? THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
MICRA | K132030 | IT WAS FURTHER NOTED THAT THE DEATH IS NOT RELATED TO THE LEADLESS IMPLANTABLE PULSE GENERATOR (IPG) OR PROCEDURE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
PATIENT INFORMATION CENTER IX | K211900 | IT WAS REPORTED NO ALARMS WERE GENERATED ON THE NURSE PHONES AND THE PATIENT PASSED AWAY. IT WAS INDICATED RED ALARMS WOULD BE DISPLAYED BRIEFLY ON THE NURSE PHONE BUT IT APPEARS THERE WAS NO AUDIBLE ALARM. | E1 REPORTING INSTITUTION PHONE#: (B)(6). IT WAS ALSO INDICATED THE DEVICE (ASCOM PHONE) WAS NOT USABLE AND A REPLACEMENT WAS REQUESTED. THE CUSTOMER ALLEGED THAT AN ASYSTOLE ALARM WAS MISSED. THE PATIENT UNDERWENT RESUSCITATION MEASURES, HOWEVER THE PATIENT PASSED AWAY. A PHILIPS PROJECT MANAGER AND APPLICATION SPECIALIST WERE CONTACTED. IT WAS NOTED THAT THE CONFIGURATION IS SET FOR DELAYED PHONE NOTIFICATION, HOWEVER THE ALLEGATION INDICATES THAT THE MESSAGE WAS RECEIVED (VERY BRIEFLY) AND DISAPPEARED. A PHILIPS REMOTE SERVICE ENGINEER (RSE) SPOKE WITH THE CUSTOMER AND NOTED THAT THE EVENT TOOK PLACE ON (B)(6) 2025 AT APPROXIMATELY 19:50 HOURS. THE ONSITE TECHNICIAN TRIGGERED A RED ALARM AND THIS WAS RECEIVED AT THE END DEVICE (MYCO 3 SMARTPHONE) WITH CAREASSIST 4.4.0. THE TEST WAS SUCCESSFUL. THE SOFTWARE WAS PROACTIVELY UPDATED TO VERSION 4.3.7. A PHILIPS FIELD SERVICE ENGINEER (FSE) WAS DISPATCHED AND THE PRELIMINARY INVESTIGATION AND FOUND THAT STARTING AT 19:55:07, AN ALARM WAS GENERATED AT THE PATIENT INFORMATION CENTER IX (PIC IX) INDICATING THAT THERE WAS NO MOBILE DEVICE AVAILABLE FOR ALERTING FOR THE AFFECTED BED LOCATION. FROM 19:55:15 AND 19:55:18, RED ALARMS WERE GENERATED ON THE MX40 MONITOR AND THE PICIX. THESE WERE ACKNOWLEDGED AT THE PICIX AT 19:55:28 AND 19:55:29, WHEREUPON THE ALARM WAS TERMINATED AT THE PICIX. ACCORDING TO THE CAREGIVER, THE NECESSARY ACTIONS WERE THEN IMMEDIATELY INITIATED FOR THE PATIENT. THE EARLY INVESTIGATION INDICATES THAT THE MOBILE PHONE WAS NOT IN USE. AS THIS IS CONTRADICTORY TO THE ORIGINAL ALLEGATION, QUALITY PERSONNEL ARE WORKING WITH THE FIELD ENGINEER TO CLARIFY THE EVENT. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A USER OF A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THE PATIENT'S WIFE REPORTS THE PATIENT WAS NOT USING THE DEVICE AT THE TIME OF PASSING. THE PATIENT HAS NOT USED THE DEVICE FOR A VERY LONG TIME. THE CAUSE OF DEATH WAS NOT PROVIDED. THERE IS NO ALLEGATION THAT THE DEVICE CAUSED THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REP DREAMSTATION AUTO CPAP, DOM - RECRT | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM UNO LEGAL LITIGATION IN REFERENCE TO A REPAIRED/RECERTIFIED DREAM STATION CPAP WITH AN ALLEGATION METASTATIC ESOPHAGEAL CARCINOMA. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
VALLEYLAB | K151649 | IT WAS REPORTED THAT AFTER AN OPEN TOTAL ABDOMINAL HYSTERECTOMY AND BILATERAL SALPINGO-OOPHORECTOMY, THE PATIENT GOT INTO HYPOVOLEMIC SHOCK THE FOLLOWING MORNING. THE BLEEDING VESSELS WERE IDENTIFIED WHEN THE PATIENT WAS TAKEN BACK TO THEATRE AS AN EMERGENCY. A MASSIVE TRANSFUSION PROTOCOL (MTP) WAS INITIATED, RESUSCITATED AND GAVE FLUIDS AND THEN TAKEN BY THE AMBULANCE CREW TO TRANSFER TO A DIFFERENT HOSPITAL WHERE THEY OPENED THE PATIENT UP AND PERFORMED HEMOSTASIS BUT THE PATIENT SUBSEQUENTLY DIED OF HYPOVOLEMIC SHOCK. AN INQUEST WAS MADE BY THE CORONER, AND IT WAS IDENTIFIED THAT THE PATIENT EXPERIENCED A DELAYED INTRA-ABDOMINAL BLEED FROM 2 OF THE 9 BLOOD VESSELS THAT WERE SEALED WITH THE LIGASURE HAD FAILED. | D10 CONCOMITANT PRODUCTS: UNKNOWN LIGASUR, UNKNOWN LIGASURE INSTRUMENT (LOT#: UNKNOWN). MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
EDWARDS ESHEATH PLUS INTRODUCER SHEATH | K200258 | AS REPORTED BY OUR GERMAN AFFILIATES, DURING IMPLANT OF A 23MM SAPIEN 3 ULTRA (S3U) VALVE IN AORTIC POSITION BY RIGHT TRANSFEMORAL APPROACH. DURING THE INSERTION OF THE COMMANDER WITH THE VALVE THROUGH THE ESHEATH+, THE VALVE COULD NOT ADVANCE. THE SYSTEM WAS REMOVED AND DISCARDED WITH NO PATIENT INJURY REPORTED AT THIS POINT. A NEW 23MM S3U VALVE WAS PREPARED AND A 16FR ESHEATH+ WAS USED THIS TIME. THE NEW 16FR ESHEATH+ WAS EASILY INSERTED BUT IT WAS ALSO DIFFICULT TO ADVANCE THE VALVE. AFTER A CONSIDERABLE EFFORT AND FORCE, THE VALVE WAS ABLE TO EXIT THE SHEATH AND THE VALVE WAS SUCCESSFULLY DEPLOYED. AFTER THE REMOVAL OF THE ESHEATH+, DESPITE THE USE OF TWO PROSTYLE SUTURES DEVICES AND 8F ANGIOSEAL, THE VESSEL WAS NOT TIGHTLY CLOSED. A NON-EDWARDS CROSSOVER SHEATH WAS INSERTED FROM THE LEFT TO THE RIGHT VESSEL. THEN, ANGIOGRAPHY REVEALED THAT THE RIGHT COMMON FEMORAL WAS PERFORATED AND A NON-EDWARDS BALLOON WAS ADVANCE TO CONTROL THE BLEEDING. THE PATIENT WAS THEN MOVED TO VASCULAR SURGERY AND THE VESSEL WAS RESECTED AND A LONG 6MM/ 4CM LONG PROSTHETIC INTERPOSITION WAS IMPLANTED. AFTER REPEATED ANGIOGRAPHY, THE EXTERNAL ILIAC ARTERY WAS SUB TOTALLY OCCLUDED BY A DISSECTED DUCT, AND A STENT WAS IMPLANTED. THE FINAL ANGIOGRAPHY SHOWED GOOD VISUALIZATION OF ALL VESSELS. THE PATIENT WAS MOVED TO ICU FOR MONITORING BUT PASSED AWAY DURING THE NIGHT DUE TO HEMORRHAGIC HYPOVOLEMIC SHOCK AFTER MASSIVE BLEEDING FROM THE COMMON FEMORAL ARTERY. | A SUPPLEMENTAL MDR IS BEING SUBMITTED DUE TO THE INVESTIGATION BEING COMPLETED. THE FOLLOWING SECTIONS OF THIS REPORT HAVE BEEN UPDATED: B4, G3, G6, H2, H6, H11. THE EVENT REPORTED IS AN ANTICIPATED IN THE RISK MANAGEMENT DOCUMENTATION FOR TRANSCATHETER HEART VALVE PROCEDURES. A PREVIOUS INVESTIGATION INTO THIS TYPE OF EVENT IS CAPTURED IN AN EDWARDS LIFESCIENCES TECHNICAL SUMMARY AND APPLIES TO THIS COMPLAINT. ADDITIONAL ASSESSMENT OF THE FAILURE MODE IS NOT REQUIRED AT THIS TIME. REVIEW OF PROVIDED 3MENSIO REVEALED THE FOLLOWING OBSERVATIONS: PATIENT'S RIGHT ACCESS VESSEL HAD PRESENCE OF TORTUOSITY, CALCIFICATION, AND UNDERSIZED DIMENSIONS. MINIMAL LUMINAL DIAMETER FOR 16F ESHEATH IS ' 6.0MM. DUE TO THE UNAVAILABILITY OF THE COMPLAINT DEVICE, ENGINEERING WAS UNABLE TO PERFORM ANY VISUAL INSPECTION, FUNCTIONAL TESTING, OR DIMENSIONAL ANALYSIS. PER THE TECHNICAL SUMMARY: THE IFU, CURRENT RISK MITIGATIONS INCLUDING DESIGN AND MANUFACTURING CONTROLS, AND TRAINING MANUALS HAVE BEEN REVIEWED, AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. THROUGH EXTENSIVE COMPLAINT INVESTIGATIONS, EVENTS REPORTING RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH AND POTENTIAL VALVE FRAME DAMAGE USING THE S3U/S3UR VALVE CONFIGURATION HAVE NOT BEEN ASSOCIATED WITH DEVICE MALFUNCTIONS OR MANUFACTURING NONCONFORMANCES. RATHER, THE ROOT CAUSE FOR THESE EVENTS HAVE HISTORICALLY BEEN DUE TO VESSEL TORTUOSITY, CALCIFICATION, UNDERSIZED VESSELS, AND/OR STEEP INSERTION ANGLE. IN ADDITION, VALVE FRAME AND/OR SHEATH DAMAGE CAN BE A RESULT OF INCREASED PUSH FORCE AND ANY EXCESSIVE DEVICE MANIPULATION OR SHEATH-VALVE INTERACTION. THE COMPLAINT FOR DIFFICULTY ADVANCING THROUGH SHEATH WAS CONFIRMED EMPIRICALLY BASED ON MEDICAL RECORD. AVAILABLE INFORMATION SUGGESTS THAT PATIENT FACTORS (CALCIFICATION, TORTUOSITY, UNDERSIZED VESSELS) MAY HAVE CONTRIBUTED TO THE REPORTED EVENT, AS CONFIRMED VIA 3MENSIO. CALCIFICATION CAN REDUCE THE VESSEL LUMEN DIAMETER AND MAY INCREASE RESTRICTION LEADING TO RESISTANCE. CALCIFICATION CAN ALSO RESULT IN THE CREATION OF SUB-OPTIMAL ANGLES DURING DELIVERY SYSTEM INSERTION THAT MAY LEAD TO RESISTANCE. SCRATCHES OBSERVED ON THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF CALCIFIED NODULES WITHIN THE ACCESS VESSEL. TORTUOUS PATIENT ANATOMY CAN CREATE SUB-OPTIMAL ANGLES THAT CAN LEAD TO NON-COAXIAL ALIGNMENT BETWEEN THE DELIVERY SYSTEM WITH CRIMPED VALVE AND SHEATH INNER LUMEN. KINKS OR CURVATURE ALONG THE SHEATH SHAFT CAN BE INDICATIVE OF THE PRESENCE OF VESSEL TORTUOSITY. UNDERSIZED VESSELS (E.G. DUE TO ANATOMICAL VARIATION, PRE-EXISTING STENT OR GRAFT, SCAR TISSUE, OR FIBROSIS) CAN CREATE A RESTRICTED PATHWAY OR CONSTRAINED CONDITION RESULTING IN DIFFICULTY DURING SHEATH EXPANSION AND INCREASED RESISTANCE DURING THE ADVANCEMENT OF THE DELIVERY SYSTEM. THE CAUSE OF THE RIGHT COMMON FEMORAL PERFORATION, WITH SUBSEQUENT BLEEDING AND DEATH IS UNKNOWN. HOWEVER, MAYBE RELATED TO THE TWO PROSTYLE SUTURES DEVICES USED FOR VESSEL CLOSURE OR RELATED TO THE RESISTANCE DURING DELIVERY SYSTEM INSERTION/ADVANCEMENT THROUGH THE SHEATH. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS ON A MONTHLY BASIS, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. SINCE NO EDWARDS PRODUCT DEFECTS OR LABELING DEFICIENCIES WERE IDENTIFIED, NO CORRECTIVE OR PREVENTATIVE ACTION IS REQUIRED.
GORE® TRI LUMEN CATHETER | K180919 | THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT USING A GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS, GORE® EXCLUDER® AAA ENDOPROSTHESIS, GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) IN ZONE 5 WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) DEVICES WERE USED AS BRANCH DEVICES IN THE VISCERAL ARTERIES. IT WAS REPORTED PRIOR TO THE PROCEDURE; THE PATIENT'S LEFT RENAL ARTERY WAS ALREADY OCCLUDED. DURING THE PROCEDURE, THERE WAS A DISTAL PERFORATION IN THE RIGHT RENAL DUE TO A 0.035" ROSEN GUIDEWIRE. THE PERFORATION WAS RESOLVED BY PLACEMENT OF A NON-GORE CORONARY COVERED STENT. THE PHYSICIAN HAD RIGHT SIDED ACCESS DUE TO THE LEFT SUBCLAVIAN BEING STENOTIC AND ALREADY HAVING AN UNKNOWN STENT IN PLACE FROM A PREVIOUS SURGERY. THE PHYSICIAN IMPLANTED ALL DEVICES WITHOUT ANY ISSUES. TWO VBX DEVICES WERE IMPLANTED IN THE SMA, THREE VBX DEVICES WERE IMPLANTED IN THE RIGHT RENAL, AND TWO VBX DEVICES WERE IMPLANTED IN THE CELIAC. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2025, IT WAS REPORTED THE PATIENT HAD A STROKE AND EXPIRED. THE PHYSICIAN STATED DISPLACEMENT OF ATHEROMA (PLAQUE) IN THE ARCH, MOVED TO THE PATIENT'S HEAD AND LEGS CAUSING THE STROKE. | ACCORDING TO THE GORE® TRI LUMEN CATHETER INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE STROKE AND DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, PATIENT UNDERWENT AN ENDOVASCULAR TREATMENT USING A GORE® EXCLUDER® ILIAC BRANCH ENDOPROSTHESIS, GORE® EXCLUDER® AAA ENDOPROSTHESIS, GORE® EXCLUDER® THORACOABDOMINAL BRANCH ENDOPROSTHESIS (TAMBE) IN ZONE 5 WHERE A GORE® VIABAHN® VBX BALLOON EXPANDABLE ENDOPROSTHESIS (VBX DEVICE) DEVICES WERE USED AS BRANCH DEVICES IN THE VISCERAL ARTERIES. IT WAS REPORTED PRIOR TO THE PROCEDURE; THE PATIENT'S LEFT RENAL ARTERY WAS ALREADY OCCLUDED. DURING THE PROCEDURE, THERE WAS A DISTAL PERFORATION IN THE RIGHT RENAL DUE TO A 0.035" ROSEN GUIDEWIRE. THE PERFORATION WAS RESOLVED BY PLACEMENT OF A NON-GORE CORONARY COVERED STENT. THE PHYSICIAN HAD RIGHT SIDED ACCESS DUE TO THE LEFT SUBCLAVIAN BEING STENOTIC AND ALREADY HAVING AN UNKNOWN STENT IN PLACE FROM A PREVIOUS SURGERY. THE PHYSICIAN IMPLANTED ALL DEVICES WITHOUT ANY ISSUES. TWO VBX DEVICES WERE IMPLANTED IN THE SMA, THREE VBX DEVICES WERE IMPLANTED IN THE RIGHT RENAL, AND TWO VBX DEVICES WERE IMPLANTED IN THE CELIAC. THE PATIENT TOLERATED THE PROCEDURE. ON (B)(6) 2025, IT WAS REPORTED THE PATIENT HAD A STROKE AND EXPIRED. THE PHYSICIAN STATED DISPLACEMENT OF ATHEROMA (PLAQUE) IN THE ARCH, MOVED TO THE PATIENT'S HEAD AND LEGS CAUSING THE STROKE. | H6: B22 AND C20 ¿ PRODUCT HISTORY REVIEW COULD NOT BE PERFORMED AS THE SERIAL# IS NOT AVAILABLE. ACCORDING TO THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE, ADVERSE EVENTS THAT MAY OCCUR AND / OR REQUIRE INTERVENTION INCLUDE DEATH. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
INTELLIVUE PATIENT MONITOR MX750 | K210906 | IT WAS REPORTED THE MONITOR DID NOT ALARM FOR SPO2 AND THE PATIENT PASSED AWAY. IN ADDITION, IT WAS NOTED THERE WERE NO ERRORS ON THE MONITOR OR CENTRAL STATION. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | A PHILIPS PRODUCT SUPPORT ENGINEER (PSE) REVIEWED PROVIDED ALARM LOGS AND AUDIT DATA FROM THE PHILIPS PATIENT INFORMATION CENTER IX (PIC IX) THAT WAS CONNECTED WITH THE MX750 MONITOR. THE PSE FOUND A LOT OF SPO2 GENERATED ALARMS IN THE LOGS FROM THE PIC IX. THE PSE NOTED THAT THERE WERE DESATURATION (DESAT) ALARMS AROUND 21:35, AND THEN AT 21:36:17 THE SPO2 ALARMS WERE SHUT OFF (FROM ON TO OFF). FOLLOWING THE TIMELINE, THE SPO2 WAS ALREADY BETWEEN 86 AND 80%, AND RELEASED BEFORE BEING SHUT OFF DESAT ALARMS (RED ALARM DESAT 82 < 84). THAT WAS CHANGED AT 22:10:29 SPO2; ALARMS FROM OFF TO ON, WHEN THE SPO2 WAS ALREADY TOO LOW BETWEEN 27% AND 23%, GENERATING AT 22:11:44 AND 22:12:37 DESAT 0<84 ALARMS (MEANING VERY CRITICAL LOW LEVEL OF SPO2); AFTER THAT, THE TIMELINE SHOWS LOWEST LEVEL SPO2 AND THE MONITOR GENERATED A LOT OF SPO2, NO PULSE, AND SPO2, AND "NOISY SIGNAL" ALARMS, AS WELL AS SPO2 SHUT FROM ON TO OFF, FOLLOWED BY AN SPO2 ALARMS FROM ON TO OFF. FURTHER INFORMATION WAS PROVIDED BY A PHILIPS REMOTE SERVICE ENGINEER (RSE) INDICATING THAT SPO2 ALARMS WERE TURNED OFF, AND PATIENT CARE GIVEN TO SOMEONE ELSE FOR BREAK RELIEF. THE PATIENT'S RESPIRATORY HEALTH DECLINED, AND ATTENTION DRAWN WHEN A RED ALARM FOR ARTERIAL BLOOD PRESURE (ABP) OCCURRED. THE PATIENT WAS DO NOT ATTEMPT RESUSCITATION (DNAR), BUT PATIENT WAS STABLE WAITING FOR A WARD BED. DUE TO LACK OF ALARM, HELP WAS NOT SOUGHT AT TIME OF INITIAL RESPIRATORY DECLINE. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS THE USERS TURNING OFF THE SPO2 ALARMS. THE REPORTED PROBLEM WAS NOT CONFIRMED. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
MAIN ASSEMBLY, VENTILATOR, COMMERCIAL, ZVENT, BASI | K162832 | COMPLAINANT ALLEGED THAT WHILE ATTEMPTING TO TREAT A PATIENT (AGE & GENDER UNKNOWN), THE DEVICE DISPLAYED A "SELF CHECK" ERROR MESSAGE. COMPLAINANT INDICATED THAT THE PATIENT SUBSEQUENTLY EXPIRED. | JUSTIFICATION FOR NO UDI: THIS DEVICE WAS MANUFACTURED BUT NOT DOMESTICALLY DISTRIBUTED; IT IS ONLY DISTRIBUTED IN THE INTENDED GEOGRAPHY. THERE IS NO EXISTING UDI REGULATION. ZOLL MEDICAL CORPORATION HAS NOT RECEIVED THE DEVICE FOR EVALUATION AND THIS COMPLAINT IS STILL UNDER INVESTIGATION.
VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | K022269 | SURGICAL OUTCOME AFTER AUTOLOGOUS BONE CHIPS REPLACEMENT IN DEPRESSED SKULL FRACTURES: A SINGLE CENTER EXPERIENCE. THE AIM OF THIS STUDY IS TO EVALUATE DIFFERENT PATTERNS OF OUTCOME INCLUDING (SAFETY, EFFICACY, COMPLICATIONS, PATIENTS¿ SATISFACTION AND NEED FOR LATE CP) OF USING AUTOLOGOUS BONE CHIPS FOR MANAGEMENT OF DEPRESSED SKULL FRACTURES (DSFS). BETWEEN MARCH 2023 AND NOVEMBER 2023, A TOTAL OF 68 TRAUMATIC BRAIN INJURY (TBI) PATIENTS WITH DSFS WERE OPERATED UPON IN OUR NEUROSURGERY DEPARTMENT. BASED ON OUR ELIGIBILITY CRITERIA, 59 PATIENTS FULFILLED OUR INCLUSION CRITERIA AND WERE ACTIVELY INCLUDED. AMONG OUR INCLUDED CASES, (89.8%) HAD MILD TBI AND THE REMAINING (10.2%) HAD MODERATE TBI. THE MEAN AGE WAS (16.505±12.42) YEARS WITH RANGE FROM 6 MONTHS TO 59 YEARS. THE VAST MAJORITY OF PATIENTS 41 CASES (69.5%) WERE =20 YEARS AND 18 PATIENTS (30.5%) WERE >20 YEARS. MALES CONSTITUTED THE MAIN BULK OF STUDY POPULATION (78%) AND FEMALE REPRESENTED ONLY (22%). RTA WAS THE MOST FREQUENT MOOD OF TRAUMA (35.6%) FOLLOWED BY ASSAULT (23.7%), DIRECT TRAUMA (23.47%) THEN: FFH (16.9%). VICRYL 0 (ETHICON) WAS USED TO MAKE A NET THROUGH EDGES OF THE GALEA THEN BONE FRAGMENTS WERE REPLACED AND FIXED USING THIS NET. REPORTED COMPLICATIONS ARE : VICRYL 0 (ETHICON) -SEIZURES (N=3.4%) TREATMENT: NOT REPORTED -CSF LEAKAGE (N=3.4%) TREATMENT : NOT REPORTED -WOUND INFECTION (N=1) TREATMENT : WE STARTED EMPIRIC BROAD-SPECTRUM ANTIBIOTIC REGIMEN WITH VANCOMYCIN (15 MG/KG IV EVERY 12 H) PLUS A THIRD A GENERATION CEPHALOSPORIN (CEFTRIAXONE 2 G IV DAILY) AND SHOWED GOOD RESPONSE TO THE ANTIBIOTIC TREATMENT VICRYL 0 (ETHICON) DEATH (N=1); THIS PATIENT HAD WOUND INFECTION AND DEVELOPED OSTEOMYELITIS TREATMENT: THE PATIENT HAD WOUND INFECTION WHICH WAS TREATED WITH EMPIRIC BROAD-SPECTRUM ANTIBIOTIC REGIMEN WITH VANCOMYCIN (15 MG/KG IV EVERY 12 H) PLUS A THIRD A GENERATION CEPHALOSPORIN (CEFTRIAXONE 2 G IV DAILY); DEVELOPED OSTEOMYELITIS AND UNDERWENT RE-SURGERY FOR WOUND DEBRIDEMENT AND REMOVAL OF BONE FRAGMENTS, LATER THIS CASE WAS EXPIRED. IN CONCLUSION, USING AUTOLOGOUS DEPRESSED BONE CHIPS FOR SKULL DEFECT RECONSTRUCTION CAN BE A SAFE AND FEASIBLE SURGICAL TECHNIQUE FOR TBI PATIENTS SUFFERING DSFS WITH GOOD AESTHETIC RESULTS, HIGH PATIENT SATISFACTION, DECREASED NEED FOR LATER CRANIOPLASTY AND CONSEQUENTLY LOW OVERALL MANAGEMENT COST. | PRODUCT COMPLAINT #: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: NEUROSURG REV. 2024 DEC 10;47(1):898. DOI: 10.1007/S10143-024-03128-Y. PMID: 39656311; PMCID: PMC11632004. HTTPS://DOI.ORG/10.1007/S10143-024-03128-Y.
DREAMSTATION | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER MAILING SENT OUT DUE TO A DEVICE RECALL. THE MANUFACTURER WAS INFORMED BY A FAMILY MEMBER THAT THE PATIENT WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DAVINCI XI | K131861 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | SECTION B5 ADDITIONAL INFORMATION: THE MOBILIZATION OF THE TUMOR WAS ALMOST COMPLETED WHEN THE BLEEDING BEGAN. A SYNCHROSEAL INSTRUMENT AND UNSPECIFIED BIPOLAR INSTRUMENT WERE BEING USED DURING THIS TASK. THE CUSTOMER SAID THERE WAS NO MALFUNCTION OF A DA VINCI PRODUCT DURING THIS EVENT. THE BLEEDING WAS COMING FROM THE LIVER VEIN, DISTAL TO THE TUMOR. CENTRALLY, THE VEIN WAS ADHERED TO THE TUMOR. THE DECISION TO CONVERT TO OPEN WAS MORE DUE TO THE VEIN BEING ADHERED TO THE TUMOR IN THE CENTRAL PART THAN THE BLEEDING. THE RESECTION OF THE TUMOR AND HEMOSTASIS WERE BOTH COMPLETED WHILE OPEN. THE BLOOD SUPPLY TO THE LIVER AND INTESTINES WERE WELL MAINTAINED AT ALL TIMES. THE VENA CAVA WAS WELL FILLED AT ALL TIMES, SO THE RETURN FLOW TO THE HEART WAS MORE VIGOROUS THAN INSUFFICIENT. THE CUSTOMER STATED THAT THE ORIGIN OF THE ASYSTOLE REMAINS UNEXPLAINED. A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT.
REMSTAR PRO C-FLEX+ | K091319 | THE REPORTER, WHO IS THE RELATIVE OF THE PATIENT, REPORTED THE PATIENT EXPIRED. THE REPORTER STATED THE PATIENT PREVIOUSLY USED; HOWEVER, THE PATIENT WAS NOT USING THE DEVICE AT THE TIME OF DEATH. THE PATIENT WAS SWITCHED TO A DIFFERENT (UNSPECIFIED) DEVICE. THE REPORTER HAD TO MAKE ANOTHER CALL AND WAS UNABLE TO ANSWER ANY FURTHER QUESTIONS. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THE CUSTOMER REQUESTED ALOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT BUT HARM WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE, THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR) NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | IT WAS REPORTED THAT THE CUSTOMER REQUESTED ALOG REVIEW TO KNOW WHAT WAS PROGRAMMED ON PATIENT ON (B)(6) 2025. IT WAS INITIALLY REPORTED THERE WAS PATIENT INVOLVEMENT BUT HARM WAS UNKNOWN. THROUGH DUE DILIGENCE ATTEMPTS, ADDITIONAL INFORMATION WAS RECEIVED. PATIENT WAS TAKEN FOR EMERGENCY SURGERY TO REPAIR A TRANSECTED POPLITEAL ARTERY AFTER SUFFERING A MECHANICAL FALL. POST-OPERATIVELY ON RETURN TO THE ICU, THE SURGEON ORDERED A FLUID BOLUS OF LACTATED RINGERS SOLUTION. THE CLINICIAN PROGRAMMED THE PUMP TO 999 ML/HOUR. UPON HEARING THE ALARM FOR AIR IN LINE, IT WAS DISCOVERED THAT THE CLINICIAN HAD PROGRAMMED THAT RATE OF 999ML/HR ON THE CHANNEL THAT HAD A FENTANYL DRIP RUNNING AT 1ML/HOUR. THE PATIENT RECEIVED THE ENTIRE BAG OF FENTANYL 2,500 MCG/100ML OVER APPROXIMATELY 20-30 MINUTES. NARCAN 2MG WAS ADMINISTERED IMMEDIATELY AFTER THE FENTANYL ADMINISTRATION. THE PATIENT WAS ALREADY ON A NEO SYNEPHRINE DRIP TO MAINTAIN BLOOD PRESSURE, THIS WAS TITRATED UP TO A MAX RATE OF 300 MCG/MINUTE. LEVOPHED DRIP WAS ADDED APPROXIMATELY 6 HOURS AFTER THE FENTANYL ADMINISTRATION. OVER THE NEXT 6 HOURS THE PATIENT¿S BLOOD PRESSURE RANGED FROM 83/46 MMHG TO 133/53MMHG. MEAN ARTERIAL PRESSURE RANGED FROM 34 MMHG TO 129 MMHG. CUSTOMER STATES THEY ARE UNABLE TO DETERMINE HOW MUCH OF THE BLOOD PRESSURE CHANGES WERE DUE TO THE FENTANYL VS POST-OPERATIVE HYPOTENSION AND KNOWN SIGNIFICANT INTRA-OP BLOOD LOSS DUE TO COAGULOPATHY, AND PRE-OP BLOOD LOSS FROM HEMATOMA. THE PATIENT CONTINUED TO EXPERIENCE POST-OP HYPOTENSION FROM LARGE VOLUME BLOOD LOSS, TREATED WITH BLOOD TRANSFUSIONS, FLUID BOLUSES AND IV PRESSORS. PATIENT ULTIMATELY SUFFERED MULTI-ORGAN FAILURE AND EXPIRED. ADDITIONAL MEDICATIONS INFUSING AT THE TIME INCLUDED: FENTANYL, 2500MCG IN 100ML NACL 0.9%, RUNNING AT 1MCG/HR (OR 1ML/HR) NEO-SYNEPHRINE, 50,000 MCG IN 250ML NACL 0.9%, RUNNING AT 60-300 MCG/MIN. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH DUE TO ALZHEIMER'S. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION, NO SECONDARY FINDINGS WAS OBSERVED. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTRE. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULD CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT CORRECTED.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. THE REPORTER HAS ALLEGED HER HUSBAND PASSED AWAY. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S PRODUCT INVESTIGATION LABORATORY FOR INVESTIGATION. DURING INVESTIGATION, THE MANUFACTURER OBSERVED AN UNKNOWN LIGHT CONTAMINATION AT THE AIR INLET AND THE BOTTOM OF THE BLOWER BOX. LIGHT BLACK SPECKS IN THE BOTTOM ENCLOSURE. UNKNOWN LIGHT WHITE DUST-LIKE CONTAMINATION WAS ON TOP AND BOTTOM OF THE BLOWER MOTOR AND THE BLOWER MOTOR SEAL. MISSING AIR INLET FILTER. DURING INVESTIGATION THE DEVICE'S EVENT LOGS WERE DOWNLOADED AND REVIEWED BY MANUFACTURER. THE MANUFACTURER FOUND NO ERROR CODES. THE MANUFACTURER APPLIED POWER SUPPLY AND POWER CORD TO THE DEVICE AND VERIFIED AIRFLOW AND CONFIRMED THAT THE DEVICE POWERS ON PROVIDES THERAPY. THE MANUFACTURER WAS NOT ABLE TO CONFIRM THE COMPLAINT. NO SPECIFIC COMPLAINT GIVEN. THE MANUFACTURER OBSERVED DUST CONTAMINATION IN THE DEVICE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH AND IS ALSO REQUESTING FOR A RETURN LABEL SENT TO THE MAILING ADDRESS DOCUMENTED ABOVE. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE MANUFACTURER'S SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE MANUFACTURER SERVICE CENTRE, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. DURING THE EVALUATION. THERE WAS 0 ERROR CODE FOUND. THE MANUFACTURER SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION.
CNS-2101 CENTRAL MONITOR | K223567 | THE CUSTOMER REPORTED THAT THE HEART RATE (HR) FOR A PATIENT DROPPED TO 30BPM; HOWEVER, THERE WERE NO ALARMS AT ALL AT THE CENTRAL NURSE'S STATION (CNS). | THE CUSTOMER REPORTED THAT THE HEART RATE (HR) FOR A PATIENT DROPPED TO 30BPM BETWEEN 1:00PM AND 1:45PM; HOWEVER, THERE WERE NO ALARMS AT ALL AT THE CENTRAL NURSE'S STATION (CNS). WHEN THEY CHECKED THE CNS, IT SHOWED THAT THE ALARMS WERE SUSPENDED. THE PATIENT EXPERIENCED BRADYCARDIA FOR APPROXIMATELY 1-2 MINUTES, WITH THE HEART RATE DROPPING INTO THE 30S. IT WAS DURING THIS TIME THEY DISCOVERED THE ALARMS WERE TURNED OFF AT THE BEDSIDE. THE PATIENT UNFORTUNATELY EXPIRED. THE CUSTOMER WANTS THE LOGS REVIEWED TO SEE IF ONE OF THE NURSES SUSPENDED OR CHANGED THE ALARMS DURING THIS TIME FRAME OR IF SOMETHING ELSE CAUSED THE ALARMS TO BE SUSPENDED. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THE FOLLOWING FIELD(S) CONTAIN NO INFORMATION (NI), AS ATTEMPTS TO OBTAIN THE INFORMATION WERE MADE, BUT NOT PROVIDED:
WATCHMAN TRUSTEER? ACCESS SYSTEM | K240018 | REPORTED VIA STUDY. IT WAS REPORTED THAT THE PATIENT DIED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WERE SELECTED TO BE USED. THE PHYSICIAN POSITIONED THE WAS IN THE LAA OF THE PATIENT AND THEN INSERTED THE WDS. THE CLOSURE DEVICE WAS DEPLOYED IN THE PATIENT BUT THE PHYSICIAN WAS UNABLE TO POSITION THE DEVICE PROPERLY IN THE LAA. THE WDS WAS PULLED BACK INTO THE RIGHT ATRIUM AND THE PHYSICIAN PERFORMED A SECOND TRANSSEPTAL PUNCTURE. SHORTLY, AFTER THIS WAS PERFORMED A PERICARDIAL EFFUSION WITH CARDIAC TAMPONADE WAS NOTED. THE PATIENTS BLOOD PRESSURE DROPPED WHICH REQUIRED VASOPRESSOR SUPPORT AND THE PATIENT RECEIVED A BLOOD TRANSFUSION. THE PROCEDURE WAS ENDED, AND THE PATIENT WAS BROUGHT TO SURGERY WHERE THEY UNDERWENT A PERICARDIAL WINDOW PROCEDURE WHICH WAS CONVERTED A FULL STERNOTOMY. THE PHYSICIAN USED SUTURE TO CLOSE A PERFORATION IN THE LAA OF THE PATIENT AND THE LAA WAS CLIPPED. THE PATIENT REMAINED IN THE HOSPITAL TO RECOVER FROM THIS EVENT. WHILE IN THE HOSPITAL THE PATIENT BECAME SEPTIC AND EXPERIENCED CARDIOGENIC SHOCK. EIGHT DAYS POST PROCEDURE THE PATIENT HAD AN INCREASED DEMAND FOR PRESSOR SUPPORT. NINE DAYS POST PROCEDURE THE PATIENT WAS EXPERIENCING PERSISTENT LACTIC ACIDOSIS, ACUTE KIDNEY INJURY RESPIRATORY FAILURE AND THEN MULTIPLE SYSTEM ORGAN FAILURES BEFORE THE PATIENT ULTIMATELY DIED. | H2 CORRECTION: THIS EVENT WAS ALSO REPORTED UNDER REPORT NUMBER 2124215-2025-34314. THIS COMPLAINT WILL BE CLOSED AS A DUPLICATE.
BIA-GEN | K182074 | COMPLAINT SUBMITTED THROUGH CLINICAL DATA COLLECTION. LIMITED INFORMATION PROVIDED. PATIENT WAS INVOLVED IN A CLINICAL STUDY FOR PSEUDOCLAUDICATION SYNDROME. INITIAL SURGERY WAS PERFORMED ON (B)(6) 2024 UTILIZING THE BIA-GEN BIOACTIVE MOLDABLE BONE GRAFT MATRIX. THE PRODUCT WAS PLACED IN VERTEBRAE LOCATIONS L2-L3-L4-L5. THERE WERE NO OTHER PRODUCTS USED IN CONJUNCTION WITH BIA-GEN DURING THE INITIAL SURGERY AND NO COMPLICATIONS NOTED. THE CLINICIAN NOTED THE PRDOUCT HAD HANDLED AND PERFORMED EXCELLENT. ON (B)(6) 2024, THE PATIENT WENT TO THE EMERGENCY DEPARTMENT FOR EMESIS, LEFT FOOT SWELLING AND EPIGASTRIC DISCOMFORT. PATIENT THEN BECAME PALE, UNRESPONSIVE, AND EXPIRED. CLINICIAN BELIEVES CAUSE OF DEATH TO BE CARDIAC ARREST, HYPOXIC ARREST, OR A MISSED MYOCARDIAL INFRACTION. CLINICIAN STATED THE SERIOUS ADVERSE EVENT IS UNRELATED TO THE BIA-GEN PRODUCT. NO FURTHER INFORMATION SURROUNDING PATIENT INFORMATION AND EVENT WILL BE PROVIDED.
PERCUFLEX URINARY DIVERSION STENT SET | K830803 | IT WAS REPORTED THAT THE FRACTURED SINGLE J STENT WAS UNABLE TO BE REMOVED AT THE TIME OF STENT CHANGE RESULTING IN RETAINED FRAGMENT WITH SUBSEQUENT SEPSIS, AND DEATH. | BLOCK H6: IMDRF DEVICE CODE A0401 CAPTURES THE REPORTABLE EVENT OF STENT SHAFT BREAK. IMDRF PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. IMDRF PATIENT CODE E2008 CAPTURES THE REPORTABLE EVENT OF FOREIGN BODY IN PATIENT. IMDRF IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF DEATH.
HT70 VENTILATOR | K111146 | IT WAS REPORTED THAT WHILE IN USE ON A CRITICALLY ILL PATIENT, THE HT70 VENTILATOR GENERATED AN ALARM, THE SCREEN WENT COMPLETELY DARK AND THE VENTILATION STOPPED. THE PATIENT HAD NO SPONTANEOUS BREATHING, WAS UNDER A DO NOT RESUSCITATE (DNR) ORDER, AND PASSED AWAY. | SECTION A PATIENT INFORMATION AND SECTION E INITIAL REPORTER CANNOT BE PROVIDED DUE TO JAPANESE PRIVACY REGULATIONS. SECTION E. JAPAN MEDTRONIC PERSONNEL REPORTED EVENT TO MANUFACTURER ON BEHALF OF THE CUSTOMER. H3: MEDTRONIC HAS NOT RECEIVED THE SUSPECT DEVICE/COMPONENT FROM THE CUSTOMER FOR EVALUATION NOR HAS THE DEVICE BEEN EVALUATED BY THE MEDTRONIC SERVICE ENGINEER. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CARDIOSAVE HYBRID, TYPE B PLUG | K112372 | IT WAS REPORTED BY THE DISTRIBUTOR THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) EXPERIENCED OVERHEATING AND POWER ON TEST FAILURE 3 DURING USE ON PATIENT. THE DEVICE WAS TURNED OFF, WAITED 10 MINUTES, AND TURNED ON AGAIN. HOWEVER, THE FAULT PERSISTED. THE PATIENT DIED LATER THAN 4 HOURS AFTER FAILURE. IT HAS NOT BEEN PROVEN THAT THE EQUIPMENT WAS RESPONSIBLE FOR THE PATIENT'S DEATH. | DUE TO CHARACTER LIMITATION IN E1, FULL INITIAL REPORTER NAME IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
MICRA | K132030 | THE PATIENT WAS REPORTED TO HAVE DIED FROM CARDIAC PERFORATION OF THE RIGHT VENTRICLE. AN ECHOCARDIOGRAM WAS PERFORMED TO CONFIRM THE PERFORATION, AFTER WHICH THE PATIENT CODED. PERICARDIOCENTESIS AND CARDIOPULMONARY RESUSCITATION WERE ALSO CONDUCTED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT WITH A DREAMSTATION AUTO CPAP DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. NO ADDITIONAL DETAILS WERE PROVIDED. THE DEVICE HAS NOT BEEN RETURNED FOR INVESTIGATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THAT HER HUSBAND PASSED AWAY AND THAT SHE NEEDED TO KNOW WHEN HE RECEIVED HIS REPLACEMENT DEVICE FOR BILLING PURPOSES FOR A LOANER DEVICE HE HAD USED. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
INTELLIVUE MX40 WLAN | K113125 | IT WAS REPORTED THE DEVICE DISPLAYED THE INCORRECT HEART RATE AND THE PATIENT PASSED AWAY. THE MX40 DEVICE MEASURED A HEART RATE OF 75BPM WHEN THE ACTUAL HEART RATE WAS 39BPM. RESUSCITATION WAS ATTEMPTED BUT THE PATIENT PASSED AWAY. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
REMSTAR PLUS W/C-FLEX | K131982 | THE MANUFACTURER RECEIVED INFORMATION FROM THE PATIENT'S SISTER STATING THAT THE PATIENT PASSED AWAY. SHE WANTED TO KNOW WHAT SHE COULD DO WITH THE REMSTAR PLUS W/ C-FLEX DEVICE. THERE WAS NO ALLEGATION OF MALFUNCTION OR THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE CALLER WAS DIRECTED TO A DME OR ELECTRONICS BIN TO DISPOSE OF THE DEVICE. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
PROTEK DUO VENOUS CANNULA | K140999 | LIVANOVA RECEIVED A MEDWATCH REPORT (B)(4) REPORT RELATED TO A PROTEK DUO CANNULA 29FR. PATIENT DEVELOPED A PROGRESSIVE DYSPNOEA FOLLOWING COVID INFECTION. A SURFACE ECHOCARDIOGRAM REVEALED PROGRESSION TO SEVERE MITRAL STENOSIS. TRANSOESOPHAGEAL ECHOCARDIOGRAM WAS CONFIRMATORY FOR SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 MMHG WITH SEVERE THICKENING OF THE POSTERIOR GREATER THAN ANTERIOR MITRAL VALVE LEAFLETS AND MITRAL ANNULAR CALCIFICATION. CARDIAC CATHETERIZATION CONFIRMED MODERATE AORTIC AND SEVERE MITRAL STENOSIS WITH MEAN GRADIENT 12 AND 35 MMHG, LEFT MAIN DISEASE EXTENDING INTO THE PROXIMAL LAD (LEFT ANTERIOR DESCENDING). PATIENT MET IN CONSULTATION WITH DR. AND HAS ELECTED TO UNDERGO SURGICAL INTERVENTION (REPLACEMENT OF IMPELLA WITH PROTEK DEVICE IN THE RIGHT VENTRICLE OF THE HEART). PATIENT WAS ADMITTED PREOPERATIVELY FOR COMPLETION OF PREOPERATIVE TESTING. PATIENT UNDERWENT THE SURGERY FOR REPLACEMENT OF IMPELLA (NON LIVANOVA DEVICE) WITH PROTEK DUO CANNULA IN THE RIGHT HEART. APPROXIMATELY ONE HOUR AFTER RETURN TO THE CVICU (CARDIOVASCULAR INTENSIVE CARE UNIT) AFTER DEVICE PLACEMENT, THE CHEST TUBE OUTPUT INCREASED DRAMATICALLY, AND PRESSOR REQUIREMENTS ALSO INCREASED. THE PATIENT WAS THEN GIVEN MULTIPLE MASSIVE TRANSFUSION PACKS WHILE THE FAMILY AND SURGICAL TEAM DETERMINED THE BEST PATH FORWARD, AND ULTIMATELY ELECTED TO HAVE THE PATIENT PROCEED WITH COMFORT MEASURES. TRANSFUSIONS WERE STOPPED, AND THE PATIENT PROMPTLY PASSED SECONDARY TO HYPOVOLEMIC SHOCK. PATIENT DEAD. | D.4. LOT NUMBER IS UNKNOWN. THIS INFORMATION WILL BE PROVIDED IN A SUPPLEMENTAL REPORT IF MADE AVAILABLE. AS THE LOT NUMBER IS UNKNOWN, UNIQUE DEVICE IDENTIFIER (UDI) NUMBER AND EXPIRY DATE ARE NOT AVAILABLE H.4. AS THE LOT NUMBER IS UNKNOWN, THE DEVICE MANUFACTURE DATE COULD NOT BE DETERMINED. THIS INFORMATION WILL BE PROVIDED IN A SUPPLEMENTAL REPORT IF MADE AVAILABLE. H11: LIVANOVA MANUFACTURES THE PROTEK DUO CANNULA. THE INCIDENT OCCURRED IN (B)(6). AS ADDITIONAL INFORMATION REPORTED IN THE RECEIVED MEDTWATCH REPORT: DUE TO PATIENT INCREASED CHEST TUBE OUTPUT (FIRST PRIMARILY FROM LEFT PLEURAL TUBE AT APPROXIMATELY 200CC/HR AND SUBSEQUENTLY FROM THE MEDIASTINAL TUBES AS WELL) A MASSIVE BLOOD TRANSFUSION WITH EQUAL PRBC (PACKED RED BLOOD CELLS), FFP (FRESH FROZEN PLASMA) AND PLATELETS WERE STARTED AND ESCALATION OF PRESSORS. IN ADDITION, TEE (TRANSESOPHAGEAL ECHOCARDIOGRAM) WAS DONE WHICH SHOWED THAT THE NEWLY PLACED PROTEK DUO CANNULA WAS IN THE RV (RIGHT VENTRICLE) WITH WHAT LOOKED LIKE AN RV INJURY PRESENT. AFTER SPEAKING WITH PATIENT'S FAMILY IT WAS AGREED NOT TO PROCEED WITH A FURTHER SURGERY. THEREFORE, ANY COAGULOPATHY WAS CORRECTED AND DESPITE THIS, THE BLEEDING CONTINUED. THEN, IT WAS AGREED TO PROCEED WITH COMFORT CARE MEASURES. THE PATIENT'S FOLLOWING CONDITION NECESSITATING CRITICAL CARE MANAGEMENT: ACIDOSIS, ACUTE BLOOD LOSS ANAEMIA, ACUTE KIDNEY INJURY, ARRHYTHMIA, CARDIOGENIC SHOCK, COAGULOPATHY, ELECTROLYTE ABNORMALITIES, HYPOTENSION, HYPOVOLEMIC SHOCK, LIVER FAILURE, MULTISYSTEM ORGAN FAILURE, RENAL FAILURE, THROMBOCYTOPENIA. THE CRITICAL CARE TREATMENT AND MANAGEMENT INCLUDED ALKALI THERAPY, ANTIBIOTICS, BLOOD PRODUCT ADMINISTRATION, CONTINUOUS RENAL REPLACEMENT THERAPY, DISCUSSION OF PATIENT'S CONDITION AND THERAPEUTIC OPTIONS WITH MEDICAL DECISION MAKERS, ELECTROLYTE AND NUTRITION MANAGEMENT, FLUID ADMINISTRATION, INVASIVE HEMODYNAMIC MONITORING, MANAGEMENT OF MECHANICAL VENTILATION, REVIEW OF MEDICAL RECORDS, SEDATION MANAGEMENT, TITRATION OF INOTROPE INFUSIONS, TITRATION OF VASOPRESSOR INFUSIONS, DISCUSSION WITH CONSULTING SERVICES, TITRATION OF ANTICOAGULANTS. AS PER LIVANOVA MEDICAL EXPERTS' OPINION, THE INJURY (BLEEDING) IS AT THE CANNULA INSERTION SITE. THIS MAY BE RELATED TO THE USED INSERTION TECHNIQUE OR TO CANNULA MIGRATION. HOWEVER, AT THE MOMENT THERE IS NO EVIDENCE, AND NO DEFINITIVE CONCLUSION CAN BE REACHED AROUND THE RELATIONSHIP BETWEEN INJURED TISSUE AND THE DEVICE. MOREOVER, BASED ON INFORMATION PROVIDED, ALSO THE SUCTION OF THE IMPELLA (NON LIVANOVA DEVICE) GENERATING EXCESSIVE NEGATIVE PRESSURE MAY HAVE CAUSED THE INJURY TO THE RIGHT VENTRICLE AS THE PROCEDURE WAS FOR IMPELLA DEVICE REPLACEMENT. REGARDING THE ULTIMATE CAUSE OF PATIENT DEATH, MEDICAL ASSESSMENT OUTCOME IS THAT IT WAS LIKELY CAUSED BY ALL OTHER SERIOUS PATIENT CONDITIONS (E.G. COVID; MYCOBACTERIA INFECTION; CORONARY HEART DISEASE; DIABETES; HEPATIC/RENAL DYSFUNCTION; HYPERTENSION; IMMUNO-COMPROMISED; MORBIDLY OBESE; OTHER) AND THAT MOST LIKELY THE BLEEDING WAS NOT THE MAIN CAUSE OF THE DEATH. INVESTIGATION IS ONGOING. IF ANY ADDITIONAL INFORMATION PERTINENT TO THE REPORTED EVENT IS RECEIVED, IT WILL BE PROVIDED IN A SUPPLEMENTAL REPORT.
MYOSURE TISSUE REMOVAL DEVICE | K142029 | IT WAS REPORTED THAT ON (B)(6) 2025, A PATIENT RECEIVED A MYOSURE XL PROCEDURE ON A 40 YEAR-OLD PATIENT, DURING THE PROCEDURE UP UNTIL 17 MINUTES EVERYTHING WAS DOING WELL AND TOTAL AMOUNT OF FLUID WAS 5 LITERS AND 3 LITERS SPILLED OUT OF THE PATIENT. THE HYSTEROSCOPE WAS REMOVED FROM THE PATIENT´S BODY AND THE PROCEDURE WAS ALMOST COMPLETE WHEN THE PATIENT SUDDENLY DEVELOPED ARRHYTHMIA AND AFTER 25 MINUTES OF RESUSCITATION FROM THE TEAM , THE PATIENT WAS TRANSFERRED TO THE ICU BUT DIED THE NEXT DAY. ECHOCARDIOGRAPHY WHICH DID NOT REVEAL ANYTHING. AN ABDOMINAL ULTRASOUND WAS PERFORMED AND THERE WERE NO PROBLEMS OBSERVED. PATIENT HAD A PREVIOUS HISTORY OF A MYOMECTOMY WITH RESECTOSCOPE A MONTH PRIOR TO THE PROCEDURE. THE PHYSICIAN BELIEVED THAT THE OPERATION WAS WELL PERFORMED AND THAT THE PATIENT´S DEATH HAD NOTHING TO DO WITH THE MYOSURE. THE FLUID MANAGEMENT SYSTEM WAS NOT A HOLOGIC PRODUCT. NO OTHER INFORMATION COULD BE OBTAINED. | HOLOGIC IS SUBMITTING THIS REPORT IN ACCORDANCE WITH 21. CFR PART 803. THE SUBMISSION IS BASED UPON THE CURRENTLY AVAILABLE INFORMATION. THE SUBMISSION OF THIS REPORT DOES NOT REPRESENT A CONFIRMATION OF DEVICE MALFUNCTION INVOLVING THE HOLOGIC PRODUCTS IN THE EVENT DESCRIBED. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED FOR THE REPORTED LOT/SERIAL NUMBER. THE DEVICE WAS RELEASED MEETING ALL QA SPECIFICATIONS. THE DEVICE INVOLVED IN THIS EVENT WAS NOT RETURNED FOR EVALUATION PURPOSES THEREFORE VISUAL AND FUNCTIONAL ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED. WE ARE UNABLE TO CONFIRM A RELATIONSHIP BETWEEN THE DEVICE AND THE ISSUE REPORTED. THE INFORMATION OBTAINED DURING COMPLAINT INVESTIGATION WILL BE INCLUDED IN OUR GLOBAL COMPLAINT TRENDING AND PRODUCT SURVEILLANCE WILL CONTINUE TO MONITOR COMPLAINTS OF THIS TYPE FOR ADVERSE TRENDS. IF THE PRODUCT IS RECEIVED OR ADDITIONAL INFORMATION IS OBTAINED, THE INVESTIGATION WILL BE REOPENED ACCORDINGLY PER STANDARD OPERATING PROCEDURE.
WATCHMAN TRUSTEER? ACCESS SYSTEM | K240018 | IT WAS REPORTED A PERICARDIAL EFFUSION (PE) WAS NOTED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. A WATCHMAN ACCESS SYSTEM (WAS) AND A 24MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. A VERSACROSS CONNECT LAAC KIT WAS SELECTED FOR THE TRANSEPTAL PUNCTURE. IT WAS A CHALLENGING SEPTUM TO CANNULATE. THE VERSACROSS WIRE WAS SEEN TENTING SEPTUM, HOWEVER ONCE THE WAS SHEATH WAS BROUGHT UP, TENTING WAS LOST. AFTER 18 MINUTES OF ATTEMPTING, THE WIRE WAS ACROSS. THE WIRE WAS LOCATED IN THE LEFT ATRIUM HOVERING OVER AORTIC VALVE. THE PHYSICIAN WAS MADE AWARE, AND SEVERAL EFFUSION SCANS DID NOT NOTE ANY CHANGES. THE PHYSICIAN CONTINUED WITH PROCEDURE. THE PATIENT BLOOD PRESSURE WAS DECLINING. THE SCAN FOR PERICARDIAL EFFUSION (PE) WAS NEGATIVE, SO THE PROCEDURE CONTINUED. WHEN THE NON-BOSTON SCIENTIFIC PIGTAIL CATHETER WAS BEING INSERTED INTO LAA, THERE WAS A CHANGE IN THE ANATOMY OF LAA AND AN EFFUSION WAS NOTED IN TRANSVERSE SINUS. THE PATIENT WAS TAPPED TO TREAT THE EFFUSION. THE PATIENT STABILIZED AND THE WATCHMAN CLOSURE DEVICE WAS IMPLANTED. ADDITIONAL BLOOD WAS DRAINED AND THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE UNIT (ICU) FOR OBSERVATION. THE PATIENT REMAINED STABLE AND WAS DISCHARGED 48 HOURS LATER. IT WAS FURTHER REPORTED THAT THE PATIENT PASSED AWAY APPROXIMATELY NINE (9) DAYS POST PROCEDURE, (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. | B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. H6: IMPACT CODE ADDED.
OMNIPOD 5 CONTROLLER | K203768 | THE PATIENT'S PHYSICIAN REPORTED THE PATIENT PASSED AWAY WHILE WEARING A POD. THE CAUSE OF DEATH WAS REPORTED TO BE UNKNOWN AND UNDER INVESTIGATION. ON A SUBSEQUENT COMMUNICATION WITH THE PATIENT'S MOTHER, IT WAS REPORTED THE PATIENT'S DEATH WAS GLUCOSE RELATED, AND SHE ALLEGED IT WAS DUE TO A DEVICE MALFUNCTION. NO ADDITIONAL INFORMATION WAS PROVIDED. IF FURTHER INFORMATION IS RECEIVED AT A LATER DATE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. | THE PHYSICAL DEVICE WAS NOT RETURNED FOR INVESTIGATION. CLOUD DATA FROM THE USER FROM (B)(6) 2025 WAS REVIEWED. THE LAST PATIENT HISTORY BUFFER (PHB) DATA POINT WAS CREATED AT 06:38:41 ON (B)(6) 2025. THE DATA SHOWED THAT THE SYSTEM WAS USED PRIMARILY IN AUTOMATED MODE AND WHILE IN AUTOMATED MODE THE AUTOMATED ALGORITHM WAS ABLE TO APPROPRIATELY ADAPT THE MICROBOLUS AMOUNTS BASED ON THE ESTIMATED GLUCOSE VALUES RECEIVED BY THE POD FROM THE SENSOR AND THE USER INPUTS. THE LAST ESTIMATED GLUCOSE VALUE RECEIVED BY THE POD WAS 106 MG/DL AT 01:08:36 ON (B)(6) 2025. ABERRANT VALUES WERE GENERATED UNTIL 02:53:36 ON (B)(6) 2025 WHEN THE SENSOR BECAME INACTIVE. THE SYSTEM RESPONDED APPROPRIATELY, PUTTING THE SYSTEM INTO AUTOMATED MODE LIMITED AFTER FOUR MISSING CYCLES UNTIL THE SYSTEM MODE WAS SWITCHED TO MANUAL MODE AT 02:08:36 ON (B)(6) 2025. THE SYSTEM WAS USED IN MANUAL MODE WITH THE POD RECEIVING NO BLOOD GLUCOSE VALUES THROUGH THE LAST DATA POINT GENERATED AT 06:38:41 ON (B)(6) 2025. THE PHB DATA SHOWED THAT THE SYSTEM WAS CORRECTLY DELIVERING THE USER'S MANUAL BASAL PROGRAM WHILE THE SYSTEM WAS IN MANUAL MODE. NO EVIDENCE OF ISSUES WITH THE SYSTEM WERE OBSERVED IN THE AVAILABLE DATA. LOCKED DOWN SMARTPHONE: LOCKDOWN. OMNIPOD SOFTWARE APP VERSION: 1.2.4. SMARTPHONE OPERATING SYSTEM: N5004L-AM-Q-MV01602-06-01.06. SMARTPHONE HARDWARE: N5004L. CGM SENSOR TYPE: UNSPECIFIED. PLEASE NOTE, THAT SECTION D IS CAPTURING THE DEVICE IDENTIFIERS AS REPORTED BY THE COMPLAINANT. THIS MAY NOT ALIGN TO THE DEVICE CONFIGURATION REPORTED IN H11, AS THIS DATA IS PULLED FROM OUR CLOUD BASED ON THE REPORTED DATE OF EVENT.
ION | K182188 | IT WAS REPORTED THAT AFTER AN ION ENDOLUMINAL LUNG BIOPSY PROCEDURE A PATIENT EXPIRED. THE BIOPSY WAS PERFORMED ON A METASTATIC MASS IN THE LEFT LUNG (PROBABLY LEFT UPPER LOBE). THE PROCEDURE WAS COMPLETED WITH NO MALFUNCTIONS OR ISSUES; HOWEVER, AT THE END OF THE PROCEDURE, WHEN THE PHYSICIAN WAS UNDOCKING THE ROBOT, BLEEDING WAS NOTICED FROM THE LUNG. THE PATIENT¿S HEART RATE THEN DROPPED TO 38BPM, ¿AND ESSENTIALLY THE PATIENT CRASHED AND SUBSEQUENTLY PASSED AWAY.¿ NO FURTHER DETAILS WERE PROVIDED. THE PHYSICIAN SPECULATED THAT THE PATIENT MAY HAVE EXPERIENCED A PULMONARY EMBOLUS. ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION WERE MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | A SYSTEM LOG REVIEW CANNOT BE PERFORMED BECAUSE THE SYSTEM LOGS ARE NOT AVAILABLE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT UNDERWENT AN ION LUNG BIOPSY FOR METASTATIC DISEASE TO THE LUNG. THE PROCEDURE WAS COMPLETED WITHOUT ISSUE. AS THE ION SYSTEM WAS UNDOCKED BLEEDING WAS NOTED ASSOCIATED WITH BRADYCARDIA INTO THE 30S. THE PATIENT DETERIORATED AND SUBSEQUENTLY DIED. THERE WAS NO MALFUNCTION OF THE ION SYSTEM, INSTRUMENTS OR ACCESSORIES. BASED ON THE AVAILABLE DATA THE EVENTS WERE LIKELY PROCEDURE RELATED AND NOT DEVICE RELATED. BRONCHOSCOPY IS A MINIMALLY INVASIVE PROCEDURE WITH A LOW RISK PROFILE FOR SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 21 EPISODES OF HEMOPTYSIS OF > 50 ML (0.38%) AND 19 EPISODES < 50 ML (0.34%). A PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 NAVIGATIONAL BRONCHOSCOPY CASES REPORTED A BLEEDING RATE OF 2.5% OVERALL AND A CTCAE GRADE 2 OR GREATER BLEEDING RATE OF 1.5%. A SINGLE CENTER RETROSPECTIVE REVIEW OF 19,017 BRONCHOSCOPIC BIOPSIES REPORTED A SEVERE BLEEDING RATE OF 0.79% WITH A HIGHER RATE IN MORE CENTRAL LESIONS. ANOTHER LARGE CASE SERIES OF 26,895 BRONCHOSCOPIES WITH BIOPSIES OF ANY TYPE REPORTED ANY BLEEDING IN 41.4% OF CASES BUT A SEVERE BLEEDING RATE OF 1.47%. A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH. BLEEDING OF ANY SEVERITY WAS REPORTED IN 2.1% OF ALL CASES. THESE DATA REFLECT VARIOUS DEFINITIONS OF ANY BLEEDING IN THE LITERATURE RANGING FROM 0.72-41.4% WITH POSSIBLY MORE CONSISTENT DEFINITIONS OF MORE SEVERE CLINICALLY SIGNIFICANT BLEEDING RANGING FROM 0.38-1.47%. THE RATE OF DEATHS ASSOCIATED WITH BLEEDING WOULD BE SOME FRACTION OF THE REPORTED BLEEDING RATES AND WILL THUS BE MORE RARE THAN SIGNIFICANT BLEEDING. A RETROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES REPORTED 4 TOTAL DEATHS (0.02%). ANOTHER PROSPECTIVE MULTICENTER INTERNATIONAL STUDY OF 1,215 REPORTED 1 DEATH (0.08%). A RECENT META-ANALYSIS OF NAVIGATIONAL BRONCHOSCOPY IN 10,381 PATIENTS REPORTED AN OVERALL ADVERSE EVENT RATE OF 5.6% WITH 1 DEATH (0.01%). A CASE SERIES OF ION ROBOTIC ASSISTED BRONCHOSCOPIES PUBLISHED AFTER THE META-ANALYSIS INCLUDING 415 CASES REPORTED NO DEATHS. FOLCH EE, PRITCHETT MA, NEAD MA, ET AL. ELECTROMAGNETIC NAVIGATION BRONCHOSCOPY FOR PERIPHERAL PULMONARY LESIONS: ONE-YEAR RESULTS OF THE PROSPECTIVE, MULTICENTER NAVIGATE STUDY. JOURNAL OF THORACIC ONCOLOGY. 2019. FACCIOLONGO N, PATELLI M, GASPARINI S, ET AL. INCIDENCE OF COMPLICATIONS IN BRONCHOSCOPY. MULTICENTRE PROSPECTIVE STUDY OF 20,986 BRONCHOSCOPIES. MONALDI ARCHIVES FOR CHEST DISEASE. 2009. KOPS SEP, HEUS P, KOREVAAR DA, ET AL. DIAGNOSTIC YIELD AND SAFETY OF NAVIGATION BRONCHOSCOPY: A SYSTEMATIC REVIEW AND META-ANALYSIS. LUNG CANCER. 2023. BO L, SHI L, JIN F, LI C. THE HEMORRHAGE RISK OF PATIENTS UNDERGOING BRONCHOSCOPIC EXAMINATIONS OR TREATMENTS. AM J TRANSL RES. 2021. BROWNLEE AR, WATSON JJJ, AKHMEROV A, ET AL. ROBOTIC NAVIGATIONAL BRONCHOSCOPY IN A THORACIC SURGICAL PRACTICE: LEVERAGING TECHNOLOGY IN THE MANAGEMENT OF PULMONARY NODULES. JTCVS. 2023.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING AN ISSUE RELATED TO A DREAMSTATION AUTO BIPAP DEVICE. THE PATIENT HAS PASSED AWAY. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
AVALON FM50 FETAL MONITOR | K140535 | PHILIPS RECEIVED A COMPLAINT ON AN AVALON FM50 FETAL MONITOR INDICATING THAT THE FETAL HEART RATE (FHR) WAS NOT ABLE TO BE MEASURED AND THE BABY PASSED AWAY AFTER BIRTH. THE MOTHER AT 37 WEEKS PLUS 4 DAYS GESTATION PRESENTED WITH AN INTERMITTENT CATEGORY 2 TRACING (NEITHER REASSURING NOR NON-REASSURING TRACE) WITH IMPROVEMENT OVER 24 HOURS WITH POSITION CHANGES. THE PATIENT THEN HAD A QUICK CHANGE TO A CATEGORY 3 TRACING (PROLONGED DECELERATIONS) AND WAS TRANSFERRED EMERGENTLY TO THE OR, WHERE THE FETAL HEART RATE IMPROVED SO A VACUUM-ASSISTED DELIVERY WAS ATTEMPTED, WHICH WAS UNSUCCESSFUL, AND THE MOTHER WAS RELUCTANT TO GO TO C-SECTION. DURING THESE ATTEMPTS THE FHR WAS LOST FOR MORE THAN 8 MINUTES SO AN EMERGENT C-SECTION WAS PERFORMED. THE BABY WAS CYANOTIC AND UNRESPONSIVE AND REQUIRED RESUSCITATION. EMERGENT INTUBATION AND TRANSFER TO NICU AND ACTIVE COOLING WAS ATTEMPTED. CARE WAS SUBSEQUENTLY WITHDRAWN AND DEATH WAS DUE TO SEVERE HYPOXIC-ISCHEMIC ENCEPHALOPATHY. THE MATERNAL HEART RATE (MHR) WAS MONITORED AT THE TIME, AND IT WAS THOUGHT BECAUSE THE MHR AND FHR WERE BEING PICKED UP INTERCHANGEABLY, THERE WAS A DELAY IN CARE, WHICH AFFECTED THE OUTCOME. BASED ON THIS INFORMATION, THE CUSTOMER BELIEVES THE DEVICE MAY HAVE CONTRIBUTED TO THE REPORTED EVENT. A PHILIPS TECHNICAL CONSULTANT (TC) WENT TO THE CUSTOMER SITE, AND WAS UNABLE TO DUPLICATE THE ISSUE THAT THE AUDIBLE HEART RATE WAS HALF THE RATE THAT SHOWED ON THE MONITOR DISPLAY. THE TC NOTED THAT THE CLINICIANS HAD DIFFICULTY DISCERNING THE MATERNAL AND FETAL HR. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
SHILEY | K150844 | IT WAS REPORTED THAT AT A HOSPITAL, THE 8.5MM AND 7.0MM TRACHEOSTOMY TUBES HAD A LEAKING CUFF DURING THE PROCEDURE BASED ON THE VIDEO PROVIDED. THERE WAS NO REPORTED PATIENT OUTCOME. | D10 CONCOMITANT PRODUCT: 8CN85R , 8CN85R 8.5MM ADT FLEX TRACH W TG CUFF X1, LOT# 202405168X MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
WATCHMAN FXD CURVE? ACCESS SYSTEM | K212228 | IT WAS REPORTED AIR EMBOLISM AND PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAA) PROCEDURE WAS BEING PERFORMED UNDER MONITORED ANESTHESIA CARE (MAC) AND INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING. A VERSACROSS CONNECT (VCC) LAAC KIT WAS SELECTED FOR USE. THE PHYSICIAN PERFORMED A SUCCESSFUL TRANSSEPTAL PUNCTURE (TSP) USING THE WATCHMAN FXD CURVE ACCESS SYSTEM (WAS). WHEN REMOVING THE VCC DILATOR, IT WAS NOTED THAT THERE WAS NO BLEED-BACK. THE END OF THE WAS WAS LIFTED UP AND HEARD A VACUUM/SUCKING SOUND, SO THE PHYSICIAN QUICKLY PLACED THE DILATOR BACK DOWN ON THE TABLE, LIFTED END OF THE WAS UP AGAIN AND HEARD THE SAME SUCKING SOUND AGAIN, QUICKLY LOWERED THE WAS DOWN. THIS IS WHEN IT WAS NOTED AIR BUBBLES ON THE ICE IMAGING, AND THE PATIENT'S HEART RATED DROPPED. THE PROCEDURE WAS ABORTED. TRANSCUTANEOUS PACING WAS INITIATED, PATIENT WAS INTUBATED, FLUIDS GIVEN AND RESUSCITATION EFFORTS STARTED ONCE PATIENT CONTINUED TO DECLINE. THE PATIENT WAS UNABLE TO BE SUCCESSFULLY RESUSCITATED. THE PATIENT WAS PRONOUNCED DEAD APPROXIMATELY AT 9:30 AM. IN THE PHYSICIAN'S OPINION, IT WAS UNCERTAIN IF THE PATIENT POSSIBLY HAD A VERY LOW OR NEGATIVE LEFT ATRIUM PRESSURE OR POSSIBLY TOOK A BREATH DURING MAC SEDATION CAUSING NEGATIVE PRESSURE RESULTING IN NO BLEED-BACK FROM THE WAS AFTER THE VCC DILATOR WAS REMOVED. THE PHYSICIANS WERE CERTAIN THE WAS/VCC SYSTEM WERE FLUSHED PROPERLY PRIOR TO INSERTION. THE WAS WAS INSPECTED AFTER THE EVENT AND APPEARED TO BE INTACT AND WORKING PROPERLY. AIR ENTRY BELIEVED TO BE AT THE PORT OF THE WAS RESULTING IN AIR IN THE LEFT SIDE OF THE HEART. THE VALVE WAS CLOSED AT THE END OF THE WAS AND ASPIRATION WAS FROM THE SIDE PORT. AIR INGRESS SPECIFICALLY TO THE LEFT SIDE OF THE HEART MAY HAVE OCCURRED WHEN THE VCC DILATOR WAS REMOVED, ULTIMATELY CAUSING CARDIAC ARREST. THE OFFICIAL CAUSE OF DEATH WAS AIR EMBOLISM RESULTING IN CARDIAC ARREST.
PATIENT INFORMATION CENTER IX | K153702 | IT WAS REPORTED THE CUSTOMER REQUESTED ASSISTANCE IN OBTAINING CLINICAL AUDIT LOGS FROM THE CENTRAL STATION FOLLOWING A PATIENT DEATH. THE CUSTOMER WANTED TO DETERMINE WHETHER THE DEVICE ALARMED APPROPRIATELY. THE CUSTOMER INDICATED THE PATIENT HAD ASPIRATED AT SOME POINT, WHICH LED TO DEATH. A LOG REVIEW WAS PERFORMED BY THE CUSTOMER, REVEALING THE ALARMS WERE GENERATED AS EXPECTED. THERE WERE MULTIPLE ALARMS WHICH WERE LATCHED BUT NOT ACKNOWLEDGED, WHICH MASKED LOWER-LEVEL ALARMS. THE LOG REVIEW PROVIDED ANSWERS TO THE CUSTOMER REQUEST. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | A PHILIPS PRODUCT SUPPORT ENGINEERS (PSE) REVIEWED THE LOGS AND DETERMINED THAT THEY REFLECT THAT THE ALARMS FOR THE BED IN QUESTION DURING THE TIME PERIOD OF INTEREST WERE NOTED. A PHILIPS CLINICAL APPLICATIONS SPECIALIST (CAS) WENT ON SITE CONFIRMED THAT THE UNIT IS WORKING AS EXPECTED, THE CLINICAL AUDIT LOGS WERE ATTACHED, AND ALARMS WERE PRESENT AS EXPECTED WITH LATCHED AND NON-LATCHED ALARMS. BASED ON THIS INFORMATION, THERE WAS NO MALFUNCTION OF THE DEVICE. IN ADDITION, THERE WAS NO ALLEGATION THE DEVICE CAUSED OR CONTRIBUTED TO THE DEATH. THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
NAVIGATOR HD | K140323 | IT WAS REPORTED THAT A NAVIGATOR ACCESS SHEATH DEVICE WAS USED DURING A URETEROSCOPY PROCEDURE. DURING THE PROCEDURE, THE NAVIGATOR ACCESS SHEATH "BROKE APART WHILE INSIDE THE PATIENT." REPORTEDLY, THE FRAGMENTS WERE REMOVED USING A GRASPER. NINE DAYS LATER, THE PATIENT "CALLED 911 COMPLAINING OF SEVERE ABDOMINAL PAIN, NAUSEA, INCLUDING ABDOMINAL DISTENSION AND FEELINGS OF CONSTIPATION." REPORTEDLY, THESE FEELINGS HAD BEEN "PRESENT SINCE HER DISCHARGE" FOLLOWING THE URETEROSCOPY PROCEDURE. UPON ARRIVAL AT THE EMERGENCY ROOM, IT WAS NOTED THE PATIENT WAS EXPERIENCING HYPOTENSION AND ATRIAL FIBRILLATION. OVER THE NEXT FIVE DAYS THE PATIENT'S "ABDOMINAL COMPLICATIONS WORSENED RAPIDLY." IT WAS NOTED THAT THE PATIENT WAS HAVING "PERITONITIS AND LIKELY AN ISCHEMIC OR INFECTIOUS BOWEL." IT WAS ALSO NOTED THAT THE URINE INPUT WAS SEVERELY LIMITED AND WAS ULTIMATELY DIAGNOSED WITH SEPSIS, SEPTIC SHOCK, AND LACTIC ACIDOSIS. AFTER SEVERAL DAYS, THE PATIENT WAS "UNRESPONSIVE IN HER HOSPITAL BED AND PRONOUNCED DEAD." IT WAS FURTHER REPORTED THAT AN AUTOPSY NOTED "HYPERTENSIVE AND ATHEROSCLEROTIC CARDIOVASCULAR DISEASE AND "COMPLICATIONS [FROM] URETEROSCOPY." FURTHER, IT WAS REPORTED THAT "THE PROCEDURE WAS COMPLICATED BY DISINTEGRATION OF THE URETEROSCOPE COVER AND SEVERE POSTOPERATIVE HYPOTENSION." | BLOCKS D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK H2: ADDITIONAL INFORMATION: BLOCK H6 (PATIENT CODES, DEVICE CODES) BLOCK A6 (RACE). BLOCK H6: DEVICE CODE A0414 CAPTURES THE REPORTABLE EVENT OF SHEATH TORN. PATIENT CODE E1002 CAPTURES THE REPORTABLE EVENT OF ABDOMINAL PAIN. PATIENT CODE E1007 CAPTURES THE REPORTABLE EVENT OF CONSTIPATION. PATIENT CODE E060102 CAPTURES THE REPORTABLE EVENT OF ATRIAL FIBRILLATION. PATIENT CODE E1024 CAPTURES THE REPORTABLE EVENT OF PERITONITIS. PATIENT CODE E0509 CAPTURES THE REPORTABLE EVENT OF ISCHEMIA. PATIENT CODE E1309 CAPTURES THE REPORTABLE EVENT OF URINARY RETENTION. PATIENT CODE E0306 CAPTURES THE REPORTABLE EVENT OF SEPSIS. PATIENT CODE E233605 CAPTURES THE REPORTABLE EVENT OF SEPSIS SHOCK. PATIENT CODE E2321 CAPTURES THE REPORTABLE EVENT OF HYPOTENSION. IMPACT CODE F23 CAPTURES THE REPORTABLE EVENT OF UNEXPECTED MEDICAL INTERVENTION. IMPACT CODE F02 CAPTURES THE REPORTABLE EVENT OF PATIENT DEATH. IMPACT CODE F1908 CAPTURES THE REPORTABLE EVENT OF PROLONGED SURGERY. BLOCK H11: CORRECTION TO FIELD D6A (IMPLANT DATE).
PROLENE POLYPROPYLENE SUTURE UNKNOWN PRODUCT | K133356 | TITLE: EFFICACY OF PROLENE VS VICRYL IN TRACHEAL RESECTION AND ANASTOMOSIS. THE AIM OF THIS STUDY IS TO COMPARE THE INCIDENCE OF MORBIDITY/MORTALITY DUE TO LEAKS AND LONG-TERM RECURRENCE OF TRACHEAL STENOSIS. BETWEEN JANUARY 2022 AND JUNE 2023, THERE WAS A TOTAL OF 66 CASES THAT REQUIRED TRACHEAL RESECTION AND ANASTOMOSIS. AMONG THEM, 32 PATIENTS UNDERWENT TRACHEAL RESECTION WITH PROLENE (GROUP P), AND THE REMAINING 34 UNDERWENT TRACHEAL RESECTION WITH VICRYL (GROUP V). REPORTED COMPLICATIONS ARE: GROUP P: N=6; LEAK FOLLOWING TRA (TRACHEAL RESECTION AND ANASTOMOSIS), TREATMENT: NOT REPORTED. GROUP V: N=1; LEAK FOLLOWING TRA (TRACHEAL RESECTION AND ANASTOMOSIS), TREATMENT: NOT REPORTED. N=3; MORTALITY DUE TO TRACHEA-RELATED COMPLICATIONS, TREATMENT: NONE. IN CONCLUSION, WE DEDUCED THAT VICRYL AND PROLENE PERFORM SIMILARLY IN REDUCING MORBIDITY (LEAK), BUT BOTH GROUPS' MORTALITY RATES REMAIN THE SAME. THEREFORE, TO ALLEVIATE THESE PROBLEMS, WE RECOMMEND USING PROLENE WHEN UNDERGOING TRACHEAL RESECTION AND ANASTOMOSIS. | PRODUCT COMPLAINT#: (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. WHAT WAS THE CAUSE OF DEATH? DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6: COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01) GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: DOI: 10.47009/JAMP.2024.6.1.330.
TRILOGY 100 | K083526 | THE MANUFACTURER RECEIVED INFORMATION CONCERNING A PRESERVATION OF EVIDENCE LETTER. THE COMPLAINT/ALLEGED ISSUE IS UNKNOWN, ONLY THAT A PRESERVATION OF EVIDENCE LETTER WAS RECEIVED WITH NO DETAILS. IT IS BELIEVED THAT THERE WAS A PATIENT DEATH, BUT NO DETAILS ON THE CAUSE OF DEATH WERE PROVIDED. THE MANUFACTURER IS TREATING THIS REPORT ARE BEING CONCERNED WITH THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE EVIDENCE PRESENT. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION CONCERNING A PRESERVATION OF EVIDENCE LETTER. THE COMPLAINT/ALLEGED ISSUE IS UNKNOWN, ONLY THAT A PRESERVATION OF EVIDENCE LETTER WAS RECEIVED WITH NO DETAILS. IT IS BELIEVED THAT THERE WAS A PATIENT DEATH, BUT NO DETAILS ON THE CAUSE OF DEATH WERE PROVIDED. THE MANUFACTURER IS TREATING THIS REPORT ARE BEING CONCERNED WITH THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. BASED ON THE INFORMATION CURRENTLY PROVIDED AND AVAILABLE, DEVICE CAUSE AND/OR CONTRIBUTION TO THE REPORTED EVENT OF DEATH CANNOT CURRENTLY BE CONFIRMED NOR REFUTED BASED ON THE EVIDENCE PRESENT. TRILOGY 100 VENTILATOR, U.S.A. BT (P/N 1054260B S/N (B)(6)) RECEIVED AT MANUFACTURER'S PRODUCT INVESTIGATION LAB (PIL) FOR INVESTIGATION. UPON INITIAL EXTERNAL VISUAL INSPECTION OF THE SUSPECT TRILOGY 100 VENTILATOR, MANUFACTURER TECH DID NOTE SOME FOREIGN MATERIALS INDICATIVE OF CONTAMINATION AT VARIOUS SITES OF THE UNIT, INCLUDING WITHIN AND AROUND THE AIRFLOW OUTLET PORT. TECH RECORDED AND REVIEWED THE UNIT¿S EVENT LOG AND SETTINGS AND PROCEEDED TO OPERATE THE UNIT ON THE PRIMARY SETTINGS WITH WHICH IT WAS RECEIVED AT PIL AFTER CONNECTING IT TO A TEST LUNG. THE UNIT OPERATED WITHOUT ISSUES OR ALARMS. THE UNIT WAS CONNECTED TO A MULTI-FUNCTION TEST STATION WHERE TECH VERIFIED FAILURES FOR POS FLOW VERIFY, ZERO FLOW VERIFY AND CONTROL PRESSURE AND MONITOR PRESSURE. TECH DISASSEMBLED THE SUSPECT UNIT FOR INTERNAL INSPECTION. TECH FOUND NO SIGNS OF DAMAGE OR EVIDENCE OF DEFECTIVE OPERATION. TECH DID NOT DETECT ANY PINCHED OR BLOCKED TUBING. TECH ALSO FOUND THE AIRPATH FOAM TO BE FIRMLY SECURED, WITHOUT ANY SIGNS OF DELAMINATION OR DEGRADATION. TECH COULD NOT CONFIRM BLACK FOAM PARTICULATES OR FOAM DEGRADATION INSIDE THE SUSPECT UNIT AFTER DISASSEMBLY. AFTER REVIEW, TECH HAS DETERMINED THAT TRILOGY 100 VENTILATOR 1054260B/(B)(6) OPERATES AS DESIGNED AND FUNCTIONS AS EXPECTED IN REGARD TO THE AIRPATH FOAM BEING FOUND TO BE FIRMLY SECURED AND THE UNIT OPERATING WITHOUT ISSUES OR ALARMS. PER THE TEST FAILURES, IT WAS DETERMINED THAT THE UNIT DRIFTED OUT OF SPEC DUE TO SEPARATE CONTAMINATION ISSUE ON THE SENSOR BOARD AND REPLACEMENT OF THE SENSOR BOARD WOULD HAVE BEEN NEEDED TO ADDRESS THE ISSUE. CORRECTION HAS BEEN MADE TO SECTION H, TYPE OF REPORTED COMPLAINT FIELD FROM PRODUCT PROBLEM TO DEATH.
PHENOM CATHETER | K151638 | IT WAS REPORTED THAT DURING AN INITIAL IMPLANT PROCEDURE USING THE PIPELINE EMBOLIZATION DEVICE 4.00MM X 30MM FOR MULTIPLE UNRUPTURED, SACCULAR ANEURYSMS LOCATED IN THE POSTERIOR COMMUNICATING SEGMENT OF THE LEFT INTERNAL CAROTID ARTERY (ICA), SEVERAL COMPLICATIONS OCCURRED. THE ANEURYSMS INCLUDED AN UPPER POSTERIOR COMMUNICATING ARTERY ANEURYSM MEASURING 7.1MM IN MAXIMUM DIAMETER WITH A 5.8MM NECK, AND ANOTHER ANEURYSM APPROXIMATELY 5MM IN SIZE BELOW. THE DISTAL LANDING ZONE WAS 3.3MM AND THE PROXIMAL LANDING ZONE WAS 3.9MM. THE VESSEL EXHIBITED SEVERE TORTUOSITY. FOLLOWING FEMORAL ARTERY PUNCTURE, A 6F 90CM LONG SHEATH, 6F 115CM TETHYS GUIDE CATHETER, PHENOM 27 MICROCATHETER, AND A 0.014 INCH 200CM GUIDEWIRE WERE USED TO ACCESS THE TARGET SITE. THE DEVICES WERE PREPARED ACCORDING TO THE INSTRUCTIONS FOR USE (IFU). 3D ROTATIONAL ANGIOGRAPHY WAS PERFORMED TO GUIDE DEVICE SELECTION AND PLACEMENT. THE STENT WAS HYDRATED AND DELIVERED, WITH THE TIP OPENED AT THE M1 SEGMENT OF THE MIDDLE CEREBRAL ARTERY. DURING PULL-BACK POSITIONING, THE STENT COULD NOT BE ANCHORED DUE TO A SHORT ANCHORING DISTANCE, RESULTING IN STENT RETRACTION. WHEN ATTEMPTING TO RETRIEVE THE STENT FOR REPOSITIONING, EXCESSIVE RESISTANCE WAS ENCOUNTERED. SUBSEQUENTLY, A RAPID RISE IN BLOOD PRESSURE WAS OBSERVED, AND ANGIOGRAPHY REVEALED RUPTURE AND BLEEDING OF THE INTERNAL CAROTID ARTERY WITH SUBARACHNOID HEMORRHAGE. PUPILLARY DILATION WAS NOTED ON THE OPERATING TABLE. CRANIOTOMY WAS PERFORMED AS A RESCUE INTERVENTION, BUT THE RESCUE WAS UNSUCCESSFUL AND THE PATIENT DIED. ADDITIONAL INFORMATION WAS RECEIVED REPORTING THAT THE PROBLEM OCCURRED THAT DAY AND THE PATIENT DIED. THE PATIENT PASSED AWAY AT (B)(6) HOSPITAL, DOCTOR INFORMED THAT THE PATIENT DIED IN THE HOSPITAL. NO AUTOPSY RECORDS WILL BE AVAILABLE. THE PATIENT'S DEATH MAY BE RELATED TO THE PIPELINE SHIELD, AND IT MAY ALSO BE RELATED TO THE OPERATOR'S OPERATION. THERE WAS INTRACRANIAL ANEURYSM LEADING UP TO THE PATIENT'S DEATH. THE CAUSE OF THE RESISTANCE DURING RETRIEVAL MAY BE VASCULAR TORTUOSITY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CARDIOSAVE HYBRID, TYPE B PLUG | K112372 | IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) WAS NOT SHOWING BALLOON PRESSURE DURING USE. THE PATIENT EXPIRED FOR REASONS OTHER THAN IABP THERAPY. | DUE TO CHARACTER LIMITATION IN E1, THE POSTAL CODE IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
PREMICATH | K954302 | PICC LINE WAS LEAKING AT THE HUB, NEEDLESS ACCESS DEVICE (NAD) CHANGED SEVERAL TIMES, T-CONNECTOR ADDED. LINE WAS LEAKING SIGNIFICANTLY. WE TRIED TROUBLE SHOOTING,USING NUMEROUS DEVICES LIKE (NADS, T-CONNECTORS, VALGUARDS) NOTHING HELPED. | ALTHOUGH THE MALFUNCTIONING DEVICE WILL NOT BE RETURNED TO VYGON, THE DETAILS OF THE MALFUNCTION WILL BE EVALUATED AS PART OF THE COMPLAINT INVESTIGATION. THE RESULTS OF THE INVESTIGATION ARE STILL PENDING AND WILL BE COMMUNICATED TO THE FDA.
REBAR | K093750 | ADDITIONAL INFORMATION RECEIVED REPORTED CONDITION BEING TREATED WAS ACUTE INTRACRANIAL LARGE ARTERY OCCLUSION. AUTOPSY NOT AVAILABLE. CAUSE OF DEATH WAS NOT DETERMINED. DEATH NOT THOUGHT TO BE RELATED TO REACT-71 CATHETER. UNKNOWN IF DEATH IS RELATED TO PROCEDURE OR PATIENT BASELINE CONDITION. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
CIRCULO | K221575 | PROCEDURE DESCRIPTION IT WAS REPORTED THAT ON (B)(6) 2025, A 27MM NAVITOR VISION VALVE WAS CHOSEN FOR A TRANSCATHETER AORTIC VALVE IMPLANTATION (TAVI) PROCEDURE USING AN UNKNOWN DELIVERY SYSTEM. DURING PROCEDURE, THE PATIENT HAD SUDDENLY NO AORTIC PRESSURE (ASYSTOLE) AND DESPITE REANIMATION THE PATIENT DID NOT RECOVER. A TRANSTHORACIC ECHOCARDIOGRAM (TTE) SHOWED MASSIVE PERICARDIAL EFFUSION (TAMPONADE AND A PERICARDIOCENTESIS WAS PERFORMED. AFTER THAT, NO RECOVERY TO THE PATIENT WAS NOTED AND HAD PERSISTENT VENTRICULAR FIBRILLATION. THEY DECIDED TO DO A STERNOTOMY; A RUPTURE OF THE LEFT VENTRICLE APPEARED TO BE THE PROBLEM DUE TO PERFORATION WITH THE TAVI GUIDEWIRE. THE PATIENT PASSED AWAY DURING THE PROCEDURE. HAS THE HEALTHCARE PROFESSIONAL STATED THAT THEY BELIEVE THE PATIENT OR USER EXPERIENCED ADVERSE HEALTH CONSEQUENCES DUE TO THE PERFORMANCE OF THE PRODUCT?: NO. PATIENT STATUS: DEAD. PRODUCT REPLACEMENT: WAS THE PRODUCT REPLACED?: NO. QUESTION: WAS THERE AN ABBOTT EMPLOYEE PRESENT FOR DURING PROCEDURE? ANSWER: YES. QUESTION: WAS THE DELIVERY SYSTEM USED AN ABBOTT PRODUCT? IF YES, PLEASE PROVIDE MODEL AND BATCH/LOT NUMBER. ANSWER: FNAV-DS-LG, LOT 10572353. QUESTION: AT WHAT POINT OF THE PROCEDURE, "THE PATIENT HAD SUDDENLY NO AORTIC PRESSURE (ASYSTOLE) AND DESPITE REANIMATION THE PATIENT DID NOT RECOVER." WAS NOTED? ANSWER: DURING DEPLOYMENT OF VALVE ( AT 80%) JUST AFTER 1ST RESHEATH QUESTION: WAS A PERICARDIAL EFFUSION NOTED PRIOR TO PROCEDURE? ANSWER: NO. QUESTION: DOES THE USER BELIEVE AN ABBOTT DEVICE CAUSED THE PERICARDIAL EFFUSION? ANSWER: YES, THE CIRCULO GUIDEWIRE HAS CAUSED A PERFORATION IN THE LV. QUESTION: IF THE USER BELIEVES AN ABBOTT DEVICE CAUSED THE PERICARDIAL EFFUSION, PLEASE SPECIFY WHICH ONE. ANSWER: CIRCULO-XS, LOT UNKNOWN. QUESTION: IF THE PERICARDIAL EFFUSION WAS NOT RELATED TO THE ABBOTT DEVICE, WHAT DOES THE USER BELIEVE CAUSED THE PERICARDIAL EFFUSION? ANSWER: PERFORATION. QUESTION: DID THE PHYSICIAN(S) CONCLUDE THAT CARDIAC TAMPONADE WAS PRESENT? ANSWER: YES, BY TTE CONFIRMED. QUESTION: WAS THE VALVE EVER RE-SHEATHED MORE THAN TWO TIMES? IF YES, IS THE PHYSICIAN AWARE OF THE IFU WARNING AGAINST THIS? ANSWER: 1 TIME. QUESTION: DURING THE PROCEDURE, WHAT WERE THE PATIENT'S ACTIVATED CLOTTING TIME (ACT) LEVELS? ANSWER: AROUND 300-350. QUESTION: DURING THE PROCEDURE, WAS HEPARIN ADMINISTERED? IF SO, HOW MUCH? ANSWER: YES, 100 IE/KG (STANDARD PROTOCOL). QUESTION: DURING THE PROCEDURE, WAS THERE ANY DIFFICULTY IMPLANTING THE DEVICE, SUCH AS MULTIPLE MANIPULATIONS? ANSWER: NO, WITH 1ST ATTEMPT JUST TOO HIGH AT THE LCC ( IN LAO VIEW), SO NEED FOR RESHEATHING. MAYBE TOO MUCH PUSH ON THE GW BY THE 2ND IMPLANTER? QUESTION: PLEASE PROVIDE THE SERIAL NUMBER OF THE 27MM NAVITOR VISION VALVE? ANSWER: (B)(6). QUESTION: PLEASE CONFIRM WAS THE 27MM NAVITOR VISION VALVE REMAIN IMPLANTED? ANSWER: YES. QUESTION: DID THEY ADMINISTER ANY PROTAMINE OR REVERSAL AGENT DURING PROCEDURE OR AFTER THE PROCEDURE? ANSWER: NO. QUESTION: WERE THERE MULTIPLE WIRE OR DELIVERY SHEATH EXCHANGES? ANSWER: NO, NOT MORE THAN USUAL PROCEDURE. QUESTION: WHAT WAS THE CAUSE OF DEATH? ANSWER: RUPTURE OF THE FREE WALL OF THE LV. QUESTION: WHAT WAS THE PATIENT'S CONDITION PRIOR TO THE IMPLANT PROCEDURE? ANSWER: GOOD. QUESTION: WHAT WAS THE PATIENT'S CONDITION AFTER THE IMPLANT PROCEDURE? ANSWER: SHE DIED JUST AFTER THE IMPLANTATION OF THE VALVE (IN THE CATHLAB). QUESTION: DOES THE IMPLANTER CLASSIFY THIS DEATH AS BEING RELATED TO THE PERFORMANCE OF THE IMPLANTED DEVICE, OR LIKELY RELATED TO THE PROCEDURE AND THE PATIENT'S COMORBIDITIES? (I.E. INQUIRE ABOUT THE FUNCTION OF THE IMPLANTED PRODUCT IF APPLICABLE). ANSWER: THE PROCEDURE (WRONG MANIPULATION GW). QUESTION: COULD YOU PLEASE PROVIDE OPERATIVE NOTES FROM THE PROCEDURE, WITH IDENTIFIABLE PATIENT INFORMATION REDACTED IF APPLICABLE? ANSWER: PREDIL WITH A 22MM BALLOON. START DEPLOYMENT IN CUSP OVERLAP VIEW. 1 RESHEATH BECAUSE OF TOO HIGH POSITION LCC IN LAA VIEW. JUST AFTER RESHEATH AND IMPLANTATION AT 80%, (B)(6) > REANIMATION, INTUBATION. TTE SHOWED TAMPONADE AND PERICARDIAL EFFUSION> PERICARDIAL PUNCTURE> NO RECOVERY> STERNOTOMY> (B)(6) DIED. QUESTION: PLEASE PROVIDE IMPLANT PROCEDURAL IMAGING, INCLUDING ECHOCARDIOGRAPHY AND ANGIOGRAPHY? DICTATED REPORTS OF ANY THERAPEUTIC INTERVENTIONS? IF NOT AVAILABLE, PLEASE PROVIDE A JUSTIFICATION. (PLEASE REDACT ANY IDENTIFIABLE PATIENT INFORMATION.) ANSWER: MEDIA WAS PROVIDED WITH THE COMPLAINT. ADDITIONAL INFORMATION RECEIVED FROM THE DISTRIBUTOR ON (B)(6) 2025: QUESTION: WHAT IS THE CONTEXT OF THE IMAGE? WAS THIS BEFORE, DURING, OR AFTER CPR? ANSWER: AFTER. QUESTION: WAS THE GUIDEWIRE LOCATION VERIFIED PRIOR TO DEPLOYMENT OF THE VALVE? ANSWER: YES. QUESTION: WAS THE DISTAL CURVE OF THE GUIDEWIRE VISUALIZED DURING DEPLOYMENT OF THE VALVE? ANSWER: YES. QUESTION: HAS ABBOTT'S MEDICAL TEAM REVIEWED THE MEDIA, AND CAN YOU PROVIDE A SUMMARY OF FINDINGS? ANSWER: GUIDEWIRE MUST HAVE PERFORATED THE LV DURING FIRST DEPLOYMENT OR RESHEATHING THE 1ST TIME. QUESTION: PLEASE LIST ALL THE OTHER DEVICES USED DURING THE PROCEDURE, INCLUDE THE MODEL, BRAND NAME, SIZES, ETC.? ANSWER: NVTR-LS-LG+: (B)(6) / NVRO-27: (B)(6)/ FNAV-DS-LG: (B)(6)/ CIRCULO-XS. QUESTION: PLEASE CONFIRM WAS THE CURVE OF THE CIRCULO GUIDEWIRE CONSTRAINED WITHIN A CATHETER DURING INSERTION INTO AND WITHDRAWAL FROM THE BODY OR TREATMENT SITE? ANSWER: YES. QUESTION: PLEASE CONFIRM WAS THE WIRE POSITION MAINTAINED WHILE TRACKING OR ADVANCING A CATHETER OR DEVICE OVER THE WIRE? ANSWER: YES, HOWEVER THERE COULD HAVE BEEN RESIDUAL PUSH. QUESTION: COULD YOU PLEASE VERIFY IF CIRCULO IS STILL AVAILABLE FOR ANALYSIS AND IF YES, WHAT IS THE PROJECTED DATE OF THE DEVICE RETURN TO INTEGER/LAKE REGION MEDICAL? ANSWER: NO. QUESTION: IF CIRCULO IS NOT AVAILABLE TO BE RETURNED, WHAT IS THE REASON FOR NO RETURN? ANSWER: IN THE PERIPROCEDURAL RUSH THE GW WAS THROWN AWAY. QUESTION: IS THIS A NEW USER OR EXPERIENCE USER? ANSWER: EXPERIENCED. QUESTIONS REGARDING THE NAVITOR VALVE: QUESTION: PLEASE CONFIRM WAS THERE ANY ISSUES NOTED ON THE NAVITOR VALVE? ANSWER: NO. QUESTION: WAS THE NAVITOR VALVE PERFORMED AS INTENDED? ANSWER: YES. | THIS MEDWATCH REPORT IS AN UPDATE TO THE PREVIOUS REPORT. ADDITIONAL INFORMATION HAS BEEN ADDED TO SECTIONS D4, H4, H6 (A), (B), (C), (D), (E), AND H11. THE DEVICE INVOLVED IN THE EVENT WAS DISCARDED THEREFORE, NO PHYSICAL OR VISUAL ANALYSIS OF THE PRODUCT COULD BE PERFORMED. A REVIEW OF THE DEVICE HISTORY RECORDS OF THE SPECIMEN DEVICE DOES NOT PRESENT ANY INDICATION OF MANUFACTURING DEFECT OR ANOMALY THAT COULD HAVE IMPACTED ON THE EVENT AS REPORTED. DURING MANUFACTURING AND PACKAGING OF THIS PRODUCT THE OPERATORS ARE INSTRUCTED TO 100% VISUALLY INSPECT FOR ANY OBVIOUS DEFECT PRIOR TO SHIPMENT. THE HISTORY RECORDS INDICATE THIS PRODUCT WAS FINAL INSPECTION TESTED AT LAKE REGION MEDICAL AND WAS DETERMINED TO BE ACCEPTABLE. A FURTHER EVALUATION OF THE AVAILABLE AND PROVIDED INFORMATION INCLUDING IMAGING WAS UNDERTAKEN BY AN INTERVENTIONAL CARDIOLOGIST ON BEHALF OF THE MANUFACTURER. SUMMARY PROVIDED BY THE INTERVENTIONAL CARDIOLOGIST: 1. FROM THE CINE IMAGES PROVIDED, IT APPEARS THAT THE LEFT VENTRICLE SIZE IS SMALL. IN PATIENTS WITH A SMALL LEFT VENTRICLE, THERE IS ALWAYS A HIGHER RISK FOR THE WIRE TO BE POSITIONED TOO DEEP IN THE LEFT VENTRICLE DURING TAVR. 2. IN ONE BAV IMAGE, THE LEFT VENTRICLE APEX IS NOT VISIBLE. THE GUIDEWIRE APPEARS TO BE POSITIONED MUCH DEEPER IN COMPARISON TO THE OTHER IMAGES PROVIDED FOR REVIEW. 3. ANOTHER IMAGE PROVIDED SHOWS THE TAVR DEVICE LOADED. FROM THE TIME OF POSITIONING THE GUIDEWIRE AND UNTIL THE TAVR VALVE GETS LOADED AND DEPLOYED, THE WIRE CAN MOVE TOO DEEP INTO THELEFT VENTRICLE DURING TAVR VALVE ADVANCEMENT WITHOUT MAINTAINING CONTROL OF THE GUIDEWIRE. THAT PARTICULAR LOSS OF CONTROL OF GUIDEWIRE AND INADVERTENT ADVANCEMENT OF THE WIRE INTO THE LEFT VENTRICLE IS NOT DOCUMENTED. WHEN THE OPERATORS REALIZE THE WIRE GOT INADVERTENTLY ADVANCED INTO THE LEFT VENTRICLE TOO DEEP, THE IMMEDIATE REACTION IS TO PULL THE WIRE BACK INTO THE IDEAL POSITION. THAT SEQUENCE, THE MAJORITY OF TIME, HAPPENS DURING REGULAR FLUORO IMAGING RATHER THAN CINE IMAGING. IT IS EXTREMELY HARD TO DOCUMENT IT. THIS INADVERTENT ADVANCEMENT OF THE GUIDEWIRE IS OBVIOUS IN THE LEFT VENTRICLE ANGIOGRAM WHERE THE PIGTAIL CATHETER HAD BEEN ADVANCED DEEP INTO THE LEFT VENTRICLE ALMOST DOUBLE THE USUAL LENGTH. 4. CONCLUSION FROM PROVIDED INFORMATION: THE LEFT VENTRICLE ANGIOGRAM INDICATES CONTRAST APPEARS TO BE LEAVING THE LEFT VENTRICLE AT THE APEX WHICH REINFORCES MY IMPRESSION THAT THERE IS A HIGH CHANCE THAT THE GUIDEWIRE, WHICH WAS POSITIONED DEEP IN THE LEFT VENTRICLE TO START WITH. DURING THE SUBSEQUENT STEPS OF THE TAVR PROCEDURE, THE WIRE WAS INADVERTENTLY ADVANCED MUCH FURTHER, CAUSING LEFT VENTRICLE PERFORATION. THE FOLLOWING INFORMATION FROM THE USER INFORMATION, WARNINGS, AND PROCEDURE SECTIONS ARE INCLUDED IN THE DEVICE INSTRUCTIONS FOR USE AND INFORM THE USERS OF STEPS TO TAKE TO PREVENT VENTRICLE PERFORATION FROM OCCURRING. AS INDICATED IN THE USER INFORMATION: THE GUIDEWIRE SHOULD BE USED ONLY BY PHYSICIANS OR CLINICIANS TRAINED IN THE INTRODUCTION AND PLACEMENT OF INTERVENTIONAL DEVICES WITHIN THE CHAMBERS OF THE HEART, INCLUDING THOSE USED DURING TRANSCATHETER AORTIC VALVE PROCEDURES. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE WARNINGS: 1. THE GUIDEWIRE SHOULD BE USED ONLY BY PHYSICIANS TRAINED IN THE INTRODUCTION AND PLACEMENT OF INTERVENTIONAL DEVICES INCLUDING THOSE USED WITHIN TRANSCATHETER AORTIC VALVE PROCEDURES. 2. CAREFULLY READ ALL INSTRUCTIONS PRIOR TO USE. OBSERVE ALL WARNINGS AND PRECAUTIONS. FAILURE TO DO SO MAY RESULT IN COMPLICATIONS. 4. MONITOR GUIDEWIRE POSITION THROUGHOUT THE PROCEDURE FOR PROPER PLACEMENT OF CURVE AND DISTAL TIP. 9. WHEN ADVANCING OR REMOVING THE GUIDEWIRE, ALWAYS USE FLUOROSCOPIC GUIDANCE WITH RADIOGRAPHIC EQUIPMENT THAT PROVIDES HIGH- RESOLUTION IMAGES. NEVER POSITION THE GUIDEWIRE BLINDLY, AS THIS MAY RESULT IN MISPLACEMENT, DISSECTION OR PERFORATION. 11. NEVER ADVANCE THE GUIDEWIRE AGAINST THE RESISTANCE WITHOUT FIRST DETERMINING THE REASON FOR RESISTANCE UNDER FLUOROSCOPY. EXCESSIVE FORCE AGAINST RESISTANCE MAY RESULT IN DAMAGE TO THE CATHETER OR VESSEL/ORGAN. CARE SHOULD BE TAKEN WHEN ADVANCING A GUIDEWIRE AFTER DEVICE DEPLOYMENT. AS INDICATED IN THE DEVICE INSTRUCTIONS FOR USE PROCEDURE: 10. ADVANCE THE INTERVENTIONAL DEVICE OVER THE GUIDEWIRE WHILE MAINTAINING GUIDEWIRE POSITION. A. VISIBLY MONITOR THE GUIDEWIRE DOUBLE CURVE WHEN ADVANCING OR WITHDRAWING A CATHETER OR INTERVENTIONAL DEVICE OVER THE WIRE. IF RESISTANCE IS MET WHILE ADVANCING THE DEVICE OVER THE WIRE STOP; EVALUATE THE CAUSE BEFORE PROCEEDING (USE FLUOROSCOPY IF NECESSARY). WE REVIEWED THE ASSOCIATED RISK DOCUMENTATION RELATIVE TO THIS PRODUCT AND CONFIRMED THAT THIS INCIDENT IS LISTED AND THAT IT IS TRENDING WITHIN THE ESTABLISHED OCCURRENCE THRESHOLD. LAKE REGION'S EFFECTIVENESS OF DESIGN VERIFICATION ACTIVITIES BRINGS THE OVERALL RISK DOWN TO AS FAR AS POSSIBLE. ARRHYTHMIA, PERICARDIAL EFFUSION, CARDIAC TAMPONADE, CARDIAC PERFORATION, AND DEATH ARE KNOWN POTENTIAL ADVERSE EVENTS AND ARE LISTED IN THE CIRCULO DEVICE INSTRUCTIONS FOR USE.
EDWARDS ESHEATH+ INTRODUCER SET | K200258 | ACCORDING TO THE TERRITORY MANAGER (TM), DURING A TRANSCATHETER AORTIC VALVE REPLACEMENT (TAVR) PROCEDURE UTILIZING A 26MM SAPIEN 3 ULTRA VALVE VIA TRANSFEMORAL APPROACH, THE ESHEATH WAS SLIGHTLY RETRACTED TO ALLOW FOR VALVE ALIGNMENT PRIOR TO THE FIRST BEND OF ABDOMINAL AORTIC TORTUOSITY. VALVE ALIGNMENT WAS ACHIEVED WITHOUT COMPLICATION. ADVANCEMENT OF THE DELIVERY SYSTEM REQUIRED INCREASED FORCE DUE TO THE TORTUOSITY, AND THE INITIAL ATTEMPT WAS UNSUCCESSFUL IN NAVIGATING THE AORTIC ARCH. A LUNDERQUIST WIRE WAS SUBSEQUENTLY INTRODUCED THROUGH THE PIGTAIL CATHETER ON THE CONTRALATERAL SIDE, WHICH FACILITATED SUCCESSFUL TRAVERSAL OF THE TORTUOUS ANATOMY AND PASSAGE OVER THE ARCH. SIGNIFICANT FLEXING OF THE DELIVERY SYSTEM WAS OBSERVED DURING ADVANCEMENT. AS THE VALVE NEARED THE ANNULUS, ANESTHESIA REPORTED A SUDDEN DROP IN THE PATIENT'S BLOOD PRESSURE. AN ECHOCARDIOGRAM REVEALED NO PERICARDIAL EFFUSION, AND PHYSICAL EXAMINATION OF THE GROIN AND ABDOMEN WAS UNREMARKABLE. AN ABDOMINAL AORTOGRAM IDENTIFIED CONTRAST EXTRAVASATION NEAR THE FIRST BEND OF THE AORTA, LIKELY AT THE PROXIMAL END OF THE SHEATH. BLOOD PRODUCTS WERE ADMINISTERED. A CODA BALLOON WAS DEPLOYED IN AN ATTEMPT TO TAMPONADE THE EXTRAVASATION, WITH PROLONGED INFLATION AND INTERMITTENT AORTOGRAPHY SHOWING REDUCED LEAKAGE. VASCULAR SURGERY WAS CONSULTED BUT DEEMED THE PATIENT INOPERABLE FOR BOTH SURGICAL AND ENDOVASCULAR REPAIR. THE PATIENT WAS STABILIZED AND TRANSFERRED TO THE ICU. ALTHOUGH INITIALLY AWAKE AND RESPONSIVE, THE PATIENT EXPERIENCED A SUDDEN DROP IN BLOOD PRESSURE AND, DESPITE SUPPORTIVE MEASURES, WAS PRONOUNCED DECEASED APPROXIMATELY ONE HOUR AFTER LEAVING THE CATHETERIZATION LAB. | PER THE INSTRUCTIONS FOR USE (IFU), CARDIOVASCULAR INJURY INCLUDING HEMATOMA, PERFORATION OR DISSECTION OF VESSELS, VENTRICLE, ATRIUM, SEPTUM, MYOCARDIUM OR VALVULAR STRUCTURES THAT MAY REQUIRE INTERVENTION ARE KNOWN POTENTIAL ADVERSE EVENTS ASSOCIATED THE OVERALL TRANSCATHETER VALVE REPLACEMENT (TVR) PROCEDURE AND MAY REQUIRE INTERVENTION. AORTIC DISSECTION OR RUPTURE MAY OCCUR WHEN MULTIPLE ATTEMPTS ARE MADE TO CROSS THE STENOTIC NATIVE VALVE, AND/OR WHEN EXCESSIVE FORCE IS USED. PHYSICIANS ARE EXTENSIVELY TRAINED BY EDWARDS BEFORE THEY ARE QUALIFIED TO USE THE SAPIEN TRANSCATHETER HEART VALVE (THV). THE THV TRAINING MANUALS GUIDE TO FACILITATE SAFE CROSSING OF THE NATIVE OR FAILING VALVE, INCLUDING CAMERA PROJECTIONS, HANDLING DURING ADVANCEMENT, AND TROUBLESHOOTING TECHNIQUES IF DIFFICULTY IS ENCOUNTERED. AS STATED, EXCESSIVE FORCE SHOULD NOT BE USED WHEN THE DEVICE HAS DIFFICULTY CROSSING THE STENOTIC VALVE. ADDING TENSION TO THE WIRE, PULLING BACK THE SYSTEM TO RE-ORIENT THE VALVE, AS NEEDED, AND TORQUING THE FLEX CATHETER MAY HELP SOLVE THE PROBLEM. IN THIS CASE, SIGNIFICANT FLEXING OF THE DELIVERY SYSTEM WAS OBSERVED DURING ADVANCEMENT. AS THE VALVE APPROACHED THE ANNULUS, ANESTHESIA REPORTED A SUDDEN DROP IN THE PATIENT'S BLOOD PRESSURE. AN ECHOCARDIOGRAM REVEALED NO PERICARDIAL EFFUSION, AND PHYSICAL EXAMINATION OF THE GROIN AND ABDOMEN WAS UNREMARKABLE. HOWEVER, AN ABDOMINAL AORTOGRAM IDENTIFIED CONTRAST EXTRAVASATION NEAR THE FIRST BEND OF THE AORTA, LIKELY AT THE PROXIMAL END OF THE SHEATH. THESE FINDINGS SUGGEST THAT MOVEMENT OF THE SHEATH TIP PARTICULARLY DURING EFFORTS TO NAVIGATE THE TORTUOUS ABDOMINAL AORTIC ANATOMY WITH THE COMMANDER DELIVERY SYSTEM MAY HAVE BEEN A CONTRIBUTING FACTOR TO THE VASCULAR INJURY. A REVIEW OF EDWARDS LIFESCIENCES RISK MANAGEMENT DOCUMENTATION WAS PERFORMED FOR THIS CASE. THE REPORTED EVENT IS AN ANTICIPATED RISK OF THE TRANSCATHETER HEART VALVE PROCEDURE, ADDITIONAL ASSESSMENT OF THIS ADVERSE EVENT IS NOT REQUIRED AT THIS TIME. THE IFU AND TRAINING MANUALS HAVE BEEN REVIEWED AND NO INADEQUACIES HAVE BEEN IDENTIFIED WITH REGARDS TO WARNINGS, CONTRAINDICATIONS, AND THE DIRECTIONS/CONDITIONS FOR THE SUCCESSFUL USE OF THE DEVICE. COMPLAINT HISTORIES FOR ALL REPORTED EVENTS ARE REVIEWED AGAINST TRENDING CONTROL LIMITS MONTHLY, AND ANY EXCURSIONS ABOVE THE CONTROL LIMITS ARE ASSESSED AND DOCUMENTED AS PART OF THIS MONTHLY REVIEW. AS SUCH, NEITHER A PRODUCT RISK ASSESSMENT, NOR CORRECTIVE OR PREVENTATIVE ACTIONS ARE REQUIRED AT THIS TIME.
UMBILICAL CATHETER | K921352 | AN UMBILICAL CATHETER WAS INSERTED AT DAY 0 ON A PREMATURE NEWBORN OF 28 WEEKS. 2 DAYS LATER AN EXTRAVASATION OF PARENTERAL WAS NOTICED. THE CATHETER WAS LOCATED IN THE LIVER. THE PREMATURE DIED ON (B)(6) 2025. | THE INVOLVED UMBILICAL CATHETER IS NOT AVAILABLE. FROM THE DESCRIPTION, THIS SERIOUS ADVERSE EVENT IS CLEARLY RELATED TO THE MALPOSITION OF THE UMBILICAL CATHETER IN THE LIVER. IT IS NOT TRACED TO A DEFECT OF OUR DEVICE. THE BATCH REVIEW DOES NOT SHOW ANY NON-CONFORMITY. THE PRODUCTS ARE IN CONFORMITY WITH THE SPECIFICATIONS. ALL CONTROLS ARE COMPLIANT, INCLUDING THE 100% TESTS CARRIED OUT ON EACH CATHETER DURING THE MANUFACTURING PROCESS. THESE TESTS ARE AS FOLLOWS: THE BLUNTNESS OF THE CATHETER TIP, MARKING THE CATHETERS, TIGHTNESS AND NON-OBSTRUCTION. THIS SERIOUS ADVERSE EVENT IS NOT TRACED TO A MALFUNCTION OR DEFECT OF OUR DEVICE.
FREESTYLE LIBRE 3 | K223435 | A LOW READINGS ISSUE WAS REPORTED WITH THE ABBOTT DIABETES CARE (ADC) DEVICE. THE CUSTOMER'S CAREGIVER REPORTED THAT SCAN RESULTS FROM THE ADC DEVICE RANGED BETWEEN 103 MG/DL AND 150 MG/DL, WHICH APPEARED WITHIN THE CUSTOMER'S NORMAL RANGE AND SHOWED A HORIZONTAL TREND ARROW, INDICATING SLOW GLUCOSE CHANGE (LESS THAN 1 MG/DL PER MINUTE). DESPITE THESE READINGS, THE CUSTOMER EXPERIENCED SYMPTOMS SUCH AS WEAKNESS, NAUSEA, AND VOMITING. NO CAPILLARY BLOOD GLUCOSE COMPARISONS WERE MADE AT THE TIME. EMERGENCY SERVICES WERE CALLED, AND DURING TRANSPORT, A HEALTHCARE PROFESSIONAL MEASURED A BLOOD GLUCOSE LEVEL OVER 400 MG/DL. UPON HOSPITAL ADMISSION, FURTHER READINGS EXCEEDED 600 MG/DL, AND A LABORATORY RESULT CONFIRMED A GLUCOSE LEVEL OVER 1000 MG/DL. THE PATIENT WAS TREATED WITH INSULIN AND MONITORED CONTINUOUSLY. WHILE HOSPITALIZED, THE PATIENT DEVELOPED PNEUMONIA AND WAS TREATED WITH ANTIBIOTICS. NO GLUCOSE READINGS WERE TAKEN FROM AN ADC DEVICE DURING HOSPITALIZATION, AS ONLY BLOOD GLUCOSE VALUES WERE USED. THE SENSOR WAS REMOVED SHORTLY AFTER ADMISSION. THE PATIENT PASSED AWAY ON (B)(6) 2025. THE CAREGIVER COULD NOT CONFIRM AN OFFICIAL CAUSE OF DEATH, NOTING THAT ALL DEVICES WERE TURNED OFF FOLLOWING THE ONSET OF PNEUMONIA. NO DEATH CERTIFICATE DETAILS WERE AVAILABLE AT THE TIME OF REPORTING. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED DEATH. | AN EXTENDED MANUFACTURING REVIEW IS PENDING AT THIS TIME. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE COMPLETED. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
REACT CATHETER | K182097 | ADDITIONAL INFORMATION RECEIVED REPORTED CONDITION BEING TREATED WAS ACUTE INTRACRANIAL LARGE ARTERY OCCLUSION. AUTOPSY NOT AVAILABLE. CAUSE OF DEATH WAS NOT DETERMINED. DEATH NOT THOUGHT TO BE RELATED TO REACT-71 CATHETER. UNKNOWN IF DEATH IS RELATED TO PROCEDURE OR PATIENT BASELINE CONDITION. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
INVACARE® PERFECTO2¿ OXYGEN CONCENTRATOR | K020386 | THE DEVICE'S PREVIOUS MANUFACTURER, INVACARE, CONTACTED VENTEC VIA EMAIL TO REPORT THE FOLLOWING EVENT: "ON (B)(6) 2025 [REDACTED], A HEALTH CARE SERVICE HAS BEEN INFORMED BY A CARER FROM THE 24/7 TEAM, THAT THE PATIENT WAS STRUGGLING TO GET OXYGEN THROUGH HIS TUBING ON THE CONCENTRATOR AND WAS STRUGGLING TO BREATH, SO THE PARAMEDIC HAD BEEN CALLED. THE PATIENT WAS WITH THE PARAMEDICS WHEN THE CARER CALLED IN AND THEY WERE ASSESSING HIM TO BE TAKEN TO HOSPITAL. THE CARER ADVISED THE PATIENT WAS GOING INTO HOSPITAL AND THEY WERE UNSURE WHEN HE WOULD BE RELEASED, AND NO ONE WOULD BE HOME TO ALLOW THE TECHNICIAN TO ATTEND. ON 30/04/2025, VIVISOL WAS INFORMED THAT THE PATIENT PASSED AWAY IN HOSPITAL. PATIENTS PRESCRIPTION ON THE CONCENTRATOR WAS 2-3LPM 20HPD. THIS CONCENTRATOR WAS INSTALLED WITH THE PATIENT ON (B)6) 2024." NO FURTHER DETAILS WERE PROVIDED. | H6: THIS DEVICE WAS MANUFACTURED AND PLACED INTO COMMERCIAL DISTRIBUTION BACK IN MAY 2011, WHICH PREDATES THE REQUIRED UDI COMPLIANCE DATE FOR CLASS II MEDICAL DEVICES. AS A RESULT, NO UDI# EXISTS FOR THIS DEVICE AND SECTION D4, UDI#, SHALL STATE ¿NONE.¿ THE REPORTER DID NOT PROVIDE ANY FURTHER INFORMATION ON THE CIRCUMSTANCES SURROUNDING THE INDIVIDUAL'S DEATH, NOR DID THEY PROVIDE THE DATE THAT THE PATIENT DIED. AS A RESULT, SECTION B2, DATE OF DEATH, SHALL BE LEFT BLANK. ALSO, SECTION D4 MODEL # AND D4 CATALOG # ARE "UNKNOWN" AS THEY WERE NOT REPORTED BY INVACARE. THE PREVIOUS DEVICE MANUFACTURER, INVACARE, IS INVESTIGATING THE REPORTED ISSUE AND HAS PROVIDED THEIR PRELIMINARY FINDINGS: "WE WOULD LIKE TO POINT OUT THAT "THE INVACARE PERFECTO2 CONCENTRATOR IS USED BY PATIENTS WITH RESPIRATORY DISORDERS WHO REQUIRE SUPPLEMENTAL OXYGEN. THE DEVICE IS NOT INTENDED TO SUSTAIN OR SUPPORT LIFE. FURTHERMORE, THE FOLLOWING SAFETY-RELEVANT INFORMATION IS STATED IN THE USER MANUAL OF THE PERFECTO2: "INVACARE RECOMMENDS AN ALTERNATE SOURCE OF SUPPLEMENTAL OXYGEN IN THE EVENT OF A POWER OUTAGE, ALARM CONDITION OR MECHANICAL FAILURE. CONSULT YOUR PHYSICIAN OR EQUIPMENT PROVIDER FOR THE TYPE OF RESERVE SYSTEM REQUIRED." THE INFORMATION WE HAVE AT THE TIME OF THIS REPORT IS VERY LIMITED. WE HAVE REQUESTED FURTHER INFORMATION AND THE PRODUCT FOR OUR INVESTIGATION. WHEN THE INVESTIGATION HAS BEEN CONCLUDED A FINAL REPORT WILL BE PROVIDED." INVACARE CONTINUES TO INVESTIGATE THE REPORTED ISSUE. ONCE COMPLETE, VENTEC WILL BE PROVIDED WITH THEIR INVESTIGATION FINDINGS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETE AS DEFINED BY 21 CFR 803.56.
DREAMSTATION BIPAP AUTOSV | K090539 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH FROM PULMONARY FIBROSIS. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. MAXIMUM ATTEMPTS WERE MADE TO HAVE THE DEVICE RETURNED FOR EVALUATION, INVESTIGATION AND TO GATHER ADDITIONAL INFORMATION BUT WERE UNSUCCESSFUL. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED.
M.R.I IMPLANTABLE PORT, HICKMAN SINGLE-LUMEN, 9.6F | K873213 | IT WAS REPORTED THROUGH THE LITIGATION PROCESS THAT SOMETIME AFTER PORT PLACEMENT PROCEDURE; THE PORT CATHETER HAD FRACTURED. IT WAS FURTHER REPORTED THAT PATIENT WAS EXPIRED, AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
ALARIS SYSTEM | K211218 | IT WAS REPORTED THAT THE PUMP WAS INTENDED TO INFUSE AT A RATE OF 4 ML/HOUR. HOWEVER, MEDICATION BAG WAS FOUND TO BE EMPTY. FURTHER INFORMATION WAS RECEIVED THROUGH DUE DILIGENCE ATTEMPTS. IT WAS REPORTED THAT MORPHINE 100 ML WAS PROGRAMMED ON (B)(6) 2025 TO INFUSE AT 4 ML/HOUR AND DOUBLE VERIFIED. THE CLINICIAN NOTED AN AIR-IN-LINE ALARM AND EMPTY MEDICATION BAG APPROXIMATELY ONE (1) HOUR AND TWENTY (20) MINUTES LATER. THERE WAS NO REPORTED CHANGE IN THE PATIENT'S CONDITION, AND NO INTERVENTIONS WERE PROVIDED. HOWEVER, IT WAS REPORTED THE PATIENT PASSED AWAY ON (B)(6) 2025. CUSTOMER RISK MANAGEMENT STATED, "THE EVENT REPORTED WAS EQUIPMENT MALFUNCTION; HOWEVER, AT THIS POINT NOT FELT TO HAVE CONTRIBUTED TO THE DEATH." | BD TECHNICAL SUPPORT TROUBLESHOOT WITH CUSTOMER OVER THE PHONE. A FOLLOW UP REPORT WILL BE SUBMITTED ONCE THE FAILURE INVESTIGATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
TRILOGY 100 | K083526 | THE MANUFACTURER RECEIVED INFORMATION THAT "A PATIENT PASSED AWAY IN THE UNIT" WHILE ON A TRILOGY 100 DEVICE. THE CUSTOMER REQUESTED THE ACTION LOG OF ALARMS DURING THE PATIENT'S USE OF THE DEVICE. DURING THE HOSPITALIZATION PERIOD, THE DEVICE ISSUED DISCONNECT AND DESATURATION ALARMS, AND THE PATIENT WAS ON A MONITOR FOR VITAL SIGNS. DESPITE RESUSCITATION EFFORTS, THE PATIENT DID NOT RECOVER. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THIS ISSUE THE CUSTOMER REPORTING THE PATIENT WAS ADMITTED TO THE HOSPITAL ON (B)(6) 2025, FOR A STROKE WHICH OCCURRED DAYS PRIOR TO THE PATIENT'S PASSING ON (B)(6) 2025. THE TIME OF THE PATIENT'S PASSING REMAINS UNCERTAIN. THE CAUSE OF DEATH IS UNKNOWN, HAS BEEN REVIEWED BY THE CLINICAL EXPERT AND DETERMINED THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE REPORTED EVENT. ADDITIONALLY, THERE IS NO OTHER CLINICAL INFORMATION PROVIDED TO INDICATE ANY CAUSAL RELATIONSHIP BETWEEN THE DEVICE AND THE REPORTED DEATH. THE VENTILATOR SETTINGS REPORTED WERE MODE: CIRCUIT PASSIVE AVAP, PRESSURE SETTINGS: 45, MINIMUM OF 35, APAP SETTINGS: 25, BREATH RATE: 13 BREATHS PER MINUTE, INSPIRATORY TIME: 1.8 SECONDS, TRIGGER TYPE: AUTO TRACK, RISE TIME: 4. THE CUSTOMER STATES THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THE REASON FOR USING THE DEVICE IS OAS AND SEVERE OBESITY. THEREFORE, BASED ON AVAILABLE INFORMATION THE REPORTED DEATH IS ASSESSED AS NOT RELATED TO THE DEVICE IN THIS CASE. THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
ENDO ANCHOR SYSTEM - HELI-FX AAA | K140036 | LITERATURE WAS REVIEWED REGARDING ¿A WIDER GRIP: ENDOANCHORS INCREASE PROXIMAL SEALING SURFACE DURING ENDOVASCULAR ANEURYSM REPAIR¿ THE TIME FRAME OF THIS STUDY WAS OVER A SEVEN YEAR PERIOD. 221 PATIENTS UNDERWENT EVAR WITH ENDURANT STENT GRAFTS. IN ANOTHER 31 PATIENTS, ENDOANCHORS WERE IMPLANTED ALONG WITH THE ENDURANT STENT GRAFTS. PATIENT MORTALITY WAS REPORTED IN THE PATIENT POPULATION (TOTAL 28.4%). IT WAS REPORTED ONE DEATH OCCURRED WITHIN 30 DAYS OF THE EVAR PROCEDURE. THE CAUSE OF DEATH WAS MULTIPLE ORGAN FAILURE IN A PATIENT WHO HAD A RUPTURED ANEURYSM. NO FURTHER INFORMATION WAS PROVIDED PERTAINING TO MEDTRONIC PRODUCTS | MEDTRONIC RECEIVED THE FOLLOWING INFORMATION FROM A JOURNAL ARTICLE ENTITLED; ¿A WIDER GRIP: ENDOANCHORS INCREASE PROXIMAL SEALING SURFACE DURING ENDOVASCULAR ANEURYSM REPAIR¿ ACCARINO G, ESPOSITO D, CAMPOLMI M, TURCHINO D, SERRA R, MARCELLO BRACALE U ANN VASC SURG 2025; 119: 37¿45, HTTPS://DOI.ORG/10.1016/J.AVSG.2025.04.119. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
LIBERTY CYCLER SET, DUAL PATIENT CONNECT | K173718 | A REGISTERED NURSE (RN) REPORTED TO FRESENIUS THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY ON (B)(6) 2025. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PATIENT¿S PERITONEAL DIALYSIS REGISTERED NURSE (PDRN). THE PATIENT WAS SEEN IN THE PD CLINIC ON (B)(6) 2025 DUE TO ABDOMINAL PAIN AND CLOUDY PD EFFLUENT. PD CULTURES WERE OBTAINED. THE PATIENT WAS DIAGNOSED WITH PERITONITIS. THE PATIENT WAS INITIATED ON EMPIRIC ANTIBIOTIC THERAPY WITH INTRAPERITONEAL (IP) VANCOMYCIN AND IP CEFEPIME (DOSE, FREQUENCY, AND DURATION UNKNOWN). THE CULTURES WERE POSITIVE FOR PSEUDOMONAS (NO SPECIES PROVIDED). THE VANCOMYCIN WAS DISCONTINUED AND THE CEFEPIME WAS TO BE CONTINUED FOR 3 WEEKS. THE PATIENT¿S SYMPTOMS WERE NOT IMPROVING. ON (B)(6) 2025 THE PATIENT WAS ADMITTED TO THE HOSPITAL DUE TO SHORTNESS OF BREATH AND SWELLING (LOCATION NOT PROVIDED). THE PATIENT¿S PD CATHETER WAS REMOVED. THE PATIENT TRANSITIONED TO HEMODIALYSIS (HD). THE PATIENT CONTINUED TO WORSEN AND NOT MAKE ANY IMPROVEMENTS. THE PATIENT PASSED AWAY ON (B)(6) 2025 IN THE HOSPITAL. THE HOSPITAL DID NOT SEND OVER THE DEATH REPORT, BUT THE DISCHARGE RECORDS STATED THE CAUSE OF DEATH WAS SEPTICEMIA AND RESPIRATORY DISTRESS SYNDROME. NO ADDITIONAL INFORMATION WAS AVAILABLE AS THE PATIENT¿S RECORDS HAD BEEN CLOSED OUT OF THE CLINIC¿S SYSTEM. THE CAUSE OF THE PERITONITIS WAS NOT ESTABLISHED. | CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PD THERAPY AND UTILIZATION OF THE CYCLER SET AND THE PATIENT EVENT OF PSEUDOMONAS PERITONITIS WITH HOSPITALIZATION AND SUBSEQUENT DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY, OR CYCLER SET DEFECT REPORTED FOR THE EVENT. THE PATIENT WAS INITIALLY DIAGNOSED WITH PSEUDOMONAS PERITONITIS. PERITONITIS CAUSED BY PSEUDOMONAS SPECIES IS A PARTICULARLY COMMON AND SERIOUS COMPLICATION AND IS OFTEN ASSOCIATED WITH A POOR RESPONSE TO ANTIBIOTICS AS WELL AS A HIGH RATE OF CATHETER REMOVAL. ALTHOUGH THE CAUSE OF THE PERITONITIS WAS NOT CONFIRMED, PSEUDOMONAS IS A SPECIES THAT NORMALLY INHABITS SOIL, WATER, AND VEGETATION AND CAN BE ISOLATED FROM THE SKIN, THROAT, AND STOOL OF HEALTHY PERSONS. THE PATIENT¿S CONDITION WORSENED OVER TIME DESPITE BEING ON ANTIBIOTIC THERAPY. THE PD CATHETER WAS EVENTUALLY REMOVED APPROXIMATELY TWO WEEKS AFTER THE PERITONITIS DIAGNOSIS. ADDITIONALLY, THE PATIENT¿S ANTIBIOTICS WERE INITIALLY CHANGED TO ONE ANTIBIOTIC WITH CEFEPIME AFTER THE CULTURE RESULTS OF PSEUDOMONAS. THE INTERNATIONAL SOCIETY OF PERITONEAL DIALYSIS (ISPD) GUIDELINES FOR MANAGEMENT OF PD-RELATED PERITONITIS FOR PSEUDOMONAS RECOMMENDS THE USE OF DUAL ANTIBIOTIC THERAPY WITH DIFFERENT MECHANISMS OF ACTION AND TREATMENT FOR 3 WEEKS. WHEN THERE IS NO CLINICAL RESPONSE AFTER 4 DAYS OF ANTIBIOTIC TREATMENT, EARLY CATHETER REMOVAL SHOULD BE PREFERRED. THE PATIENT¿S WORSENING SYMPTOMS AND EVENTUAL PD CATHETER REMOVAL SUGGESTS THAT THE BACTERIA POSSIBLY DEVELOPED A BIOFILM WITHIN THE PD CATHETER AND THAT THE BACTERIA INVADED TISSUES SPREADING TO CAUSE THE EVENTUAL SEPTICEMIA WITH RESPIRATORY DISTRESS SYNDROME AND SUBSEQUENT DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | K022269 | TIMING OF COMPLICATIONS IN OPEN AIRWAY RECONSTRUCTION. THE AIM OF THIS STUDY IS TO INVESTIGATE THE INCIDENCE AND TIMING OF COMPLICATIONS THAT OCCUR AFTER TRACHEAL RESECTION (TR)/CRICOTRACHEAL RESECTION (CTR), TO BETTER CHARACTERIZE WHEN CLINICAL SUSPICION SHOULD REMAIN HIGHEST AND GUIDE POST-OPERATIVE MANAGEMENT. BETWEEN 2015 TO 2022, SIXTY-NINE CASES OF TR/CTR WERE REVIEWED. ALL PATIENTS WHO UNDERWENT TR/CTR DURING THIS TIME PERIOD WERE ANALYZED FOR TIMING OF COMPLICATIONS IN DAYS. DEMOGRAPHIC DATA AND CLINICAL FACTORS WERE COLLECTED FOR A SECONDARY AIM OF ASSESSING RISK FACTORS. A SUBSET OF NINE PATIENTS WERE EXCLUDED FROM RISK FACTOR ANALYSIS GIVEN INCOMPLETE MEDICAL RECORDS. CASES WERE PERFORMED BY FOUR SURGEONS, ALL USING THE SAME ANASTOMOSIS TECHNIQUE WITH FULL-THICKNESS INTERRUPTED VICRYL SUTURES. REPORTED COMPLICATIONS ARE THE FOLLOWING: VICRYL SUTURES (ETHICON) POST-SURGICAL NECK HEMATOMA OCCURRED BETWEEN 1 AND 5 DAYS AFTER SURGERY (N=4) TREATMENT: TWO WERE MANAGED WITH BEDSIDE INCISION AND DRAINAGE AND TWO UNDERWENT EVACUATION IN THE OPERATING ROOM. ANASTOMOTIC DEHISCENCE (WITH DEFINITION RANGING FROM SUBCUTANEOUS EMPHYSEMA TO BRONCHOSCOPICALLY-VISUALIZED AIR LEAK) OCCURRED ON AVERAGE OF 7.5 DAYS AFTER SURGERY (N=4) TREATMENT: TWO PATIENTS WERE MANAGED CONSERVATIVELY WITH ANTIBIOTICS AND LOCAL WOUND PACKING, AND TWO PATIENTS HAD SIGNIFICANT DEHISCENCE REQUIRING REVISION TRACHEOSTOMY. UNKNOWN EVENT ON FIRST OCCURRING 4.6 DAYS AFTER SURGERY ON AVERAGE (N=8) TREATMENT: NOT SPECIFIED BUT THERE WERE PATIENT HAD REQUIRED RE-INTUBATION AT LEAST ONCE WITH UNKNOWN CAUSE EARLY RESTENOSIS (N=1), ANASTOMOTIC DEHISCENCE (N=2), PROLONGED REINTUBATION IN THE SETTING OF ARDS (N=1) ON AVERAGE OCCURRING 10.3 DAYS POST-OP TREATMENT: REQUIRED REVISION TRACHEOSTOMY RESTENOSIS AT THE ANASTOMOTIC LINE DEVELOPED THIS DURING THEIR POST-OPERATIVE HOSPITALIZATION AT DAY 8 (N=1) TREATMENT: MANAGED WITH REINTUBATION FOLLOWED BY TRACHEOSTOMY IN THE ACUTE SETTING AND HAS GONE ON TO DEVELOP NEAR-COMPLETE SUPRASTOMAL/ANASTOMOTIC STENOSIS. RESTENOSIS AT THE ANASTOMOTIC LINE DEVELOPED BETWEEN 25 AND 92 DAYS POST-OP (N=6) TREATMENT: ALL SEEN DURING ROUTINE FOLLOW-UP APPOINTMENTS ON WHICH ONE PATIENT HAD ONLY MILD STENOSIS AND WAS OBSERVED WITHOUT INTERVENTION. THE OTHER FIVE ALL REQUIRED AT LEAST TWO FURTHER OPERATIVE ENDOSCOPIC DILATIONS AND INTRALESIONAL STEROID INJECTIONS WITH CONTROL OF THEIR SYMPTOMS. VICRYL SUTURES (ETHICON): DEATH (N=3) DUE TO SEVERE ARDS (N=2) AND ADVANCED AIRWAY INVOLVEMENT OF THYROID CARCINOMA DEVELOPED ANASTOMOTIC DEHISCENCE LIKELY DUE TO LOCAL INFECTION, AND ONE SUFFERED AIRWAY LOSS AND CARDIAC ARREST IN THE OPERATING ROOM AND WITHDRAWAL OF CARE 7 DAYS POST OPERATIVELY (N=1). TREATMENT: NONE. IN CONCLUSION, IN THIS SINGLE-CENTER STUDY, THE MAJORITY OF POST OPERATIVE COMPLICATIONS AFTER TR OR CTR, INCLUDING HEMATOMA AND ANASTOMOTIC DEHISCENCE, OCCURRED WITHIN THE FIRST 8 DAYS. THUS, PATIENTS WITH SIGNIFICANT MEDICAL COMORBIDITIES OR LONGER-SEGMENT TRACHEAL STENOSIS MAY BENEFIT FROM A PERIOD OF LONGER OBSERVATION IN THE HOSPITAL. MOST CASES OF RESTENOSIS WERE NOTED APPROXIMATELY 1¿ 3 MONTHS AFTER SURGERY AT ROUTINE FOLLOW-UP VISITS. THIS MAY INFORM CLINICAL DECISION-MAKING REGARDING PATIENT MONITORING IN THE ACUTE SETTING AS WELL AS POST-OPERATIVE SURVEILLANCE AFTER OPEN AIRWAY SURGERY. | PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: LARYNGOSCOPE. 2024 AUG;134(8):3527-3531. DOI: 10.1002/LARY.31362. EPUB 2024 FEB 23. PMID: 38390693. HTTPS://DOI.ORG/10.1002/LARY.31362.
DREAMSTATION AUTO BIPAP | K131982 | A RELATIVE OF THE PATIENT REPORTED THE PATIENT EXPIRED. THE REPORTER STATED THE PATIENT WAS NOT USING THE DEVICE AT THE TIME THE PATIENT PASSED AWAY. THE REPORTER STATED THE PATIENT WAS USING ANOTHER UNSPECIFIED DEVICE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
REMSTAR SYSTEM ONE 60 SERIES | K113068 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THE PATIENT'S RELATIVE IS RETURNING THEIR DEVICE. THE CAUSE OF DEATH WAS NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
INTELLIVUE MX550 PATIENT MONITOR | K131872 | IT WAS REPORTED THE MONITOR DID NOT GENERATE ALARMS FOR APPROXIMATELY 25 MINUTES AND THE PATIENT PASSED AWAY. ENGINEERING REVIEWED THE CLINICAL AUDIT LOGS, REVEALING BOTH VISUAL AND AUDIBLE RED AND YELLOW ALARMS HAD OCCURRED AT BOTH THE MONITOR AND CENTRAL STATION. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE WAS RETURNED TO THE THIRD-PARTY SERVICE CENTER FOR FURTHER EVALUATION. DURING THE EVALUATION OF THE DEVICE AT THE THIRD-PARTY SERVICE CENTER, THE DEVICE WAS VISUALLY INSPECTED AND FOUND NO EVIDENCE OF FOAM DEGRADATION. THE DEVICE POWERED ON AND AIRFLOW WAS CONFIRMED. THE DEVICE'S DOWNLOADED LOGS WERE REVIEWED BY THE THIRD-PARTY SERVICE CENTER. THERE WAS NO ERROR CODE FOUND. THE THIRD-PARTY SERVICE CENTER CONCLUDES THAT THEY COULDN'T CONFIRM THE CUSTOMER'S ALLEGATION AND THERE WAS NO VISIBLE FOAM DEGRADATION AND UNIT GOT SCRAPPED.
TEMPUS LS | K200849 | THIS REPORT IS BASED ON INFORMATION PROVIDED BY REMOTE SERVICE ENGINEER RSE AND IS CURRENTLY UNDER INVESTIGATION BY THE PHILIPS COMPLAINT HANDLING TEAM. PHILIPS RECEIVED A COMPLAINT ON THE TEMPUS LS MANUAL INDICATING THE SCREEN FAILED TO PROPERLY SHOW ECG (ELECTROCARDIOGRAM) RHYTHM WHILE ON A CALL. THE SCREEN WAS CLEANED AFTER RETURNING TO STATION, AND AS SOON AS PRESSURE WAS APPLIED, SCREEN BEGAN TO FUNCTION PROPERLY. THE PATIENT WAS IN CARDIAC ARREST AND BEING MONITORED THROUGH THE DEFIBRILLATOR PADS. THE ECG (ELECTROCARDIOGRAM) RHYTHM WAS SUCCESSFULLY DISPLAYED ON THE TEMPUS PRO, BUT NOT THE TEMPUS LS MANUAL, AS SUCH, DID NOT IMPACT THE CARE OR OUTCOME OF THIS PATIENT.
ALARIS SYSTEM | K133532 | IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON (B)(6) 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
ALARIS SYSTEM | K133532 | IT WAS INITIALLY REPORTED AN OVER INFUSION EVENT THAT OCCURRED ON 27 MAY 2025. PER REPORT, THE PATIENT "DIED." THE CUSTOMER IS REQUESTING THE MANUFACTURER TO REVIEW THE DEVICE LOGS TO IDENTIFY WHICH MEDICATIONS WERE INFUSED FROM 3:00 PM ON (B)(6) 2025 TO 1:30 AM ON (B)(6) 2025. NO FURTHER INFORMATION WAS PROVIDED. BD RECEIVED ADDITIONAL INFORMATION. IT WAS REPORTED THAT "THE REASON TO DO THIS (REVIEW OF DEVICE LOGS) ONLY TO CHECK THE MEDICATIONS AND THE RATE OF IT GIVEN BY THE MACHINE. I¿M NOT BELIEVE THAT THE DEVICE CAUSED THE DEATH BUT MAYBE HUMAN MISTAKE." MULTIPOLE ATTEMPTS HAVE BEEN MADE REQUESTING FOR THE DETAILS OF THE EVENT. HOWEVER, NO FURTHER INFORMATION HAS BEEN PROVIDED TO DATE. | THE ACTUAL DATE OF DEATH IS UNKNOWN. THE EVENT LOGS HAVE BEEN RECEIVED AND THE EVALUATION IS PENDING. A FOLLOW UP REPORT WILL BE SUBMITTED WITH INVESTIGATION RESULTS ONCE THE EVALUATION HAS BEEN COMPLETED. PER 803.52(F)(11)(III) THE INFORMATION PROVIDED REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. THE COMPLAINANT OR REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO THE MANUFACTURER.
DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | K213919 | IT WAS REPORTED VIA WEB SELF-SERVICE THAT A PATIENT PASSED AWAY. THE REPORTER STATED, "MY BOYFRIEND DIED BECAUSE HIS DEXCOM WAS HOOKED TO A PUMP.". NO OTHER EVENT DETAILS WERE PROVIDED AND FOLLOW UP ATTEMPTS TO THE REPORTER FOR ADDITIONAL EVENT DETAILS WERE UNSUCCESSFUL. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. THE ALLEGATION AND PROBABLE CAUSE COULD NOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. | (B)(4) H6 MEDICAL DEVICE PROBLEM CODE - 3191 - PATIENT WAS UNABLE TO IDENTIFY DEVICE ISSUE THEREFORE A MORE SPECIFIC CODE COULD NOT BE SELECTED.
BEAD BLOCK | K150876 | IT WAS REPORTED VIA LITERATURE THAT DEATH OCCURRED. PURPOSE: ABLATIVE THERAPIES FOR PRIMARY AND SECONDARY LIVER MALIGNANCIES ARE INCREASINGLY ADOPTED IN CURRENT GUIDELINES. NEVERTHELESS, SURGICAL RESECTION REMAINS THE GOLD STANDARD IN MOST CURATIVE THERAPY SETTINGS. HOWEVER, THERE ARE INDICATIONS FOR ABLATIVE TREATMENTS SUCH AS THERMAL ABLATION, HIGH-CONFORMAL RADIATION, AND LOCO-REGIONAL THERAPIES. WE INVESTIGATED COMPLICATIONS AND MORTALITY AFTER ABLATIVE TREATMENT IN A LARGE, UNSELECTED STUDY COHORT. THIS STUDY EVALUATES TREATMENT DATA FROM THE GERMAN DRG-BASED HOSPITAL REIMBURSEMENT SYSTEM OF ABLATIVE PROCEDURES OF PRIMARY AND SECONDARY LIVER MALIGNANCIES PERFORMED BETWEEN 2005 AND 2020. OVER 15 YEARS, THERE WERE 4374 ABLATIVE PROCEDURES OF THE LIVER PERFORMED. PATIENTS UNDERWENT 3085 HIGH-DOSE-RATE BRACHYTHERAPY (HDR- BT) (70.5%), 296 THERMAL ABLATIONS (6.8%), 111 TRANSARTERIAL CHEMOEMBOLIZATION (TACE)(2.5%), AND 882 SELECTIVE INTERNAL RADIATION THERAPY (SIRT) (20.2%). EVALUATION OF THE TREATMENT DATA FOCUSED ON PRE-EXISTING MEDICAL CONDITIONS, LENGTH OF HOSPITALIZATION, MAJOR COMPLICATIONS DURING THE HOSPITALIZATION, AND IN-HOSPITAL DEATH. PROCEDURE SUMMARY - TRANSARTERIAL CHEMOEMBOLIZATION (TACE): WE INSERTED A 5F DIAGNOSTIC ANGIOGRAPHIC CATHETER USING TRANSFEMORAL ACCESS AND A COAXIAL NON-BOSTON SCIENTIFIC 2.0F MICROCATHETER TO DISPLAY TUMOR-SUPPLYING ARTERIES SELECTIVELY. WE INJECTED DC BEADS M1 70 -150 MICROMETRE PARTICLES, WHICH WERE LOADED WITH 50 -75 MG OF DOXORUBICIN (DEBDOX). NONLOADED MICROPARTICLES OF DIFFERENT SIZES FROM BEAD BLOCK (100-300 MICRONS AND 300-500 MICRONS) WERE USED TO REACH FLOW STAGNATION IF THE INJECTION OF DEBDOX WAS INSUFFICIENT. EVENT SUMMARY THERE WERE 5 IN-HOSPITAL DEATHS IN TREATMENT CASES WITH TACE. | MARCH, CHRISTINE, ET AL. "MORTALITY AND POSTINTERVENTIONAL COMPLICATIONS AFTER ABLATIVE TREATMENT OF LIVER MALIGNANCIES: A COHORT STUDY OF 4374 PATIENTS." BRACHYTHERAPY 23.6 (2024): 743-750. B3: DATE OF EVENT: USED FIRST DAY IN THE FIRST MONTH OF THE YEAR OF THE ARTICLE. D2B: PRO CODE (PRODUCT CODE): HCG. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC AS THE EVENT WAS REPORTED THROUGH LITERATURE. BASED ON THE NATURE OF THE INFORMATION PROVIDED TO BSC, IT IS NOT POSSIBLE TO PERFORM A GOOD FAITH EFFORT TO OBTAIN ADDITIONAL INFORMATION. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WE ARE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | A PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) REPORTED TO FRESENIUS CUSTOMER SUPPORT THAT A PATIENT PASSED AWAY ON THE CYCLER. THE CAUSE OF DEATH IS UNKNOWN. A RED BAG PICKUP FROM THE CLINIC PROCESS WAS REQUESTED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT DISASSOCIATE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED WHILE UNDERGOING TREATMENT ON (B)(6) 2025. THE PATIENT REPORTEDLY AWOKE GASPING FOR AIR AND PASSED SHORTLY THEREAFTER. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY. CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS.
PANORAMA CENTRAL STATION WITH TELEMETRY | K150082 | IT WAS REPORTED THAT ON (B)(6) 2025, BETWEEN 6:30 AM AND 7:30 AM, CLINICAL STAFF DID NOT HEAR AN ALARM AT THE WORKSTATION (PANORAMA CENTRAL STATION). IN ADDITION, IT WAS REPORTED THE PATIENT WAS ONLY CONNECTED TO PATIENT MONITORING (MINDRAY V12). THE PATIENT EXPIRED. | A SIMULATOR WAS CONNECTED TO AN EMERGENCY DEPARTMENT MONITOR, TRIGGERED AN ALARM, AND CONFIRMED THAT THERE WAS NO AUDIO COMING OUT OF THE WORKSTATION. SYSTEM LOGS WERE NOT ABLE TO BE COLLECTED AS THE SYSTEM INDICATED A BAD HDD. CUSTOMER BIOMED, WITH THE ASSISTANCE OF MINDRAY FIELD SERVICE, INVESTIGATED THE AUDIO PROBLEMS OF THE CENTRAL STATION AND IT WAS FOUND THAT THE STAFF NOT HEARING THE AUDIBLE ALARM WAS CAUSED BY THE INCORRECT CONNECTION OF THE SPEAKER CABLES TO THE KVM IN THE SERVER ROOM. THIS LED TO THE AUDIO OF THE ALARM NOT SOUNDING AT THE NURSE'S STATION. IT IS LIKELY THE EQUIPMENT WAS ALTERED AFTER THE MANUFACTURER INSTALLATION. THE MINDRAY REPRESENTATIVE ADJUSTED THE CABLES THAT WERE INCORRECTLY CONNECTED. HE MOVED THE TWO CABLES MISTAKENLY INSERTED INTO THE MICROPHONE PORTS TO THE CORRESPONDING SPEAKER PORTS. AFTER THE ADJUSTMENT WAS COMPLETED, THE AUDIO ALARMS FUNCTIONED PER SPECIFICATIONS. MINDRAY PANORAMA CENTRAL STATION PERFORMED PER SPECIFICATIONS.
DAVINCI X | K171294 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED UMBILICAL HERNIA TAPP PROCEDURE, THE COLON WAS INADVERTENTLY INJURED, LEADING TO AN INFECTION AND A SUBSEQUENT PROCEDURE, AFTER WHICH THE PATIENT EXPIRED. THE INJURY OCCURRED DURING THE INITIAL PORT INSERTION USING THE BLUNT OBTURATOR AND CANNULA [TROCAR] ACCESSORY. THERE WAS NO BLEEDING, AND THE INJURY WAS IMMEDIATELY IDENTIFIED AND CLOSED WITH SUTURES. THE PROCEDURE WAS COMPLETED ROBOTICALLY WITHOUT FURTHER COMPLICATIONS AND THE SURGEON DOES NOT BELIEVE THAT A MALFUNCTION OF A DA VINCI SYSTEM, INSTRUMENT, OR ACCESSORY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE PATIENT WAS REPORTEDLY MORBIDLY OBESE AND THE HERNIA WAS LARGE AND OBSTRUCTED. SHORTLY AFTER THE PROCEDURE, THE PATIENT DEVELOPED ABDOMINAL PAIN, SWELLING, AND DIFFICULTY BREATHING, AND WAS ADMITTED TO THE ICU WITH SIGNS OF PERITONITIS. ON POSTOPERATIVE DAY THREE, AN EXPLORATORY LAPAROTOMY REVEALED AN INTESTINAL INFECTION CAUSED BY LEAKAGE AT THE SITE OF THE INITIAL INJURY. ACCORDING TO THE HOSPITAL¿S STATEMENT, A TEAM OF "SPECIALIST DOCTORS" PROVIDED ALL POSSIBLE MEDICAL CARE. HOWEVER, THE PATIENT¿S CONDITION DETERIORATED DUE TO INFECTION AND MULTI-ORGAN FAILURE, ULTIMATELY RESULTING IN DEATH. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, MEGA SUTURECUT NEEDLE DRIVER, FORCE BIPOLAR WITH DUAL GRIP. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT SUFFERED AN INJURY TO THE BOWEL UPON PLACING A BLUNT OBTURATOR AND PORT. HOW THE INJURY OCCURRED AND HOW THE SURGEON¿S TECHNIQUE MAY HAVE CONTRIBUTED TO THIS EVENT WERE NOT PROVIDED. THIS WAS RECOGNIZED AND REPAIRED AT THE TIME OF THE INITIAL INJURY. HOW THE REPAIR WAS DONE OR ANY SAFEGUARDS TAKEN TO CONFIRM THE REPAIR WAS ADEQUATE WERE NOT PROVIDED. THE PATIENT SUBSEQUENTLY DEVELOPED A LEAK AND REQUIRED RE-EXPLORATION. THEY ULTIMATELY DIED AFTER ADDITIONAL COMPLICATIONS. NO AUTOPSY WAS PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. .
DAVINCI XI | K131861 | IT WAS REPORTED THAT THE PATIENT UNDERWENT A DA VINCI-ASSISTED ROUX-EN-Y GASTRIC BYPASS PROCEDURE. THE PATIENT SUBSEQUENTLY EXPIRED ON POST-OPERATIVE DAY 6. NO INFORMATION REGARDING THE EVENTS LEADING UP TO THE DEATH OR THE CAUSE OF DEATH WERE PROVIDED. THERE WAS NO REPORT OF ANY INTRAOPERATIVE COMPLICATIONS. THE SITE VICE PRESIDENT OF BARIATRICS REPORTED THAT NO SPECIFIC INFORMATION REGARDING THE POST-OPERATIVE COURSE WOULD BE AVAILABLE; HOWEVER, THERE WAS A POTENTIAL THAT THE PATIENT'S CAUSE OF DEATH COULD BE ATTRIBUTED TO A POST-OPERATIVE OVERDOSE. NO ADDITIONAL INFORMATION WAS PROVIDED. THERE WAS NO ALLEGATION THAT A DA VINCI SYSTEM, INSTRUMENT, OR ACCESSORY MALFUNCTION OCCURRED DURING THE PROCEDURE. | THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, PERMANENT CAUTERY HOOK, TIP-UP FENESTRATED GRASPER, CADIERE FORCEPS, AND MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. DHR REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. REVIEW OF THE SUREFORM 60 STAPLER LOGS FOUND USED IN THIS EVENT FOUND THE INSTRUMENT WAS INSTALLED ON THE SYSTEM 8 TIMES AND FIRED 7 WHITE RELOADS. ON INSTALL 1, A WHITE RELOAD WAS LOADED, HOWEVER NO CLAMPING OR FIRING WAS ATTEMPTED. ON INSTALL 2, THE FIRST CLAMP ATTEMPT STOPPED AT APPROXIMATELY 94% COMPLETION. THE NEXT TWO CLAMPS WERE SUCCESSFUL, AND THE FIRING WAS COMPLETED WITH 2-3 PAUSES FOR COMPRESSION. ON INSTALLS 3-5, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE COMPLETED WITH 1 PAUSE FOR COMPRESSION EACH. ON INSTALLS 6-8, THE FIRST CLAMP WAS SUCCESSFUL, AND THE FIRINGS WERE COMPLETED WITH NO PAUSES FOR COMPRESSION EACH. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS, AND THERE WAS NO REPORT OF ANY ISSUES FROM THE OR TEAM REGARDING THE STAPLER INSTRUMENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT PATIENT IN THIS REPORT DIED 6 DAYS FOLLOWING A REPORTEDLY UNEVENTFUL GASTRIC BYPASS PROCEDURE. HOW THEY DIED AND WHAT EVENTS LED TO THEIR DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER | K203365 | A SHOCKWAVE M5+ PERIPHERAL INTRAVASCULAR LITHOTRIPSY (IVL) CATHETER WAS USED WAS USED TO TREAT THE SUPERFICIAL FEMORAL ARTERY (SFA) AND POPLITEAL ARTERY. THE SHOCKWAVE CATHETER DELIVERED 300 PULSES AT A PULSE PRESSURE OF 4ATM SUCCESSFULLY. IT WAS REPORTED THAT THE PATIENT EXPERIENCED COMPLICATIONS FOLLOWING STENT DELIVERY. IT WAS REPORTED THAT THE SHOCKWAVE DEVICE HAD NO IMPACT ON THE PATIENT'S COMPLICATIONS. ACCORDING TO THE MANAGEMENT TEAM, THE PATIENT DEATH OCCURRED POST PROCEDURE, AFTER BEING TRANSFERRED TO THE HEART CENTER. THERE WAS NO ADDITIONAL INFORMATION KNOWN AND/OR PROVIDED REGARDING THE EVENT. BASED ON THE INFORMATION PROVIDED, THE PROCEDURE WAS PERFORMED ON (B)(6) 2025 AND PATIENT DEATH OCCURRED ON THE SAME DAY. | THE DEVICE REFERENCED IN THIS REPORT WAS NOT RETURNED TO SHOCKWAVE MEDICAL, INC. THEREFORE, A PHYSICAL EXAMINATION CANNOT BE PERFORMED. BASED ON THE REPORTED INFORMATION, THERE WAS NO IVL MALFUNCTION AND NO INFORMATION SUGGESTING THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. MULTIPLE ATTEMPTS WERE MADE TO GATHER ADDITIONAL INFORMATION RELATED TO THE EVENT WITH LIMITED INFORMATION TO DATE. SHOCKWAVE MEDICAL HAS CONTROLS IN PLACE TO ENSURE DEVICES ARE BUILT TO APPROVED PROCEDURES AND MEET LOT RELEASE ACCEPTANCE CRITERIA PRIOR TO BEING DISTRIBUTED. A REVIEW OF THE MANUFACTURING AND TEST DOCUMENTATION FOR SUBJECT LOT DOES NOT REVEAL ANY ISSUES WITH THE MANUFACTURING OF THE DEVICE. THE DEVICE PASSED ALL OF SHOCKWAVE MEDICAL, INC. ACCEPTANCE CRITERIA PRIOR TO SHIPPING. COMPLAINT TRENDS WILL CONTINUE TO BE MONITORED.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER ALLEGING VISUALIZATION OF PARTICLES RELATED TO A CPAP DEVICE'S SOUND ABATEMENT FOAM. NO MEDICAL INTERVENTION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
VICRYL POLYGLACTIN 910 SUTURE UNKNOWN PRODUCT | K022269 | TITLE: INSTITUTIONAL EXPERIENCE ON COMPARISON OF ALLOGRAFTS WITH MULTIPLE-RENAL ARTERY VS. SINGLE RENAL ARTERY IN LIVING AND DECEASED DONOR TRANSPLANTATION THE AIM OF THIS STUDY IS TO COMPARE UROLOGICAL (URINE LEAK AND STENOSIS) AND VASCULAR COMPLICATIONS (RENAL ARTERY AND VEIN THROMBOSIS) IN RENAL ALLOGRAFTS WITH A SINGLE ARTERY VS. MULTIPLE ARTERIES. ADDITIONALLY, THE SECONDARY OBJECTIVES INCLUDED ASSESSING DELAYED GRAFT FUNCTION AND GRAFT SURVIVAL AT 3- AND 6-MONTHS POST-TRANSPLANTATION. BETWEEN APRIL 2022 TO APRIL 2024, A TOTAL OF E 95 PATIENTS, 67.4% (64) WERE MALE, AND 32.6% (31) WERE FEMALE. THE MEAN AGE OF THE PARTICIPANTS WAS 34.65 ± 10.41 YEARS. FOR MRA GRAFTS A POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) WAS USED FOR THE VASCULAR CONNECTIONS. AT THE SAME TIME IN ALL CASE, MODIFIED LICH-GREGOIR URETERONEOCYSTOSTOMY WAS PERFORMED USING 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). REPORTED COMPLICATIONS ARE : POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON). -ACUTE TUBULAR NECROSIS (N=10). TREATMENT: NOT PROVIDED. -VASCULAR COMPLICATIONS (N=3). TREATMENT: NOT PROVIDED. -UROLOGICAL COMPLICATION (N=5). TREATMENT: NOT PROVIDED. 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). -ACUTE TUBULAR NECROSIS (N=10). TREATMENT: NOT PROVIDED. -VASCULAR COMPLICATIONS (N=3). TREATMENT: NOT PROVIDED. -EXPERIENCING GRAFT FAILURE (N=1). TREATMENT: NOT PROVIDED. -UROLOGICAL COMPLICATION (N=5). TREATMENT: NOT PROVIDED. POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3). TREATMENT: NOT PROVIDED. 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3). TREATMENT: NOT PROVIDED. IN CONCLUSION, THIS STUDY DEMONSTRATES THAT RENAL TRANSPLANTATION OUTCOMES ARE SUPERIOR TO THOSE OF SINGLE RENAL ARTERY ALLOGRAFTS COMPARED TO THOSE OF MULTIPLE RENAL ARTERIES. IMMEDIATE GRAFT FUNCTION PLAYS A CRITICAL ROLE IN ENSURING GRAFT SURVIVAL AND MULTIPLE RENAL ARTERIES ARE ASSOCIATED WITH HIGHER VASCULAR COMPLICATIONS AND MORTALITY. THESE FINDINGS HIGHLIGHT THE IMPORTANCE OF SURGICAL EXPERTISE, METICULOUS PERIOPERATIVE CARE, AND TARGETED POSTOPERATIVE MANAGEMENT OF PATIENTS WITH ANATOMICALLY COMPLEX ALLOGRAFTS. | PRODUCT COMPLAINT # (B)(4). THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON INC, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON INC, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. ATTEMPTS ARE BEING MADE TO OBTAIN THE FOLLOWING INFORMATION. TO DATE NO RESPONSE HAS BEEN PROVIDED. IF FURTHER DETAILS ARE RECEIVED AT A LATER DATE A SUPPLEMENTAL MEDWATCH WILL BE SENT. DOES THE SURGEON BELIEVE THAT ANY OF THE ETHICON PRODUCTS INVOLVED CAUSED AND/OR CONTRIBUTED TO THE POST-OPERATIVE COMPLICATIONS/DEATH DESCRIBED IN THE ARTICLE? WHICH SPECIFIC ETHICON PRODUCTS HAVE BEEN USED DURING THE PROCEDURES (PRODUCT CODE, LOT NUMBER)? DOES THE SURGEON BELIEVE THERE WAS ANY DEFICIENCY WITH ANY OF THE ETHICON PRODUCTS USED IN THIS PROCEDURE? IF SO, PLEASE PROVIDE DETAILS. WERE THE CASES DISCUSSED IN THIS ARTICLE PREVIOUSLY REPORTED TO ETHICON? IF YES, PLEASE PROVIDE A COMPLAINT REFERENCE NUMBER. PATIENT DEMOGRAPHICS? HAS AN AUTOPSY BEEN PERFORMED FOR DEATH CASE? IF SO, CAN THE RESULTS BE SHARED? WOULD THE JOURNAL AUTHOR / PHYSICIAN LIKE TO SPEAK WITH ETHICON MEDICAL SAFETY AND ENGINEERING VIA SCHEDULED CONFERENCE CALL REGARDING THE PRODUCT INVOLVED IN THIS EVENT? THIS REPORT IS RELATED TO A JOURNAL ARTICLE; THEREFORE, NO PRODUCT WILL BE RETURNED FOR ANALYSIS AND THE BATCH HISTORY RECORDS CANNOT BE REVIEWED AS THE LOT NUMBER HAS NOT BEEN PROVIDED. H6 COMPONENT CODE: G07002 ¿ DEVICE NOT RETURNED. D4: UDI: AS THE CATALOG/MODEL NUMBER WAS NOT PROVIDED, THE (01)GTIN IS NOT AVAILABLE. THE SINGLE COMPLAINT WAS REPORTED WITH MULTIPLE EVENTS. THERE ARE NO ADDITIONAL DETAILS REGARDING THE ADDITIONAL EVENTS. CITATION: INT J ACAD MED PHARM 2025; 7 (1); 6-10. DOI: 10.47009/JAMP.2025.7.1.2.
PROLENE POLYPROPYLENE SUTURE UNKNOWN PRODUCT | K133356 | TITLE: INSTITUTIONAL EXPERIENCE ON COMPARISON OF ALLOGRAFTS WITH MULTIPLE-RENAL ARTERY VS. SINGLE RENAL ARTERY IN LIVING AND DECEASED DONOR TRANSPLANTATION THE AIM OF THIS STUDY IS TO COMPARE UROLOGICAL (URINE LEAK AND STENOSIS) AND VASCULAR COMPLICATIONS (RENAL ARTERY AND VEIN THROMBOSIS) IN RENAL ALLOGRAFTS WITH A SINGLE ARTERY VS. MULTIPLE ARTERIES. ADDITIONALLY, THE SECONDARY OBJECTIVES INCLUDED ASSESSING DELAYED GRAFT FUNCTION AND GRAFT SURVIVAL AT 3- AND 6-MONTHS POST-TRANSPLANTATION. BETWEEN APRIL 2022 TO APRIL 2024, A TOTAL OF E 95 PATIENTS, 67.4% (64) WERE MALE, AND 32.6% (31) WERE FEMALE. THE MEAN AGE OF THE PARTICIPANTS WAS 34.65 ± 10.41 YEARS. FOR MRA GRAFTS A POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) WAS USED FOR THE VASCULAR CONNECTIONS. AT THE SAME TIME IN ALL CASE, MODIFIED LICH-GREGOIR URETERONEOCYSTOSTOMY WAS PERFORMED USING 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON). REPORTED COMPLICATIONS ARE: POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) ACUTE TUBULAR NECROSIS (N=10). TREATMENT: NOT PROVIDED VASCULAR COMPLICATIONS (N=3). TREATMENT: NOT PROVIDED UROLOGICAL COMPLICATION (N=5). TREATMENT: NOT PROVIDED 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON) . -ACUTE TUBULAR NECROSIS (N=10). TREATMENT : NOT PROVIDED -VASCULAR COMPLICATIONS (N=3). TREATMENT : NOT PROVIDED -EXPERIENCING GRAFT FAILURE (N=1). TREATMENT : NOT PROVIDED -UROLOGICAL COMPLICATION (N=5). TREATMENT : NOT PROVIDED POLYPROPYLENE SUTURES (PROLENE 6-0; ETHICON) -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3) TREATMENT : NOT PROVIDED. 5-0 POLYGLACTIN SUTURES (VICRYL; ETHICON) -MORTALITY WAS OBSERVED, ALL OF WHICH WERE ATTRIBUTED TO SEPSIS (N=3) TREATMENT : NOT PROVIDED . IN CONCLUSION, THIS STUDY DEMONSTRATES THAT RENAL TRANSPLANTATION OUTCOMES ARE SUPERIOR TO THOSE OF SINGLE RENAL ARTERY ALLOGRAFTS COMPARED TO THOSE OF MULTIPLE RENAL ARTERIES. IMMEDIATE GRAFT FUNCTION PLAYS A CRITICAL ROLE IN ENSURING GRAFT SURVIVAL AND MULTIPLE RENAL ARTERIES ARE ASSOCIATED WITH HIGHER VASCULAR COMPLICATIONS AND MORTALITY. THESE FINDINGS HIGHLIGHT THE IMPORTANCE OF SURGICAL EXPERTISE, METICULOUS PERIOPERATIVE CARE, AND TARGETED POSTOPERATIVE MANAGEMENT OF PATIENTS WITH ANATOMICALLY COMPLEX ALLOGRAFTS. | PRODUCT COMPLAINT # (B)(4). CORRECTED INFORMATION: B1, B2, H1 - ADDITIONAL INFORMATION WAS RECEIVED AND IT WAS REPORTED THAT THIS EVENT NO LONGER MEETS REPORTABLE CRITERIA. THEREFORE, THIS MEDWATCH REPORT IS BEING VOIDED.
DREAMSTATION ST30 | K102465 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE PREVIOUS REPORT HAD THE INCORRECT HEALTH IMPACT CODE. THIS HAS BEEN UPDATED TO DEATH UNDER THE HEALTH IMPACT GRID.
CAPIOX®FX | K071494 | THE USER FACILITY REPORTED THAT A SMALL BLOOD LEAK ARE DETECTED AT THE ARTERIAL CONNECTION PORT OUTLET DURING THE SETUP OF THE EXTRACORPOREAL CIRCUIT FOR THE PATIENT'S SURGERY. THE ISSUE WAS IDENTIFIED WHEN THE CIRCUIT WAS PRIMED. SINCE THE PATIENT WAS ALREADY CONNECTED TO THE CARDIOPULMONARY BYPASS (CPB) AND THERE WAS INSUFFICIENT TIME TO REPLACE THE CIRCUIT, THE LEAK WAS RESOLVED BY APPLYING A TEGADERM TO PREVENT A HIGH-RISK SITUATION. THE PATIENT EXPERIENCED HYPOTENSION AND AIRWAY BLEEDING, WHICH WAS THE CAUSE OF DEATH. | D6A: IMPLANTED DATE: DEVICE WAS NOT IMPLANTED D6B: EXPLANTED DATE: DEVICE WAS NOT EXPLANTED E3: OCCUPATION: PERFUSIONIST G4: 510K: K1300520. THE ACTUAL DEVICE IS NOT AVAILABLE FOR RETURN; THEREFORE, AN EVALUATION OF THE ACTUAL DEVICE WAS UNABLE TO BE CONDUCTED. THE MANUFACTURING RECORD AND THE SHIPPING INSPECTION RECORD OF THE ACTUAL DEVICE - NO ANOMALY WAS FOUND. PAST COMPLAINT FILE - NO OTHER SIMILAR REPORT OF THE PRODUCT WITH THE INVOLVED PRODUCT CODE/LOT NUMBER WAS FOUND. CAUSE OF OCCURRENCE/CONCLUSION BASED ON THE INVESTIGATION RESULT, NO ANOMALY WAS FOUND IN THE MANUFACTURING RECORDS. SINCE THE ACTUAL DEVICE COULD NOT BE INVESTIGATED, IT WAS NOT POSSIBLE TO CLARIFY THE CAUSE OF OCCURRENCE. RELEVANT IFU REFERENCE: "DO NOT USE AN OXYGENATOR AND RESERVOIR THAT LEAKS. REPLACE IT WITH ANOTHER CAPIOX FX25 OXYGENATOR AND RESERVOIR." TERUMO MEDICAL CORPORATION (TMC) (IMPORTER) REGISTRATION NO. 2243441 IS SUBMITTING THIS REPORT ON BEHALF OF ASHITAKA FACTORY OF TERUMO CORPORATION (MANUFACTURER) REGISTRATION NO. 9681834.
CARDIOHELP | K133598 | MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-06-25: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2025-05-15 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING: RETROPERITONEAL HEMORRHAGE" AND THE CONSEQUENCE FOR THE PATIENT WAS "PATIENT PASSED AWAY¿ NO ADDITIONAL INFORMATION WAS RECEIVED. FURTHER INFORMATION ABOUT THE EVENT AND DEVICE WAS REQUESTED BUT IS STILL PENDING. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. COMPLAINT ID (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE SPAIN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT A TEVAR PROCEDURE TO TREAT AN ANEURYSM UTILIZING A GORE® DRYSEAL FLEX INTRODUCER SHEATH. REPORTEDLY, WHEN GETTING READY TO PUT DEVICE IN THE 24FR SHEATH, PHYSICIAN TOOK THE DILATOR OUT AND IN DOING SO RUPTURED THE COMMON AND INTERNAL ILIAC ARTERY LEADING TO PATIENT DEATH. THERE WERE NO ANATOMICAL CHALLENGES WITH TORTUOSITY OR CALCIUM BUT ONE SIDE HAD SMALLER ACCESS SO PHYSICIAN USED OTHER SIDE ACCESS AND PATIENT DEATH IS DUE TO PROCEDURAL ISSUE WITH THE DILATOR. FURTHER INFORMATION IS NOT AVAILABLE. | C20- A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. IFU STATEMENTS: THE DRYSEAL SHEATH DEVICE INSTRUCTIONS FOR USE (IFU) WERE REVIEWED WITH RESPECT TO THE COMPLAINT DETAIL FOR THE APPLICABLE REGION AND TIME PERIOD, AND THE FOLLOWING IFU STATEMENTS WERE IDENTIFIED IN RELATION TO THE COMPLAINT. WARNINGS: ¿ ADEQUATE VESSEL ACCESS IS REQUIRED TO INTRODUCE THE SHEATH INTO THE VASCULATURE. CAREFUL EVALUATION OF VESSEL SIZE, ANATOMY, TORTUOSITY, AND DISEASE STATE (INCLUDING CALCIFICATION, PLAQUE, AND THROMBUS) IS REQUIRED TO ENSURE SUCCESSFUL SHEATH INTRODUCTION AND SUBSEQUENT WITHDRAWAL. IF VESSEL IS NOT ADEQUATE FOR ACCESS, MAJOR BLEEDING, VESSEL DAMAGE, OR SERIOUS INJURY TO THE PATIENT, INCLUDING DEATH, MAY RESULT. ¿ IF VESSEL SIZE IS SMALLER THAN THE NOMINAL BODY OD (TABLE 1), MAJOR BLEEDING, VESSEL DAMAGE, OR SERIOUS INJURY TO THE PATIENT, INCLUDING DEATH, MAY RESULT. INSTRUCTIONS FOR OPERATION OR USE: SHEATH PREPARATION - 2. VERIFY THE VESSEL IS OF ADEQUATE DIAMETER AND TORTUOSITY TO ACCOMMODATE THE INTRODUCER SHEATH. 8. INSERT THE DILATOR TIP THROUGH THE VALVE AND INTO THE SHEATH UNTIL THE LOCKING CAP ON THE DILATOR IS IN CONTACT WITH THE MATING SURFACE OF THE VALVE. LOCK THE DILATOR TO THE VALVE BODY BY TWISTING THE LOCKING CAP (CLOCKWISE) ON THE DILATOR UNTIL THE POINTER ON THE DILATOR CAP ALIGNS WITH THE ¿LOCK¿ ICON ON THE VALVE BODY. THIS ENSURES THAT THE TAPERED PORTION OF THE DILATOR IS BEYOND THE LEADING END OF THE INTRODUCER SHEATH TIP. THIS WILL OPTIMIZE THE FLEXIBILITY OF THE LEADING END OF THE DILATOR AND ENSURE THAT THE DILATOR STAYS IN PLACE WHILE ADVANCING THE INTRODUCER SHEATH WITH DILATOR INTO THE PATIENT¿S ACCESS VESSEL. 9. HYDROPHILIC COATING ACTIVATION: WET THE ENTIRE OUTER SURFACE OF THE SHEATH TUBE WITH EITHER STERILE SALINE OR WATER TO ACTIVATE THE HYDROPHILIC COATING. NOTE: IT IS IMPORTANT TO KEEP THE SHEATH TUBE OUTER SURFACE WET/SLIPPERY TO TOUCH THROUGHOUT THE PROCEDURE AND DURING INSERTION OF THE INTRODUCER SHEATH. ONCE THE HYDROPHILIC COATING IS ACTIVATED, THE SHEATH OUTER TUBE WILL BE WET AND SLIPPERY. ACTIVATE THE INTRODUCER SHEATH COATING IMMEDIATELY PRIOR TO INSERTING INTO THE VASCULATURE. FOR PROCEDURES OF EXTENDED DURATION, IT MAY BE NECESSARY TO REACTIVATE THE HYDROPHILIC COATING. THIS CAN BE ACHIEVED THROUGH MINOR ROTATIONAL MOVEMENT OF THE SHEATH TO ALLOW BLOOD TO REACTIVATE THE COATING. ENSURE THE DILATOR IS LOCKED TO THE INTRODUCER SHEATH VALVE PRIOR TO INSERTION INTO THE ACCESS VESSEL OR PRIOR TO ANY SHEATH ADVANCEMENT WHILE IN THE PATIENT¿S VASCULATURE. POTENTIAL DEVICE AND PROCEDURE-RELATED ADVERSE EVENTS: POSSIBLE ADVERSE EVENTS AND COMPLICATIONS THAT MAY OCCUR WITH THE USE OF GORE® DRYSEAL FLEX INTRODUCER SHEATH OR ANY INTRODUCER SHEATH OR IN ANY ENDOVASCULAR DEVICE INSERTION PROCEDURE AND WHICH MAY OR MAY NOT REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO (SHOWN IN ENGLISH ALPHABETICAL ORDER): ¿ BLOOD LOSS, BLEEDING, HEMATOMA. ¿ DEATH. ¿ VASCULAR TRAUMA (I.E., DISSECTION, RUPTURE, PERFORATION, TEAR, ETC.) W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
CARDIOHELP | K133598 | MAQUET CARDIOPULMONARY GMBH WAS RECEIVED THE FOLLOWING INFORMATION FROM THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES ON 2025-06-25: THE VIGILANCE AREA OF MEDICAL DEVICES OF THE SPANISH AGENCY OF MEDICINES AND MEDICAL DEVICES HAS BEEN INFORMED THROUGH RENACER REGISTRY ABOUT AN INCIDENT INVOLVING THE PRODUCT ¿CARDIOHELP SYSTEM¿, MANUFACTURED BY GETINGE AND DISTRIBUTED BY, OCCURRED ON 2025-05-15 WITH REPORT NUMBER (B)(4). THE INFORMATION WAS PROVIDED THAT THE PATIENT HAD THE INCIDENT "MAJOR BLEEDING: RETROPERITONEAL HEMORRHAGE" AND THE CONSEQUENCE FOR THE PATIENT WAS "PATIENT PASSED AWAY¿ NO ADDITIONAL INFORMATION WAS RECEIVED. FURTHER INFORMATION ABOUT THE EVENT AND DEVICE WAS REQUESTED BUT IS STILL PENDING. AS THE PATIENT PASSED AWAY, A REPORT IS REQUIRED. COMPLAINT ID (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE SPAIN MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, CARDIOHELP¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K133598. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR CARDIOHELP WITH CATALOG NUMBER 701048012.
SWARTZ¿ BRAIDED TRANSSEPTAL GUIDING INTRODUCER, SL1¿, 63 CM LENGTH, 8.5F | K052644 | DURING A LEFT ATRIAL PROCEDURE, AN AIR EMBOLISM OCCURRED WHICH RESULTED IN A NEUROLOGICAL EVENT. THE VALVE OF THE SHEATH WAS REPORTEDLY LEAKING AND THE PATIENT EXPERIENCED ST-ELEVATION IN ALL LEADS FOLLOWED BY ELECTRO-MECHANICAL DISSOCIATION. REANIMATION MEASURES WERE INITIATED. AFTER THE PATIENT WAS STABILIZED AND ST-ELEVATION RETURNED TO NORMAL, A CORONARY ANGIOGRAM WAS PERFORMED WHICH EXCLUDED CAD AND THROMBOEMBOLISM. THE PATIENT WAS TRANSFERRED TO INTENSIVE CARE AND CURRENT NEUROLOGICAL STATUS IS NOT CLEAR. | THE RESULTS OF THE INVESTIGATION ARE INCONCLUSIVE SINCE THE DEVICE WAS NOT RETURNED FOR ANALYSIS. THE DEVICE HISTORY RECORD WAS REVIEWED TO ENSURE THAT EACH MANUFACTURING AND INSPECTION OPERATION WAS PERFORMED. BASED ON THE INFORMATION RECEIVED, THE CAUSE OF THE REPORTED AIR EMBOLISM AND STROKE REMAINS UNKNOWN.
ERBE VIO 3 | K190823 | IT WAS REPORTED THAT A PATIENT INCIDENT OCCURRED WITH THE ELECTROSURGICAL UNIT (ESU/GENERATOR) DURING A PARTIAL GASTRECTOMY AND HERNIA REPAIR OF THE ABDOMEN. THE ESU WAS USED WITH AN ERBE LAP BISECT [PART NUMBER (P/N) 20195-204, LOT NUMBER (L/N): INFORMATION NOT PROVIDED (NI) WHICH INCLUDED AN ERBE LAP BISECT INSERT (P/N 20195-206, L/N 2304793)] (NOTE: THE ERBE LAP BISECT AND INSERT ARE NOT CLEARED FOR DISTRIBUTION IN THE U.S. AND ARE NOT DISTRIBUTED IN THE U.S.). NO INFORMATION WAS PROVIDED REGARDING ANY OTHER ACCESSORY USED IN THE PROCEDURE. ALSO, THE SPECIFIC ESU SETTINGS USED WERE NOT PROVIDED. DURING THE OPERATION AND WHEN ACTIVATED IN A CUTTING MODE, THERMALLY INDUCED PERFORATIONS OCCURRED ON A BOWEL LOOP DUE TO TISSUE COMING INTO CONTACT WITH THE NON-INSULATED JOINT OF THE BISECT INSTRUMENT. SUBSEQUENTLY, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT OF ANOTHER HOSPITAL FOR FURTHER SURGICAL TREATMENT OF THE PERFORATION SITES. HOWEVER, THE PATIENT'S CONDITION DETERIORATED, AND HE DIED ON (B)(6) 2025. | THE ESU AND LAP BISECT INSERT WERE EVALUATED. THE FINDINGS WERE AS FOLLOWS: ESU.THE GENERATOR WAS INSPECTED/TESTED. THE UNIT WAS FOUND TO BE FUNCTIONING AS INTENDED. THE EVALUATION INCLUDED AN ELECTRICAL SAFETY CHECK, A FUNCTIONAL CHECK OF EACH OF THE EQUIPMENT'S FEATURES AND A POWER OUTPUT CHECK. THE ESU WAS/IS WITHIN SPECIFICATIONS AND ALL FEATURES WERE/ARE FUNCTIONING PROPERLY. IN ADDITION, NO ANOMALIES WERE FOUND IN THE DEVICE HISTORY RECORD (DHR) OF THE ESU. LAP BISECT INSERT AN INSPECTION OF THE INSERT SHOWED SIGNS OF DAMAGE DURING USE. ITS BLACK PTFE INSULATION WAS BADLY DAMAGED IN THE JOINT AREA. TRACES OF MELTING CAN BE SEEN ON THE WHITE PTFE INSULATION IN THE REAR AREA. THE SHAFT END OF THE INSERT WAS HEAVILY SCALED. A MATERIAL/MANUFACTURING DEFECT WAS NOT EVIDENT. IN CONCLUSION, NO ERBE EQUIPMENT PROBLEM WAS FOUND THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE INCIDENT. THERE WERE MANY FACTORS INVOLVED IN THE EVENT. THE THERMAL LESIONS OF THE BOWEL WERE INITIALLY PRESUMABLY CAUSED BY (UNINTENTIONAL) OR INCORRECT CONTACT OF THE PROXIMAL NON-INSULATED BRANCH PARTS OF THE INSTRUMENT WITH THE BOWEL. THESE AREAS CAN ALSO BECOME VERY HOT, AND CARE MUST BE TAKEN TO ENSURE THAT THERE IS NO UNINTENTIONAL CONTACT WITH TISSUE IN THIS AREA. THE NOTES ON USE FOR THE INSTRUMENT EXPLICITLY STATES THAT THE PRODUCT, INCLUDING ALL INSULATION, MUST BE CHECKED FOR MECHANICAL DAMAGE BEFORE EACH USE AND THAT THE PRODUCT MUST NOT BE USED IN THE EVENT OF DAMAGE. THERE IS ALSO ANOTHER WARNING THAT UNINTENTIONAL TISSUE CONTACT WITH LIVE, NON-INSULATED INSTRUMENT PARTS CAN LEAD TO UNINTENTIONAL COAGULATION AND IRREVERSIBLE TISSUE DAMAGE. THIS ALSO APPLIES TO THE JOINT AREA BEHIND THE INSTRUMENT JAWS. THE INSTRUMENT WAS ALSO CLEANED INTRAOPERATIVELY WITH A SALINE SOLUTION, WHICH SHOULD ALSO BE AVOIDED ACCORDING TO THE ACCESSORY'S INSTRUCTIONS. IN SUMMARY PER THE REPORTED INFORMATION AND THE EVALUATION OF ERBE'S PRODUCTS, THE INCIDENT IS RELATED TO USER ERROR. NO TRENDS HAVE BEEN IDENTIFIED AND ERBE USA, INC. IS NOW CLOSING THE FILE ON THIS EVENT.
TRILOGY 202 VENTILATOR, INTERNATIONAL | K093905 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING ACUTE NIV (NON-INVASIVE VENTILATION) CIRCUIT HAD DISCONNECTED DUE TO PATIENT PULLING OFF AND STAFF NOTICED ALARMS WERE OFF ON ACUTE NIV MACHINE, THEREFORE DID NOT ALARM TO ALERT STAFF. PATIENT FOUND UNRESPONSIVE WITH VENTILATOR DISCONNECTED. NO ALARMS TO ALERT STAFF OF DISCONNECTION. MET (MEDICAL EMERGENCY TEAM) CALL PUT OUT, UNFORTUNATELY PATIENT PASSED AWAY ON MET TEAM ARRIVAL. VENTILATOR CHECKED WITH PHYSIO NEXT MORNING, PATIENT HAD CHANGED VENTILATION MODES DURING DAY, NO ALARMS HAD BEEN SET, THE PREVIOUS ALARMS DON'T FOLLOW THROUGH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION ALLEGING ACUTE NIV (NON-INVASIVE VENTILATION) CIRCUIT HAD DISCONNECTED DUE TO PATIENT PULLING OFF AND STAFF NOTICED ALARMS WERE OFF ON ACUTE NIV MACHINE, THEREFORE DID NOT ALARM TO ALERT STAFF. PATIENT FOUND UNRESPONSIVE WITH VENTILATOR DISCONNECTED. NO ALARMS TO ALERT STAFF OF DISCONNECTION. MET (MEDICAL EMERGENCY TEAM) CALL PUT OUT, UNFORTUNATELY PATIENT PASSED AWAY ON MET TEAM ARRIVAL. VENTILATOR CHECKED WITH PHYSIO NEXT MORNING, PATIENT HAD CHANGED VENTILATION MODES DURING DAY, NO ALARMS HAD BEEN SET, THE PREVIOUS ALARMS DON'T FOLLOW THROUGH. THE VENTILATOR WAS RETURNED TO A THIRD-PARTY SERVICE CENTER FOR EVALUATION. NO COSMETIC DAMAGE TO DEVICE WAS FOUND AND NO ACTIONABLE ERRORS WERE FOUND IN THE EVENT LOG THAT OCCURRED WHILE IN POSSESSION OF THE CUSTOMER. THE DEVICE WAS TESTED AND AN ERROR CODE RELATED TO A COMMUNICATION ISSUE WAS OBSERVED. ERR_DSP_LOSS_OF_RIPC_COMMUNICATION SIGNIFIES THAT A BRIEF SYSTEM RESET OCCURRED DUE TO DETECTION OF AN ISSUE. THIS IS A SYSTEM RELIABILITY FUNCTION TO KEEP THE SYSTEM WORKING THROUGH TRANSIENT UPSETS. THE SYSTEM BOARD WAS REPLACED TO ADDRESS THE ISSUE. THIS EVENT DID NOT OCCUR WHILE IN USE BY THE CUSTOMER AND NO EVIDENCE WAS FOUND REGARDING THIS EVENT BEING RELATED TO THE REPORTED COMPLAINT. THE SUBJECT DEVICE WAS BENCH TESTED AND FUNCTIONED AS DESIGNED. THE FULL DATA EXPORT THAT WAS GIVEN TO PHILIPS RESPIRONICS BY THE CUSTOMER INCLUDED THE ENTIRETY OF THE EVENT LOG EVENTS FROM JULY 21, 2024 TO JULY 1, 2025. THIS DATA IS RECORDED TO THE VENTILATOR INTERNALLY AND DOES NOT REQUIRE AN SD CARD TO BE IN USE DURING EVENTS. DURING AN ANALYSIS OF THE EXPORTED EVENT LOG, IT WAS DETERMINED THAT NONE OF THE ALARM SETTINGS WERE CHANGED BETWEEN JUNE 3, 2025 AND JUNE 25, 2025. ALL USER-SETTABLE ALARMS WERE TURNED OFF (WITH THE EXCEPTION OF THE LOW SPO2 ALARM), INCLUDING THE CIRCUIT DISCONNECT ALARM. SYSTEM ALARMS EXIST ON THE TRILOGY 202 THAT ARE NOT ABLE TO BE TURNED OFF OR HAVE THEIR LIMITS ADJUSTED. IN THE PC MODE THAT WAS BEING USED DURING THE EVENT, ONE OF THESE ALARMS IS LOW PEAK INSPIRATORY PRESSURE (PIP), WHICH CAN OCCUR DURING LARGE LEAKS AND CIRCUIT DISCONNECTS. WHILE THE LOW PIP ALARM CAN BE SEEN IN THE EVENT LOG FOR PRIOR TIMES ON (B)(6) 2025 (THE LAST LOW PIP ALARM ON (B)(6) 2025 OCCURRED AT APPROXIMATELY 14:59 BST), INSTANCES OF THIS ALARM DO NOT APPEAR IN THE TIMELINE IMMEDIATELY SURROUNDING THE REPORTED EVENT. NO EVIDENCE WAS FOUND OF THE REPORTED CIRCUIT DISCONNECT IMMEDIATELY PRIOR TO THE ESTIMATED TIME THE PATIENT WAS FOUND UNRESPONSIVE (21:33 BST ON (B)(6) 2025). ACCORDING TO THE USER MANUAL FOR THE TRILOGY 202, WARNINGS AND CAUTIONS, PAGE 3 STATES: FOR VENTILATOR DEPENDENT PATIENTS, DO NOT RELY ON ANY SINGLE ALARM TO DETECT A CIRCUIT DISCONNECT CONDITION. THE LOW TIDAL VOLUME, LOW MINUTE VENTILATION, LOW RESPIRATORY RATE, AND APNEA ALARMS SHOULD BE USED IN CONJUNCTION WITH THE CIRCUIT DISCONNECT AND LOW PEAK INSPIRATORY PRESSURE ALARMS. TEST THE OPERATION OF THE CIRCUIT DISCONNECT FUNCTION DAILY AND WHENEVER A CHANGE IS MADE TO THE PATIENT CIRCUIT. AN INCREASE IN CIRCUIT RESISTANCE CAN PREVENT PROPER OPERATION OF SOME ALARMS. DO NOT SET THE LOW PEAK INSPIRATORY PRESSURE ALARM TOO LOW, OR THE SYSTEM MAY NOT DETECT LARGE CIRCUIT LEAKS, OR A PATIENT DISCONNECT. BASED ON THE EVALUATION OF THE DEVICE AND DOWNLOADED DEVICE EVENT DATA, PHILIPS RESPIRONICS CONCLUDES THAT THE DEVICE OPERATED AS DESIGNED. THE CUSTOMER REPORTED THAT THE SD CARD WAS NOT FULLY INSERTED INTO THE VENT RESULTING IN MISSING THERAPY DATA AROUND THE REPORTED EVENT TIME, THEREFORE DEVICE CONTRIBUTION TO A USE ERROR THAT MAY HAVE LED TO THE PATIENT DEATH CANNOT BE CONFIRMED OR REFUTED AT THIS TIME.
REMSTAR SYSTEM ONE 60 SERIES | K131982 | THE MANUFACTURER ATTEMPTED TO CONTACT THE PATIENT IN REFERENCE TO A REMSTAR AUTO WITH HEATED HUMIDIFIER, SYSTEM ONE 60 SERIES DEVICE REPLACEMENT. THE PATIENT'S SPOUSE ANSWERED THE CALL AND STATED THE PATIENT WAS DECEASED. THE PATIENT'S SPOUSE ENDED THE CALL PRIOR TO ANY ADDITIONAL INFORMATION BEING OBTAINED. DETAILS SURROUNDING THE PATIENT'S DEATH WERE NOT PROVIDED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
RADICAL-7 | K212161 | PER MEDWATCH MW5170850 THE CUSTOMER REPORTED "A SIX-MONTH-OLD INFANT DIED DURING PSG WHEN A HIDDEN FREEZE MASKED A DROP IN OXYGEN SATURATION." THERE WAS A PATIENT DEATH REPORTED AS A PART OF THIS COMPLAINT. | OTHER, OTHER TEXT: DEVICE DETAILS FOR THE RADICAL-7 WERE OMITTED FROM THE MEDWATCH REPORT, AND THE DEVICE HAS NOT BEEN RETURNED TO MASIMO FOR INVESTIGATION. AS A RESULT, THE UDI, DEVICE SERIAL NUMBER, AND MANUFACTURE DATE ARE UNKNOWN. AS THE INITIAL REPORTER REQUESTED TO REMAIN ANONYMOUS, MASIMO IS UNABLE TO GATHER ADDITIONAL INFORMATION FROM THE REPORTER. BASED ON THE INFORMATION PROVIDED IN THE MEDWATCH REPORT, MASIMO DOES NOT BELIEVE THAT THIS IS A PROBLEM WITH THE RADICAL-7 DEVICE AND ITS HARDWARE, BUT RATHER AN ISSUE WITH THE USE OF AN UNVALIDATED DIY RS-232 TO USB CABLE USED TO STREAM SPO2 VALUES INTO THE PSG SYSTEM. IT WAS REPORTED THAT DURING PSG TESTING, THE UNVALIDATED CABLE EXPERIENCED OVER 120 COMMUNICATION DROPOUTS. WHEN THESE DROPOUTS OCCURRED, THE PSG TRACE FROZE ON THE LAST VALUE RESULTING IN HIDDEN DESATURATION EVENTS. MASIMO HAS DETERMINED THAT THE USE OF AN UNVALIDATED DIY CABLE, COMBINED WITH INAPPROPRIATE DEVICE SETTINGS, IMPROPER SYSTEM INTEGRATION INTO THE PSG PLATFORM AND A LACK OF INVESTIGATION BY THE CUSTOMER DESPITE KNOWN AND DOCUMENTED ISSUES CONTRIBUTED TO THE REPORTED EVENT.
DAVINCI X | K171294 | IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED BILATERAL INGUINAL HERNIA REPAIR, THE PATIENT EXPERIENCED SUDDEN, UNEXPLAINED PAIN AND RETURNED TO THE HOSPITAL IN CARDIAC ARREST. THE SITE ROBOTICS COORDINATOR REPORTED THAT THE SURGICAL PROCEDURE WAS COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS. THE PATIENT WOKE UP WELL, WAS EXTUBATED, AND WAS TRANSFERRED TO THE ICU AS A PRECAUTION DUE TO COMORBIDITIES AND ADVANCED AGE. ON POSTOPERATIVE DAY (POD) 3, THE PATIENT WAS TRANSFERRED OUT OF THE ICU AND REMAINED HOSPITALIZED FOR AN ADDITIONAL DAY BEFORE BEING DISCHARGED HOME ON POD 4 IN STABLE CONDITION. ON POD 8, THE PATIENT EXPERIENCED A SUDDEN ONSET OF UNSPECIFIED PAIN AND RETURNED TO THE HOSPITAL IN CARDIAC ARREST. AFTER APPROXIMATELY 40 MINUTES OF RESUSCITATION EFFORTS, THE PATIENT EXPIRED. THE CAUSE OF DEATH IS SUSPECTED TO BE EITHER A HEART ATTACK OR A PULMONARY EMBOLISM. AN AUTOPSY WAS NOT PERFORMED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, MONOPOLAR CURVED SCISSORS, LARGE NEEDLE DRIVER, FENESTRATED BIPOLAR FORCEPS. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED 8 DAYS FOLLOWING A REPORTEDLY UNEVENTFUL ROBOTIC HERNIA REPAIR. THE PATIENT RETURNED WITH PAIN BUT THE LOCATION AND TYPE OF PAIN IS NOT PROVIDED. A PULMONARY EMBOLISM OR CARDIAC EVENT WAS SPECULATED TO HAVE OCCURRED BUT NO SUPPORTING INFORMATION FOR THOSE POTENTIAL CAUSES WAS PROVIDED. NO ALLEGATION AGAINST ANY INTUITIVE PRODUCT OR INSTRUMENT WAS MADE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
RENTAL, DREAMSTATION ST30, GB | K102465 | MANUFACTURER RECEIVED A CALL FROM THE CLINIC THAT THE PATIENT PASSED AWAY. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
PHILIPS CPAP | K131982 | MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S SISTER STATING THAT HER BROTHER WAS ON THE DEVICE FOR YEARS BUT PASSED AWAY IN 2012 AND WAS REQUESTING FOR THE SERIAL NUMBER OF HER BROTHER'S DEVICE, THE CALLER WAS TRANSFERRED TO PATIENT RECALL SUPPORT. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER A DA VINCI-ASSISTED RADICAL CYSTECTOMY AND URINARY DIVERSION PROCEDURE, THE PATIENT EXPIRED POST-OPERATIVELY DUE TO BLEEDING. THE PATIENT REPORTEDLY HAD SEVERE ADHESIONS, WAS UNDERGOING CHEMOTHERAPY, AND HAD FRAGILE BLOOD VESSELS. THE SURGEONS CONVERTED TO LAPAROTOMY, BUT COULD NOT IDENTIFY THE SOURCE OF THE BLEEDING. NO ADDITIONAL INFORMATION WAS PROVIDED. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION WERE MADE; NO RESPONSE WAS RECEIVED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO SYSTEM ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 8MM 30 DEGREE ENDOSCOPE PLUS, 0 DEGREE ENDOSCOPE, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, LARGE SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND X2. A SITE HISTORY REVIEW FOUND THAT ALL MULTIPLE USE INSTRUMENTS WERE USED IN SUBSEQUENT PROCEDURES WITH NO REPORTED ISSUES THAT WOULD BE RELATED TO THIS EVENT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AFTER A RADICAL CYSTECTOMY. BLEEDING IS SUGGESTED AS THE CAUSE OF DEATH AND POTENTIAL CONTRIBUTING FACTORS ARE SUGGESTED BUT HOW AND WHEN THE BLEEDING OCCURRED IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
K-PACK II | K984576 | THE USER FACILITY REPORTED AN INCIDENT TO THE HEALTH AUTHORITY OF POLAND FOLLOWING THE DEATH OF A CHILD AFTER A VACCINATION. TO ENSURE THE SAFETY OF PATIENTS, AN IMMEDIATE REVIEW OF THE BATCH RECORDS AND A THOROUGH INVESTIGATION FOCUSING ON THE STERILITY OF THE TWO NEEDLE LOTS IN QUESTION HAVE BEEN REQUESTED. THE EVENT OCCURRED POST-TREATMENT, AND THE FINAL PATIENT IMPACT WAS DEATH. | D4: LOT#: 2306025 OR 2311027. D6A: IMPLANTED DATE: DEVICE WAS NOT IMPLANTED. D6B: EXPLANTED DATE: DEVICE WAS NOT EXPLANTED. E3: OCCUPATION: SPECIALIST, SPU IM OPS BE, PRIMARY PACK. H4: DEVICE MANUFACTURE DATE: 2023-06-25 OR 11-27-2023. THE ACTUAL DEVICE IS NOT AVAILABLE FOR RETURN. THE BATCH RECORDS FOR LOT NUMBERS: 2306025 AND 2311027 (KN2516RB05) WERE REVIEWED. DURING NEEDLE ASSEMBLY, AS PART OF IN-PROCESS CONTROL, 40 ASSEMBLED NEEDLES (NEEDLE ASSEMBLED WITH CASE) WERE VISUALLY INSPECTED TWICE PER SHIFT (8-HOUR SHIFTS DURING THE WEEK AND 12-HOUR SHIFTS DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGE TO THE CASE. ONE NON-CONFORMITY WAS RECORDED FOR CASE DAMAGE; HOWEVER, THE DAMAGE WAS SUPERFICIAL AND POSED NO RISK OF CRACKING THE CASE. DURING PACKAGING, 40 LABELED PRODUCTS WERE VISUALLY INSPECTED TWICE PER SHIFT (8-HOUR SHIFTS DURING THE WEEK AND 12-HOUR SHIFTS DURING THE WEEKEND) FOR FOREIGN MATTER AND DAMAGE TO THE CASE AND CAP. NO DEVIATIONS OR ABNORMALITIES RELATED TO THE REPORTED COMPLAINT WERE OBSERVED DURING PRODUCTION. DURING FINAL INSPECTION, WHICH IS PART OF RELEASE TESTING, 1,250 UNITS WERE EXAMINED AND NO DEFECTS WERE DETECTED. NO RELEVANT DEVIATIONS OR ABNORMALITIES WERE NOTED IN THE BATCH RECORDS OR DURING THE INSPECTIONS DESCRIBED ABOVE THAT WOULD RELATE TO THE REPORTED COMPLAINT. LIMULUS AMEBOCYTE LYSATE (LAL) TESTING AND STERILIZATION CYCLES FOR BOTH LOTS WERE PERFORMED ACCORDING TO SPECIFICATIONS, WITH NO OBSERVATIONS NOTED. THE 2023 BIOBURDEN MONITORING PROGRAM WAS REVIEWED, AND RESULTS WERE CONSISTENT WITH PREVIOUS YEARS. MONITORING SAMPLES FROM JUNE 2023 AND NOVEMBER 2023 (CLOSEST TO THE MANUFACTURING DATES) DID NOT SHOW ANY DEVIATIONS. ADDITIONALLY, NO MATERIALS OF ANIMAL ORIGIN WERE USED. ALL REMAINING RETENTION SAMPLES WERE VISUALLY INSPECTED FOR FOREIGN MATTER AND BROKEN OR DAMAGED CASES WITH THE NAKED EYE ON 06/05/2025. NO DEFECTS WERE FOUND. NO COMPLAINT SAMPLE WAS RECEIVED; THEREFORE, NO DIRECT INVESTIGATION OF A SAMPLE COULD BE PERFORMED. NON-CONFORMITIES RELATED TO THE FINISHED PRODUCT BATCHES AND SEMI-FINISHED NEEDLE ASSEMBLY BATCHES WERE REVIEWED. NO RELATED NON-CONFORMITIES WERE IDENTIFIED, EXCEPT FOR ONE: A NON-CONFORMITY WAS RECORDED FOR COSMETIC DAMAGE (NO CRACK) TO THE CASE DURING NEEDLE ASSEMBLY ON LINE NA7. THE AFFECTED PRODUCTION WAS CONTAINED AND SCRAPPED. NOTIFICATIONS RELATED TO THE FINISHED PRODUCT AND SEMI-FINISHED NEEDLE ASSEMBLY BATCHES WERE ALSO REVIEWED. NO RELATED NOTIFICATIONS WERE FOUND. THERE IS NO CORRECTIVE AND PREVENTIVE ACTION (CAPA) ASSOCIATED WITH ISSUES INVOLVING THE STERILE BARRIER SYSTEM OF THE PRODUCTS. BASED ON THE INVESTIGATION, INCLUDING BATCH RECORD REVIEW AND RETENTION SAMPLE INSPECTION, NO ROOT CAUSE RELATED TO THE PRODUCTION PROCESS WAS IDENTIFIED. FOR INFORMATION REGARDING THE IMPORTER REPORT FOR THIS EVENT, PLEASE REFER TO SECTION H10.
ILET BIONIC PANCREAS | K231485 | ON 27-MAY-2025, A CERTIFIED DIABETES CARE AND EDUCATION SPECIALIST (CDCES) NOTIFIED BETA BIONICS THAT THE USER HAD PASSED AWAY OVER THE WEEKEND. FOLLOW-UP ON (B)(6) 2025 CONFIRMED THE USER PASSED AWAY ON (B)(6) 2025. AT THE TIME OF DISCOVERY, THE USER WAS WEARING THE ILET, BUT THE INFUSION SET WAS FOUND TO BE PARTIALLY DISLODGED. THE REPORTED CAUSE OF DEATH WAS NATURAL CAUSES AND HIGH BLOOD PRESSURE. NO AUTOPSY WAS PERFORMED. REVIEW OF DEVICE DATA CONFIRMS THE ILET WAS IN USE THROUGH (B)(6) 2025. A DEVICE INVESTIGATION IS PENDING. A FOLLOW-UP REPORT WILL BE SUBMITTED ONCE EVALUATION IS COMPLETE. | NO PRODUCT WAS RETURNED FOR EVALUATION. THE DEVICE HISTORY RECORD (DHR) REVIEW WAS COMPLETED, AND THIS DEVICE PASSED ALL MANUFACTURING RELEASE CRITERIA FOR DISTRIBUTION. THERE WERE NO ISSUES IDENTIFIED THAT WOULD HAVE IMPACTED THIS EVENT. AT THE TIME OF THIS REPORT, THE ILET EVALUATION IS STILL IN PROGRESS. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE INVESTIGATION IS COMPLETED.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE WAS BEING RETURNED BECAUSE THE PATIENT HAD PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K220414 | IT WAS REPORTED A PATIENT DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A WATCHMAN LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE PHYSICIAN PERFORMED A SUCCESSFUL TRANSSEPTAL PUNCTURE (TSP) USING WATCHMAN DOUBLE CURVE SHEATH (WAS). WHEN REMOVING THE DILATOR, IT WAS NOTED THAT THERE WAS NO BLEED. THE END OF THE SHEATH WAS LIFTED UP AND HEARD A VACUUM/SUCKING SOUND, QUICKLY PLACED THE DILATOR BACK DOWN ON THE TABLE, LIFTED END OF THE SHEATH UP AGAIN AND HEARD THE SAME SUCKING SOUND AGAIN, QUICKLY LOWERED THE SHEATH DOWN. THIS IS WHEN IT WAS NOTED AIR BUBBLES ON THE INTRACARDIAC ECHOCARDIOGRAPHY (ICE) IMAGES AND THE PATIENT'S HEART RATED DROPPED. TRANSCUTANEOUS PACING WAS INITIATED, PATIENT WAS INTUBATED, FLUIDS GIVEN AND CODE STARTED ONCE PATIENT CONTINUED TO DECLINE. CATH TEAM WAS CALLED IN. THE PATIENT WAS UNABLE TO BE SUCCESSFULLY RESUSCITATED. THE CODE ENDED AND PATIENT PRONOUNCED DEAD APPROXIMATELY AT 9:30 AM. IN THE PHYSICIAN OPINION, HE WAS UNCERTAIN IF THE PATIENT POSSIBLY HAD A VERY LOW OR NEGATIVE LEFT ATRIUM PRESSURE OR POSSIBLY TAKING A BREATH DURING MAC ANESTHESIA CAUSING NEGATIVE PRESSURE RESULTING IN NO BLEED-BACK FROM THE SHEATH AFTER THE DILATOR WAS REMOVED. THE PHYSICIANS WERE CERTAIN THE SHEATH/DILATOR WERE FLUSHED PROPERLY PRIOR TO INSERTION. SHEATH WAS INSPECTED AFTER EVENT AND APPEARS TO BE INTACT AND WORKING PROPERLY. THE DILATOR IS EXPECTED TO BE RETURNED FOR ANALYSIS. AIR ENTRY BELIEVED TO BE AT THE PORT OF THE DOUBLE CURVE SHEATH RESULTING IN AIR IN THE LEFT SIDE OF THE HEART. THE VALVE WAS CLOSED AT THE END OF THE SHEATH. ASPIRATION FROM THE SIDE PORT. FLUID AND ANESTHESIA SUPPORT. AIR INGRESS WHEN REMOVED DILATOR TO LEFT SIDE OF THE HEART. PATIENT STATUS DECLINED, CODE INITIATED, THEN DEATH. AIR INGRESS TO LEFT SIDE OF HEART CAUSING CARDIAC ARREST. NO DIFFICULTY WITH THE TSP WORKFLOW THAT WAS APPARENT DURING THE PROCEDURE. THE PHYSICIAN DID NOTE THAT THE FEMORAL VEIN ANATOMY WAS TORTUOUS BUT DID NOT APPEAR TO HAVE DIFFICULTY ADVANCING THE SHEATH. THE PHYSICIAN DID NOT MENTION ANYTHING ABOUT A DEFECTIVE PRODUCT OR THAT THE SHEATH OR VERSACROSS CONNECT WAS THE ISSUE. IN THE PHYSICIAN OPINION, IT WAS MORE OF A LOW OR NEGATIVE LEFT ATRIAL PRESSURE AND THE USE OF MAC THAT MAY HAVE CONTRIBUTED TO THE PROBLEM. | THE DEVICE HAS NOT BEEN RECEIVED FOR ANALYSIS. UPON RECEIPT AND COMPLETION OF THE FAILURE ANALYSIS OF THE COMPLAINT DEVICE, IF THERE IS ANY FURTHER RELEVANT INFORMATION FROM THAT REVIEW, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
MEDTRONIC NAVIGATION | K081656 | DOI: 10.3171/2025.1.PEDS24387. SUMMARY: OBJECTIVE STEREOTACTIC LASER ABLATION (SLA) OR LASER INTERSTITIAL THERMAL THERAPY (LITT) IS AN EMERGING ALTERNATIVE TO CONVENTIONAL SURGERY FOR CNS TUMORS. FURTHER CHARACTERIZATION OF ITS EFFECTIVENESS AND SAFETY IN THE PEDIATRIC POPULATION IS NEEDED. METHODS A REVIEW WAS CONDUCTED ACCORDING TO PRISMA GUIDELINES USING THE PUBMED/MEDLINE, SCOPUS, EMBASE, GOOGLE SCHOLAR, AND SCIENCE DIRECT DATABASES TO INVESTIGATE THE EFFECTIVENESS AND SAFETY IN THE USE OF LITT TO TREAT PEDIATRIC CNS TUMORS IN CHILDREN. RESULTS A TOTAL OF 24 ARTICLES MET THE INCLUSION CRITERIA. FROM THE 4 RETROSPECTIVE STUDIES COMPRISING 97 PATIENTS (MEAN AGE 11.4 YEARS, 55.0% MALES IN THE POOLED PREVALENCE), THE MOST COMMON TUMOR TYPES WERE PILOCYTIC ASTROCYTOMA (31.5%) AND SUBEPENDYMAL GIANT CELL ASTROCYTOMA (SEGA; 68.5%). TUMORS WERE PRIMARILY LOCATED IN THE FRONTAL LOBE (29.7%) AND THALAMUS (24.4%). POSTOPERATIVE COMPLICATIONS INCLUDED TRANSIENT NEUROLOGICAL DEFICITS IN 12.1% AND PERMANENT DEFICITS IN 6.0% OF PATIENTS. AT A MEAN FOLLOW-UP OF 43.9 MONTHS, MASS REDUCTION WAS OBSERVED IN 68.8% OF PATIENTS AND OVERALL DISEASE IMPROVEMENT IN 91.9% OF PATIENTS, AND THE MORTALITY RATE WAS 2.6%. FROM THE 20 CASE REPORTS/SERIES INVOLVING 67 PATIENTS (MEAN AGE 10.8 ± 4.7 YEARS, 52.2% MALES), THE MEAN TUMOR SIZE WAS 15.7 ± 8.7 CM3. PREDOMINANT TUMOR SUBTYPES WERE PILOCYTIC ASTROCYTOMA (29.9%) AND SEGA (16.4%). TUMORS WERE LOCATED IN ELOQUENT AREAS IN 61.3% OF CASES, NOTABLY THE THALAMUS (24.2%) AND VENTRICULAR SYSTEM (24.2%). PRIOR TREATMENTS INCLUDED SURGERY (78.8%), CHEMOTHERAPY (51.5%), AND IMMUNOTHERAPY (27.3%). KEY LITT PARAMETERS WERE DURATION (7.2 ± 8.8 MINUTES), DOSE (10.2 ± 2.4 W), AND EXTENT OF TUMOR VOLUME DECREASE (68.3% ± 30.4%). THE MOST COMMONLY USED LITT SYSTEM WAS VISUALASE (95.7%). POSTOPERATIVE COMPLICATIONS WERE REPORTED IN 26.9% OF PATIENTS, INCLUDING TRANSIENT NEUROLOGICAL DEFICITS (55.6%), PERILESIONAL EDEMA (22.2%), AND HYDROCEPHALUS (22.2%). HOSPITAL STAYS WERE = 3 DAYS IN 92.7% OF PATIENTS. TUMOR SIZE REDUCTION WAS ACHIEVED IN 86.7% OF PATIENTS, AND 78.9% EXPERIENCED DISEASE IMPROVEMENT. COMPARATIVE ANALYSIS SHOWED THAT A GREATER EXTENT OF ABLATION WAS ASSOCIATED WITH A REDUCED NEED FOR SECONDARY SURGERY (P = 0.038, OR 0.94) AND IMPROVED DISEASE OUTCOMES (P = 0.023, OR 1.05). LONGER LITT DURATION WAS SIGNIFICANTLY ASSOCIATED WITH POSTOPERATIVE COMPLICATIONS (P = 0.050). CONCLUSIONS LITT APPEARS EFFECTIVE IN REDUCING TUMOR SIZE AND IMPROVING DISEASE OUTCOMES IN PEDIATRIC CNS TUMORS. THE LONG-TERM EFFECTIVENESS OF LITT IN PEDIATRIC BRAIN TUMORS REQUIRES FURTHER RANDOMIZED PROSPECTIVE INVESTIGATION. REPORTED EVENT(S): THIS ARTICLE DOCUMENTED REVIEW OF 24 ARTICLES THAT DOCUMENT 20 CASE REPORTS / CASE SERIES AND 4 RETROSPECTIVE STUDIES. THESE REPORTS WERE COMPRISED OF 97 PATIENTS (54.99% MALE, 45.01% FEMALE) WITH A MEAN AGE OF 11.43 YEARS. THE MORTALITY RATE WAS 2.55%. | H3, H6: NO PRODUCTS WERE RETURNED TO MEDTRONIC FOR ANALYSIS. CODES B17, C20, AND D15 ARE APPLICABLE. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REP DREAMSTATION AUTO BIPAP, DOM-RECRT | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO A CUSTOMER INQUIRING ABOUT A DREAMSTATION DEVICE RECEIVED FROM THEIR DME. THE MANUFACTURER WAS INFORMED THAT THE PREVIOUS USER OF THE DEVICE WAS DECEASED. THERE IS NO ALLEGATION OF DEVICE MALFUNCTION. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE PATIENT DEATH. NO OTHER INFORMATION IS AVAILABLE AT THIS TIME. THIS REPORT IS BEING SUBMITTED FOR AN ALLEGATION OF A PATIENT DEATH ONLY DUE TO INSUFFICIENT INFORMATION RELATED TO THE DEVICE. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. THE MANUFACTURER INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
WECK VISTA | K152663 | REFER TO H11 FOR FOLLOW-UP INFORMATION. | ADDITIONAL INFORMATION SECTION B5: THE SURGEON REPORTED THAT THIS EVENT WAS UNRELATED TO THE DA VINCI SYSTEM OR INSTRUMENTS, AND WAS ADVISED BY THE HOSPITAL ADMINISTRATION TO NOT DISCLOSE ALL THE DETAILS OF THE EVENT AS IT WAS NOT ROBOTIC RELATED. SECTION H11: A REVIEW OF THE SYSTEM LOGS FOUND THAT NO ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM SHOWED THAT THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE FOUND.
WECK VISTA | K152663 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED OVARIAN CYSTECTOMY AND ENDOMETRIOSIS RESECTION PROCEDURE, THE PATIENT'S AORTA WAS INADVERTENTLY PERFORATED DURING THE INITIAL TROCAR INSERTION, PRIOR TO DOCKING THE ROBOTIC SYSTEM. ACCESS INTO THE PATIENT'S ABDOMEN WAS PERFORMED USING A THIRD-PARTY 5MM LAPAROSCOPIC ENDOSCOPE AND A DA VINCI 8MM CANNULA WITH AN 8MM BLADELESS OPTICAL OBTURATOR. BLEEDING WAS IDENTIFIED WITHIN MINUTES, AND THE PROCEDURE WAS CONVERTED TO LAPAROTOMY TO ATTEMPT HEMOSTASIS. THE PATIENT WAS PLACED ON A MASS TRANSFUSION PROTOCOL. A VASCULAR SURGEON WAS CALLED TO ASSIST; HOWEVER, THE PATIENT EXPIRED ON THE OPERATING TABLE BEFORE THE VASCULAR SURGEON WAS ABLE TO INTERVENE. | THE REPORTED EVENT OCCURRED PRIOR TO DOCKING THE DA VINCI SYSTEM; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF A TROCAR INJURY TO THE AORTA WHILE USING A BLADELESS OPTICAL OBTURATOR AND A 3RD PARTY SCOPE. ALTHOUGH NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT, HOW THE INJURY OCCURRED IS NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.
T:SLIM X2 INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | K203234 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025. CAUSE OF DEATH WAS DUE TO HYPERGLYCEMIA RESULTING IN DIABETIC KETOACIDOSIS WITH A BLOOD GLUCOSE (BG) LEVEL OF 600 MG/DL. THE PUMP HISTORY WAS REVIEWED BY THE HEALTHCARE PROVIDER AND MULTIPLE HIGH BG ALERTS WERE OBSERVED TO HAVE OCCURRED BETWEEN (B)(6). THE CUSTOMER INITIALLY ARRIVED AT THE EMERGENCY ROOM ON (B)(6) AND WAS TRANSFERRED BY EMERGENCY SERVICES TO ANOTHER HOSPITAL LATER THAT SAME DAY. EMERGENCY SERVICES DISCARDED THE INFUSION SET USED AT THE TIME OF THE EVENT, AND NO INFORMATION WAS PROVIDED ON THE CONTINUOUS GLUCOSE MONITOR (CGM) SENSOR USED AT THE TIME OF THE EVENT. NO ADDITIONAL INFORMATION COULD BE CONFIRMED AT THE TIME OF THIS REPORT, AS PUMP DATA IS NOT AVAILABLE FOR REVIEW. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS RECEIVED. | D4 (SERIAL NUMBER), H4.
ARROW CVC SET: 4-LUMEN 8.5FR X 16CM | K862056 | IT WAS REPORTED "AFTER REMOVAL OF THE CENTRAL VENOUS CATHETER (CVC) PLACED IN THE RIGHT SUBCLAVIAN VEIN, THE PATIENT BEGAN TO SHOW CHANGES IN CONSCIOUSNESS AND MYOCLONUS - PROBABLY GAS EMBOLISM. TREATMENT: HYPERBARIC CHAMBER TREATMENT, + VENTILATION INVASIVE MECHANICS + VASOPRESSOR. " THE PATIENT'S CURRENT CONDITION IS REPORTED AS "DECEASED". ADDITIONAL INFORMATION REPORTS "ROOT CAUSE ANALYSIS WAS CARRIED OUT ON THIS INCIDENT AND THERE WAS NO EVIDENCE OF PROBLEMS WITH THE DEVICE USED, HOWEVER, THE DISTRIBUTOR WAS MADE AWARE OF THE OCCURENCE." | (B)(4)
PU-681RA | K102376 | THE HOSPITAL'S RISK MANAGER AT THE HOSPITAL REPORTED THAT AFTER RETURNING FROM A PROCEDURE, THE PATIENT EXPIRED WHILE BEING MONITORED AT THE CENTRAL NURSE'S STATION (CNS). THIS OCCURRED ON (B)(6) 2025 AT 1200 HOURS. THE LOGS ARE CURRENTLY BEING PULLED FROM THE DEVICES. SHE WAS CONCERNED BECAUSE IN THE PAST, MONITOR TECHS HAD BEEN SILENCING THE ALARMS, AND THEY WANT TO VERIFY THAT THIS IS NOT THE CASE WITH THIS EVENT. SHE ALSO WANTED TO SEE WHAT HAPPENED WITH THE PATIENT. | THE HOSPITAL'S RISK MANAGER AT THE HOSPITAL REPORTED THAT AFTER RETURNING FROM A PROCEDURE, THE PATIENT EXPIRED WHILE BEING MONITORED AT THE CENTRAL NURSE'S STATION (CNS). THIS OCCURRED ON (B)(6) 2025 AT 1200 HOURS. THE LOGS ARE CURRENTLY BEING PULLED FROM THE DEVICES. SHE WAS CONCERNED BECAUSE IN THE PAST, MONITOR TECHS HAD BEEN SILENCING THE ALARMS, AND THEY WANT TO VERIFY THAT THIS IS NOT THE CASE WITH THIS EVENT. SHE ALSO WANTED TO SEE WHAT HAPPENED WITH THE PATIENT. NIHON KOHDEN CONTINUES TO INVESTIGATE THE REPORTED EVENT. NIHON KOHDEN WILL SUBMIT A SUPPLEMENTAL REPORT IN ACCORDANCE WITH 21 CFR SECTION 803.56 WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THE FOLLOWING FIELD CONTAINS LIMITED INFORMATION, AS ATTEMPTS TO OBTAIN THE INFORMATION WERE MADE, BUT NOT PROVIDED: D10 ATTEMPT # 1: 06/12/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ATTEMPT # 2: 06/25/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ATTEMPT # 3: 07/07/2025 EMAILED THE CUSTOMER VIA MICROSOFT OUTLOOK FOR ALL INFORMATION IN THE NI LIST ABOVE: NO REPLY WAS RECEIVED. ADDITIONAL DEVICE INFORMATION: D10 CONCOMITANT MEDICAL DEVICE: THE FOLLOWING DEVICES WERE USED IN CONJUNCTION WITH THE CNS: MULTIPLE PATIENT RECEIVER (ORG): MODEL #: NI. SERIAL #: NI. DEVICE MANUFACTURER DATA: NI. UNIQUE IDENTIFIER (UDI) #: NI. RETURNED TO NIHON KOHDEN: NA. ZM TRANSMITTER: MODEL #: ZM-531PA. SERIAL #: (B)(6). DEVICE MANUFACTURER DATA: 09/25/2020. UNIQUE IDENTIFIER (UDI) #: (B)(4). RETURNED TO NIHON KOHDEN: NA.
WATCHMAN FXD CURVE? ACCESS SYSTEM | K212228 | IT WAS REPORTED A PERICARDIAL EFFUSION OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. PRIOR TO THE PROCEDURE, A TRACE PERICARDIAL EFFUSION WAS NOTED ON BASELINE TRANSESOPHAGEAL ECHOCARDIOGRAPHY (TEE). THE TRANS-SEPTAL PUNCTURE (TSP) WAS PERFORMED USING AN NRG RF NEEDLE AND A FIXED CURVE SHEATH UNDER INTRACARDIAC ECHOCARDIOGRAPHY (ICE) GUIDANCE. THE IMPLANTING PHYSICIAN EXPERIENCED DIFFICULTY OBTAINING A CLEAR ANTERIOR-POSTERIOR (AP) VIEW ON ICE BUT OBSERVED A GOOD TENTING OF THE SEPTUM IN THE INFERIOR/SUPERIOR VIEW. THE NEEDLE WAS INITIALLY ADVANCED ACROSS THE SEPTUM, THEN RETRACTED TO THE RIGHT ATRIUM. AFTER REVIEWING TEE IMAGING, THE PHYSICIAN PROCEEDED WITH SEPTAL CROSSING USING RF ENERGY. AS THE SHEATH WAS EXCHANGED FOR A WAS DOUBLE CURVE SHEATH AND THE PIGTAIL CATHETER WAS ADVANCED, THE TEE IMAGING PHYSICIAN NOTED THAT THE PREVIOUSLY OBSERVED TRACE EFFUSION HAD INCREASED IN SIZE. A PERICARDIOCENTESIS TRAY WAS PREPARED, AND CONTRAST IMAGING OF THE LEFT ATRIAL APPENDAGE (LAA) WAS PERFORMED. NO CONTRAST LEAK WAS OBSERVED FROM THE APPENDAGE OR INTO THE PERICARDIAL SPACE. PERICARDIOCENTESIS WAS COMPLETED AND THE PATIENT WAS TREATED WITH PROTAMINE, IV FLUIDS, AND HEMODYNAMIC SUPPORT PROVIDED BY THE CERTIFIED REGISTERED NURSE ANESTHETIST (CRNA). A CARDIOTHORACIC SURGEON WAS CALLED FOR CONSULTATION. THE PATIENT RECEIVED TWO UNITS OF PACKED RED BLOOD CELLS AND WAS TRANSFERRED TO THE OPERATING ROOM FOR SURGICAL MANAGEMENT. INTRAOPERATIVELY, A PERFORATION WAS IDENTIFIED IN THE LEFT ATRIAL WALL NEAR THE RIGHT INFERIOR PULMONARY VEIN, WHICH WAS DETERMINED TO BE THE SOURCE OF THE EFFUSION. THE PERFORATION WAS SURGICALLY REPAIRED, AND THE LAA WAS LIGATED. THE PATIENT TOLERATED THE PROCEDURE WELL, WITH NO FURTHER COMPLICATIONS, AND WAS TRANSFERRED TO RECOVERY IN STABLE CONDITION. IT WAS FURTHER REPORTED THAT THE PATIENT WAS EXTUBATED ONE (1) DAY POST-OP, HOWEVER THE PATIENT BEGAN TO DECLINE. THREE (3) DAYS POST-OP THE PATIENT PASSED AWAY DUE TO SEPTIC SHOCK. | B5 DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. B2: DATE OF DEATH ESTIMATED AS (B)(6) 2025 AS THE PATIENT DIED 3 DAYS POST-OPERATION. H6: IMPACT CODE ADDED.
INTELLIVUE MX40 1.4GHZ | K113125 | PHILIPS RECEIVED A COMPLAINT ON THE MX40 1.4 GHZ SMART HOPPING INDICATING A PATIENT CONNECTED TO TELEMETRY REMOVED THE LEADS, WHICH RESULTED IN THE GENERATION OF A BLUE INOP ALARM FOR 'LEADS OFF' INSTEAD OF RED, WHICH WAS EXPECTED BY THE USER. RESPIRATORY STAFF FOUND THE PATIENT IN PULSELESS ELECTRICAL ACTIVITY (PEA) WITH ECG LEADS OFF AND THE PATIENT HAD PASSED AWAY. IT WAS SUSPECTED WHEN TELEMETRY WENT OFFLINE FOR A BATTERY CHANGE, THE ALARM SETTINGS DEFAULTED BACK TO THE BEDSIDE SETTINGS; HOWEVER, THIS CANNOT BE CONFIRMED AT THIS TIME. | THE AUDIT LOGS WERE REVIEWED BY THE CLINICAL PRODUCT SPECIALIST (CPS) AND THE PRODUCT SPECIALIST ENGINEER (PSE) TO DETERMINE IF THE DEVICE FUNCTIONED AS EXPECTED. THEY EXPLAINED THAT PRIOR TO THE ECG ¿LEADS OFF¿ THAT GENERATED A CYAN COLOR ON (B)(6) 2023 AT 05:13:40, THE DEVICE WAS RECONNECTED TO THE PIC AT THE START OF MONITORING HENCE THE ¿NO DATA TELE¿ ENDED PRIOR TO THAT. IN THE INFORMATION FOR USE (IFU), IT STATES THAT THE ECG LEADS OFF WILL DISPLAY AS A CYAN INITIALLY UNTIL THE ECG SIGNAL IS OBTAINED. SO, IF THE DEVICE IS AT THE START OF MONITORING DUE TO PIC IX REBOOT OR IF IT IS OUT OF THE ACCESS POINT (AP) RANGE, THEN IT WILL GO CYAN AT THE START UNTIL THERE IS AN ECG SIGNAL. DUE TO THE DEVICE BEING RECONNECTED, THERE HAD BEEN NO ECG SIGNAL ACQUIRED; THEREFORE, THE INOP REMAINS CYAN IN COLOR. BASED ON THIS INFORMATION, IT WAS DETERMINED THAT NO PRODUCT MALFUNCTION OCCURRED. BASED ON THE INFORMATION AVAILABLE AND THE TESTING CONDUCTED, THE CAUSE OF THE REPORTED PROBLEM WAS USER ERROR. THE REPORTED PROBLEM WAS NOT CONFIRMED THE DEVICE WAS CONFIRMED TO BE OPERATING PER SPECIFICATIONS AND NO FAILURE WAS IDENTIFIED. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO DEATH OF A PATIENT. THERE IS NO ALLEGATION OF PRODUCT MALFUNCTION. NO OTHER MEDICAL INFORMATION WAS SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | ON (B)(6) 2025 A HOSPITAL NURSE REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT HAD PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. THE PATIENT WAS AT THE END OF TREATMENT (STEP 8). THE NURSE DISCONNECTED THE PATIENT AND REALIZED THE PATIENT WAS NOT WAKING UP. THE NURSE TOOK THE PATIENT¿S VITAL SIGNS, AND HE SUBSEQUENTLY WENT INTO CARDIAC ARREST AND DIED. ADDITIONAL INFORMATION WAS OBTAINED THROUGH THE DIALYSIS UNIT NURSE. THE PATIENT WAS AN IN-PATIENT REHABILITATION (REHAB) DIALYSIS PATIENT. THE PATIENT WAS IN REHAB DUE TO A LEFT BELOW KNEE AMPUTATION (BKA) WHICH WAS UNRELATED TO DIALYSIS. THE DATE OF THE ADMISSION TO THE REHAB WAS UNKNOWN. THE NURSE STATED THAT ON (B)(6) 2025, THE HOSPITAL NURSE WAS DOING ROUNDS WHEN THEY WENT TO DISCONNECT THIS PATIENT FROM THE LIBERTY SELECT CYCLER AT THE END OF TREATMENT. THE NURSE WENT TO TAKE THE PATIENT¿S VITALS AND FOUND THE PATIENT¿S PULSE TO BE THREADY. THE PATIENT SUBSEQUENTLY WENT INTO CARDIAC ARREST. THE NURSE COULD NOT CONFIRM WHAT RESUSCITATIVE EFFORTS WERE ADMINISTERED. THE EFFORTS WERE NOT SUCCESSFUL, AND THE PATIENT WAS PRONOUNCED DECEASED. THE CAUSE OF DEATH WAS CARDIAC ARREST. THE NURSE STATED THE DEATH WAS NOT RELATED TO PD THERAPY OR USE OF THE LIBERTY SELECT CYCLER OR OTHER FRESENIUS PRODUCT(S). NO FURTHER INFORMATION WAS PROVIDED. | ADDITIONAL INFORMATION: D9, H3 PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL EVALUATION. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGNS OF PHYSICAL DAMAGE. THERE WERE NOT VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT. THERE WERE VISUAL INDICATION OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NOT BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. TEACH PUMPS TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
AEROGEN SOLO | K070642 | ON NOVEMBER 16, 2024, AN ADVERSE EVENT OCCURRED DURING A TRANSPLANT INVOLVING THE USE OF A PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WITH AN AEROGEN SOLO NEBULISER. THE PATIENT, A 57-YEAR OLD MALE, EXPERIENCED PROLONGED RESPIRATORY CHALLENGES DUE TO AN INCORRECT SETUP OF THE CIRCUIT, WHICH LIKELY CAUSED THE FILTER TO CLOG. THE FILTER AEROGEN SOLO NEBULISER WAS INTENDED TO DELIVER INHALED EPOPROSTENOL AFTER THE PATIENT WAS LIBERATED FROM CARDIOPULMONARY BYPASS. THE PATIENT WAS TAKEN OFF BYPASS AND CONNECTED TO A VENTILATOR FOR THE DELIVERY OF INHALED APOPROSTENOL. HOWEVER, WITHIN 2 HOURS, THE PATIENT'S TIDAL VOLUMES DECREASED, INDICATING DIFFICULTY IN VENTILATION AND OXYGENATION, LIKELY DUE TO THE CLOGGED FILTER OBSTRUCTING AIRFLOW AND EPOPROSTENOL DELIVERY FROM THE AEROGEN SOLO NEBULIZER. THIS LED TO AN ANOXIC BRAIN INJURY AND ULITMATELY THE PATIENTS DEATH, WHICH OCCURED ON (B)(6) 2024. THE ADVERSE EVENT WAS IDENTIFIED AS POTENTIALLY LINKED TO THE SETUP OF THE FILTER WITH THE AEROGEN SOLO NEBULIZER AFTER DISCOVERING THAT THE FILTER ON THE VENTILATOR HAD CLOGGED, WHICH OBSTRUCTED THE DELIVERY OF OXYGEN AND EPOPROSTENOL. | " ON NOVEMBER 16, 2024, AN ADVERSE EVENT OCCURRED DURING A TRANSPLANT INVOLVING THE USE OF A PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WITH AN AEROGEN SOLO NEBULISER. THE PATIENT, A 57-YEAR OLD MALE, EXPERIENCED PROLONGED RESPIRATORY CHALLENGES DUE TO AN INCORRECT SETUP OF THE CIRCUIT, WHICH LIKELY CAUSED THE FILTER TO CLOG. THE FILTER AEROGEN SOLO NEBULISER WAS INTENDED TO DELIVER INHALED EPOPROSTENOL AFTER THE PATIENT WAS LIBERATED FROM CARDIOPULMONARY BYPASS. THE PATIENT WAS TAKEN OFF BYPASSAND CONNECTED TO A VENTILATOR FOR THE DELIVERY OF INHALED APOPROSTENOL. HOWEVER, WITHIN 2 HOURS, THE PATIENT'S TIDAL VOLUMES DECREASED, INDICATING DIFFICULTY IN VENTILATION AND OXYGENATION, LIKELY DUE TO THE CLOGGED FILTER OBSTRUCTING AIRFLOW AND EPOPROSTENOL DELIVERY FROM THE AEROGEN SOLO NEBULIZER. THIS LED TO AN ANOXIC BRAIN INJURY AND ULITMATELY THE PATIENTS DEATH, WHICH OCCURED ON (B)(6) 2024. THE ADVERSE EVENT WAS IDENTIFIED AS POTENTIALLY LINKED TO THE SETUP OF THE FILTER WITH THE AEROGEN SOLO NEBULIZER AFTER DISCOVERING THAT THE FILTER ON THE VENTILATOR HAD CLOGGED, WHICH OBSTRUCTED THE DELIVERY OF OXYGEN AND EPOPROSTENOL." THIS IS BEING INVESTIGATED AS A NON-RETURN. THE DEVICE REPORTED WAS NOT RETURNED TO AEROGEN FOR INVESTIGATION AS IT WAS DISCARDED BY THE CUSTOMER AND THERE WAS INSUFFICIENT EVIDENCE/INFORMATION SUPPLIED BY THE CUSTOMER TO DETERMINE A ROOT CAUSE. CLINICAL CONCLUSION BASED ON THE INFORMATION AVAILABLE THE DEATH REPORTED WAS LIKELY ATTRIBUTABLE SOLELY TO USER ERROR. THIS EVENT DID NOT OCCUR, AS CONFIRMED BY THE USER FACILITY, AS A RESULT OF DEVICE [21 CFR 803.3]: - FAILURE; - MALFUNCTION; - IMPROPER OR INADEQUATE DESIGN; - MANUFACTURE; OR - LABELLING. THE USER FACILITY REPRESENTATIVE CONFIRMED THAT "USER ERROR", IN THIS INSTANCE A MISTAKE MADE BY AN ATTENDING HEALTH CARE PROFESSIONAL IN THE CONFIGURATION IN WHICH THE PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WAS PLACED IN RELATION TO THE AEROGEN® SOLO NEBULISER WAS THE CAUSE OF THIS EVENT12. THERE WAS NO PERFORMANCE ISSUE RELATED TO THE AEROGEN® SOLO NEBULISER REPORTED BY THE USER FACILITY. SUMMARY NO PERFORMANCE ISSUE WAS REPORTED WITH THE AEROGEN MEDICAL DEVICE INVOLVED. THE USER FACILITY CONFIRMED THAT AN "INCORRECT" CONFIGURATION OF THE MEDICAL DEVICES BY AN ATTENDING CLINICAL STAFF MEMBER WAS THE CAUSE OF THE EVENT. BASED ON THE INFORMATION AVAILABLE THE FOLLOWING IMDRF ANNEX A, E & F CODES MAY BE CONSIDERED. APPROPRIATE: A24 - ADVERSE EVENT WITHOUT IDENTIFIED DEVICE OR USE PROBLEM. E0726 - HYPOXIA. F22 - UNEXPECTED DETERIORATION. F23 - UNEXPECTED MEDICAL INTERVENTION. F02 - DEATH. THE FOLLOWING QUESTIONS (BOLD ITALICS) WERE PROVIDED TO THE USER FACILITY BY AEROGEN LTD. ON THE XXX. A WRITTEN RESPONSE (NORMAL TEXT) WAS PROVIDED TO QUESTION 1 ONLY. VERBAL RESPONSES (REPRODUCED AS FAITHFULLY AS POSSIBLE - NORMAL TEXT) WERE PROVIDED TO OTHER QUESTIONS IN THE FACE-TO-FACE INTERVIEW BY (B)(6) REPRESENTATIVES L PACKER ON THE 27TH JUNE 2025. DUE TO ONGOING LITIGATION (B)(6) DECLINE TO RESPOND TO SOME QUESTIONS ASKED. THESE ARE MARKED [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] BELOW. 1. COULD YOU DESCRIBE THE FORM OF RESPIRATORY SUPPORT (OR MECHANICAL VENTILATION) THE PATIENT WAS RECEIVING AT THE TIME, INCLUDING CONFIGURATION (ACTIVE OR PASSIVE HUMIDIFICATION, HUMIDIFIER USED, DUAL/SINGLE LIMB ETC.) AND INTERFACE TO THE PATIENT (E.G., ETT VS. TRACHEOSTOMY TUBE). MECHANICAL VENTILATION WAS ACHIEVED USING AN GE DATEX-OHMEDA AISYS CS2 ANESTHESIA SYSTEM, CONNECTED VIA A SINGLE LIMB ANESTHESIA CIRCUIT TO THE ETT OF THE PATIENT. THE PATIENT-SIDE HME FILTER WAS REMOVED FROM THE CIRCUIT. THERE WERE VENTILATOR-SIDE PLEATED HYDROPHOBIC MEMBRANE FILTERS (PALL ULTIPOR 55) PLACED BOTH IN THE INSPIRATORY AND EXPIRATORY PORT. THE AEROGEN SOLO NEBULIZER WAS PLACED IN THE INSPIRATORY PORT. THE INCORRECT SETUP WAS SUCH THAT THE AEROGEN SOLO NEBULIZER WAS PLACED IN THE INSPIRATORY PORT OF THE VENTILATOR FIRST, FOLLOWED BY THE PALL ULTIPOR FILTER, FOLLOWED BY THE ANESTHESIA CIRCUIT, LIKELY CAUSING THE HYDROPHOBIC FILTER TO CLOG WITH AERSOSOLIZED EPOPROSTANOL. 2. WAS THE PALL ULTIPOR 55 FILTER BEING EMPLOYED AS A FILTER (ON THE MACHINE SIDE PROTECTING THE EXPIRATORY PORT) OR AS A HME (AND POSITIONED PROXIMAL TO THE PATIENT)? A. IF THE PALL ULTIPOR 55 FILTER WAS PLACED ON THE EXPIRATORY LIMB, WAS ANOTHER HME EMPLOYED OR WAS ACTIVE HUMIDIFICATION USED? PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WAS PLACED AT BOTH THE INSPIRATORY AND EXPIRATORY POT OF THE VENTILATOR. THIS IS NORMAL PRACTICE AT THIS INSTITUTION. 3. COULD YOU DESCRIBE WHERE THE AEROGEN® SOLO NEBULISER WAS PLACED IN THIS CIRCUIT IN RELATION TO THE PALL ULTIPOR 55 FILTER? N/A - CONSIDERED ADDRESSED IN PREVIOUS ANSWERS TO Q1 & Q2. 4. WAS THE PALL ULTIPOR 55 FILTER PLACED [?]UPSTREAM' OF THE AEROGEN® SOLO NEBULISER, I.E., PROXIMAL TO THE PATIENT WITH RESPECT TO THE POSITION OF NEBULISER POSITION? OR WAS THE AEROGEN® SOLO NEBULISER BETWEEN THE PALL ULTIPOR 55 FILTER AND THE PATIENT? YES. THE FILTER WAS PLACED BETWEEN THE NEBULISER AND THE PATIENT. 5. WHAT WAS THE INDICATION FOR AEROSOL DRUG DELIVERY USING THE AEROGEN® SOLO NEBULISER? REDUCTION OF PULMONARY VASCULAR RESISTANCE (PVR), RELIEF OF RIGHT VENTRICULAR AFTERLOAD. INSTITUTIONAL PRACTICE FOR PREVENTION OF PRIMARY GRAFT DYSFUNCTION (REFERENCED TO RCT CONDUCED AT (B)(6) IN 2022 - SEE BIBLIOGRAPHY). 6. WHEN WAS THE AEROGEN SOLO FIRST USED FOLLOWING THE PATIENT'S SEPARATION FROM BYPASS? 2 HOURS. 7. HOW LONG WAS THE NEBULISER IN USE DURING THE POST-OPERATIVE PERIOD? [CONSIDERED ADDRESSED BY RESPONSE TO Q6] 8. WERE THE STAFF INVOLVED FAMILIAR WITH AEROGEN'S IFU AND THE CORRECT SETUP OF THE DEVICE? YES, (B)(6) STAFF ARE VERY FAMILIAR WITH AEROGEN DEVICES, THE IFU AND USE AEROGEN NEBULISERS FOR IEPO AS STANDARD PRACTICE. 9. HAS THIS TYPE OF CONFIGURATION (AEROGEN SOLO WITH IN-LINE FILTER) BEEN USED BEFORE AT YOUR FACILITY, AND IF SO, HAVE ANY SIMILAR ISSUES OCCURRED? THIS SPECIFIC CONFIGURATION, AS OCCURRED IN THIS INCIDENT - NO. THE CORRECT CONFIGURATION, WITH A FILTER ON THE INSPIRATORY PORT & EXPIRATORY PORT, WITH AN AEROGEN NEBULISER PROXIMAL TO THE PATIENT IS STANDARD IN (B)(6). PAGE 11 OF 16 RELATED TO THE FILTER OBSTRUCTION: 10. IN THIS SPECIFIC CASE HOW LONG WAS THE PALL ULTIPOR 55 FILTER USED? AND A. AT WHAT FREQUENCY ARE THESE FILTERS TYPICALLY CHECKED ACCORDING TO YOUR INTERNAL SOPS, AND CHANGED/REPLACED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 11. SPECIFICALLY, WAS THE FILTER SATURATED OR [?]CLOGGED'/'OBSTRUCTED' WITH PARTICULATE LIKE CRYSTALS? YES. THE FILTER WAS NOTED TO BE SATURATED BY THE HCPS. NO - THERE WAS NO PARTICULATE VISIBLE. 12. WAS THERE VISIBLE CONDENSATION OR SATURATION NOTED IN THE FILTER OR CIRCUIT? NO, THE SATURATION WAS [?]NON-OBVIOUS' TO THE ATTENDING CLINICAL TEAM. 13. AT WHAT POINT WAS A CLOGGED OR OBSTRUCTED FILTER SUSPECTED OR CONFIRMED? THE PROFOUND DROP IN TIDAL VOLUME WAS THE FIRST INDICATION TO THE CLINICAL STAFF OF A POTENTIAL ISSUE. IT WAS DISCOVERED SHORTLY AFTER THIS. 14. AT WHAT POINT WAS A CLOGGED OR OBSTRUCTED FILTER SUSPECTED OR CONFIRMED? CONSIDERED ADDRESSED BY RESPONSE TO Q13. RELATED TO THE AE: 15. WAS A PRE-USE CHECK OR INSPECTION OF THE BREATHING CIRCUIT AND FILTER CONDUCTED PRIOR TO INITIATION? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 16. WAS A PRE-USE CHECK OF THE VENTILATOR PERFORMED TO ENSURE THE PATIENT SETTINGS, ALARMS AND MONITORING DEVICES ARE ACCURATELY CONNECTED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 17. THE REPORT REFERS TO A PROLONGED RESPIRATORY CHALLENGE[?]COULD YOU PROVIDE MORE DETAIL AROUND THIS? HOW LONG DID IT PERSIST AND WHAT SYMPTOMS WERE NOTICED/OBSERVED? AFTER THE DROP IN TIDAL VOLUME THE [?]RESPIRATORY CHALLENGE' OCCURRED FOR BETWEEN 15-20 MINS UNTIL THE SATURATED FILTER WAS DISCOVERED. WHEN THIS WAS REPLACED, AND CONFIGURATION CORRECT, THE PATIENT WAS VENTILATED SUCCESSFULLY. 18. WHAT WERE THE PATIENT'S OXYGENATION LEVELS (SPO2 OR SIMILAR) BEFORE THE AE. [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 19. RELATED TO THE "DECREASED TIDAL VOLUMES" AND "DIFFICULTY IN VENTILATION" WERE ANY OF THE FOLLOWING ALARMS RECORDED: A. HIGH PEEP, OR B. HIGH AIRWAY PRESSURE ALARM, OR C. LOW EXHALED TIDAL VOLUME ALARM, OR D. APNEA OR LOW MINUTE VENTILATION, OR E. FLOW SENSOR OR EXPIRATORY VALVE ERROR. NO - NO ALARMS SOUNDED TO ALERT THE CLINICAL TEAM. THE FIRST INDICATION OF AN ISSUE WAS THE DROP IN TIDAL VOLUME. 20. WERE KEY RESPIRATORY PARAMETERS SUCH AS TIDAL VOLUME, PEAK PRESSURES, RESISTANCE, OR ETCO CONTINUOUSLY MONITORED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 21. WHAT CLINICAL MONITORING WAS IN PLACE (E.G., SPO[?], ETCO[?], FIO[?] TRENDS) [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 22. WHAT INTERVENTIONS WERE MADE IN RESPONSE TO THE OBSERVED DETERIORATION IN VENTILATION AND OXYGENATION? PAGE 12 OF 16 A BRONCHOSCOPY WAS ATTEMPTED TO DETERMINE IF AIRWAY OBSTRUCTION WAS PRESENT. UPON DETECTION THE FILTER WAS REMOVED - AT WHICH POINT THE PATIENT COULD BE VENTILATED SUCCESSFULLY. 23. WHEN WAS THE ONSET OF ANOXIC BRAIN INJURY FIRST SUSPECTED OR CONFIRMED? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 24. WERE OTHER POSSIBLE CAUSES OF HYPOXIA (E.G., AIRWAY OBSTRUCTION, PULMONARY EMBOLISM, BRONCHOSPASM) CONSIDERED OR RULED OUT? CONSIDERED ADDRESSED BY RESPONSE TO Q22 25. THE AE REPORT STATES THE AE WAS "LIKELY DUE TO A CLOGGED FILTER OBSTRUCTING AIRFLOW" AND REFERS TO A 2 HOUR TIME PERIOD OF WHICH PERIOD "DECREASED TIDAL VOLUME, DIFFICULTY IN VENTILATION AND OXYGENATION OCCURRED". WAS THE "CLOGGED" FILTER NOTED DURING THIS PERIOD OR AFTER? CONSIDERED ADDRESSED BY PREVIOUS ANSWERS. FINALLY: 26. HAS ANY CHANGE BEEN MADE TO LOCAL PRACTICE OR PROTOCOLS AS A RESULT OF THIS EVENT? [DECLINED TO ANSWER - NO RESPONSE AVAILABLE] 27. IS THE AEROGEN DEVICE AVAILABLE TO BE SENT BACK FOR INVESTIGATION. NO - THE AEROGEN DEVICE WAS NOT RETAINED. CLINICAL STAFF REPORTED NO ISSUE WITH THE AEROGEN DEVICE, AND IT WAS NOT RETAINED. THE PALL ULTIPOR 55 BREATHING CIRCUIT FILTER WAS RETAINED.
CARDIOSAVE HYBRID, 3.1 EDITION | K112372 | IT WAS REPORTED THAT CARDIOSAVE INTRA-AORTIC BALLOON PUMP (IABP) PRESSURE WAVEFORM AND BLOOD PRESSURE VALUES WERE NO LONGER DISPLAYED DURING CLINICAL USE WITH SENSOR BALLOON. THE PATIENT DIED ON (B)(6) 2025. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
WATCHMAN TRUSTEER? ACCESS SYSTEM | K240018 | IT WAS REPORTED PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED. A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND A 35MM WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. THE FIRST TRANSEPTAL PUNCTURE (TSP) WAS PERFORMED WITHOUT FLUOROSCOPIC GUIDANCE USING THE VERSACROSS CONNECT. IT WAS NOTED THAT THE TSP WAS NOT AT AN IDEAL POSITION. DESPITE THE SUBOPTIMAL LOCATION, THE PHYSICIAN PROCEEDED WITH ATTEMPTS TO DEPLOY THE WDS HOPING THE WAS TRUSTEER SHEATH WOULD HELP COMPENSATE FOR THE INADEQUATE TSP. MULTIPLE ATTEMPTS TO POSITION THE DEVICE CORRECTLY IN THE APPENDAGE WERE MADE, BUT THE POSITIONING WAS DETERMINED NOT ACCEPTABLE. THE PHYSICIAN RECAPTURED THE CLOSURE DEVICE AND REMOVED BOTH THE WDS AND WAS TRUSTEER SHEATH IN ORDER TO PERFORM A SECOND, BETTER-POSITIONED TSP. UPON REMOVAL OF THE DEVICES, A PERICARDIAL EFFUSION WAS NOTED. PERICARDIOCENTESIS WAS INITIATED, AND 2-3 SYRINGES (60CC EACH) OF BLOOD WERE DRAINED, TEMPORARILY STABILIZING THE PATIENT. A SECOND TSP WAS SUCCESSFULLY PERFORMED, AND THE TRUSTEER SHEATH AND 35MM DEVICE WERE RE-INTRODUCED AND THE CLOSURE DEVICE WAS REDEPLOYED. INITIALLY, THE TEAM BELIEVED THE CLOSURE DEVICE HELPED TO STOP THE EFFUSION, BUT THE EFFUSION REACCUMULATED. THE PHYSICIAN DECIDED TO RELEASE THE 35MM CLOSURE DEVICE WHICH WAS WELL POSITIONED IN THE APPENDAGE. ADDITIONAL DRAINAGE WAS THEN PERFORMED, BUT THE PATIENT WENT INTO CARDIAC ARREST. THE NURSES PERFORMED HEART MASSAGE. THE PATIENT EXPERIENCED SEVERE BRADYCARDIA, AND HYPOTENSION (ARTERIAL TENSION LOW 16 HIGH 34). A DEFIBRILLATOR (200 JOULESX1) WAS USED, AND LIFE CELL WAS USED TO GIVE ADDITIONAL BLOOD UNITS. A CODE WAS CALLED IN THE HOSPITAL TO ADVISE THE SURGERY TEAM. PROTAMINE WAS GIVEN TO THE PATIENT. THE PATIENT SEEMED TO STABILIZE SLIGHTLY, BUT THE EFFUSION PERSISTED. THE PHYSICIAN KEPT REMOVING 60CC SYRINGES OF BLOOD FROM THE PATIENT. THE PATIENT REMAINED IN THE INTERVENTIONAL CARDIOLOGY LAB UNDER ACTIVE RESUSCITATION AND WAS REPORTED TO HAVE DIED. IT WAS FURTHER REPORTED THAT THE PHYSICIAN SUSPECTED THAT THE INADEQUATE TRANSEPTAL LOCATION CAUSED THE WATCHMAN TRUSTEER ACCESS SYSTEM TO RUB AGAINST THE RIDGE AND POSSIBLY CAUSED A SMALL TEAR IN THE APPENDAGE. THERE WAS NO PERFORATION NOTED ON ANY IMAGING. THE CAUSE OF DEATH WAS DETERMINED TO BE TAMPONADE. | B5: DESCRIBE EVENT OR PROBLEM: UPDATED WITH ADDITIONAL INFORMATION RECEIVED. D1: BRAND NAME CORRECTED FROM WATCHMAN FLX PRO TO WATCHMAN TRUSTEER ACCESS SYSTEM. D4: LOT NUMBER UPDATED. H6: PATIENT CODE ADDED.
TRANS-RAY PLUS 7.5FR 35CC IAB WITH ACCESSORIES | K122628 | IT WAS REPORTED THAT DURING INTRA-AORTIC BALLOON (IAB) THERAPY, THE DEVICE WAS USED PER STANDARD PROCEDURE. HOWEVER, IN FULL AUTO MODE WITH ECG TRIGGER, THE TRP SENSOR PRESSURE WAVEFORM AND NUMERICAL VALUES WERE NOT DISPLAYED ON THE CSAVE MONITOR, DESPITE CONTINUED IAB CATHETER ASSISTANCE. RECONNECTING THE SENSOR CONNECTOR HAD NO EFFECT. THE TRANSDUCER WAS THEN CONNECTED TO THE Y-FITTING, ALLOWING ARTERIAL PRESSURE INPUT TO CONTINUE IABP THERAPY. THE PATIENT LATER PASSED AWAY ON (B)(6) 2025. SUBSEQUENT TESTING OF THE CARDIOSAVE UNIT WITH A TEST BALLOON AND SYSTEM TRAINER SHOWED NORMAL ARTERIAL PRESSURE WAVEFORM AND VALUES. | UPDATED FIELDS: B4, D8, D9, E1 (INITIAL REPORTER), G1 (CONTACT PERSON MFG SITE), G3, G6, H2, H3, H6 (TYPE OF INVESTIGATION, INVESTIGATION FINDINGS, COMPONENT CODES AND INVESTIGATION CONCLUSIONS), H11. NO DECON OR VISUAL INSPECTION PERFORMED SINCE PRODUCT WAS NOT RETURNED AND COULD NOT BE EVALUATED. THEREFORE, WE WERE UNABLE TO CONFIRM OR DETERMINE A ROOT CAUSE FOR THE REPORTED FAILURE. THE FAILURE MODE IS ADDRESSED IN THE RISK FILE AND IS OPERATING WITHIN ITS RISK PROFILE. THE IFU ADDRESSES THE REPORTED FAILURE. THERE WERE NO NCMRS IDENTIFIED WHICH COULD CAUSE OR CONTRIBUTE TO THE REPORTED FAILURE. THE INVESTIGATION DOES NOT INDICATE THAT THE DEVICE WAS INADVERTENTLY RELEASED AS NON-CONFORMING OR AN ADULTERATED PRODUCT OR WAS A COUNTERFEIT. THE COMPLAINT HISTORY REVIEW DID NOT IDENTIFY AN ADVERSE TREND (INCREASE IN NUMBER OF COMPLAINTS OVER PAST THREE (3) MONTHS). BASED ON THE RATIONAL PROVIDED ABOVE, NO ESCALATION TO THE CAPA PROCESS IS REQUIRED.
TRILOGY EVO, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION THAT A PATIENT USING A FISHER AND PAYKEL VITERA MASK HAD DETACHED FROM THE CIRCUIT AT THE SWIVEL CONNECTOR AND A CIRCUIT DISCONNECT ALARM SOUNDED FROM THE TRILOGY EVO VENTILATOR. IT WAS REPORTED THAT THE PATIENT WAS FOUND LYING FACE DOWN IN RESPIRATORY ARREST. THE DOCTOR WAS CONTACTED, AND MANUAL VENTILATION WAS INITIATED, HOWEVER THE PATIENT EXPIRED. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | THE MANUFACTURER PREVIOUSLY REPORTED INFORMATION RECEIVED THAT A PATIENT USING A FISHER AND PAYKEL VITERA MASK HAD DETACHED FROM THE CIRCUIT AT THE SWIVEL CONNECTOR AND A CIRCUIT DISCONNECT ALARM SOUNDED FROM THE TRILOGY EVO VENTILATOR. IT WAS REPORTED THAT THE PATIENT WAS FOUND LYING FACE DOWN IN RESPIRATORY ARREST. THE DOCTOR WAS CONTACTED, AND MANUAL VENTILATION WAS INITIATED, HOWEVER THE PATIENT EXPIRED. THERE IS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. THE COMPLAINT REVIEW HAS DETERMINED THAT DUE TO INSUFFICIENT INFORMATION THIS WILL BE REPORTED. ALTHOUGH THERE IS NOT AN ALLEGATION THAT THE DEVICE CAUSED OR CONTRIBUTED TO THE MENTIONED "DEATH". DURING THE EVALUATION OF THE DEVICE'S EVENT LOG NO ERROR INDICATING A DEVICE FAILURE FOR THE REPORTED DATE AND TIME WAS CONFIRMED. NO DEVICE FAILURE RELATED TO THE REPORTED ISSUE WAS CONFIRMED BY ANY OTHER INVESTIGATIONS EITHER. SINCE DIRT WAS CONFIRMED, INLINE PROXIMAL FILTER WAS REPLACED. SINCE USB EXTENSION CABLE WAS FOUND DAMAGED DUE TO HAVING BEEN PINCHED, IT WAS REPLACED. (THERE WAS NEITHER THERMAL DAMAGE NOR EXPOSED CONDUCTING WIRES.) PERFORMED FINAL TEST, BATTERY CHECK, VALVE CHECK, RUN IN TESTS AND CLEANING THEN CONFIRMED THE UNIT OPERATED PROPERLY. CONFIRMED THE SOFTWARE VERSION IS 1.06.10.00.
AMPLATZER VASCULAR PLUG | K031810 | THE ARTICLE, "IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO EVALUATE THE EFFECT OF TRANSCATHETER PATENT DUCTUS ARTERIOSUS CLOSURE (TCDC) IN PREMATURE INFANTS ON PULMONARY EDEMA AND RESPIRATORY STATUS, AND MORE SPECIFICALLY TO CHARACTERIZE THE SUBSET OF PATIENTS WHO HAVE SIGNIFICANTLY IMPROVED CHEST X-RAY (CXR) IMMEDIATELY AFTER TCDC. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER PICCOLO OCCLUDER, MEDTRONIC MICROVASCULAR PLUG, COVIDIEN MICROVASCULAR PLUG, AMPLATZER VASCULAR PLUG, AND AMPLATZER VASCULAR PLUG II. THE ARTICLE CONCLUDED CHEST X-RAY HAZINESS IMPROVED AFTER TCDC, WITH SMALLER INFANTS AND EARLIER CLOSURE HAVING MORE IMPROVEMENT. INFANTS WITH LUNG DISEASE HAD LESS NOTICEABLE IMPROVED EDEMA, INDICATING THE DIFFICULTY TO ASSESS THE HEMODYNAMIC SIGNIFICANCE OF THEIR PDA PRIOR TO CLOSURE. FURTHER STUDIES ARE NEEDED TO IDENTIFY WHICH NEONATES BENEFIT MOST FROM TCDC. [THE PRIMARY AND CORRESPONDING AUTHOR WAS AHMAD CHMAISSE, DIVISION OF CARDIOLOGY, RILEY HOSPITAL FOR CHILDREN, INDIANA UNIVERSITY SCHOOL OF MEDICINE, INDIANAPOLIS, IN, UNITED STATES, WITH CORRESPONDING EMAIL: ACHMAIS@IU.EDU]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2017 TO JUNE 2021. A TOTAL OF 68 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 59.5% RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE, WEIGHT, AND MAJORITY GENDER WAS NOT REPORTED. COMORBIDITIES INCLUDED PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. | SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. ONE OF THE COMPLICATION REPORTED WAS DEATH; THIS COMPLICATION IS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS.
AMPLATZER VASCULAR PLUG II | K071699 | THE ARTICLE, "IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS", WAS REVIEWED. THE ARTICLE PRESENTED A RETROSPECTIVE, SINGLE CENTER STUDY TO EVALUATE THE EFFECT OF TRANSCATHETER PATENT DUCTUS ARTERIOSUS CLOSURE (TCDC) IN PREMATURE INFANTS ON PULMONARY EDEMA AND RESPIRATORY STATUS, AND MORE SPECIFICALLY TO CHARACTERIZE THE SUBSET OF PATIENTS WHO HAVE SIGNIFICANTLY IMPROVED CHEST X-RAY (CXR) IMMEDIATELY AFTER TCDC. DEVICES INCLUDED IN THE STUDY WERE AMPLATZER PICCOLO OCCLUDER, MEDTRONIC MICROVASCULAR PLUG, COVIDIEN MICROVASCULAR PLUG, AMPLATZER VASCULAR PLUG, AND AMPLATZER VASCULAR PLUG II. THE ARTICLE CONCLUDED CHEST X-RAY HAZINESS IMPROVED AFTER TCDC, WITH SMALLER INFANTS AND EARLIER CLOSURE HAVING MORE IMPROVEMENT. INFANTS WITH LUNG DISEASE HAD LESS NOTICEABLE IMPROVED EDEMA, INDICATING THE DIFFICULTY TO ASSESS THE HEMODYNAMIC SIGNIFICANCE OF THEIR PDA PRIOR TO CLOSURE. FURTHER STUDIES ARE NEEDED TO IDENTIFY WHICH NEONATES BENEFIT MOST FROM TCDC. [THE PRIMARY AND CORRESPONDING AUTHOR WAS AHMAD CHMAISSE, DIVISION OF CARDIOLOGY, RILEY HOSPITAL FOR CHILDREN, INDIANA UNIVERSITY SCHOOL OF MEDICINE, INDIANAPOLIS, IN, UNITED STATES, WITH CORRESPONDING EMAIL: ACHMAIS@IU.EDU]. THE TIME FRAME OF THE STUDY WAS FROM JANUARY 2017 TO JUNE 2021. A TOTAL OF 68 PATIENTS WERE INCLUDED IN THE STUDY, OF WHICH 59.5% RECEIVED AN ABBOTT DEVICE. THE AVERAGE AGE, WEIGHT, AND MAJORITY GENDER WAS NOT REPORTED. COMORBIDITIES INCLUDED PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. | SUMMARIZED PATIENT OUTCOMES/COMPLICATIONS OF IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS WERE REPORTED IN A RESEARCH ARTICLE IN A SUBJECT POPULATION WITH MULTIPLE CO-MORBIDITIES INCLUDING PULMONARY HYPERTENSION, PULMONARY EDEMA, PRETERM BIRTH, AND PATENT DUCTUS ARTERIOSUS. ONEE OF THE COMPLICATION REPORTED WAS DEATH; THIS COMPLICATION IS ANTICIPATED FOR THE PROCEDURE AND SUBJECT POPULATION. A MORE COMPREHENSIVE ASSESSMENT COULD NOT BE PERFORMED AS THE EVENT WAS NON-CONTEMPORANEOUSLY REPORTED THROUGH A LITERATURE REVIEW AND NO DEVICE OR INDIVIDUAL PATIENT INFORMATION WAS RECEIVED FOR ANALYSIS. B2: DATE OF DEATH IS ESTIMATED. B3: DATE OF EVENT IS ESTIMATED. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED. LITERATURE: IMPROVEMENT OF PULMONARY EDEMA AND RESPIRATORY STATUS AFTER TRANSCATHETER PDA CLOSURE IN THE SMALLEST AND MOST PREMATURE INFANTS.
VIZIGO | K170997 | IT WAS REPORTED A PATIENT UNDERWENT A CARDIAC ABLATION PROCEDURE WITH A VIZIGO AND A QDOT MICRO CATHETER. THE PATIENT EXPERIENCED ACCUMULATION OF BLOOD IN THE ABDOMEN TREATED WITH BLOOD TRANSFUSION FOR THE BLEED AND LATER THE PATIENT EXPERIENCED CARDIAC TAMPONADE AND EXPIRED TWO DAYS AFTER THE PROCEDURE. IT WAS INITIALLY REPORTED THAT AFTER THE PROCEDURE, AN ACCUMULATION OF BLOOD IN THE ABDOMEN WAS FOUND WHEN THE DRAPE WAS REMOVED. INTRACARDIAC WAS CHECKED BY ECHOCARDIOGRAPHY, AND NO PERICARDIAL EFFUSION WAS CONFIRMED. AFTER THE PATIENT LEFT THE PROCEDURE ROOM A BLOOD TRANSFUSION WAS PERFORMED. THE PHYSICIAN'S OPINION WAS THAT THERE WAS AN ISSUE WITH SOME ASPECT OF THE WORKFLOW DURING THE PUNCTURE BEFORE THE USE OF BWI PRODUCTS. ON 23-JUN-2025, ADDITIONAL INFORMATION WAS RECEIVED INDICATING THAT IN ADDITION TO THE ABDOMINAL BLOOD ACCUMULATION OBSERVED IMMEDIATELY AFTER THE PROCEDURE, CARDIAC TAMPONADE WAS CONFIRMED AT A LATER DATE AND THE PATIENT PASSED AWAY IN THE HOSPITAL 2 DAYS AFTER THE INDEX PROCEDURE. THE DETAILS WERE UNCLEAR, BUT THE PHYSICIAN SUSPECTED THAT THE ISSUE OCCURRED DURING THE PUNCTURE PRIOR TO THE USE OF BWI DEVICES. THE PHYSICIAN'S ASSESSMENT OF THE HEALTH HAZARD WAS SERIOUS. THE PHYSICIAN'S OPINIONS ON THE RELATIONSHIP BETWEEN THE EVENT AND THE PRODUCT WAS IMMEDIATELY AFTER THE PROCEDURE, INTRACARDIAC WAS CHECKED BY ECHOCARDIOGRAPHY, NO PERICARDIAL EFFUSION WAS CONFIRMED. NO ABNORMALITIES WERE OBSERVED WITH THE PRODUCT PRIOR TO OR DURING USE OF THE PRODUCTS. | ON 22-JUL-2025, ADDITIONAL INFORMATION ABOUT THE PATIENT AND EVENT WERE RECEIVED. IT WAS REPORTED THE PROCEDURAL ACTIVATED CLOTTING TIME (ACT) WAS 300 SECONDS. SEVERAL CATHETER EXCHANGES OCCURRED DURING THE PROCEDURE. THE ENERGY DELIVERED WAS RADIOFREQUENCY (RF). THERE WAS NO EVIDENCE OF STEAM POP. TRANSSEPTAL PUNCTURE WAS PERFORMED WITH UNKNOWN NEEDLE. THE VIZIGO SHEATH WAS NOT USED FOR TRANSSEPTAL ACCESS. NO ERROR MESSAGES OBSERVED ON BIOSENSE WEBSTER EQUIPMENT DURING THE PROCEDURE. THE CORRECT CATHETER SETTINGS WERE SELECTED ON THE GENERATOR. THERE WERE NO ISSUES WITH FLOW RATE CHANGES AT THE BEGINNING OF THE PROCEDURE. FORCE VISUALIZATION FEATURES USED INCLUDED VECTOR, REALTIME GRAPH. VISITAG MODULE PARAMETERS FOR STABILITY WERE RANGE: 3MM, TIME: 3SEC, FORCE OVER TIME (FOT): 25%, TAG SIZE: 2MM. NO ADDITIONAL FILTERS WRE USED. COLOR OPTIONS USED WERE TAG INDEX. THE PHYSICIAN'S OPINION ON THE CAUSE OF DEATH WAS CARDIAC TAMPONADE OR HEMORRHAGIC SHOCK BEFORE THE CARDIAC TAMPONADE WAS CONFIRMED. IT WAS REPORTED THAT DURING THE PROCEDURE, A DECREASE IN BLOOD PRESSURE WAS OBSERVED; HOWEVER, AT THAT TIME, NO PERICARDIAL EFFUSION WAS CONFIRMED ON ULTRASOUND. AFTER THE PROCEDURE, WHEN THE COVERING CLOTH WAS REMOVED, A SUBCUTANEOUS BLOOD ACCUMULATION IN THE ABDOMEN WAS NOTED, AND DRAINAGE AND TRANSFUSION WERE PERFORMED. THE PHYSICIAN SUSPECTED THAT VASCULAR DISSECTION DURING THE PUNCTURE MAY HAVE BEEN THE CAUSE. ON THE MORNING OF 2 DAYS AFTER THE PROCEDURE, CARDIAC TAMPONADE WAS FOUND, AND THE PATIENT SUBSEQUENTLY DIED. BASED ON THE ADDITIONAL INFORMATION RECEIVED, THE VIZIGO SHEATH IS THE MOST SUSPECTED DEVICE FOR VASCULAR HEMORRHAGE / DISSECTION SINCE THE VIZIGO COMES INTO DIRECT CONTACT WITH VASCULAR TISSUE. PLEASE CONSIDER THE CARDIAC TAMPONADE (E0605) AS REMOVED FROM FIELD H6. HEALTH EFFECT - CLINICAL CODE AND VASCULAR DISSECTION (E0515) AND HEMORRHAGE/BLEEDING (E0506) HAVE BEEN ADDED TO FIELD H6. HEALTH EFFECT - CLINICAL CODE. THE MOST SUSPECTED DEVICE FOR THE CARDIAC TAMPONADE CONTINUES TO BE QDOT MICRO ABLATION CATHETER AS IT DELIVERS ENERGY DIRECTLY TO CARDIAC TISSUE. BECAUSE THE CAUSE OF DEATH IS UNCLEAR WHETHER IT IS FROM HEMORRHAGE OR CARDIAC TAMPONADE, BOTH THE VIZIGO AND THE QDOT MICRO ABLATION CATHETER WILL CONTINUE TO CAPTURE THE DEATH. IF ADDITIONAL INFORMATION IS RECEIVED REGARDING THIS EVENT, A SUPPLEMENTAL 3500A REPORT WILL BE SUBMITTED TO THE FDA. MANUFACTURER'S REF. # (B)(4).
PORT-A-CATH IMPLANTABLE ACCESS SYSTEM | K072657 | ICU MEDICAL LEARNED OF A CIVIL LAWSUIT FILED IMPLICATING A PORT-A-CATH II, 21-4455-24, LOT 3829680. THE LAWSUIT ALLEGES THE FOLLOWING. THE PATIENT, (B)(6), UNDERWENT PLACEMENT OF A PORT IN THE LEFT CHEST, ON (B)(6) 2021 FOR VEIN ACCESS. THE PATIENT PRESENTED TO (B)(6) WITH SHORTNESS OF BREATH. DOCTORS DETERMINED SHE HAD AN INFECTION WITH ACCUMULATION OF FLUID IN THE LUNGS. THE PATIENT'S PORT WAS EXPLANTED IMMEDIATELY AFTER THIS DETERMINATION ON OR ABOUT (B)(6) 2023. THE PATIENT PASSED AWAY ON (B)(6) 2023. | NO DEVICE WAS RETURNED FOR INVESTIGATION. DUE TO THIS, NO DEVICE ISSUE WAS DETERMINED. IF THE DEVICE IS RETURNED FOR INVESTIGATION, THE DEVICE WILL BE EXAMINED WITH THE RESULTS SENT IN A SUPPLEMENTAL REPORT.
AVEIR¿ RETRIEVAL SYSTEM CATHETER | K213494 | IT WAS REPORTED THAT AFTER THE LEADLESS (LP) PACEMAKER IMPLANT WAS COMPLETED, THE PATIENT EXPERIENCED CARDIAC ARREST RESULTING IN PULSELESS HEART ACTIVITY. A THORACOTOMY AND CARDIAC MASSAGE WERE PERFORMED. DURING THE CARDIAC MASSAGE, THE LP DISLODGED AND MIGRATED TO THE PULMONARY ARTERY. THE LP WAS UNABLE TO BE RETRIEVED WITH A RETRIEVAL CATHETER DUE TO A STRETCHED HELIX. THE PACEMAKER WAS SURGICALLY REMOVED. THE PATIENT WAS ALERT WITH NORMAL VITAL SIGNS FOLLOWING THE PROCEDURE, HOWEVER, PASSED AWAY A FEW DAYS LATER. THE PHYSICIAN ALLEGED THE PATIENT PASSED AWAY DUE TO A PULMONARY EMBOLISM AND THE PROCEDURE TIME MAY HAVE CONTRIBUTED TO THE PATIENT'S PASSING. THE PHYSICIAN ALLEGED THAT THE ABBOTT PRODUCTS USED WERE NOT RELATED TO THE CARDIAC ARREST OR THE PATIENT DEATH. ADDITIONALLY, IT WAS STATED THAT THE PATIENT EXPERIENCED A PLEURAL EFFUSION. IT IS UNKNOWN IF THE PULMONARY EMBOLISM OR PLEURAL EFFUSION WAS RELATED TO THE ABBOTT PRODUCTS OR THE PROCEDURE. | FURTHER INFORMATION WAS REQUESTED BUT NOT RECEIVED. D4 - PRIMARY UNIQUE DEVICE IDENTIFICATION (UDI) NUMBER CANNOT BE PROVIDED AS A PRODUCT IDENTIFIER COULD NOT BE OBTAINED.
INTRODUCER KIT FOR GASTROSTOMY FEEDING TUBE | K080253 | IT WAS REPORTED ON (B)(6) 2025, A PATIENT DIED SHORTLY AFTER GASTRONOMY. THE FACILITY BELIEVE IT IS UNLIKELY THAT AN ISSUE WITH THE KIT CONTRIBUTED TO THE SUTURES FAILING. PER ADDITIONAL INFORMATION RECEIVED ON REPORTABLE ALERT DATE (B)(6) 2025, "PATIENT UNDERWENT ELECTIVE INSERTION OF 14F MICKEY BUTTON USING THE INTRODUCER KIT ON (B)(6). COMMENCED ON ROUTINE FEEDS POST-OPERATIVELY AND DISCHARGED HOME 48 HOURS LATER. REPRESENTED TO LOCAL HOSPITAL AFTER 4 DAYS WITH SUSPECTED PNEUMONIA. UNDERWENT CHANGE IN WATER IN THE BALLOON ON (B)(6) (WAS RECOMMENDED TO NOT PERFORM BEFORE (B)(6)) - UNSURE IF ANY CHECKS TOOK PLACE LOCALLY TO CONFIRM POSITION OF BALLOON - RECOMMENCED ON FEEDS AND DETERIORATED CLINICALLY - CT SCAN SUSPICIOUS OF BALLOON INFLATION OUTSIDE THE STOMACH - LAPAROTOMY ON (B)(6) - CONFIRMED BALLOON OUTSIDE STOMACH - T-FASTENERS CAME AWAY EASILY FROM ABDOMINAL WALL WHEN DRESSING WAS REMOVED AT TIME OF LAPAROTOMY. AT LAPAROTOMY, DEFECT IN STOMACH CLOSED AND FREKA GASTROSTOMY SITED. PATIENT REMAINED IN PICU/HDU ENVIRONMENT FOR 6-7 WEEKS POST-PROCEDURE. DEVELOPED AN ASPIRATION PNEUMONITIS FOLLOWING A FOLLOW-THROUGH STUDY TO ASSESS FOR AN OBSTRUCTION - SUBSEQUENTLY DEMISED 24 HOURS LATER ON (B)(6) 2025. INCIDENT REVIEW MEETING COMPLETED AT (B)(6) WITH LOCAL HOSPITAL." ENDOSCOPY WAS USED AT THE TIME OF INITIAL SURGERY TO VERIFY TUBE PLACEMENT. THE BALLOON INFLATION VOLUME WAS 5ML. | THE ACTUAL COMPLAINT PRODUCT WAS NOT RETURNED FOR EVALUATION. A REVIEW OF THE DEVICE HISTORY RECORD IS NOT POSSIBLE AS NO LOT NUMBER WAS PROVIDED; THEREFORE, THE UDI-PI IS UNAVAILABLE. ALL INFORMATION REASONABLY KNOWN AS OF 09 JUL 2025 HAS BEEN INCLUDED IN THIS HEALTH AUTHORITY REPORT. SHOULD ADDITIONAL INFORMATION BE OBTAINED, A FOLLOW-UP HEALTH AUTHORITY REPORT WILL BE PROVIDED. THE INFORMATION PROVIDED BY AVANOS MEDICAL INC. REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO AVANOS MEDICAL INC. AVANOS MEDICAL INC. HAS NO INDEPENDENT KNOWLEDGE OF THE REPORTED EVENT BUT IS RELAYING THE INFORMATION PROVIDED BY THE USER FACILITY WHERE THE INCIDENT OCCURRED. THIS PRODUCT INCIDENT IS DOCUMENTED IN THE AVANOS MEDICAL COMPLAINT DATABASE AND IDENTIFIED AS COMPLAINT (B)(4). THIS INFORMATION IS SUBMITTED PURSUANT TO 21CFR803, IN COMPLIANCE WITH THE MEDICAL DEVICE REPORTING REQUIREMENT AND SHOULD NOT BE CONSIDERED TO BE AN ADMISSION THAT AN AVANOS MEDICAL INC. PRODUCT IS DEFECTIVE OR CAUSED SERIOUS INJURY.
INTELLIVUE MULTI-MEASUREMENT MODULE X3 | K171801 | IT WAS REPORTED FOLLOWING A PATIENT INCIDENT, CLINICAL LEADERSHIP REQUESTED ASSISTANCE RETRIEVING AND REVIEWING THE CLINICAL AUDIT LOGS. THERE WAS A REQUEST FOR AN ON-SITE ENGINEER TO CONFIRM THE DEVICE WAS WORKING AS EXPECTED. IT WAS INDICATED THE PATIENT PASSED AWAY A FEW DAYS LATER BUT BIOMED WAS NOT SURE IF THE DEATH WAS RELATED TO THE INCIDENT. NO FURTHER DETAILS WERE PROVIDED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | THE FOLLOWING FUNCTIONAL/DIAGNOSTIC TESTING AND COMMUNICATIONS WERE PERFORMED; THE PHILIPS REMOTE CLINICAL APPLICATION SPECIALIST (CAS) SPOKE WITH THE BIOMED WHO CONFIRMED THAT THE X3 MONITOR IS CURRENTLY IN USE AND THE DEVICE IS PERFORMING AS EXPECTED. THE CAS REQUESTED AN ONSITE VISIT TO RETRIEVE THE CLINICAL AUDIT, TECHNICAL LOGS, AND CONFIRM THE MONITOR IS PERFORMING AS EXPECTED. A PHILIPS TECHNICAL CONSULTANT WENT TO THE CUSTOMER'S SITE AND PULLED LOGS. THE WORK ORDER INDICATES THE SYSTEM MEETS SPECIFICATION FOR THE PERFORMED SERVICE AND IS RETURNED TO USE. GOOD FAITH EFFORTS WERE MADE TO OBTAIN ADDITIONAL INFORMATION; HOWEVER, NO RESPONSE WAS RECEIVED FROM THE CUSTOMER CONTACT. ALTHOUGH DEVICE LOGS WERE AVAILABLE, NO DETAILS ABOUT THE PATIENT INCIDENT OR DEATH WERE PROVIDED AND NO DEVICE PROBLEM WAS ALLEGED; THEREFORE, THERE IS INSUFFICIENT INFORMATION FOR AN INVESTIGATION. BASED ON THE INFORMATION AVAILABLE THE CAUSE OF THE REPORTED PROBLEM WAS UNABLE TO BE DETERMINED. THE DEVICE REMAINS IN USE AT THE CUSTOMER'S SITE. THE INVESTIGATION CONCLUDES THAT NO FURTHER ACTION IS REQUIRED AT THIS TIME. IF ADDITIONAL INFORMATION IS RECEIVED THE COMPLAINT FILE WILL BE REOPENED.
CG FUTURE BAND | K061127 | IT WAS REPORTED THAT A PATIENT WHO HAD BEEN TREATED FOR MITRAL VALVE DISEASE WITH THE IMPLANTATION OF A CG FUTURE ANNULOPLASTY HEART RING IN THE MITRAL POSITION, PASSED AWAY. THE DEVICE WAS IMPLANTED ON (B)(6) 2025, AND THE DEATH OCCURRED ON (B)(6) 2025. THE CAUSE OF DEATH WAS NOT RECEIVED. THEREFORE, RELATEDNESS OF THE DEATH TO THE PRODUCT COULD NOT BE EXCLUDED. IT IS UNKNOWN WHETHER AN AUTOPSY WAS PERFORMED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
PATIENT INFORMATION CENTER IX | K183387 | IT WAS REPORTED THE ALARMS FOR APNEA, BRADYCARDIA, AND ASYSTOLE WERE NOT GENERATED. THE ICU PATIENT WAS FOUND PASSED AWAY IN THEIR BED AND RESUSCITATION ATTEMPTS WERE UNSUCCESSFUL. TESTING OF THE DEVICE ONSITE REVEALED NO TECHNICAL ANOMALIES. IT WAS CLARIFIED NONE OF THE ALARMS, AUDIBLE OR VISUAL, WERE GENERATED AT THE CENTRAL STATION. THE CONDITION OF THE PATIENT PRIOR AND THE CAUSE OF DEATH WERE UNKNOWN, BUT IT WAS INDICATED THE PATIENT WAS ELDERLY. | A PHILIPS FIELD SERVICE ENGINEER (FSE) WENT ONSITE AND COLLECTED THE LOGS OF THE PIC IX AND LOGS AND CONFIGURATION OF THE MP5. THE FSE DID NOT FIND ANY TECHNICAL DEFECTS WHILE ONSITE. INITIAL REVIEW OF THE LOGS INDICATES MANUAL DEACTIVATION OF THE PULSE ALARMS. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. SECTION E REPORTING INSTITUTION PHONE #: (B)(6).
SIGNIA | K191070 | IT WAS REPORTED THAT WHEN CLAMPING A SEGMENT OF THE LEFT PULMONARY ARTERY (A1/3), BLEEDING OCCURRED AFTER REMOVAL OF THE MAGAZINE AT THE PROXIMAL CLAMP LINE, WHICH RAPIDLY INCREASED AND NECESSITATED AN EMERGENCY THORACOTOMY. DUE TO MASSIVE BLEEDING WITH CARDIAC ARREST, THE PATIENT DIED. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
MESH - VENTRALEX | K132441 | PER INFORMATION PROVIDED: ON (B)(6) 2008 ¿ PATIENT WAS DIAGNOSED WITH RECURRENT VENTRAL HERNIA ABOVE UMBILICUS THEREBY UNDERWENT OPEN REPAIR WITH VENTRALEX MESH (DEVICE #1, #2). PER OP NOTES, ¿A VENTRALEX PATCH (DEVICE #1) WAS PLACED INTO PREPERITONEAL POCKET AND SUTURED. THEN, PRUDENT TROCAR SITE WAS REPAIRED USING VENTRALEX PATCH (DEVICE #2) PLACED BETWEEN THE FASCIA AND THE PERITONEUM AND SECURED USING SUTURES CATCHING A PORTION OF UNDERLYING MESH (DEVICE #1).¿ ON (B)(6) 2015 ¿ PATIENT WAS DIAGNOSED WITH CELLULITIS AND ABSCESS THEREBY UNDERWENT INCISION & DRAINAGE OF ABSCESS. PER NOTES, ¿A SMALL AMOUNT OF SEROUS FLUID WAS OBTAINED, AND ANTIBIOTICS WERE PROVIDED.¿ ON (B)(6) 2016 ¿ PATIENT WAS DIAGNOSED WITH INFECTED UMBILICAL MESH WITH CUTANEOUS FISTULA THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF MESHES (DEVICE #1, #2). PER OP NOTES, ¿A OLD CUTANEOUS DEFECTS ERODED THROUGH THE ABDOMINAL WALL AND MESH (DEVICE#1, #2) WERE NOTED. THE MESH (DEVICE #1, #2) IMBEDDED INTO FASCIA WAS REMOVED COMPLETELY. PERITONEUM HAD CHRONIC GRANULATION TISSUE WAS EXCISED.¿ ATTORNEY ALLEGES THAT THE PATIENT PASSED AWAY ON (B)(6) 2017. | BASED ON THE MEDICAL RECORD REVIEW, POST IMPLANT OF VENTRALEX MESHES (DEVICE #1, #2), PATIENT HAD BACTERIAL INFECTION, CUTANEOUS FISTULA, CELLULITIS, ABSCESS, THEREBY UNDERWENT LAPAROSCOPIC REPAIR WITH THE REMOVAL OF MESHES (DEVICE #1, #2). PER PLAINTIFF PROFILE FORM, PATIENT HAD ABSCESS, SEROMA, CHRONIC INFECTION, CELLULITIS, ABDOMINAL DRAINING AND ALLEGES PATIENT DEATH. NO INFORMATION REGARDING PATIENT¿S DEATH IN THE MEDICAL RECORDS PROVIDED TO DATE AND NO AUTOPSY REPORT, OR DEATH CERTIFICATE HAVE BEEN PROVIDED." ROOT CAUSE IS INCONCLUSIVE; NO CONCLUSION CAN BE MADE AS TO THE EXTENT TO WHICH THE IMPLANT MAY HAVE CAUSED OR CONTRIBUTED TO THE PATIENT'S POSTOPERATIVE COURSE. NOTE, THE MANUFACTURING DATE (15-FEB-2008) IS CONSIDERED TO BE A BEST ESTIMATE. THIS EMDR REPRESENTS THE BARD MESH - VENTRALEX (DEVICE #2). AN ADDITIONAL SUPPLEMENTAL EMDR WAS SUBMITTED TO REPRESENT THE BARD MESH - VENTRALEX (DEVICE #1). NOTE: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
BENEVISION DISTRIBUTED MONITORING SYSTEM (DMS) | K220058 | THE CUSTOMER IS NOT REPORTING A PRODUCT MALFUNCTION OF A MINDRAY DEVICE. THE CUSTOMER INDICATED THAT THEY WERE TRYING TO RETRIEVE DISCHARGED PATIENT INFORMATION BUT COULD NOT LOCATE IT AND IS REQUESTING HELP IN RETRIEVING PATIENT'S DISCHARGED DATA AS THE PATIENT EXPIRED ON (B)(6) 2025. | MINDRAY FIELD SERVICE REPRESENTATIVE VISITED THE FACILITY AND COLLECTED BENEVISION DMS LOGS, BUT DISCHARGED PATIENT'S DATA WAS NOT AVAILABLE. MINDRAY PERFORMED A REVIEW OF THE AVAILABLE SYSTEM LOGS, AND NO DEVICE PERFORMANCE ISSUES WERE IDENTIFIED. THE SYSTEM PERFORMED PER SPECIFICATIONS.
FREESTYLE LIBRE 3 PLUS | K223435 | AN ANSM USER REPORT WAS RECEIVED THAT OBTAINED THE FOLLOWING INFORMATION: A HEALTH CARE PROVIDER (HCP) THAT A CUSTOMER, WHO WAS PREGNANT, WAS RECEIVING FLAT AND LOW SENSOR VALUES FOR SEVERAL DAYS FROM AN ADC DEVICE. THE HCP FURTHER REPORTED THAT THE CUSTOMER EXPERIENCED SEVERE KETOACIDOSIS RESULTING IN FETAL DEATH. ON (B)(6) 2025, THE CUSTOMER WAS CONTACTED BY ADC AND REPORTED THAT ON (B)(6) 2025 AROUND 8:00 PM THEY EXPERIENCED SYMPTOMS OF DISCOMFORT, DRY MOUTH, DIZZINESS AND THE FIRE DEPARTMENT WERE CALLED WHERE A BLOOD GLUCOSE READING OF ¿BETWEEN 400 TO 600 MG/DL¿ WAS OBTAINED DURING TRANSPORT TO THE HOSPITAL. REPORTEDLY, THE CUSTOMER WAS TREATED BY THE FIREMEN AND AT THE HOSPITAL IN "REANIMATION SERVICE" (CRITICAL CARE) FOR A DIAGNOSIS OF HYPERGLYCEMIA; HOWEVER, TREATMENT DETAILS OR ADDITIONAL GLUCOSE READINGS WERE NOT PROVIDED. IT IS UNKNOWN WHETHER THE CUSTOMER NOTED A TREND ARROW ON THE DEVICE AT THE TIME OF THE EVENT. THE CUSTOMER INDICATED THE DEATH OF HER BABY OCCURRED ON (B)(6) 2025 WHILE SHE WAS IN "REANIMATION SERVICE" OF THE HOSPITAL. NO AUTOPSY REPORT OR DEATH CERTIFICATE HAS BEEN PROVIDED AT THIS TIME. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED FETAL DEATH. | AN EXTENDED MANUFACTURING REVIEW OF THE REPORTED PRODUCT HAS BEEN PERFORMED. LOT 1001001995 PERFORMED WITHIN SPECIFICATION FOR ALL MEASUREMENTS DURING THE SENSOR MANUFACTURING AND RELEASE TESTING PROCESS. THERE WERE NO NON-CONFORMANCES IN THE SAME TIME PERIOD THAT WAS RELATED TO MANUFACTURE OF LOT 1001001995. ALL MEASUREMENTS REVIEWED WERE WITHIN SPECIFICATIONS. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. THE REPORTED PRODUCT IS NOT EXPECTED TO BE RETURNED AS THE CUSTOMER INDICATED THE DEVICE WAS DISCARDED. IN THE EVENT THAT PRODUCT IS RECEIVED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ABBOTT DIABETES CARE RECEIVED AN ANSM USER REPORT (R2518288) IN REGARD TO THE SAME INCIDENT. PLEASE REFER TO ADC CASE ID (B)(4) FOR ALL SUBSEQUENT REPORTS RELATED TO THIS INCIDENT. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
TANDEM MOBI INSULIN PUMP WITH CONTROL-IQ TECHNOLOGY | K233044 | IT WAS REPORTED THAT THE CUSTOMER PASSED AWAY ON (B)(6) 2025 DUE TO A HEART ATTACK AND DIABETIC KETOACIDOSIS. THE CUSTOMER WAS WEARING THE PUMP AT THE TIME OF DEATH. A REVIEW OF PUMP DATA WILL BE PERFORMED, AND THE RESULTS WILL BE SUBMITTED IN A SUPPLEMENTAL REPORT.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A NURSE REPORTED A PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT). DURING THE INTAKE PROCESS, IT WAS REPORTED THE PATIENT WAS UNDERGOING CCPD THERAPY WHEN PATIENT EXPIRED. FOLLOW-UP WITH THE PATIENTS¿ HOME PROGRAM MANAGER (HPM) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. THE PATIENT REPORTEDLY EXPIRED WHILE SLEEPING AND WAS DISCOVERED BY THE PATIENT¿S SPOUSE. THE HPM REPORTED THE PATIENT¿S PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST AND WAS UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. | ADDITIONAL INFORMATION: D.9.,H.3. PLANT INVESTIGATION: THE ACTUAL DEVICE WAS RETURNED TO THE MANUFACTURER FOR PHYSICAL ANALYSIS. A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR SHOWED NO SIGN OF PHYSICAL DAMAGE.THERE WERE VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE COMPARTMENT.THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA.HERE WERE NOT BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE.SYSTEM AIR LEAK TEST PASSED.VALVE ACTUATION TEST PASSED.A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE VISUAL INDICATIONS OF DRIED FLUID UNDER THE PUMP ASSEMBLY ON THE BOTTOM COVER. THE CAUSE OF THE OBSERVED DRIED FLUID COULD NOT BE DETERMINED.THERE WERE NOT DISCREPANCIES ENCOUNTERED WITH THE MUSHROOM HEADS. THE DEVICE HISTORY RECORD DID NOT REVEAL ANY ISSUES OR PROBLEMS RELATED TO THE REPORTED SYMPTOM CODE(S). UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER AN UNSPECIFIED DA VINCI-ASSISTED PROCEDURE, THE PATIENT LATER PRESENTED FOR UNSPECIFIED REASONS, UNDERWENT AN EXPLORATORY LAPAROTOMY, AND SUBSEQUENTLY EXPIRED. ADDITIONAL INFORMATION HAS BEEN REQUESTED, BUT NO RESPONSE HAS BEEN RECEIVED. | INSUFFICIENT INFORMATION IS AVAILABLE FOR PROCEDURE VERIFICATION; THEREFORE, NO DA VINCI SYSTEM OR INSTRUMENT LOGS ARE AVAILABLE FOR REVIEW. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN UNKNOWN ROBOTIC PROCEDURE. NO ADDITIONAL INFORMATION IS PROVIDED ABOUT THE PROCEDURE OR ANY INTRAOPERATIVE ISSUES. THE EVENTS THAT LED TO THEIR REOPERATION, THE FINDINGS OF THE REOPERATION, AND THE SUBSEQUENT ISSUES THAT LED TO THEIR DEATH AND CAUSE OF DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI XI | K131861 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED LIVER RESECTION SURGICAL PROCEDURE, THE PATIENT EXPIRED DUE TO AN INJURY TO THE VENA CAVA, WHICH RESULTED IN A SIGNIFICANT AMOUNT OF BLOOD LOSS. MULTIPLE REQUESTS FOR ADDITIONAL INFORMATION HAVE BEEN MADE; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO ERRORS OCCURRED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MEGA SUTURECUT NEEDLE DRIVER, MONOPOLAR CURVED SCISSORS, FENESTRATED BIPOLAR FORCEPS, PROGRASP FORCEPS, MARYLAND BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. STAPLER LOG REVIEW SHOWS THE SUREFORM 60 STAPLER INSTRUMENT WAS INSTALLED ON THE SYSTEM TWICE AND SUCCESSFULLY FIRED 2 BLUE RELOADS. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. A REVIEW OF THE VESSEL SEALER EXTEND INSTRUMENT AND THE E-100 GENERATOR LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THE REPORT DIED AS THE RESULT OF BLEEDING FROM AN INJURY TO THE VENA CAVA DURING A LIVER RESECTION. HOW THIS EVENT OCCURRED OR DETAILS OF WHAT MAY HAVE LED TO THE INJURY ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. AFTER THE FIRST ATTEMPT TO HAVE THE DEVICE AND COMPONENTS EVALUATED, THE CUSTOMER RESPONDED, BUT UNABLE TO GATHER ADDITIONAL INFORMATION RELATED DEVICE RETURN. THE MANUFACTURER IS SUBMITTING AN UPDATED REPORT AT THIS TIME. IF PERTINENT INFORMATION BECOMES AVAILABLE TO THE MANUFACTURER AT A LATER DATE, AN ADDENDUM TO THIS FINAL REPORT WILL BE FILED.
POWERPORT CLEARVUE ISP IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K122899 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT SOMETIME LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED WITH UNSPECIFIED INFECTION AND RESULTED IN DEATH. THE CURRENT STATUS OF THE PATIENT WAS UNKNOWN. | H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, HENCE DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS NOT PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS COULD NOT BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
DRAGONFLY OPSTAR¿ | K192019 | IT WAS REPORTED THAT THE PROCEDURE WAS TO TREAT THE RIGHT CORONARY ARTERY WITH HEAVY CALCIFICATION, HEAVY TORTUOSITY AND 80% STENOSIS. THE DRAGONFLY OPSTAR IMAGING CATHETER WAS ADVANCED TO THE LESION WITH RESISTANCE NOTED BUT WAS ABLE TO CROSS THE LESION AND PERFORM PULLBACK. A DISSECTION WAS SUSPECTED TO BE CAUSED BY THE DRAGONFLY AND ORBITAL ATHERECTOMY WAS PERFORMED. THERE WAS A CLINICALLY SIGNIFICANT DELAY IN THE PROCEDURE DUE TO THE DISSECTION AND UNSPECIFIED TREATMENT AND THE PATIENT WAS UNSTABLE AND SYMPTOMATIC THROUGHOUT THE PROCEDURE. THE PATIENT EXPERIENCED CARDIAC ARREST AND EXPIRED. NO ADDITIONAL INFORMATION WAS PROVIDED. | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4 - A PARTIAL UDI WAS PROVIDED AS THE LOT NUMBER IS NOT KNOWN.
TRILOGY EVO, O2, JAPAN | K181166 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING WHEN CHANGING SETTINGS THAT THE VALUE HAD BECOME THE MAXIMUM VALUE ON A TRILOGY EVO O2 DEVICE. THE SALES REPRESENTATIVE VISITED THE CUSTOMER SITE AND THAT "AFTER TURNING THE POWER OFF, IT WAS AFTER A WHILE." WHEN YOU TOUCH AN ITEM THAT YOU WANT TO CHANGE BUT IT BECAME A MAXIMUM VALUE WHEN TOUCHING AN ITEM THAT WANTED TO BE CHANGED." THE SALES REPRESENTATIVE GOES ON TO STATE, "IT DID NOT RESPOND WHEN TOUCHING IT, BUT IT WAS IN A SITUATION LIKE "+" PR "+- WAS BEING PRESSED ALL THE TIME". THE SALES REPRESENTATIVE VISITED THE ME, "THERE WAS AN INQUIRY REGARDING THE CARE ORCHESTRATOR DATA. SETTINGS CHANGED ON 8TH, BUT IT WAS NOT RECORDED IN THE ELECTRONIC MEDICAL RECORDS." THE PATIENT HAD PASSED AWAY. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, WE ARE UNABLE TO CONFIRM THE ALLEGED MALFUNCTION. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. | UPDATED H10/H11, PATIENT IDENTIFIER, DATE OF BIRTH, AND RELEVANT TEST/LAB DATA FIELDS: THE MANUFACTURER RECEIVED INFORMATION ALLEGING WHEN CHANGING SETTINGS THAT THE VALUE HAD BECOME THE MAXIMUM VALUE ON A TRILOGY EVO O2 DEVICE. THE SALES REPRESENTATIVE VISITED THE CUSTOMER SITE AND THAT "AFTER TURNING THE POWER OFF, IT WAS AFTER A WHILE." WHEN YOU TOUCH AN ITEM THAT YOU WANT TO CHANGE BUT IT BECAME A MAXIMUM VALUE WHEN TOUCHING AN ITEM THAT WANTED TO BE CHANGED." THE SALES REPRESENTATIVE GOES ON TO STATE, "IT DID NOT RESPOND WHEN TOUCHING IT, BUT IT WAS IN A SITUATION LIKE "+" PR "+- WAS BEING PRESSED ALL THE TIME". THE SALES REPRESENTATIVE VISITED THE ME, "THERE WAS AN INQUIRY REGARDING THE CARE ORCHESTRATOR DATA. SETTINGS CHANGED ON 8TH, BUT IT WAS NOT RECORDED IN THE ELECTRONIC MEDICAL RECORDS." THE PATIENT HAD PASSED AWAY. THE VENTILATOR WAS RETURNED TO THE MANUFACTURER FOR EVALUATION AND THE CUSTOMER¿S COMPLAINT WAS NOT CONFIRMED DURING AN OPERATIONAL CHECK OF THE DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN WAS ABLE TO CONFIRM "THAT DURING CHANGES TO EACH SETTING, ADJUSTMENTS COULD BE MADE BOTH "+"AND "-" ACCORDING TO THE TOUCH OPERATIONS ON THE DISPLAY, WITHOUT ANY FORCED CHANGES TO THE SETTINGS. THE MANUFACTURER'S SERVICE TECHNICIAN REVIEWED THE DEVICE'S ERROR LOG AND OBSERVED NO ERRORS INDICATING ABNORMALITIES FOR THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN DISASSEMBLED THE DEVICE AND PERFORMED AN INTERNAL INSPECTION OF THE DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN CONFIRMED THERE WERE NO ABNORMALITIES FOR THIS DEVICE. THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED A FUNCTIONAL TEST ON THE DEVICE. THE DEVICE PASSED THE FUNCTIONAL TEST. THE MANUFACTURER'S SERVICE TECHNICIAN "REQUEST FOR SOFTWARE IMPROVEMENT HAS BEEN MADE. AS FOR THE CHANGE IN SETTING VALUES AT THE SPECIFIC DATE AND TIME, IT IS BELIEVED THAT THE CHANGE WAS MADE MANUALLY, AS THE CHANGE IS NOT APPLIED UNLESS THE CONFIRM BUTTON IS PRESSED EVEN IF THE ISSUE OCCURS. ADDITIONAL FINDINGS NOT RELATED TO THE MALFUNCTION/ISSUE THE SILVER STICKER AROUND THE SILENCE BUTTON WAS GONE. THE MANUFACTURER'S SERVICE TECHNICIAN REPLACED THE VANITY COVER TO ADDRESS THIS ISSUE. THE MANUFACTURER'S SERVICE TECHNICIAN PROACTIVELY REPLACED THE AIR INLET FOAM FILTER AND PARTICLE FILTER FOR THIS DEVICE. AFTERWARDS, THE MANUFACTURER'S SERVICE TECHNICIAN PERFORMED THE FINAL AND RUN-IN TESTS, ALONG WITH THE BATTERY AND VALVE CHECKS. THE MANUFACTURER'S SERVICE TECHNICIAN CONFIRMED THE DEVICE WAS OPERATIONAL, AND IT WAS CLEANED.
FREESTYLE LIBRE 3 PLUS | K223435 | ON 02-JULY-2025, ADC RECEIVED AN EMAIL FROM A CAREGIVER, WHO REPORTED A LOW READINGS ISSUE WITH THE ADC DEVICE. IT WAS REPORTED THAT THE CUSTOMER HAD BEEN RECEIVING UNSPECIFIED ¿NORMAL¿ READINGS ON THE ADC DEVICE LEADING UP TO THE EVENT, AND THEREFORE THE CUSTOMER DID NOT ADMINISTER INSULIN FOR SEVERAL DAYS AND SUBSEQUENTLY LOST CONSCIOUSNESS. THE CUSTOMER WAS ADMITTED TO THE HOSPITAL WITH A LABORATORY BLOOD GLUCOSE RESULT OF 1400 MG/DL AND DIAGNOSED WITH DIABETIC KETOACIDOSIS. THE CAREGIVER STATED THAT DURING THE HOSPITALIZATION, THE CUSTOMER PASSED AWAY ON (B)(6) 2025. NO INFORMATION WAS PROVIDED AS TO WHAT TREATMENT WAS PROVIDED DURING THE HOSPITALIZATION NOR THE CAUSE OF DEATH. NO AUTOPSY REPORT OR DEATH CERTIFICATE HAS BEEN PROVIDED AT THIS TIME. GIVEN THE INFORMATION PRESENTED AT THIS TIME, IT IS INCONCLUSIVE THAT THE ADC DEVICE HAS LED TO OR CONTRIBUTED TO THE REPORTED DEATH. ADC CUSTOMER SERVICE IS CURRENTLY IN THE PROCESS OF MAKING ADDITIONAL ATTEMPTS TO GATHER FURTHER INFORMATION REGARDING THE REPORTED DEATH. A FOLLOW-UP REPORT WILL BE PROVIDED IF ADC OBTAINS ANY ADDITIONAL INFORMATION REGARDING THIS EVENT. ON 15-JULY-2025, ADC RECEIVED ADDITIONAL INFORMATION REGARDING THIS EVENT. THE CAREGIVER REPORTED A GLUCOSE READING OF 128 MG/DL OBTAINED FROM THE ADC DEVICE WHILE THE CUSTOMER WAS IN THE HOSPITAL¿S INTENSIVE CARE UNIT. IT IS UNKNOWN WHETHER THE CAREGIVER NOTED A TREND ARROW ON THE DEVICE. ADDITIONALLY, IT IS UNKNOWN IF ANY COMPARISON READINGS WERE TAKEN WITHIN 10 MINUTES AT THE TIME OF THE MEDICAL EVENT. THE CUSTOMER REPORTEDLY DID NOT RECEIVE THIRD-PARTY TREATMENT FROM ANYONE OTHER THAN THEMSELVES. A DEATH CERTIFICATE HAS BEEN ISSUED; HOWEVER THE DEATH CERTIFICATE DOES NOT SPECIFY AN OFFICIAL CAUSE OF DEATH, AND THE CAREGIVER IS NOT AWARE OF WHAT THE OFFICIAL CAUSE MAY BE. | THE FOLLOWING SECTION HAS BEEN UPDATED: B5 - DESCRIBE EVENT OR PROBLEM THIS REPORT IS RELATED TO FDA MFR REPORT # 2954323-2025-18852. ADC RECEIVED ADDITIONAL INFORMATION ON 02-JULY-2025 IN WHICH IT WAS REPORTED THAT THE ADVERSE EVENT WAS RELATED TO A LOW READINGS ISSUE WITH THE ADC DEVICE. IT WAS PREVIOUSLY REPORTED TO ADC THAT THE ADVERSE EVENT WAS RELATED TO A SENSOR ERROR MESSAGE; THIS EVENT WAS SUBMITTED TO THE FDA UNDER MFR REPORT # 2954323-2025-18852. PER THE ADDITIONAL INFORMATION RECEIVED, NO FURTHER REPORTS WILL BE SUBMITTED UNDER FDA MFR REPORT # 2954323-2025-18852. THE DEVICE IN QUESTION WAS REMOVED AT THE HOSPITAL AND ITS CURRENT STATUS IS UNKNOWN. AN EXTENDED MANUFACTURING REVIEW OF THE REPORTED PRODUCT HAS BEEN PERFORMED. LOT 1001015875 PERFORMED WITHIN SPECIFICATION FOR ALL MEASUREMENTS DURING THE SENSOR MANUFACTURING AND RELEASE TESTING PROCESS. THERE WERE NO NON-CONFORMANCES IN THE SAME TIME PERIOD THAT WAS RELATED TO MANUFACTURE OF LOT 1001015875. ALL MEASUREMENTS REVIEWED WERE WITHIN SPECIFICATIONS. THE DHRS (DEVICE HISTORY RECORD) FOR THE PRODUCT WAS REVIEWED AND THE DHRS SHOWED THE PRODUCT MET SPECIFICATIONS PRIOR TO RELEASE TO DISTRIBUTION. THE MOST PROBABLE ROOT CAUSES ASSOCIATED WITH THIS FAILURE MODE ARE DISCONNECTED, FAULTY OR DAMAGED COMPONENTS, SOFTWARE/DATA CORRUPTION, OR MISUSE. HOWEVER, MITIGATIONS ARE IN PLACE TO REDUCE AND PREVENT SUCH ISSUES. ALL VITAL MANUFACTURING STEPS ARE VALIDATED, MONITORED, AND VERIFIED DURING MANUFACTURING TO ENSURE THE SYSTEM IS IN CONFORMANCE WITH THE VERIFIED DESIGN. ALL VITAL FUNCTIONS ARE MONITORED BY THE SYSTEM AND, WHEN NECESSARY, FUNCTION IS SUSPENDED TO SAFEGUARD AGAINST INACCURATE RESULTS. LABELING IS PROVIDED TO INSTRUCT THE USER ON THE INTENDED USE OF ALL VITAL PARTS OF THE SYSTEM TO MINIMIZE MISUSE. ALL COMPLAINTS AND COMPLAINT TRENDS ARE INVESTIGATED TO DETERMINE IF THERE IS A PRODUCT DEFECT/ DEFICIENCY. IF A PRODUCT DEFECT/ DEFICIENCY IS IDENTIFIED, A RISK EVALUATION IS COMPLETED AND COMPARED TO THE RISK MANAGEMENT REPORT, TO ENSURE THE RISK PROFILE HAS NOT CHANGED. ADDITIONALLY, AS A PART OF ABBOTT¿S POST-MARKET SURVEILLANCE PROCESS, ALL RISK EVALUATIONS WITH ASSOCIATED COMPLAINT DATA ARE REVIEWED ANNUALLY TO DETERMINE IF THE RISK PROFILES HAVE CHANGED AS COMPARED TO THE PRODUCT RISK MANAGEMENT REPORTS. THESE MONITORING PROCESSES ENSURE THAT ALL PRODUCT RISK PROFILES REMAIN ACCEPTABLE AND HAVE A POSITIVE BENEFIT/ RISK RATIO. IF PRODUCT IS RETURNED, A PHYSICAL INVESTIGATION WILL BE PERFORMED PER ADC'S ESTABLISHED PROCESSES AND PROCEDURES AND A REPORT WILL BE SUBMITTED UPON COMPLETION OF INVESTIGATION. ALL PERTINENT INFORMATION AVAILABLE TO ABBOTT DIABETES CARE HAS BEEN SUBMITTED.
NELLCOR | K142865 | IT WAS REPORTED THAT DURING USE, THE MONITOR FAILED TO ALARM, RESULTING IN THE PATIENT'S CARDIAC ARREST AND SUBSEQUENT DEATH. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
MITRACLIP | K190167 | IT WAS REPORTED THIS WAS A MITRACLIP PROCEDURE TO TREAT MITRAL REGURGITATION (MR) WITH A GRADE OF 4. AFTER A SUCCESSFUL CLIP DEPLOYMENT AND REMOVAL OF THE STEERABLE GUIDE CATHETER (SGC), THE PATIENT'S BLOOD PRESSURE DROPPED WITHIN MINUTES, LEADING TO CPR. A VEIN RUPTURE WAS SUSPECTED, AND EMERGENCY SURGICAL CONVERSION WAS PERFORMED, HOWEVER THE PATIENT DID NOT SURVIVE. THE PATIENT PASSED AWAY ON THE SAME DATE. | THE DEVICE WAS NOT RETURNED FOR ANALYSIS. A REVIEW OF THE LOT HISTORY RECORD REVEALED NO MANUFACTURING NONCONFORMITIES ISSUED TO THE REPORTED LOT THAT WOULD HAVE CONTRIBUTED TO THIS EVENT. THE INVESTIGATION WAS UNABLE TO DETERMINE A CAUSE FOR THE REPORTED PERFORATION, HYPOTENSION, AND DEATH. PERFORATION, HYPOTENSION, AND DEATH, LISTED IN THE INSTRUCTIONS FOR USE, ARE KNOWN POSSIBLE COMPLICATION ASSOCIATED WITH MITRACLIP PROCEDURES. THE REPORTED UNEXPECTED MEDICAL INTERVENTION AND SURGICAL INTERVENTION WERE RESULTS OF CASE SPECIFIC CIRCUMSTANCES. THERE IS NO INDICATION OF A PRODUCT ISSUE WITH RESPECT TO MANUFACTURE, DESIGN, OR LABELING.
DREAMSTATION | K131982 | MANUFACTURER RECEIVED A CALL FROM THE PATIENT'S DAUGHTER IN LAW, REQUESTING REPLACEMENT DEVICE SENT TO HER AS HER MOTHER-IN-LAW HAS PASSED AWAY. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A SUPPLEMENTAL FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K222318 | IT WAS REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT WAS RECEIVING DIALYSIS AND EXPIRED WHILE IN TREATMENT. THE NURSE WAS ADVISED TO DISCONTINUE THE USE OF THE CYCLER. A REPLACEMENT CYCLER WAS ISSUED. SUBSEQUENT ATTEMPTS TO OBTAIN ADDITIONAL INFORMATION (E.G., DISCHARGE SUMMARY, HOSPITAL RECORDS, ESRD DEATH NOTIFICATION, DEATH CERTIFICATE, TIMELINE OF EVENTS, PATIENT DEMOGRAPHICS) WERE UNSUCCESSFUL. | CLINICAL INVESTIGATION: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER, AND THE SERIOUS ADVERSE EVENT OF DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING TREATMENT WHEN THE SERIOUS ADVERSE EVENT OCCURRED. THE ETIOLOGY OF THE SERIOUS ADVERSE EVENTS IS UNKNOWN; THEREFORE, CAUSALITY COULD NOT BE ESTABLISHED. IT SHOULD BE NOTED THAT LIMITED CLINICAL FOLLOW-UP PRECLUDED A MORE COMPREHENSIVE ASSESSMENT/INVESTIGATION. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. OF THESE DEATHS, CARDIOVASCULAR DISEASE (OR SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CANNOT BE DISASSOCIATED FROM HAVING A POSSIBLE CONTRIBUTORY ROLE IN THE PATIENT¿S EXPIRATION. THERE IS CURRENTLY NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION OCCURRED. HOWEVER, GIVEN THE UNKNOWN CAUSE OF DEATH, PAST MEDICAL HISTORY, SUMMARY/TIMELINE OF EVENTS, DISCHARGE SUMMARY, AND NO MANUFACTURER EVALUATION OF THE PATIENT¿S DEVICE, THIS CLINICAL INVESTIGATION CANNOT EXCLUDE THE PATIENT¿S LIBERTY SELECT CYCLER FROM PLAYING A POSSIBLE ROLE IN THE SERIOUS ADVERSE EVENTS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
ILET BIONIC PANCREAS | K231485 | ON 14-APR-2025, BETA BIONICS WAS CONTACTED BY (B)(6) REGARDING THE DEATH OF A 56-YEAR-OLD MALE WHO HAD BEEN USING THE ILET. THE MEDICAL EXAMINER REQUESTED DEVICE EVALUATION, AS THE USER WAS WEARING THE DEVICE AT THE TIME OF PASSING. THE ILET WAS SUBSEQUENTLY RETURNED TO BETA BIONICS FOR INVESTIGATION. | NO DEVICE MALFUNCTIONS WERE IDENTIFIED; THE ILET FUNCTIONED WITHIN SPECIFICATION. INSULIN DELIVERY AND SUSPENSION OCCURRED AS EXPECTED IN RESPONSE TO CGM VALUES AND MEAL ANNOUNCEMENTS. PREDICTIVE AND THRESHOLD-BASED HYPOGLYCEMIA ALARMS WERE APPROPRIATELY TRIGGERED BUT NOT ACKNOWLEDGED. DESPITE AUTOMATIC SUSPENSION OF INSULIN DELIVERY IN RESPONSE TO FALLING GLUCOSE LEVELS, SUSTAINED HYPOGLYCEMIA (CGM <54MG/DL) WAS OBSERVED BEGINNING AT APPROXIMATELY 00:19 ON (B)(6) 2025. A CGM SENSOR ISSUE WAS SUBSEQUENTLY REPORTED, FOLLOWED BY SENSOR FAILURE AND LOSS OF GLUCOSE DATA. BASED ON ALL AVAILABLE INFORMATION, THE ILET SYSTEM OPERATED AS INTENDED. THERE IS NO EVIDENCE TO SUGGEST THAT THE ILET CAUSED OR CONTRIBUTED TO THE USER'S DEATH. THE DOCUMENTED CAUSE OF DEATH WAS HYPERTENSIVE AND ATHEROSCLEROTIC CORONARY HEART DISEASE.
DAVINCI XI | K131861 | IT WAS REPORTED THAT A PATIENT EXPIRED DURING A DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA SURGICAL PROCEDURE. THE SURGEON REPORTED THAT THE PATIENT HAD A HISTORY OF LUNG CANCER WITH RADIATION AND CHEMOTHERAPY TREATMENTS RESULTING IN DIFFICULT, SCARRED, AND WEAK TISSUE ANATOMY. WHILE DISSECTING TISSUE NEAR THE VAGUS NERVE, CONTROLLED BLEEDING OCCURRED AND PRESSURE WAS APPLIED. THE PATIENT THEN EXPERIENCED BRADYCARDIA FROM A VASOVAGAL RESPONSE THAT DETERIORATED TO ASYSTOLE. THE BLEEDING SITE WAS PACKED, THE ROBOT WAS EMERGENTLY UNDOCKED, AND THE PROCEDURE WAS CONVERTED TO OPEN. CHEST COMPRESSIONS WERE INITIATED. DESPITE RESUSCITATION EFFORTS THE PATIENT REMAINED IN ASYSTOLE; NO CARDIAC RHYTHM WAS RESTORED AND THE PATIENT EXPIRED. THE SURGEON STATED THAT NO MALFUNCTIONS OF ANY INTUITIVE PRODUCTS OCCURRED DURING THE PROCEDURE. NO ADDITIONAL INFORMATION WAS PROVIDED. | A REVIEW OF THE SYSTEM LOGS FOR THE REPORTED PROCEDURE DATE FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM SHOW THE SYSTEM FUNCTIONED NORMALLY AND THERE WERE NO INDICATIONS RELATING TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, TWO CADIERE FORCEPS, PERMANENT CAUTERY HOOK, VESSEL SEALER EXTEND, SUCTION IRRIGATOR, 8MM 0 DEGREE ENDOSCOPE. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) INVOLVED WITH THE REPORTED EVENT FOUND NO NON-CONFORMANCES WERE IDENTIFIED TO BE RELATED TO THIS COMPLAINT. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A BRADYCARDIC EVENT LIKELY FROM A VAGAL RESPONSE DURING A HIATAL HERNIA REPAIR. THIS LED TO ASYSTOLE AND THE PATIENT EXPIRED DESPITE RESUSCITATIVE EFFORTS. NO ALLEGATION IS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. BLANK MDR FIELDS: THE MISSING PATIENT INFORMATION IN SECTIONS A AND B WAS EITHER UNKNOWN, UNAVAILABLE, NOT PROVIDED, OR NOT APPLICABLE. THE EXPIRATION DATE FOR SECTION D4 IS NOT APPLICABLE. FIELD D6 IS BLANK BECAUSE THE PRODUCT IS NOT IMPLANTABLE. INFORMATION FOR THE BLANK FIELDS IN SECTION E1 IS NOT AVAILABLE. FIELD E4 IS BLANK BECAUSE IT IS UNKNOWN IF THE INITIAL REPORTER SUBMITTED A REPORT TO THE FDA. FIELDS G5 AND G7 ARE NOT APPLICABLE.
DAVINCI 5 | K232610 | IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED PANCREATICODUODENECTOMY SURGICAL PROCEDURE FOR A BENIGN TUMOR, THE PATIENT EXPIRED POSTOPERATIVELY. THE ROBOTIC PROCEDURE WAS COMPLETED AS PLANNED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS. BLOOD LOSS WAS MINIMAL AND THE SURGICAL FIELD WAS "DRY" AT THE END OF THE PROCEDURE. THE SURGEON REPORTED THAT THE PATIENT REQUIRED A SUBSEQUENT OPEN LAPAROTOMY A FEW HOURS POSTOPERATIVELY DUE TO BLOOD PRESSURE COMPLICATIONS INDICATING BLEEDING. AN OPEN LAPAROTOMY WAS PERFORMED AND THE BLEEDING WAS IDENTIFIED FROM A SMALL BLOOD VESSEL. THE BLEEDING WAS STOPPED WITH UNSPECIFIED INTERVENTIONS, AND THE ANASTOMOSIS WAS REDONE. THE PATIENT ULTIMATELY EXPIRED THREE DAYS POSTOPERATIVELY IN THE INTENSIVE CARE UNIT DUE TO MULTISYSTEM ORGAN FAILURE. THE SURGEON REPORTED THAT THE PATIENT¿S POSTOPERATIVE COURSE WAS RELATED TO THE MANAGEMENT OF MEDICAL CARE AND DOES NOT ATTRIBUTE THE PATIENT'S POSTOPERATIVE COURSE TO BE RELATED TO THE INITIAL DA VINCI PROCEDURE. NO SPECIFIC DETAILS WERE PROVIDED. | REVIEW OF THE SYSTEM LOGS FOUND NO ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON DETERMINED THAT THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE 30 DEGREE, TIP-UP FENESTRATED GRASPER, FORCE BIPOLAR, TWO CLIP APPLIERS, MONOPOLAR CURVED SCISSORS, PERMANENT CAUTERY HOOK, LARGE SUTURECUT NEEDLE DRIVER, FORCE FEEDBACK MARYLAND BIPOLAR FORCEPS, VESSEL SEALER EXTEND, SUREFORM 60 STAPLER, SUREFORM 30 CURVED-TIP STAPLER, SUREFORM 45 STAPLER, SUCTION IRRIGATOR. SITE HISTORY REVIEW FOUND NO COMPLAINTS FOR THE INSTRUMENTS USED IN THE PROCEDURE, AND THE MULTIPLE-USE DEVICES WERE USED IN SUBSEQUENT PROCEDURES AFTER THE EVENT DATE WITH NO REPORTED COMPLAINTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. A REVIEW OF THE VESSEL SEALER EXTEND LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. STAPLER LOG REVIEWS OF THE STAPLER INSTRUMENTS FOUND THE SUREFORM 60 STAPLER WAS INSTALLED ONCE AND FIRED ONE BLUE RELOAD. THE SUREFORM 30 CURVED-TIP STAPLER INSTRUMENT WAS INSTALLED FIVE TIMES AND FIRED THREE WHITE RELOADS. THE SUREFORM 45 STAPLER INSTRUMENT WAS INSTALLED FIVE TIMES AND FIRED THREE BLUE RELOADS AND ONE WHITE RELOAD. THERE WERE NO STAPLER RELATED ERRORS IN THE LOGS. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT CONCLUDED THAT THE PATIENT IN THIS REPORT DEVELOPED A DELAYED BLEED AFTER A ROBOTIC PANCREATICODUODENECTOMY. WHAT SPECIFIC VESSEL BLED AND HOW OR WHY IT MAY HAVE BLED WAS NOT PROVIDED. THE AMOUNT OF BLEEDING AND THE ISSUES THAT RESULTED IN THE BLEEDING WERE NOT PROVIDED. THE SPECIFIC ANASTOMOSIS THAT WAS REDONE WAS NOT PROVIDED. PERFORMING THE ANASTOMOSIS REVISION MAY INDICATE THAT IT WAS THE SITE OF BLEEDING; HOWEVER, THE SPECIFIC SITE OF BLEEDING WAS NOT DEFINED. THE PATIENT DEVELOPED MULTISYSTEM ORGAN FAILURE AND EXPIRED 3 DAYS LATER; HOW THE ORGAN FAILURE DEVELOPED WAS NOT PROVIDED. NO ALLEGATION WAS MADE ABOUT ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT.
DAVINCI 5 | K232610 | IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED PARAESOPHAGEAL HIATAL HERNIA SURGICAL PROCEDURE, THE PATIENT EXPIRED DUE TO COMPLICATIONS OF AN ASPIRATION EVENT. THE SURGEON STATED THAT THE HIATAL HERNIA WAS LARGE, BUT THE ROBOTIC PROCEDURE WAS COMPLETED SUCCESSFULLY WITHOUT ANY INTRAOPERATIVE COMPLICATIONS, AND WITHOUT ERRORS OR ISSUES RELATED TO THE DA VINCI SYSTEM. THE PATIENT WAS RECOVERING; HOWEVER, REQUIRED PROLONGED INTUBATION DUE TO PULMONARY EDEMA. THREE WEEKS POST-PROCEDURE, THE PATIENT WAS STILL INTUBATED, AND DURING AN ENDOTRACHEAL TUBE EXCHANGE, THE PATIENT ASPIRATED AND ULTIMATELY EXPIRED DUE TO COMPLICATIONS FROM THE ASPIRATION EVENT. THE SURGEON BELIEVES THE PROLONGED HOSPITALIZATION AND INTUBATION WERE RELATED TO THE PATIENT'S LARGE HIATAL HERNIA AND PULMONARY EDEMA. THE PATIENT¿S MEDICAL HISTORY AND SPECIFIC ORIGIN OF THE PULMONARY EDEMA WERE NOT PROVIDED. | REVIEW OF THE SYSTEM LOGS FOUND THE SYSTEM FUNCTIONED NORMALLY AND THERE WERE NO ERRORS FOUND RELATED TO THIS EVENT. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM DETERMINED IT FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: ENDOSCOPE 30 DEGREE, FORCE FEEDBACK CADIERE FORCEPS, FORCE FEEDBACK FENESTRATED BIPOLAR FORCEPS, FORCE FEEDBACK MEGA SUTURECUT NEEDLE DRIVER, VESSEL SEALER EXTEND, SUCTION IRRIGATOR. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. A REVIEW OF THE VESSEL SEALER EXTEND LOGS FOUND THE INSTRUMENT WAS INSTALLED ONCE AND PASSED HOMING. THE CUT COMPLETE AND SEAL EVENTS WERE COMPLETED WITH NO RELEVANT ERRORS NOTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THIS REPORT DIED AS A RESULT OF AN ASPIRATION EVENT WHILE IN THE ICU SEVERAL WEEKS AFTER A ROBOTIC HIATAL HERNIA REPAIR. THIS TYPE OF POST OPERATIVE COMPLICATION CAN OCCUR WITH INTUBATED PATIENTS AND HAS NO RELATIONSHIP TO THE ORIGINAL PROCEDURE. THERE IS NO ALLEGATION AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, NO INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION B2 DATE OF DEATH IS ESTIMATED. THE SURGEON STATED THAT THE PATIENT'S DATE OF DEATH WAS UNKNOWN, APPROXIMATELY THREE WEEKS POST-PROCEDURE.
CS100 INTRA-AORTIC BALLOON PUMP, ENGLISH, 220V | K031636 | IT WAS REPORTED ON (B)(6) 2025, THAT CS100 INTRA-AORTIC BALLOON PUMP (IABP) SHOWED ELECTRICAL TEST FAILD CODE #118 DURING USE ON PATIENT. THE PATIENT HAS DIED. CUSTOMER ALSO REPORTED THAT A SYSTEM FAILURE ERROR OCCURRED ON (B)(6) 2025. | DUE TO CHARACTER LIMITATION IN E1, 1. FULL EVENT SITE NAME IS (B)(6). A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
EVERFLO | K061261 | THE RELATIVE OF A PATIENT INQUIRED ABOUT RETURNING A PATIENT'S DEVICE REPORTING THE PATIENT PASSED AWAY. THE PATIENT'S CAUSE OF DEATH WAS NOT PROVIDED. ADDITIONAL INFORMATION IS BEING REQUESTED FROM THE REPORTER SURROUNDING THE DETAILS OF THE PATIENT'S DEATH, AND IF THERE WAS ANY KNOWN DEVICE ISSUE. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
EMBOTRAP | K193063 | ON 25-JUN-2025, THE HEALTHCARE PROFESSIONAL REPORTED THAT SHE WAS MADE AWARE OF AN EVENT THAT OCCURRED ON 23-DEC-2024. THE PROCEDURE WAS A MECHANICAL THROMBECTOMY. THE 86-YEAR-OLD PATIENT PRESENTED WITH ¿LOTS OF COMORBIDITIES¿ PRESENTED WITH AN OCCLUSION IN THE DISTAL M2 SEGMENT OF THE MIDDLE CEREBRAL ARTERY (MCA). IT WAS REPORTED THAT ¿IT WAS CONSIDERED WHETHER PATIENT WAS A SUITABLE CANDIDATE FOR THROMBECTOMY. [THE] PATIENT HAD LOTS OF VESSEL ATROPHIES. VESSELS [DID NOT] LOOK LIKE THEY WOULD BE SUITABLE FOR USING A STENT RETRIEVER, SO THE FIRST PASS WAS PERFORMED WITH ASPIRATION ONLY. THE FIRST PASS WAS PERFORMED USING A RED 62 REPERFUSION CATHETER (PENUMBRA) WAS NOT SUCCESSFUL, NO [CLOT] WAS REMOVED. THE SECOND PASS WAS PERFORMED VIA CO-ASPIRATION, USING A 5MM X 22MM EMBOTRAP III REVASCULARIZATION DEVICE (ET307522 / 24J087AV) DELIVERED VIA THE PHENOM¿ 021 MICROCATHETER (MEDTRONIC), THE RED 62 REPERFUSION CATHETER.¿ IT WAS REPORTED THAT ONE PASS OF THE EMBOTRAP III DEVICE AND THE CLOT WAS REMOVED, AND THE VESSEL WAS OPENED, ¿BUT THE VESSEL EVULSED AND A BLEED WAS SEEN JUST PAST THE BEND OF THE M1/M2 SEGMENTS.¿ THE PHYSICIAN PERFORMED COIL EMBOLIZATION TO STOP THE BLEED, BUT THE PATIENT EXPIRED FOLLOWING THE PROCEDURE. THE NEUROINTERVENTIONALIST COMMENTED THAT ¿THE OUTCOME WAS NOT CAUSED BY EMBOTRAP AND THAT HE WOULD HAVE HAD THE SAME OUTCOME WITH ANY STENT RETRIEVER THAT WAS USED. THE OUTCOME WAS PATIENT-RELATED AND NOT DEVICE-RELATED. THE PATIENT WAS PROBABLY NOT SUITABLE FOR PROCEDURE, ELDERLY, VERY UNWELL WITH LOTS OF COMORBIDITIES AND FRIABLE VESSELS.¿ | MANUFACTURER¿S REF. NO: (B)(4). INFORMATION REGARDING PATIENT IDENTIFIER, DATE OF BIRTH, SEX, GENDER, WEIGHT, RACE, AND ETHNICITY WERE NOT PROVIDED. SECTION B.2: DATE OF DEATH: THE DATE OF DEATH IS NOT KNOWN / REPORTED. SECTION E.1: THE INITIAL REPORTER PHONE IS NOT AVAILABLE / REPORTED. BASED ON COMPLAINT INFORMATION, THE DEVICE IS NOT AVAILABLE TO BE RETURNED FOR ANALYSIS. A REVIEW OF THE MANUFACTURING DOCUMENTATION ASSOCIATED WITH THIS LOT (24J087AV) TOP AND SUB ASSEMBLY PRESENTED NO ISSUES DURING THE MANUFACTURING OR INSPECTION PROCESSES RELATED TO THE REPORTED COMPLAINT. THERE WERE NO INTERNAL ACTIONS RELATED TO DEVICE MANUFACTURE OR INSPECTION. PRODUCT ANALYSIS CANNOT BE CONDUCTED AS THE PRODUCT WAS NOT RETURNED FOR ANALYSIS. NO DETERMINATION OF CAUSES AND POSSIBLE CONTRIBUTING FACTORS COULD BE MADE. AS SUCH, THE INVESTIGATION WILL BE CLOSED. THERE WAS NO DEVICE PERFORMANCE ISSUE OR DEVICE MALFUNCTION REPORTED DURING THE PROCEDURE. BASED ON THE MANUFACTURING DOCUMENTATION REVIEW, THERE IS NO INDICATION THAT THE EVENT IS RELATED TO THE DEVICE MANUFACTURING PROCESS. AS PART OF THE POST MARKET SURVEILLANCE PROGRAM, INFORMATION FROM THIS COMPLAINT IS TRENDED FOR STATISTICAL SIGNALS AND CORRECTIVE / PREVENTIVE ACTION MAY BE TRIGGERED LATER. SINCE THERE WAS NO EVIDENCE TO SUGGEST THE EVENT WAS RELATED TO A MANUFACTURING OR DESIGN ISSUE, NO CORRECTIVE ACTIONS WILL BE TAKEN AT THIS TIME. VESSEL INJURY AND CEREBRAL HEMORRHAGE ARE KNOWN POTENTIAL COMPLICATIONS ASSOCIATED WITH THE EMBOTRAP III DEVICE AND ARE MENTIONED IN THE INSTRUCTIONS FOR USE (IFU) AS SUCH. THERE WAS NO ALLEGED QUALITY ISSUE RELATED TO THE USED DEVICES, AS THE DEVICES PERFORMED AS INTENDED. THERE ARE PATIENT AND PROCEDURAL FACTORS, INCLUDING VESSEL CHARACTERISTICS (I.E., ATROPHY, FRIABLE VESSELS), ADVANCED AGE WITH MULTIPLE COMORBIDITIES, DEVICE INTERACTION, AND OPERATOR TECHNIQUES, THAT MAY HAVE CONTRIBUTED TO THE REPORTED EVENTS RATHER THAN THE DESIGN OR MANUFACTURE OF THE DEVICE. ADDITIONALLY, THE PHYSICIAN STATED THE EVENTS WERE ATTRIBUTED TO PATIENT-RELATED FACTORS AND THE PATIENT WAS PROBABLY NOT SUITABLE FOR PROCEDURE. HOWEVER, SINCE THE EVENTS OCCURRED DURING THE USE OF THE EMBOTRAP DEVICE, THE CORRELATION BETWEEN EVENTS TO THE DEVICE USED AS A CONTRIBUTING FACTOR CANNOT BE RULED OUT. ADDITIONALLY, THE EVENTS REQUIRED A SURGICAL INTERVENTION, AND THE PATIENT ULTIMATELY EXPIRED SHORTLY AFTER THE PROCEDURE. BASED ON THIS INFORMATION, THIS EVENT MEETS US FDA REPORTING CRITERIA UNDER 21 CFR 803 WITH A CLASSIFICATION OF ¿DEATH.¿ THE FILE WILL BE RE-REVIEWED IF ADDITIONAL INFORMATION IS RECEIVED AT A LATER DATE. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, JOHNSON & JOHNSON MEDTECH, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE. THE MANUFACTURER WILL SUBMIT A SUPPLEMENTAL REPORT IF NEW FACTS ARISE WHICH MATERIALLY ALTER INFORMATION SUBMITTED IN A PREVIOUS MDR REPORT. ADDITIONAL INFORMATION WILL BE SUBMITTED WITHIN 30 DAYS OF RECEIPT.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED THAT THIS PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY SUDDENLY. THE CAUSE OF DEATH WAS UNKNOWN. IT WAS NOT REPORTED THAT THE PATIENT WAS IN ACTIVE TREATMENT AT THE TIME OF DEATH. ADDITIONAL INFORMATION REQUESTED BUT HAS NOT BEEN OBTAINED. | ADDITIONAL INFORMATION PROVIDED IN D9 AND H3. CORRECTION PROVIDED IN G4. PLANT INVESTIGATION: A VISUAL INSPECTION OF THE RETURNED CYCLER EXTERIOR NO SHOWED SIGNS OF PHYSICAL DAMAGE. THERE WERE NO VISUAL INDICATIONS OF DRIED FLUID WITHIN THE CASSETTE. THERE WERE VISUAL INDICATIONS OF PARTICULATES WITHIN THE CASSETTE AREA. THERE WERE NO BURRS OR SHARP EDGES IN CASSETTE AREA THAT MAY HAVE PUNCTURED A CASSETTE MEMBRANE. SYSTEM AIR LEAK TEST PASSED. VALVE ACTUATION TEST PASSED. A SIMULATED TREATMENT USING (AS-RECEIVED) TREATMENT SETTINGS WITH REDUCED DWELL TIMES WAS PERFORMED AND COMPLETED WITHOUT FAILURES. NO FLUID LEAKS IN THE TEST CASSETTE DURING THE TREATMENT TEST. THERE WERE NO VISUAL DISCREPANCIES ENCOUNTERED DURING THE INTERNAL INSPECTION. A REVIEW OF THE DEVICE MANUFACTURING RECORDS WAS CONDUCTED BY THE MANUFACTURER. THERE WERE NO DEVIATIONS OR NON-CONFORMANCES DURING THE MANUFACTURING PROCESS. IN ADDITION, A DEVICE HISTORY RECORD (DHR) REVIEW WAS PERFORMED AND VERIFIED THAT THE RESULTS OF THE IN-PROGRESS AND FINAL QUALITY CONTROL (QC) TESTING MET ALL REQUIREMENTS. UPON COMPLETION OF THE EVALUATION, THERE WERE NO MALFUNCTIONS THAT COULD HAVE CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. THE CYCLER PERFORMED AS DESIGNED AND AN ASSOCIATED CAUSE COULD NOT BE DETERMINED.
INTELLIVUE MX40 WLAN | K113125 | IT WAS REPORTED THE PATIENT PASSED AWAY WHILE BEING MONITORED ON THE MX40 TELEMETRY DEVICE AND THE CUSTOMER REQUESTED ASSISTANCE WITH AUDIT LOG INTERPRETATION. THE PATIENT WAS ADMITTED FOR OBSERVATION POST-SYNCOPAL EPISODE WITH A FALL. THE PATIENT THEN PASSED AWAY WHILE ON TELEMETRY AND THE CUSTOMER REQUESTED INTERPRETATION OF THE CLINICAL AUDIT LOGS. | PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE. REPORTING INSTITUTION PHONE # (B)(6).
GORE® MOLDING & OCCLUSION BALLOON CATHETER | K172567 | THE FIELD SALES ASSOCIATE (FSA) REPORTED THE FOLLOWING TO GORE: ON (B)(6) 2025, THE PATIENT UNDERWENT AN ENDOVASCULAR PROCEDURE TO TREAT AN ANEURYSM - RUPTURE UTILIZING THE GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM. IT WAS REPORTED THAT THE ANEURYSM WAS EMERGENT, IT HAD NOT YET RUPTURED HOWEVER, IT WAS EXTREMELY ENLARGED. IT WAS OVER 10CM. AFTER IMPLANTATION OF THE TRUNK IPSILATERAL COMPONENT, AN AORTIC OCCLUSION BALLOON WAS USED TO DILATE THE MAIN BODY PORTION OF THE DEVICE. AFTER THE BALLOON WAS DEFLATED, THE PATIENTS PRESSURE MARKEDLY DROPPED. AN ANGIOGRAM WAS PERFORMED WHICH SHOWED A LARGE AMOUNT OF EXTRALUMINAL CONTRAST ORIGINATING FROM THE DISTAL AORTA NECK. AORTIC EXTENDERS WERE DEPLOYED PROXIMALLY TO TRY AND SEAL THE LARGE TYPE 1A ENDOLEAK WHICH WAS PRESENT DUE TO CALCIUM IN THE PROXIMAL NECK. THE PATIENT WAS NOT A CANDIDATE FOR CONVERSION TO OPEN REPAIR SO COMFORT CARE WAS CHOSEN, AND THE PATIENT EXPIRED A COUPLE HOURS AFTER THE PROCEDURE. THE PHYSICIAN BELIEVES THAT CAUSE OF THE REPORTED EVENT WAS DUE TO THE DISTAL AORTIC NECK RUPTURE DURING BALLOON DILATION OF THE PROXIMAL INFRARENAL AORTIC NECK, CALCIUM PRESENT IN THE PROXIMAL SEALING ZONE, AND A LARGE 1A ENDOLEAK THAT COULDN'T BE RESOLVED THAT LED TO EXSANGUINATION. | W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
SOLITAIRE FR4 | K183022 | ADDITIONAL INFORMATION RECEIVED REPORTED AN AUTOPSY WAS NOT PERFORMED. BASELINE CONDITION TICI 0; NO BLOOD FLOW PERFUSION. POST-THROMBECTOMY CONDITION: TICI 1: PARTIAL REPERFUSION, WITH LIMITED BLOOD FLOW RESTORATION TO THE AFFECTED AREA. BASED ON THE INFORMATION FROM PHYSICIAN, THERE ARE MANY FACTOR THAT CAN BE THE CAUSE OF THE DEATH, NOT ONLY BECAUSE THE DEFECT OF THE ITEM. THE CAUSE OF THE SOLITAIRE RESISTANCE WAS NOT DEFINITELY KNOWN; POSSIBLY DUE TO EXTREMELY HARD CLOT. CONTINUOUS SALINE FLUSH WAS ADMINISTERED AND MAINTAINED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THERE WAS NO DAMAGE OBSERVED TO THE REBAR CATHETER. THE SOLITAIRE BROKE IN THE FIRST PASS. THE SOLITAIRE DEVICE WAS NOT TORQUED OR REPOSITIONED DURING DELIVERY OR RETRIEVAL. THE PATIENT DID NOT HAVE VESSEL STENOSIS PROXIMAL TO THE THROMBUS SITE. THE MICROCATHETER TIP COVERED THE SOLITAIRE DEVICE PROXIMAL MARKER DURING THE ATTEMPTED RETRIEVAL. DURING THE PROCEDURE, THE PATIENT REMAINED STABLE, WITH VITAL SIGNS INCLUDING BLOOD PRESSURE WITHIN NORMAL LIMITS. UPON COMPLETION OF THE PROCEDURE AT APPROXIMATELY 2:00 AM, THE PATIENT CONTINUED TO EXHIBIT STABLE VITAL SIGNS. IN ACCORDANCE WITH STANDARD HOSPITAL PROTOCOL, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR POST-PROCEDURAL RECOVERY. AT APPROXIMATELY 10:00 AM, THE PATIENT WAS REPORTED TO HAVE PASSED AWAY IN THE ICU DUE TO SHOCK. THE SOLITAIREX DEVICE WAS UTILIZED DURING THE PROCEDURE. THE ATTENDING PHYSICIAN HAS CONFIRMED THAT THE PATIENT'S DEATH WAS NOT ATTRIBUTABLE TO ANY PRODUCT DEFECT OR DAMAGE RELATED TO SOLITAIREX. THE DAMAGED COMPONENT OF THE SOLITAIREX DEVICE WAS SUCCESSFULLY RETRIEVED DURING THE PROCEDURE, AND NO REMNANTS WERE LEFT IN THE PATIENT'S BODY. BASED ON THE ATTENDING PHYSICIAN¿S ASSESSMENT AND PROCEDURAL DOCUMENTATION, THE PATIENT¿S DEATH WAS NOT CAUSED BY DEVICE MALFUNCTION OR PROCEDURAL ERROR. ALL STANDARD PROTOCOLS WERE FOLLOWED, AND THE PATIENT WAS MONITORED APPROPRIATELY THROUGHOUT THE PROCESS. IT WAS CLARIFIED THE PATIENT WAS ALREADY UNCONSCIOUS PRIOR TO THE THROMBECTOMY PROCEDURE, ALTHOUGH VITAL SIGNS REMAINED STABLE. THE FAMILY REQUESTED THAT THE THROMBECTOMY STILL BE PERFORMED AND AGREED TO ALL ASSOCIATED RISKS, AS THE PATIENT'S CONDITION WOULD HAVE FURTHER DETERIORATED WITHOUT THE PROCEDURE. AFTER THE THROMBECTOMY PROCEDURE WAS COMPLETED, THE PATIENT'S CONDITION REMAINED THE SAME AS IT WAS PRIOR TO THE OPERATION. THE PATIENT DID NOT PASS AWAY DUE TO SHOCK, BUT DUE TO INSUFFICIENT RESTORATION OF BLOOD FLOW IN THE DISTAL MIDDLE CEREBRAL ARTERY. THE ATTENDING PHYSICIAN HAS CLINICALLY CONFIRMED THAT THERE WAS NO BLEEDING OR VESSEL RUPTURE RESULTING FROM THE THROMBECTOMY PROCEDURE. WHAT WAS MEANT BY "SHOCK" WAS ACTUALLY "DETERIORATION". | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
REBAR | K093750 | ADDITIONAL INFORMATION RECEIVED REPORTED AN AUTOPSY WAS NOT PERFORMED. BASELINE CONDITION TICI 0; NO BLOOD FLOW PERFUSION. POST-THROMBECTOMY CONDITION: TICI 1: PARTIAL REPERFUSION, WITH LIMITED BLOOD FLOW RESTORATION TO THE AFFECTED AREA. BASED ON THE INFORMATION FROM PHYSICIAN, THERE ARE MANY FACTOR THAT CAN BE THE CAUSE OF THE DEATH, NOT ONLY BECAUSE THE DEFECT OF THE ITEM. THE CAUSE OF THE SOLITAIRE RESISTANCE WAS NOT DEFINITELY KNOWN; POSSIBLY DUE TO EXTREMELY HARD CLOT. CONTINUOUS SALINE FLUSH WAS ADMINISTERED AND MAINTAINED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THERE WAS NO DAMAGE OBSERVED TO THE REBAR CATHETER. THE SOLITAIRE BROKE IN THE FIRST PASS. THE SOLITAIRE DEVICE WAS NOT TORQUED OR REPOSITIONED DURING DELIVERY OR RETRIEVAL. THE PATIENT DID NOT HAVE VESSEL STENOSIS PROXIMAL TO THE THROMBUS SITE. THE MICROCATHETER TIP COVERED THE SOLITAIRE DEVICE PROXIMAL MARKER DURING THE ATTEMPTED RETRIEVAL. DURING THE PROCEDURE, THE PATIENT REMAINED STABLE, WITH VITAL SIGNS INCLUDING BLOOD PRESSURE WITHIN NORMAL LIMITS. UPON COMPLETION OF THE PROCEDURE AT APPROXIMATELY 2:00 AM, THE PATIENT CONTINUED TO EXHIBIT STABLE VITAL SIGNS. IN ACCORDANCE WITH STANDARD HOSPITAL PROTOCOL, THE PATIENT WAS TRANSFERRED TO THE INTENSIVE CARE UNIT (ICU) FOR POST-PROCEDURAL RECOVERY. AT APPROXIMATELY 10:00 AM, THE PATIENT WAS REPORTED TO HAVE PASSED AWAY IN THE ICU DUE TO SHOCK. THE SOLITAIREX DEVICE WAS UTILIZED DURING THE PROCEDURE. THE ATTENDING PHYSICIAN HAS CONFIRMED THAT THE PATIENT'S DEATH WAS NOT ATTRIBUTABLE TO ANY PRODUCT DEFECT OR DAMAGE RELATED TO SOLITAIREX. THE DAMAGED COMPONENT OF THE SOLITAIREX DEVICE WAS SUCCESSFULLY RETRIEVED DURING THE PROCEDURE, AND NO REMNANTS WERE LEFT IN THE PATIENT'S BODY. BASED ON THE ATTENDING PHYSICIAN¿S ASSESSMENT AND PROCEDURAL DOCUMENTATION, THE PATIENT¿S DEATH WAS NOT CAUSED BY DEVICE MALFUNCTION OR PROCEDURAL ERROR. ALL STANDARD PROTOCOLS WERE FOLLOWED, AND THE PATIENT WAS MONITORED APPROPRIATELY THROUGHOUT THE PROCESS. IT WAS CLARIFIED THE PATIENT WAS ALREADY UNCONSCIOUS PRIOR TO THE THROMBECTOMY PROCEDURE, ALTHOUGH VITAL SIGNS REMAINED STABLE. THE FAMILY REQUESTED THAT THE THROMBECTOMY STILL BE PERFORMED AND AGREED TO ALL ASSOCIATED RISKS, AS THE PATIENT'S CONDITION WOULD HAVE FURTHER DETERIORATED WITHOUT THE PROCEDURE. AFTER THE THROMBECTOMY PROCEDURE WAS COMPLETED, THE PATIENT'S CONDITION REMAINED THE SAME AS IT WAS PRIOR TO THE OPERATION. THE PATIENT DID NOT PASS AWAY DUE TO SHOCK, BUT DUE TO INSUFFICIENT RESTORATION OF BLOOD FLOW IN THE DISTAL MIDDLE CEREBRAL ARTERY. THE ATTENDING PHYSICIAN HAS CLINICALLY CONFIRMED THAT THERE WAS NO BLEEDING OR VESSEL RUPTURE RESULTING FROM THE THROMBECTOMY PROCEDURE. WHAT WAS MEANT BY "SHOCK" WAS ACTUALLY "DETERIORATION". | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
0.14 GUIDEWIRES | K113454 | MEDTRONIC RECEIVED INFORMATION REGARDING A SOLITAIRE X STENT RETRIEVER (SFR4-6-40-10, LOT: B791192) THAT BROKE DURING A PROCEDURE. IT WAS REPORTED THAT PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT A NON-HEMORRHAGIC STROKE. THE SOLITAIRE AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE SOLITAIRE WAS DEPLOYED FOR CLOT RETRIEVAL. HOWEVER, DURING RETRIEVAL, THERE WAS RESISTANCE IN THE DISTAL END OF THE CATHETER (REBAR, 105-5081-153, LOT: B813957) AND THE SOLITAIRE BROKE AND REMAINED IN THE CLOT. THE SURGEON USED ANOTHER SOLITAIRE X (SFR4-4-40-10, LOT: B777662) TO RETRIEVE THE BROKEN SOLITAIRE ALONG WITH THE CLOT. IT APPEARS THE BROKEN DEVICE WAS SUCCESSFULLY RETRIEVED. HOWEVER, IT WAS NOTED THAT THE PATIENT ULTIMATELY HAD AN OUTCOME OF DEATH. | INFORMATION RECEIVED INDICATED PLI30 AND PLI40 ARE NO LONGER REPORTABLE. SINCE A REGULATORY REPORT HAS BEEN SENT GCH FUNCTIONALITY DOES NOT ALLOW THE COMPLAINT FLAG TO BE CHANGED TO NO. SHOULD FURTHER INFORMATION MAKE THE EVENT REPORTABLE AGAIN THE EVENT WILL BE REOPENED AND A REGULATORY REPORT WILL BE SENT. SEE SECTION H.6. FOR CORRECTED CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
SOLITAIRE FR4 | K183022 | MEDTRONIC RECEIVED INFORMATION REGARDING A SOLITAIRE X STENT RETRIEVER (SFR4-6-40-10, LOT: B791192) THAT BROKE DURING A PROCEDURE. IT WAS REPORTED THAT PATIENT WAS UNDERGOING A MECHANICAL THROMBECTOMY PROCEDURE TO TREAT A NON-HEMORRHAGIC STROKE. THE SOLITAIRE AND ALL ACCESSORY DEVICES WERE PREPARED AS INDICATED IN THE INSTRUCTIONS FOR USE (IFU). THE SOLITAIRE WAS DEPLOYED FOR CLOT RETRIEVAL. HOWEVER, DURING RETRIEVAL, THERE WAS RESISTANCE IN THE DISTAL END OF THE CATHETER (REBAR, 105-5081-153, LOT: B813957) AND THE SOLITAIRE BROKE AND REMAINED IN THE CLOT. THE SURGEON USED ANOTHER SOLITAIRE X (SFR4-4-40-10, LOT: B777662) TO RETRIEVE THE BROKEN SOLITAIRE ALONG WITH THE CLOT. IT APPEARS THE BROKEN DEVICE WAS SUCCESSFULLY RETRIEVED. HOWEVER, IT WAS NOTED THAT THE PATIENT ULTIMATELY HAD AN OUTCOME OF DEATH. | INFORMATION RECEIVED INDICATED PLI30 AND PLI40 ARE NO LONGER REPORTABLE. SINCE A REGULATORY REPORT HAS BEEN SENT GCH FUNCTIONALITY DOES NOT ALLOW THE COMPLAINT FLAG TO BE CHANGED TO NO. SHOULD FURTHER INFORMATION MAKE THE EVENT REPORTABLE AGAIN THE EVENT WILL BE REOPENED AND A REGULATORY REPORT WILL BE SENT. SEE SECTION H.6. FOR CORRECTED CODES. MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION AUTO CPAP DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION ST30 DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
PATIENT INFORMATION CENTER IX | K211900 | IT WAS REPORTED THE USER DID NOT HEAR OR RESPOND TO THE ALARMS AND THE PATIENT PASSED WAY; HOWEVER, THERE WAS NO ALLEGATION OF ALARM FAILURE. IT WAS INDICATED THE USER DID NOT HEAR THE ECG LEADS OFF ALARMS AND DID NOT RESPOND TO THE ECG LEADS OFF YELLOW TECHNICAL INOP ALARMS. THE CUSTOMER REQUESTED ASSISTANCE OBTAINING THE UNFILTERED AUDIT LOGS, WHICH REVEALED THERE WERE TWO HOURS OF ECG LEADS OFF ALARMS WITH NO ALARM ACKNOWLEDGEMENT AND THE ALARMS WERE PAGED APPROPRIATELY. RESUSCITATION WAS ATTEMPTED BUT UNSUCCESSFUL AND THE PATIENT PASSED AWAY. THE NURSE MANAGER INDICATED THE ALARM VOLUME FOR YELLOW ALARMS IS SET LOWER THAN RED ALARMS AND THE STAFF IS NOT RESPONDING TO THE YELLOW ALARMS. THE CUSTOMER WILL BE REQUESTING THE ECG LEADS OFF TECHNICAL INOP BE CHANGED TO A RED ALARM VIA THE INTERNAL ALARM COUNCIL. | THE CLINICAL AUDIT LOGS WERE PULLED FOR INVESTIGATION. PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
WATCHMAN TRUSTEER? ACCESS SYSTEM | K240018 | IT WAS REPORTED THAT PERICARDIAL EFFUSION (PE), CARDIAC TAMPONADE, AND DEATH OCCURRED. A LEFT ATRIAL APPENDAGE (LAA) CLOSURE PROCEDURE WAS BEING PERFORMED. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING NOTED THERE WAS NO BASELINE PE. GROIN ACCESS OBTAINED AND NON-BOSTON SCIENTIFIC VASCULAR PRE-CLOSURE DEVICE WAS USED. A VERSACROSS RADIOFREQUENCY (RF) WIRE WAS INSERTED UP TO THE SUPERIOR VENA CAVA (SVC), A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND VERSACROSS CONNECT DILATOR TO FOLLOW. NO TRANSSEPTAL PUNCTURE (TSP) WAS ACHIEVED ON THE FIRST DROP DOWN. ON THE SECOND DROPDOWN TSP WAS PERFORMED WITHOUT RF DELIVERY REQUIRED. TSP LOCATION APPEARED TO BE MID/MID-POSTERIOR BASED ON ICE IMAGING. NO PE WAS NOTED AT THIS POINT IN THE PROCEDURE AND PATIENT VITALS WERE STABLE. TSP LOCATION WAS DILATED WITH THE WAS AND THEN WAS AND VCC DILATOR RETRACTED FOR LEFT ATRIUM (LA) ACCESS OF THE ICE CATHETER. TSP ACCESS OF ICE CATHETER WAS STILL UNSUCCESSFUL. A PE WAS THEN DISCOVERED FOUR (4) MINUTES LATER. WAS REMOVED FROM THE GROIN AND ICE VISUALIZED THE PE LOCATION ON THE RIGHT SIDE OF THE HEART. THE PROCEDURE WAS ABORTED. CARDIAC TAMPONADE OCCURRED, WHICH CAUSED THE NEED FOR CARDIOPULMONARY RESUSCITATION (CPR), BLOOD TRANSFUSION, PERICARDIOCENTESIS REMOVING BRIGHT RED BLOOD OF OVER THREE (3) LITERS, FRESH FROZEN PLASMA (FFP), AND CRYOPRECIPITATE. THE FAMILY WAS CONTACTED A LITTLE OVER TWO (2) HOURS LATER AND AS THE PATIENT WAS A DO NOT RESUSCITATE (DNR) STATUS AND THEY REQUESTED NO ADDITIONAL LIFE-SAVING MEASURES, SO TIME OF DEATH WAS CALLED BY THE PHYSICIAN. THE OFFICIAL CAUSE OF DEATH IS PE THAT RESULTED IN DEATH. IT WAS NOTED THE PATIENT WAS ON ELIQUIS AND STOPPED TAKING ELIQUIS TWO (2) DAYS PRIOR TO THE PROCEDURE. LAST RECORDED INR LAB RESULT WAS 1.4.
DREAMSTATION 2 ADVANCED AUTO CPAP | K200480 | A RELATIVE OF THE PATIENT INQUIRED ABOUT THE OWNER OF THE DEVICE, A DREAMSTATION 2 ADVANCED AUTO CPAP. THE RELATIVE OF THE PATIENT REPORTED THE PATIENT PASSED AWAY. THE REPORTER STATED THE PATIENT WAS NOT ON THE DEVICE AT THE TIME OF DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
LIBERATOR | K800742 | AS REPORTED: AN ADDITIONAL OXYGEN TANK WAS DELIVERED ON (B)(6) 2025 TO A PATIENT'S HOME DUE TO INCREASED OXYGEN NEEDS. THE TANK WAS FIRST CONNECTED AND USED ON FRIDAY, ON (B)(6) 2025. IT FUNCTIONED NORMALLY: THE LED INDICATORS WERE FULLY GREEN, AND OXYGEN WAS BEING DELIVERED PROPERLY. ON SATURDAY, ON (B)(6) 2025, AROUND 11:00 P.M., A ROUTINE CHECK WAS PERFORMED. ALL CONNECTIONS WERE IN ORDER, THE LED INDICATORS REMAINED GREEN, AND OXYGEN APPEARED TO BE AVAILABLE IN SUFFICIENT QUANTITY. THE PATIENT WAS SLEEPING IN A HOSPITAL BED IN THE LIVING ROOM, WEARING A NASAL CANNULA. OTHER HOUSEHOLD MEMBERS WERE ASLEEP IN THEIR RESPECTIVE ROOMS. ON THE MORNING OF SUNDAY, ON (B)(6) 2025, AT APPROXIMATELY 6:40 A.M., THE PATIENT WAS FOUND DECEASED IN BED. THE NASAL CANNULA WAS STILL IN PLACE AND CORRECTLY CONNECTED TO THE TANK, BUT NO OXYGEN WAS FLOWING, DESPITE THE LED INDICATORS STILL SHOWING GREEN. ANOTHER TANK WAS IMMEDIATELY CONNECTED, AND OXYGEN FLOWED NORMALLY. CPR WAS ATTEMPTED BUT WAS UNSUCCESSFUL. THE PATIENT'S MOBILE PHONE, USUALLY KEPT WITHIN REACH TO CALL FOR HELP, WAS FOUND ON THE FLOOR, SUGGESTING AN ATTEMPT TO CALL DURING THE NIGHT. ON 29/06/2025, A CALL WAS MADE TO THE SUPPLIER. AS IT WAS A SUNDAY, A TECHNICIAN ANSWERED AND GAVE A CONFUSED EXPLANATION, MENTIONING THAT AN ALARM SHOULD HAVE SOUNDED IN CASE OF OXYGEN FAILURE. HOWEVER, THE TECHNICIAN WAS REFERRING TO A DIFFERENT DEVICE. THE CALLER WAS ADVISED TO CONTACT THE MAIN OFFICE THE NEXT MORNING. ON 30/06/2025, THE MAIN OFFICE WAS CONTACTED. IT WAS CONFIRMED THAT THE MODEL IN QUESTION, THE LIBERATOR, DOES NOT HAVE A BUILT-IN ALARM SYSTEM, WHICH THE FAMILY WAS ALREADY AWARE OF AFTER SEVERAL YEARS OF USE. THAT SAME DAY, THE PATIENT'S GENERAL PRACTITIONER INITIATED A JUDICIAL INVESTIGATION. THE PUBLIC PROSECUTOR'S OFFICE WAS NOTIFIED. THE OXYGEN TANKS WERE SEALED ON-SITE BY THE POLICE AND PLACED IN A SEPARATE ROOM IN THE HOME. THE OFFICERS ATTEMPTED TO WEIGH THE TANKS USING A HOUSEHOLD SCALE, WHICH WAS NOT SUITABLE FOR SUCH HEAVY EQUIPMENT. ONE OF THE OFFICERS EXPRESSED DOUBT ABOUT THE RELIABILITY OF THIS METHOD, BUT THE WEIGHING WAS CARRIED OUT FOR REPORTING PURPOSES. THE TANKS REMAIN SEALED AT THE RESIDENCE, WITH A POLICE NOTICE ATTACHED INDICATING A "SUSPICIOUS DEATH." AN AUTOPSY WAS PERFORMED ON (B)(6) 2025 AT (B)(6) HOSPITAL. THE CASE IS REGISTERED UNDER A PV NUMBER WITH THE PUBLIC PROSECUTOR'S OFFICE, UNDER THE SUPERVISION OF A DEPUTY PUBLIC PROSECUTOR. CONTACT DETAILS ARE AVAILABLE UPON REQUEST FOR OFFICIAL FOLLOW-UP. THE INCIDENT WAS ALSO REPORTED TO THE FEDERAL PUBLIC SERVICE (FPS) HEALTH VIA EMAIL ON 08/07/2025. A RESPONSE WAS RECEIVED FROM THE DEPUTY PROSECUTOR, BUT IT IS RECOMMENDED TO CONTACT HER DIRECTLY FOR FURTHER INFORMATION. IT IS UNDERSTOOD THAT THE PUBLIC PROSECUTOR'S OFFICE WILL NOT CONDUCT A TECHNICAL INVESTIGATION. THEREFORE, AN INDEPENDENT TECHNICAL ANALYSIS IS NECESSARY TO DETERMINE WHETHER A MALFUNCTION OCCURRED THAT CAUSED AN INTERRUPTION IN OXYGEN DELIVERY, DESPITE THE LED INDICATORS SHOWING NORMAL STATUS. SUCH AN INVESTIGATION IS CRUCIAL TO PREVENT SIMILAR INCIDENTS IN THE FUTURE. THE SEALED TANKS REMAIN IN THE HOME, WHICH IS A DAILY EMOTIONAL BURDEN FOR THE FAMILY. A THOROUGH TECHNICAL EXAMINATION OF THE TANK IS STRONGLY REQUESTED. | PURSUANT TO TITLE 21 - FOOD AND DRUGS, CHAPTER I - FOOD AND DRUG ADMINISTRATION DEPARTMENT OF HEALTH AND HUMAN SERVICES, SUBCHAPTER H -0 MEDICAL DEVICE, PART 803 - MEDICAL DEVICE REPORTING, SUBPART A - GENERAL PROVISIONS, SECTION 803.16, NEITHER THIS REPORT NOR ANY INFORMATION SUBMITTED HEREIN CONSTITUTES AN ADMISSION BY CAIRE INC. THAT THE DEVICE STATED IN THIS REPORT, CAIRE INC., OR CAIRE INC.'S EMPLOYEES, CAUSED OR CONTRIBUTED TO THE REPORTABLE EVENT STATED HEREIN. CAIRE CUSTOMER SERVICE CONTACTED THE DISTRIBUTOR, OXYSPHAIR, TO GATHER ADDITIONAL INFORMATION REGARDING ALL CAIRE DEVICES IN THE PATIENT'S HOME. OXYSPHAIR AND CAIRE WILL REQUEST THAT THE DEVICE BE RELEASED FROM POLICE CUSTODY AND SENT TO CAIRE'S LANGENFELD, GERMANY FACILITY FOR EVALUATION. CAIRE WILL CONDUCT AN EVALUATION OF THE DEVICE, IF POSSIBLE. A FOLLOW UP AND FINAL REPORT WILL BE SUBMITTED. NOTE: THE INTENDED USE OF LIBERATOR LIQUID OXYGEN BASE IS SUPPLEMENTAL OXYGEN TO INCREASE BLOOD OXYGEN SATURATION. IT IS NOT INDICATED FOR LIFE SUSTAINING OR LIFE SUPPORT ACTIVITIES. LIBERATOR DOES NOT HAVE PATIENT SURVEILLANCE CAPABILITIES. PER THE INITIAL REPORT, "THE PATIENT'S MOBILE PHONE, USUALLY KEPT WITHIN REACH TO CALL FOR HELP, WAS FOUND ON THE FLOOR, SUGGESTING AN ATTEMPT TO CALL DURING THE NIGHT." A MOBILE PHONE IS NOT AN ADEQUATE PATIENT SURVEILLANCE SYSTEM. SECTION D4 PRIMARY UNIQUE DEVICE IDENTIFIER (UDIS) # IS INTENTIONALLY BLANK PER RULE 21 CFR 801.30(8). THE DEVICE WAS INTENDED FOR EXPORT ONLY.
RENTAL-DREAMSTATION AUTO CPAP, GB | K131982 | THE MANUFACTURER RECEIVED INFORMATION THAT A RENTAL DREAMSTATION AUTO CPAP DEVICE IS RETURNING BECAUSE THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE EXAMINATION HAS NOT BEGUN BECAUSE THE DEVICE HAS NOT YET BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION AUTO CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING DEATH. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
EPOC BLOOD ANALYSIS SYSTEM | K200107 | THE CUSTOMER REPORTED THAT THEY RECEIVED A DISCREPANT HIGH HCO3 RESULT OF 12.8 ON A SINCE DECEASED PATIENT THAT DID NOT MATCH CLINICAL PICTURE OF THE PATIENT. THE CUSTOMER DOES NOT KNOW IF THE PATIENT WAS TREATED ON THE DISCREPANT RESULT. REPEATED TESTING OF A NEW SAMPLE GAVE A RESULT OF -17.3. THE CUSTOMER HAS NOT PROVIDED WHAT THEIR EXPECTED RESULT WAS FOR THIS PATIENT. | SIEMENS HAS REQUESTED MORE INFORMATION AND INSTRUMENT FILES FOR FURTHER INVESTIGATION. INVESTIGATION WILL COMMENCE ONCE INFORMATION HAS BEEN RECEIVED. THE CUSTOMER HAS NOT ALLEGED THAT THE DISCREPANT RESULT CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. HOWEVER, SIEMENS IS FILING THIS REPORT IN AN ABUNDANCE OF CAUTION DUE TO THE PASSING OF THE PATIENT. THE CAUSE OF THIS EVENT IS UNKNOWN.
EPOC BLOOD ANALYSIS SYSTEM | K200107 | THE CUSTOMER REPORTED THAT THEY RECEIVED A DISCREPANT LOW HCO3 RESULT OF -17.3 ON A SINCE DECEASED PATIENT THAT DID NOT MATCH CLINICAL PICTURE OF THE PATIENT. THE CUSTOMER DOES NOT KNOW IF THE PATIENT WAS TREATED ON THE DISCREPANT RESULT. REPEATED TESTING OF A NEW SAMPLE GAVE A RESULT OF 12.8. THE CUSTOMER HAS NOT PROVIDED WHAT THEIR EXPECTED RESULT WAS FOR THIS PATIENT. NOTE THIS IS THE SAME PATIENT AS MDR 3002637618-2025-00055. THIS IS FOR THE OTHER INSTRUMENT. | SIEMENS HAS REQUESTED MORE INFORMATION AND INSTRUMENT FILES FOR FURTHER INVESTIGATION. INVESTIGATION WILL COMMENCE ONCE INFORMATION HAS BEEN RECEIVED. THE CUSTOMER HAS NOT ALLEGED THAT THE DISCREPANT RESULT CAUSED OR CONTRIBUTED TO THE PATIENT DEATH. HOWEVER, SIEMENS IS FILING THIS REPORT IN AN ABUNDANCE OF CAUTION DUE TO THE PASSING OF THE PATIENT. THE CAUSE OF THIS EVENT IS UNKNOWN. NOTE THIS IS THE SAME PATIENT AS MDR 3002637618-2025-00055.
BIPAP PRO BIFLEX, W/SD CARD, US-RENTAL | K091319 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING THE PATIENT PASSED AWAY 4 YEARS AGO. MEDICAL INTERVENTION WAS NOT SPECIFIED. THERE IS NO REPORT OF ANY DEVICE MALFUNCTION RELATED TO THE PATIENT¿S DEATH. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
SUREFORM | K173721 | IT WAS REPORTED THAT DURING A DA VINCI-ASSISTED TRANSTHORACIC ESOPHAGECTOMY WITH CHEST ANASTOMOSIS PROCEDURE, THE SUREFORM 60 STAPLER INSTRUMENT WAS IN THE PROCESS OF FIRING WHEN COINCIDENTALLY, THE PATIENT REQUIRED IMMEDIATE RESUSCITATION DUE TO SYMPTOMS OF A HEART ATTACK. AT THAT TIME, THE STAPLING CYCLE HAD ALREADY BEEN INITIATED AND COULD NOT BE INTERRUPTED. ALTHOUGH THE ROBOTIC SYSTEM WAS UNDOCKED, BOTH THE STAPLER INSTRUMENT AND THE ENDOSCOPE REMAINED INSIDE THE PATIENT THROUGHOUT RESUSCITATION EFFORTS. AN INTUITIVE TECHNICAL SUPPORT ENGINEER (TSE) WAS CONTACTED DURING THE EVENT AND ADVISED THE SURGICAL TEAM TO OBTAIN VISUAL ACCESS TO THE STAPLER INSTRUMENT INSIDE THE ABDOMEN TO ASSESS WHETHER THE EMERGENCY RELEASE FUNCTION COULD BE SAFELY ACTIVATED. MULTIPLE UNSUCCESSFUL ATTEMPTS WERE MADE TO UNCLAMP AND RELEASE THE STAPLER INSTRUMENT FROM THE SMALL INTESTINE. THE TSE ADVISED THE CUSTOMER TO CONTINUE MOVING THE RELEASE KNOB TO ALLOW THE INSTRUMENT TO BE UNCLAMPED FROM THE TISSUE. RELEASE OF THE STAPLER INSTRUMENT WAS THEN SUCCESSFUL. IT WAS REPORTED THAT THE PATIENT ULTIMATELY EXPIRED. | A REVIEW OF THE SYSTEM LOGS FOUND THAT A FORCE FIRE FAILURE OCCURRED, RESULTING IN THE SUREFORM 60 STAPLER INSTRUMENT BEING FORCE EXPIRED. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE REPORTED EVENT. A DEVICE HISTORY RECORD (DHR) REVIEW FOUND NO NON-CONFORMANCES DOCUMENTED THAT WOULD BE RELATED TO THIS EVENT. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, TIP-UP FENESTRATED GRASPER, FENESTRATED BIPOLAR FORCEPS, VESSEL SEALER EXTEND. A SITE HISTORY REVIEW FOUND NO COMPLAINTS WERE MADE FOR THESE INSTRUMENTS USED DURING THIS PROCEDURE. THE SUREFORM 60 STAPLER INSTRUMENT WAS RECEIVED FOR EVALUATION; THE SUREFORM 60 RELOAD WAS NOT RETURNED. THE INSTRUMENT WAS FOUND TO HAVE A BLOCKED RADIO FREQUENCY IDENTIFIER (RFID) CHIP DUE TO THE USE OF THE INSTRUMENT¿S MANUAL RELEASE KNOB. AFTER RESETTING THE RFID CHIP, THE INSTRUMENT WAS PLACED AND RECOGNIZED BY THE IN-HOUSE ROBOT, THE INSTRUMENT WAS DRIVEN AND PASSED THE INITIALIZATION, RECOGNITION, AND ENGAGEMENT TESTS MULTIPLE TIMES. THE INSTRUMENT MOVED INTUITIVELY WITH FULL RANGE OF MOTION IN ALL DIRECTIONS, AND OPERATED AS EXPECTED: THE JAWS OPENED AND CLOSED PROPERLY, AND THE INSTRUMENT CLAMPED, FIRED A NEW RELOAD, AND UNCLAMPED WITHOUT ISSUE. A VISUAL INSPECTION REVEALED NO SIGNS OF DAMAGE. THE INSTRUMENT WAS FULLY FUNCTIONAL, AND NO PRODUCT ISSUE WAS IDENTIFIED. REVIEW OF THE SUREFORM 60 STAPLER LOG SHOWS THAT THE INSTRUMENT WAS INSTALLED 8 TIMES AND FIRED 8 RELOADS. EACH INSTALL CLAMPED AND FIRED SUCCESSFULLY, WITH THE EXCEPTION OF THE LAST INSTALL, WHERE FIRING FAILED, STOPPING AT APPROXIMATELY 94% COMPLETION AFTER APPROXIMATELY 51 SECONDS. THE LOGS SHOW AN ERROR REPRESENTING THE FAILURE, AND NO UNCLAMPING WAS ATTEMPTED. APPROXIMATELY 10 SECONDS LATER, THE GRIP RELEASE TOOL WAS DETECTED, THOUGH THE E-STOP HAD NOT YET BEEN PRESSED, AND THE STAPLER WAS THEN FORCE EXPIRED BY THE SYSTEM AS EXPECTED. APPROXIMATELY 18 MINUTES LATER, THE INSTRUMENT WAS RE-INSTALLED, AND LESS THAN A MINUTE LATER, THE E-STOP BUTTON WAS PRESSED ON THE SURGEON SIDE CONSOLE. THERE WERE NO FURTHER ERRORS OR EVENTS RELATED TO THE STAPLER FOUND IN THE LOGS. PER THE CUSTOMER AND TSE REPORT, THE INSTRUMENT WAS THEN SUCCESSFULLY RELEASED WITH THE GRIP RELEASE TOOL. THE DA VINCI SYSTEM AND INSTRUMENT MANUALS PROVIDE THE FOLLOWING WARNINGS: **WARNING: THE SYSTEM NEEDS TO BE IN A FAULTED STATE IN ORDER TO PERFORM MANUAL STAPLER RELEASE TO PREVENT PATIENT INJURY. **WARNING: DO NOT PERFORM GRIP RELEASE ON A NON-FAULTED SYSTEM WITHOUT FIRST PRESSING THE EMERGENCY STOP BUTTON. FAILURE TO OBSERVE THIS WARNING MAY RESULT IN UNINTENDED INSTRUMENT MOTION OR DAMAGE TO THE GRIP RELEASE MECHANISM. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN INTRAOPERATIVE EVENT. THE STAPLER WAS IN THE PROCESS OF BEING USED ON THE ESOPHAGUS WHEN THE PATIENT REQUIRED EMERGENT RESUSCITATION. THE DETAILS OF WHAT PRECIPITATED THAT EVENT ARE NOT PROVIDED. IT IS UNKNOWN IF THE USE OF THE STAPLER AT THAT MOMENT CONTRIBUTED TO THE EVENT. THE SURGEON COULD NOT RELEASE THE STAPLER AT THAT TIME. HOW THAT MAY HAVE AFFECTED EITHER THE SITUATION OR RESUSCITATIVE EFFORTS IS NOT PROVIDED. THE PATIENT ULTIMATELY EXPIRED. HOW THEY DIED AND THE REMAINDER OF THE EVENTS THAT LED TO THE DEATH ARE NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCTS OR INSTRUMENTS CONTRIBUTED TO THIS EVENT. SECTION B2: DATE OF DEATH IS ESTIMATED TO BE THE SAME DATE AS THE PROCEDURE, BUT WAS NOT CONFIRMED. .
DAVINCI XI | K131861 | IT WAS REPORTED THAT AFTER AN UNEVENTFUL DA VINCI-ASSISTED INGUINAL HERNIA REPAIR, THE PATIENT WAS DISCHARGED HOME. AN INTUITIVE CLINICAL TERRITORY ASSOCIATE WAS LATER INFORMED BY A HOSPITAL NURSE THAT THE PATIENT HAD PRESENTED TO THE EMERGENCY DEPARTMENT EIGHT HOURS LATER WITH UNSPECIFIED SYMPTOMS AND EXPIRED ON POST-OPERATIVE DAY 1. NO INFORMATION REGARDING THE CAUSE OF DEATH OR TREATMENT RENDERED WAS PROVIDED. THE ROBOTIC PROCEDURE WAS REPORTEDLY COMPLETED WITHOUT ANY INTRAOPERATIVE COMPLICATIONS OR DA VINCI SYSTEM ERRORS. ADDITIONAL INFORMATION WAS REQUESTED; HOWEVER, NO RESPONSE HAS BEEN RECEIVED. | REVIEW OF THE SYSTEM LOGS FOUND NO RELATED ERRORS WERE LOGGED DURING THE PROCEDURE. LOG REVIEW OF THE 5 SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM ALSO FOUND NO RELATED ERRORS OCCURRED. THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES FOUND. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THE PROCEDURE: ENDOSCOPE PLUS 30 DEGREE, MONOPOLAR CURVED SCISSORS, FORCE BIPOLAR, FENESTRATED BIPOLAR FORCEPS, MEGA SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO COMPLAINTS HAVE BEEN REPORTED FOR ANY OF THESE PRODUCTS. DEVICE HISTORY RECORD (DHR) REVIEW FOR THE INSTRUMENTS CONFIRMED THAT THERE WERE NO RELATED NON-CONFORMANCES DOCUMENTED. AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) REVIEW OF THE EVENT WAS PERFORMED AND CONCLUDED THAT THE PATIENT IN THIS REPORT DIED FOLLOWING AN INGUINAL HERNIA REPAIR. ALTHOUGH NO INTRAOPERATIVE ISSUES OR COMPLICATIONS WERE REPORTED, THE REASON FOR THE RETURN TO THE EMERGENCY ROOM JUST 8 HOURS AFTER SURGERY DISCHARGE IS NOT PROVIDED. THE CAUSE OF DEATH IS ALSO NOT PROVIDED. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS AVAILABLE TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT. SECTION A2: PATIENT AGE IS ESTIMATED; REPORTED TO BE "EARLY TO MID 70'S."
PORTEX BLUSELECT SUCTIONAID CUFFED TRACHEOSTOMY | K173384 | IT WAS REPORTED THAT THE PATIENT DIED DUE TO PROBLEMS WITH THE DEVICE. THE BALLOON WAS DEFLATING, AND THE PATIENT WAS MICROASPIRATING, WHICH CONTRIBUTED TO THE DEATH. ACCORDING TO THE CAREGIVER, THE PATIENT ASPIRATED MULTIPLE TIMES, ULTIMATELY RESULTING IN DEATH BY BRONCHOASPIRATION CAUSED BY A CANNULA OBSTRUCTING THE AIRWAY. | NO PRODUCT WAS RETURNED. THE REPORTED ISSUE CANNOT BE CONFIRMED AS NO PRODUCT WAS RETURNED FOR INVESTIGATION. IF THE PRODUCT IS RETURNED, THIS COMPLAINT WILL BE REOPENED FOR FURTHER INVESTIGATION. A DEVICE HISTORY RECORD (DHR) REVIEW WAS CONDUCTED WHICH INDICATED ALL INSPECTIONS WERE COMPLETED AND NO ISSUES WERE NOTED DURING MANUFACTURE.
ERBE VIO 3 | K190823 | IT WAS REPORTED THAT A PATIENT INCIDENT OCCURRED WITH THE ELECTROSURGICAL UNIT (ESU/GENERATOR) AND AN ERBE TRISECT RAPIDE INSTRUMENT (PART NUMBER 21150-402, LOT NUMBER WO433553.) DURING A TOTAL PELVIC EXENTERATION (TPE) PROCEDURE (NOTE: THE ERBE TRISECT RAPIDE INSTRUMENT IS NOT CLEARED FOR DISTRIBUTION IN THE U.S. AND IS NOT DISTRIBUTED IN THE U.S.). NO INFORMATION WAS PROVIDED REGARDING ANY OTHER ACCESSORIES USED IN THE OPERATION. ALSO, THE ESU SETTINGS EMPLOYED WERE NOT CONVEYED TO ERBE. DURING THE POSTOPERATIVE PERIOD, A DELAYED PERFORATION OF THE AORTA WAS DISCOVERED WHICH RESULTED IN MASSIVE BLEEDING AND THE PATIENT DIED (NOTE: IT APPEARS THAT A THERMAL LESION ON THE AORTA OCCURRED DURING THE TPE PROCEDURE WHILE USING THE ERBE TRISECT RAPIDE INSTRUMENT.). | THE ESU WAS THOROUGHLY INSPECTED/TESTED (NOTE: THE ERBE TRISECT RAPIDE INSTRUMENT WAS DISPOSED OF AFTER THE PROCEDURE AND WAS NOT AVAILABLE FOR AN EVALUATION.). THE GENERATOR WAS FOUND TO BE FUNCTIONING AS INTENDED. THE EVALUATION INCLUDED AN ELECTRICAL SAFETY CHECK, A FUNCTIONAL CHECK OF EACH OF THE EQUIPMENT'S FEATURES, AND A POWER OUTPUT CHECK. THE UNIT WAS/IS WITHIN SPECIFICATIONS AND ALL FEATURES WERE/ARE WORKING PROPERLY. IN ADDITION, NO ANOMALIES WERE FOUND IN THE DEVICE HISTORY RECORD (DHR) OF THE ESU OR THE TRISECT RAPIDE INSTRUMENT.). IN CONCLUSION, NO ERBE EQUIPMENT PROBLEM WAS FOUND THAT WOULD HAVE CAUSED OR CONTRIBUTED TO THE INCIDENT. BASED UPON THE REPORTED INFORMATION, MOST LIKELY THE THERMAL INSULT TO THE AORTA WAS CAUSED BY INADVERTENT CONTACT BETWEEN THE BRANCHES OF THE TRISECT RAPIDE INSTRUMENT AND THE AORTA. THE USER MANUAL OF THE ESU WARNS THAT BURNS OR FIRES MAY OCCUR DUE TO ACCIDENTALLY ACTIVATING AN INSTRUMENT. ALSO, THE MANUAL WARNS THAT UNWANTED TISSUE TRAUMA CAN OCCUR IF TISSUE COMES INTO CONTRACT WITH A INSTRUMENT THAT IS STILL HOT. THE TRISECT RAPIDE INSTRUMENT NOTES ON USE WARNS THAT TISSUE IN THE JOINT AREA MAY BE THERMALLY DAMAGED WHEN THE INSTRUMENT IS ACTIVATED WITH THE BRANCHES WIDE OPEN AND THAT TISSUE CONTACT WITH THE JOINT AREA OF THE INSTRUMENT SHOULD BE AVOIDED IN THIS CASE. ERBE USA, INC. IS NOW CLOSING THE FILE ON THIS EVENT.
RENTAL, DREAMSTATION ST30, GB | K102465 | THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE RENTAL DREAMSTATION ST30 DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. NO MEDICAL INTERVENTION WAS REPORTED. THE MANUFACTURER IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT, AND THE COMPLAINT IS STILL UNDER INVESTIGATION. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
EVIS EXERA III DUODENOVIDEOSCOPE | K250701 | OLYMPUS MEDICAL SYSTEMS CORP. (OMSC) RECEIVED A LITERATURE TITLED ""DIRECT COMPARISON OF SIMULTANEOUS AND SEQUENTIAL ENDOSCOPIC METALLIC BILATERAL STENTING IN MALIGNANT HILAR BILIARY OBSTRUCTION." THE AIM OF THE STUDY IS TO EVALUATE A POSSIBLE SUPERIORITY OF SIMULTANEOUS ¿SIDE BY SIDE¿ (SBS) BILIARY DRAINAGE IN UNRESECTABLE MHBO. BACKGROUND: ENDOSCOPIC BILATERAL BILIARY DRAINAGE IS A FIRST LINE PALLIATIVE TREATMENT FOR UNRESECTABLE MALIGNANT HILAR BILIARY OBSTRUCTION (MHBO) BUT REMAINS TECHNICALLY CHALLENGING. THE EMERGENCE OF SELF-EXPANDABLE METALLIC STENTS CARRIED BY AN ULTRATHIN (6 FR OR SMALLER) DELIVERY SYSTEM NOW PERMITS SIMULTANEOUS BILATERAL STENT PLACEMENT. TO DATE, ONLY A FEW STUDIES HAVE COMPARED THIS NEW METHOD WITH CONVENTIONAL SEQUENTIAL BILATERAL STENTING. METHODS THIS STUDY IDENTIFIED 135 PATIENTS WHO BENEFITED FROM BILATERAL DRAINAGE USING UNCOVERED SELF-EXPANDABLE METALLIC STENTS BETWEEN 2010 AND 2023. AMONG THEM, 62 BENEFITED FROM SIMULTANEOUS SBS BILATERAL DRAINAGE BETWEEN 2017 AND 2023, AND 73 BENEFITED FROM SEQUENTIAL BILATERAL DRAINAGE [38 USING ¿STENT IN STENT¿ (SIS) TECHNIQUE AND 35 USING SBS TECHNIQUE BETWEEN 2010 AND 2017]. RESULTS: TECHNICAL SUCCESS WAS SIGNIFICANTLY INCREASED IN SIMULTANEOUS DRAINAGE COMPARED WITH SEQUENTIAL DRAINAGE (94% VS 75%, P = 0.008). HOWEVER, SIMULTANEOUS SBS DRAINAGE AND SEQUENTIAL SIS DRAINAGE HAD A SIMILAR TECHNICAL SUCCESS (94% VS 95%). WE OBSERVED NO DIFFERENCES REGARDING CLINICAL SUCCESS, PROCEDURE DURATION AND RECURRENT BILIARY OBSTRUCTION RATE. STENT PATENCY WAS SHORTER IN THE SIS GROUP COMPARED WITH THE SIMULTANEOUS GROUP (103 DAYS VS 144 DAYS). EARLY ADVERSE EVENTS WERE MORE FREQUENT IN THE SEQUENTIAL GROUP (31% VS 21%, P = 0.205), WITH NO DIFFERENCES REGARDING SIS OR SBS TECHNIQUE. TECHNICAL FAILURE WAS ASSOCIATED WITH A HIGHER RATE OF INFECTIOUS FATAL ADVERSE EVENTS (9.5% VS 1.7%, P = 0.02). REINTERVENTION AFTER RECURRENT BILIARY OBSTRUCTION SEEMS TO BE MORE SUCCESSFUL AFTER USING SBS RATHER THAN SIS TECHNIQUES (83% VS 75%, P = 0.53). CONCLUSION: SIMULTANEOUS SBS METALLIC STENT PLACEMENT USING AN ULTRA-THIN DELIVERY SYSTEM WAS TECHNICALLY EASIER AND AS EFFICIENT AS SEQUENTIAL BILATERAL STENTING IN UNRESECTABLE MHBO TO ACHIEVE BILATERAL DRAINAGE. THE SIS PROCEDURE REMAINS A GOOD OPTION IN UNRESECTABLE MHBO."" TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS FOUR PATIENTS (3%) DIED DUE TO INFECTIOUS COMPLICATIONS (CHOLANGITIS) LEADING TO MULTIVISCERAL FAILURE AND SEPTIC SHOCK. THESE DEATHS OCCURRED AFTER TECHNICAL FAILURE OF THE STENTING PROCEDURE, WHICH WAS SIGNIFICANTLY ASSOCIATED WITH A HIGHER RISK OF FATAL OUTCOMES. THE SEQUENTIAL GROUP HAD THREE FATAL COMPLICATIONS, WHILE THE SIMULTANEOUS GROUP HAD ONE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE. | THE DEVICE HAS NOT BEEN RETURNED TO OLYMPUS FOR EVALUATION. BASED ON THE RESULTS OF THE INVESTIGATION, THE RELATIONSHIP BETWEEN THE DEVICE AND THE ADVERSE EVENT CANNOT BE CONFIRMED. THERE IS NO EVIDENCE OF AN OLYMPUS DEVICE MALFUNCTION IN THE LITERATURE. THEREFORE, THE ROOT CAUSE CANNOT BE DETERMINED. THE COMPLAINT IS RELATED TO THE FOLLOWING PATIENT IDENTIFIERS: (B)(6). THE LITERATURE DESCRIBED "TJF-190" AND OLYMPUS SELECTED "TJF-Q190V" AS THE REPRESENTATIVE MODEL. HECC- MULTIVISCERAL FAILURE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
EVIS EXERA DUODENOVIDEOSCOPE | K024033 | OLYMPUS MEDICAL SYSTEMS CORP. (OMSC) RECEIVED A LITERATURE TITLED "DIRECT COMPARISON OF SIMULTANEOUS AND SEQUENTIAL ENDOSCOPIC METALLIC BILATERAL STENTING IN MALIGNANT HILAR BILIARY OBSTRUCTION." THE AIM OF THE STUDY IS TO EVALUATE A POSSIBLE SUPERIORITY OF SIMULTANEOUS ¿SIDE BY SIDE¿ (SBS) BILIARY DRAINAGE IN UNRESECTABLE MHBO. BACKGROUND: ENDOSCOPIC BILATERAL BILIARY DRAINAGE IS A FIRST LINE PALLIATIVE TREATMENT FOR UNRESECTABLE MALIGNANT HILAR BILIARY OBSTRUCTION (MHBO) BUT REMAINS TECHNICALLY CHALLENGING. THE EMERGENCE OF SELF-EXPANDABLE METALLIC STENTS CARRIED BY AN ULTRATHIN (6 FR OR SMALLER) DELIVERY SYSTEM NOW PERMITS SIMULTANEOUS BILATERAL STENT PLACEMENT. TO DATE, ONLY A FEW STUDIES HAVE COMPARED THIS NEW METHOD WITH CONVENTIONAL SEQUENTIAL BILATERAL STENTING. METHODS THIS STUDY IDENTIFIED 135 PATIENTS WHO BENEFITED FROM BILATERAL DRAINAGE USING UNCOVERED SELF-EXPANDABLE METALLIC STENTS BETWEEN 2010 AND 2023. AMONG THEM, 62 BENEFITED FROM SIMULTANEOUS SBS BILATERAL DRAINAGE BETWEEN 2017 AND 2023, AND 73 BENEFITED FROM SEQUENTIAL BILATERAL DRAINAGE [38 USING ¿STENT IN STENT¿ (SIS) TECHNIQUE AND 35 USING SBS TECHNIQUE BETWEEN 2010 AND 2017]. RESULTS: TECHNICAL SUCCESS WAS SIGNIFICANTLY INCREASED IN SIMULTANEOUS DRAINAGE COMPARED WITH SEQUENTIAL DRAINAGE (94% VS 75%, P = 0.008). HOWEVER, SIMULTANEOUS SBS DRAINAGE AND SEQUENTIAL SIS DRAINAGE HAD A SIMILAR TECHNICAL SUCCESS (94% VS 95%). WE OBSERVED NO DIFFERENCES REGARDING CLINICAL SUCCESS, PROCEDURE DURATION AND RECURRENT BILIARY OBSTRUCTION RATE. STENT PATENCY WAS SHORTER IN THE SIS GROUP COMPARED WITH THE SIMULTANEOUS GROUP (103 DAYS VS 144 DAYS). EARLY ADVERSE EVENTS WERE MORE FREQUENT IN THE SEQUENTIAL GROUP (31% VS 21%, P = 0.205), WITH NO DIFFERENCES REGARDING SIS OR SBS TECHNIQUE. TECHNICAL FAILURE WAS ASSOCIATED WITH A HIGHER RATE OF INFECTIOUS FATAL ADVERSE EVENTS (9.5% VS 1.7%, P = 0.02). REINTERVENTION AFTER RECURRENT BILIARY OBSTRUCTION SEEMS TO BE MORE SUCCESSFUL AFTER USING SBS RATHER THAN SIS TECHNIQUES (83% VS 75%, P = 0.53). CONCLUSION: SIMULTANEOUS SBS METALLIC STENT PLACEMENT USING AN ULTRA-THIN DELIVERY SYSTEM WAS TECHNICALLY EASIER AND AS EFFICIENT AS SEQUENTIAL BILATERAL STENTING IN UNRESECTABLE MHBO TO ACHIEVE BILATERAL DRAINAGE. THE SIS PROCEDURE REMAINS A GOOD OPTION IN UNRESECTABLE MHBO." TYPE OF ADVERSE EVENTS/NUMBER OF PATIENTS. FOUR PATIENTS (3%) DIED DUE TO INFECTIOUS COMPLICATIONS (CHOLANGITIS) LEADING TO MULTIVISCERAL FAILURE AND SEPTIC SHOCK. THESE DEATHS OCCURRED AFTER TECHNICAL FAILURE OF THE STENTING PROCEDURE, WHICH WAS SIGNIFICANTLY ASSOCIATED WITH A HIGHER RISK OF FATAL OUTCOMES. THE SEQUENTIAL GROUP HAD THREE FATAL COMPLICATIONS, WHILE THE SIMULTANEOUS GROUP HAD ONE. THERE WAS NO ALLEGATION OF DEVICE MALFUNCTION IDENTIFIED IN THE ARTICLE. | THE DEVICE HAS NOT BEEN RETURNED TO OLYMPUS FOR EVALUATION. BASED ON THE RESULTS OF THE INVESTIGATION, THE RELATIONSHIP BETWEEN THE DEVICE AND THE ADVERSE EVENT CANNOT BE CONFIRMED. THERE IS NO EVIDENCE OF AN OLYMPUS DEVICE MALFUNCTION IN THE LITERATURE. THEREFORE, THE ROOT CAUSE CANNOT BE DETERMINED. THE COMPLAINT IS RELATED TO THE FOLLOWING PATIENT IDENTIFIERS: (B)(6). THE LITERATURE DESCRIBED "TJF-160" AND SO OLYMPUS SELECTED ¿TJF-160VR " AS THE REPRESENTATIVE MODEL. HECC- MULTIVISCERAL FAILURE. SHOULD ADDITIONAL RELEVANT INFORMATION BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. OLYMPUS WILL CONTINUE TO MONITOR FIELD PERFORMANCE FOR THIS DEVICE.
AGILIA SP MC WIFI GB | K121613 | THE FOLLOWING HAS BEEN REPORTED BY CUSTOMER: EMAIL (NOT FORMAL) CARDIFF COMPLAINT - NNU LIPIDS - ATTENDED A CALL WITH [NURSE] FROM CARDIFF TODAY TO ASSIST WITH CONFIGURATION QUERY. SHE RAISED THAT THIS WAS REQUIRED DUE TO AN INCIDENT WITH A NEONATAL LIPID INFUSION, WHERE THE LIPID WAS GIVEN AT AN INCORRECT FLOW RATE. [NURSE] CONFIRMED THAT THIS WAS A CLINICAL ERROR, AS THE LIPID WAS INFUSED AT THE RATE OF THE SECOND INFUSION (AQUEOUS SOLUTION). THE TRUST HAVE CARRIED OUT THEIR REVIEW OF THE DEVICE, DATIX INFO AND PUMP DATA AND HAVE STATED THEY DO NOT REQUIRE SUPPORT WITH THIS AS IT WAS DEEMED A CLINICAL MISTAKE AND DUE TO IT BEING A FEW MONTHS BACK, HAS BEEN RESOLVED. UNFORTUNATELY, THE PATIENT DIED AS A RESULT OF THIS INCIDENT. REPORTING AS A CONSERVATIVE MEASURE. ADVERSE EFFECTS WERE REPORTED. MORE INFORMATION IS NEEDED TO COMPLETE THE INVESTIGATION. AS DEVICE SERIAL NUMBER WAS NOT PROVIDED, DEVICE HISTORY RECORD COULD NOT BE REVIEWED. DEVICE LOG WAS NOT PROVIDED; THEREFORE, NO REVIEW COULD BE PERFORMED. THE DEVICE WAS NOT RECEIVED IN OUR LABORATORY; HOWEVER, THE INVESTIGATION WAS CONDUCTED BASED ON THE DESCRIPTION PROVIDED. BASED ON THIS DESCRIPTION, THERE HAS BEEN AN INCIDENT WITH A NEONATAL LIPID INFUSION, WHERE THE LIPID WAS GIVEN AT AN INCORRECT FLOW RATE. [NURSE] CONFIRMED THAT THIS WAS A CLINICAL ERROR, AS THE LIPID WAS INFUSED AT THE RATE OF THE SECOND INFUSION (AQUEOUS SOLUTION). FOR THIS COMPLAINT, THE ROOT CAUSE HAS BEEN IDENTIFIED AS AN INCORRECT FLOW RATE SETTING BY THE USER, WHICH CONSTITUTES A USE ERROR. THIS COMPLAINT HAS BEEN INCLUDED IN A TREND ANALYSIS. THIS COMPLAINT IS CONSIDERED AS: USE ERROR. THE TREND IS: NORMAL.
DREAMSTATION | K131982 | IT WAS REPORTED THAT THE DREAMSTATION AUTO CPAP DEVICE WAS BEING RETURNED DUE TO THE USER PASSING AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. REPORTABLE SINCE DEVICE ISSUE/EVENT HAS OR MAY HAVE CAUSED OR CONTRIBUTED TO A DEATH.
DREAMSTATION AUTO | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE / RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP, AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION ALLEGING A DEATH HAS OCCURRED. WIFE CALLED IN STATING HER HUSBAND GOT LUNG CANCER FROM THE DEVICE & PASSED AWAY. MEDICAL INTERVENTION WAS NOT SPECIFIED. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
EVERFLO 230V, OPI, AUS, ULTRAFILL | K061261 | THE MANUFACTURER WAS CONTACTED ABOUT A DECEASED PATIENT. LAW ENFORCEMENT INFORMED THEM THAT THE DEVICE IS INVOLVED IN A HOMICIDE INVESTIGATION. THE DEVICE IS NOW IN LAW ENFORCEMENT CUSTODY. CURRENTLY, THERE IS NO EVIDENCE SUGGESTING THE DEVICE CAUSED OR CONTRIBUTED TO THE PATIENT'S DEATH. THE MANUFACTURER'S INVESTIGATION IS STILL IN PROGRESS, AND A FOLLOW-UP REPORT WILL BE PROVIDED ONCE IT IS COMPLETED.
COLLECTIVE CONCENTRATE- NATURALYTE | K192017 | ON 24/JUN/2025 FRESENIUS WAS MADE AWARE OF AN ARTICLE TITLED ¿UNFAMILIAR PRODUCT AND BARCODE SCANNING ERROR LEADS TO PATIENT DEATH¿ WHICH WAS PUBLISHED IN THE AMERICAN JOURNAL OF NURSING AUGUST 2024, VOLUME 124, NUMBER 8. THE ARTICLE REVIEWED A RECENT ALERT FROM THE INSTITUTE FOR SAFE MEDICATION PRACTICES (ISMP) WHICH HIGHLIGHTED A CASE IN WHICH AN INTENSIVE CARE NURSE (ICU) ACCIDENTALLY ADMINISTERED NATURALYTE® (A HEMODIALYSIS PRODUCT) INSTEAD OF THE PRESCRIBED BOWEL PREPARATION PRODUCT (SUPREP BOWEL PREP) TO A PATIENT. THE PATIENT LATER DIED. ADDITIONAL DETAILS WERE OBTAINED FROM A NEWS ARTICLE WRITTEN BY LEIGH SEARCY AND MORGAN EADS ON 25/JUL/2023 FOR WLEX18, LEXINGTON, KENTUCKY. THE ARTICLE PRESENTED A TIMELINE OF EVENTS WITH DATES AND PATIENT INFORMATION. AN 81-YEAR-OLD MALE PATIENT WAS TRANSFERRED TO CHI ST. JOSEPH¿S HEATH MAIN IN LEXINGTON, KY FOR HIGHER LEVEL CARE DUE TO GASTROINTESTINAL (GI) BLEED (DATE UNKNOWN). THIS PATIENT WAS ADMITTED TO THE ICU. ON 30/JUN/2022 A BOWEL PREP WAS ORDERED FOR THIS PATIENT. THE DIALYSIS TEAM HAD PREVIOUSLY LEFT BEHIND A LARGE PLASTIC CONTAINER OF NATURALYTE®, WHICH HAD BEEN PRESCRIBED TO A DIFFERENT PATIENT. THE CONTAINER HAD BEEN LEFT IN THE ICU PHARMACY DELIVERY AREA. THE CONTAINER REPORTEDLY COULD HAVE BEEN THERE FOR UP TO THREE DAYS PER DOCUMENTS CITED IN THE NEWS ARTICLE. THIS PATIENT¿S NURSE SAW THE NATURALYTE® CONTAINER AND ASSUMED THAT IT WAS THE SAME AS GOLYTELY. GOLYTELY IS AN ELECTROLYTE-BASED BOWEL PREP WITH SIMILAR PACKAGING AS NATURALYTE®. ADDITIONALLY, THEY CONTAIN SIMILAR ELECTROLYTES SUCH AS MAGNESIUM, POTASSIUM, AND SODIUM. FURTHERMORE, THIS NURSE ASSUMED THAT GOLYTELY WAS THE SAME GENERIC PRODUCT AS SUPREP BOWEL PREP. THE NURSE DID NOT READ THE LABELING CAREFULLY AS THE NATURALYTE® LABEL DOES STATE THAT IT IS INTENDED FOR HEMODIALYSIS. THE NURSE DID NOT CONFIRM THAT THE PRODUCTS WERE IDENTICAL WITH THE PHARMACY OR BY USING AN ONLINE INDEX. THE NURSE ATTEMPTED TO SCAN THE PHARMACY BARCODE ON THE NATURALYTE® CONTAINER (THINKING IT WAS THE GOLYTELY). THE BARCODE DID NOT SCAN. THE PHARMACY BARCODE DID NOT CONTAIN A PATIENT NAME OR IDENTIFYING FACTOR. THE NURSE CALLED THE PHARMACY AT 5:35PM AND STATED THAT THE CONTAINER WOULD NOT SCAN. THE PHARMACY DID NOT SEND THE CORRECT PRODUCT (SUPREP BOWEL PREP) WITH A LABEL FOR THAT SPECIFIC PATIENT NOR COME UP TO THE ICU TO INSPECT THE MEDICATION. INSTEAD, THE PHARMACY SENT A NEW BARCODE LABEL TO THE ICU THROUGH THE TUBE DELIVERY SYSTEM. THE NEW BARCODE LABEL ARRIVED APPROXIMATELY FIVE MINUTES AFTER THE CALL TO THE PHARMACY. THE NURSE PROCEEDED TO ADMINISTER APPROXIMATELY 8OZ OF THE NATURALYTE® TO THE PATIENT, ORALLY, BEFORE THE END OF HER SHIFT. THE PATIENT WAS UNABLE TO TOLERATE THE LIQUID AS THE TASTE WAS UNPLEASANT AND IT PRODUCED NAUSEA. DOCUMENTED NOTES ANNOTATED THAT THE PHYSICIAN WHO INITIALLY ORDERED THE COLONOSCOPY PREP SAID THE PATIENT MUST TAKE THE FULL AMOUNT OF THE CONTAINER. IT WAS ORDERED THAT THE PREP BE ADMINISTERED VIA FEEDING TUBE. AFTER THE INITIAL NURSE ENDED HER SHIFT, ANOTHER NURSE ATTENDED TO THE PATIENT. THAT NURSE DID NOT CAREFULLY READ THE PREPRINTED LABEL ON THE CONTAINER OF NATURALYTE® AND ADMINISTERED THE REMAINDER OF THE LIQUID THROUGH A FEEDING BAG. THE PATIENT BEGAN TO DETERIORATE. ANOTHER PHYSICIAN ASSESSED THE PATIENT AND MISTOOK THE NATURALYTE® FOR GOLYTELY. SIGNIFICANT ELECTROCARDIOGRAM CHANGES DEVELOPED. THE MEDICATION ERROR WAS CAUGHT AT ABOUT MIDNIGHT. THE PATIENT PASSED AWAY AT 7:35AM ON (B)(6) 2022. THE CORONER¿S REPORT DOCUMENTED THE CAUSE OF DEATH AS COMPLICATIONS OF INADVERTENT ADMINISTRATION OF NATURALYTE® IN THE SETTING OF GI HEMORRHAGE. THE INITIAL NURSE CARING FOR THE PATIENT WAS ASSIGNED TO ADDITIONAL PATIENTS IN THE ICU THAT NIGHT WHICH IS SAID TO HAVE BEEN A FACTOR IN THE MEDICATION ERROR. | CLINICAL INVESTIGATION: THERE IS A TEMPORAL RELATIONSHIP BETWEEN THE ERRONEOUS ADMINISTRATION OF NATURALYTE® TO THIS ICU PATIENT PRESCRIBED A COLONOSCOPY BOWEL PREP WITH SUPREP BOWEL PREP AND ADVERSE OUTCOME OF DEATH. THERE IS NO EVIDENCE OF NATURALYTE® PRODUCT DEFICIENCY OR MANUFACTURING DEFICIENCY AND WAS THE RESULT OF MEDICATION ERROR. NURSES DID NOT PERFORM RIGHTS OF MEDICATION ADMINISTRATION PRIOR TO ADMINISTERING NATURALYTE®. EACH NURSE IS RESPONSIBLE TO CHECK FOR RIGHT PATIENT, DRUG, ROUTE, TIME, AND DOSE PRIOR TO ADMINISTERING MEDICATION. THERE WAS NO LABEL ON THE CONTAINER WHICH IDENTIFIED NATURALYTE® AS BEING PRESCRIBED FOR THAT PATIENT; THERE WAS ONLY A PHARMACY BARCODE. NURSES DID NOT READ MANUFACTURER'S LABEL ON NATURALYTE® CONTAINER WHICH INDICATES WHAT PRODUCT IS IN CONTAINER: ¿NATURALYTE® LIQUID ACID CONCENTRATE FOR BICARBONATE HEMODIALYSIS.¿ INDICATIONS FOR USE STATES: ¿NATURALYTE LIQUID ACID CONCENTRATE IS INDICATED FOR USE IN PATIENTS UNDERGOING EXTRACORPOREAL BICARBONATE HEMODIALYSIS FOR ACUTE AND CHRONIC RENAL FAILURE. NATURALYTE LIQUID ACID CONCENTRATE IS INTENDED TO BE USED AS ONE COMPONENT IN THE PREPARATION OF DIALYSATE IN A THREE-STREAM PROPORTIONING HEMODIALYSIS MACHINE ACCORDING TO A PHYSICIAN¿S PRESCRIPTION.¿ THERE IS A WARNING ON LABEL: ¿WARNING FAILURE TO FOLLOW THESE INSTRUCTIONS FOR USE MAY RESULT IN PATIENT INJURY OR DEATH.¿ GOLYTELY PREP AND SUPREP BOWEL PREPS REQUIRE RECONSTITUTING POWDER OR LIQUID WITH ANOTHER LIQUID WHICH NURSES WOULD HAVE HAD TO DO THEMSELVES OR CHECK IT WAS COMPLETED. HAD NURSES STOPPED TO CHECK THE CORRECTLY LABELED MANUFACTURER LABEL ON NATURALYTE®, COMPLETED FIVE RIGHTS OF MEDICATION ADMINISTRATION AND NOT RELIED ON BARCODE LABEL FROM PHARMACY, THE MEDICATION ERROR WOULD NOT HAVE BEEN MADE. BASED ON AVAILABLE INFORMATION, IT CAN BE CONCLUDED THE PATIENT DEATH RESULTED FROM ADMINISTRATION ERROR BY MEDICAL STAFF AND NOT BY ANY NATURALYTE® PRODUCT OR MANUFACTURING DEFICIENCY. PLANT INVESTIGATION IN PROCESS.
AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEM | K150692 | BOSTON SCIENTIFIC CORPORATION BECAME AWARE OF AN EVENT INVOLVING AXIOS STENT AND ELECTROCAUTERY ENHANCED DELIVERY SYSTEMS AND WALLFLEX BILIARY RX FULLY COVERED STENT SYSTEMS THROUGH THE ARTICLE TITLED, "OPTIMIZING EUS-GUIDED CHOLEDOCHODUODENOSTOMY WITH LUMEN-APPOSING METAL STENTS FOR PRIMARY DRAINAGE OF MALIGNANT DISTAL BILIARY OBSTRUCTION (SCORPION-IIP): A PROSPECTIVE PILOT STUDY", BY JESKA A. FRITZSCHE ET AL. PER THE ARTICLE, BETWEEN AUGUST 2023 TO APRIL 2024, 27 PATIENTS SUFFERING FROM MALIGNANT DISTAL BILE DUCT OBSTRUCTION UNDERWENT EUS-GUIDED CHOLEDOCHODUODENOSTOMY (EUS-CDS) USING AXIOS AS LUMEN-APPOSING METAL STENTS (LAMS) WITH A PLACEMENT OF A WALLFLEX FULLY COVERED SELF-EXPANDING METAL STENTS THROUGH IT. ONE MORTALITY OCCURRED DURING THE TEST PERIOD OF THESE PROCEDURES POTENTIALLY RELATED TO INGROWTH IN THE WALLFLEX DEVICE AND CHOLECYSTITIS. THE PATIENT ALSO HAD CONCOMITANT RENAL FAILURE. | BLOCK B2: THE EXACT DATE OF THE PATIENTS' DEATH WAS NOT REPORTED. THE RETROSPECTIVE STUDY START DATE OF AUGUST 1, 2023, IS USED FOR THE ESTIMATED DATE OF DEATH. BLOCK B3: THE EXACT DATE OF EVENT WAS NOT REPORTED. THE RETROSPECTIVE STUDY START DATE OF AUGUST 1, 2023, IS USED FOR THE ESTIMATED DATE OF EVENT. BLOCK D4, H4: DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. BECAUSE THE PRODUCT IS UNKNOWN AT THIS TIME, WERE UNABLE TO PROVIDE THE COMPLETE UNIQUE IDENTIFIER (UDI) # AND OTHER PRODUCT SPECIFIC INFORMATION. IF ADDITIONAL DETAILS BECOME AVAILABLE, A SUPPLEMENTAL REPORT WILL BE SUBMITTED. BLOCK G2: LITERATURE SOURCE: JESKA A. FRITZSCHE; PAUL FOCKENS; MARC G. BESSELINK; OLIVER R. BUSCH; FREEK DAAMS; MATTHEUS C. B. WIELENGA; JOHANNA W. WILMINK; ROGIER P. VOERMANS; ROY L. J. VAN WANROOIJ. "OPTIMIZING EUS-GUIDED CHOLEDOCHODUODENOSTOMY WITH LUMEN-APPOSING METAL STENTS FOR PRIMARY DRAINAGE OF MALIGNANT DISTAL BILIARY OBSTRUCTION (SCORPION-IIP): A PROSPECTIVE PILOT STUDY." GASTROINTESTINAL ENDOSCOPY 2025; VOLUME 101, NO. 5: 1009-1016. BLOCK H6: IMDRF PATIENT CODE E2326 CAPTURES THE REPORTABLE EVENT OF CHOLECYSTITIS. IMDRF IMPACT CODE F02 CAPTURES THE PATIENT'S DEATH.
QUANTUM WORKSTATION 12.1" ELITE | K202557 | ADVERSE EVENT OCCURRED THAT RESULTED IN PATIENT DEATH WHILE PATIENT WAS CONNECTED TO THE MEDICAL DEVICE. USER REPORTED THAT BLOOD FLOW WAS UNABLE TO BE RE-ESTABLISHED FOLLOWING CPR AND PATIENT EXPERIENCED HEMORRHAGE BENEATH THE HEART. THE INVESTIGATION DETERMINED THAT THE DEVICE DID NOT MALFUNCTION DURING USE AND WAS OPERATING AS INTENDED. | THE INVESTIGATION CONFIRMED THE ADVERSE EVENT OCCURRED AS THE USER DESCRIBED; HOWEVER, THERE WAS NO EVIDENCE OF A PROBLEM WITH THE QUANTUM SYSTEM THAT WOULD HAVE RESULTED IN SUCH EVENT. IT WAS CONCLUDED THAT THE QUANTUM SYSTEM WORKED AS INTENDED AND THE ADVERSE EVENT OCCURRED DUE TO EXTERNAL FACTORS FROM THE QUANTUM SYSTEM.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | IT WAS REPORTED BY THE PATIENT'S SPOUSE THAT THIS PATIENT WITH END STAGE RENAL DISEASE (ESRD) ON CONTINUOUS CYCLIC PERITONEAL DIALYSIS [CC(PD)] UTILIZING A LIBERTY SELECT CYCLER FOR RENAL REPLACEMENT THERAPY (RRT) EXPIRED ON (B)(6) 2025 WHILE UNDERGOING CCPD THERAPY. FOLLOW-UP WITH THE PATIENT¿S PD REGISTERED NURSE (PDRN) CONFIRMED THE PATIENT EXPIRED AT HOME WHILE UNDERGOING CCPD THERAPY ON (B)(6) 2025. THE PATIENT WAS DISCOVERED BY HER SPOUSE WHEN HE AWOKE IN THE MORNING. THE PDRN REPORTED THE PATIENT PASSED AWAY PEACEFULLY WHILE SLEEPING WITH NO APPARENT SIGNS OF DISTRESS. THE PATIENT¿S TREATMENT DATA FROM (B)(6) 2025 INDICATED THE PATIENT COMPLETED TREATMENT, AND NO ALARMS WERE NOTED. A NON-URGENT CALL WAS PLACED TO EMERGENCY MEDICAL SERVICES WHO CONFIRMED THE PATIENT¿S PASSING IN THE HOME AND TRANSPORTED THE PATIENT TO THE MORGUE. THE PDRN STATED THE PRIMARY CAUSE OF DEATH WAS CARDIAC ARREST SECONDARY TO HER EXTENSIVE CARDIAC COMORBID CONDITIONS (E.G., ATRIAL FIBRILLATION, CORONARY ARTERY DISEASE, UNDERWENT MULTIPLE CARDIAC CATHETERIZATIONS WITH STENTING. THE PDRN STATED THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) DEFICIENCY OR MALFUNCTION. | CLINICAL REVIEW: A TEMPORAL RELATIONSHIP EXISTS BETWEEN CCPD THERAPY UTILIZING A LIBERTY SELECT CYCLER AND THE SERIOUS ADVERSE EVENTS OF CARDIAC ARREST AND DEATH, AS THE PATIENT WAS ACTIVELY UNDERGOING CCPD THERAPY WHEN HE EXPIRED. THE PDRN ATTRIBUTED CAUSALITY TO THE PATIENTS¿ SIGNIFICANT CARDIAC COMORBID CONDITIONS. THE ESRD POPULATION CONTINUES TO HAVE SIGNIFICANTLY HIGHER MORTALITY, AND FEWER EXPECTED YEARS OF LIFE WHEN COMPARED TO THE GENERAL POPULATION. CARDIOVASCULAR DISEASE (INCLUDING SUDDEN CARDIAC DEATH) ACCOUNTS FOR THE MOST DEATHS AMONG THE ESRD POPULATION. ADDITIONALLY, ADULTS WITH ESRD HAVE MORTALITY RATES UP TO 30-FOLD HIGHER THAN THE GENERAL POPULATION. PER THE PDRN, THE SERIOUS ADVERSE EVENTS WERE UNRELATED TO ANY FRESENIUS DEVICE(S) AND/OR PRODUCT(S) MALFUNCTION OR DEFICIENCY. BASED ON THE INFORMATION AVAILABLE, THE LIBERTY SELECT CYCLER CAN BE DISASSOCIATED FROM THE SERIOUS ADVERSE EVENTS. THERE IS NO ALLEGATION OR OBJECTIVE EVIDENCE INDICATING A FRESENIUS DEVICE(S) AND/OR PRODUCT(S) CAUSED OR CONTRIBUTED TO THE SERIOUS ADVERSE EVENTS. FURTHERMORE, THERE WAS NO REPORT A FRESENIUS PRODUCT(S) AND/OR DEVICE(S) FAILED TO MEET THE USERS¿ EXPECTATIONS AND/OR THE MANUFACTURERS¿ SPECIFICATIONS. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DS2ADV AUTO CPAP | K200480 | THE MANUFACTURER BECAME AWARE OF A DS2ADV AUTO CPAP STATING THAT THE PATIENT'S SIGNIFICANT OTHER ADVISED THAT THE PATIENT SADLY PASSED AWAY AND THAT THERE IS NO ALLEGATION THAT THE DEVICE HAS CONTRIBUTED TO THE DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. NO ADDITIONAL INFORMATION CAN BE REQUESTED AT THIS TIME. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE PERFORMED. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A FOLLOW UP REPORT WILL BE FILED.
HEART LUNG MACHINE | K991864 | THE EVENT OCCURRED IN JAPAN. IT WAS REPORTED THAT THE ERROR MESSAGE ¿REFERENC ERROR¿ OCCURRED ON THE ROTAFLOW CONSOLE AND THE PUMP STOPPED DUE TO THE "REFERENC ERROR" DURING PATIENT USE. THE CUSTOMER USED THE EMERGENCY DRIVE AND REPLACED THE AFFECTED ROTAFLOW CONSOLE WITH ANOTHER ROTAFLOW CONSOLE. THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER PASSED AWAY, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR ACCORDING TO THE CUSTOMER. DUE TO THE REPORTED DEATH OF THE PATIENT A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). | THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE JAPANESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW CONSOLE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW CONSOLE WITH CATALOG NUMBER 701051711.
MEGA 7.5FR. 40CC IAB | K120868 | IT WAS REPORTED DURING USE, THROUGH THE EMERGENCY SUPPORT LINE, THAT THE PATIENT HAD LOW BLOOD PRESSURE WITH HIGH AUGMENTATION. AN X-RAY IMAGE WAS REVIEWED WHICH SHOWED THE IAB MARKER AT THE 8TH INTERCOSTAL SPACE, INDICATING THE BALLOON WAS POSITIONED LOW. THE USER WAS ADVISED TO NOTIFY THE PROVIDER. APPROXIMATELY ONE HOUR LATER, THE USER CALLED BACK AND SHARED ANOTHER X-RAY SCREENSHOT, WHICH SHOWED THAT THE BALLOON POSITION REMAINED RELATIVELY UNCHANGED AND WAS NOTED TO BE ALMOST AT THE HUB. IT WAS RECOMMENDED TO RETURN THE PATIENT TO THE CATH LAB FOR FURTHER EVALUATION UNDER FLUOROSCOPY, AS THE RADIOPAQUE TIP SHOULD BE POSITIONED BETWEEN THE 2ND AND 3RD INTERCOSTAL SPACE. THE PATIENT EXPIRED. | A SUPPLEMENTAL REPORT WILL BE SUBMITTED UPON COMPLETION OF OUR INVESTIGATION.
M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, KIT, 8F | K873213 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT APPROXIMATELY TWO YEARS, SIX MONTHS AND TWENTY-FIVE DAYS LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED INFECTION, BACTEREMIA, SEPSIS AND ORGAN DYSFUNCTION. THE PORT WAS REMOVED AND WAS REPLACED WITH A NEW DEVICE. LATER THE PATIENT PASSED AWAY AND CAUSE OF DEATH WAS NOT PROVIDED. | DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS ON (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. H11: AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F ¿ THE INFORMATION PROVIDE BY BD REPRESENTS ALL THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT/REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
GORE® DRYSEAL FLEX INTRODUCER SHEATH | K160254 | THE FOLLOWING WAS REPORTED TO GORE: ON (B)(6) 2025, A PATIENT WAS TREATED FOR A PSEUDOANEURYSM IN THE ASCENDING AORTA USING A GORE® TAG® CONFORMABLE THORACIC STENT GRAFT WITH ACTIVE CONTROL SYSTEM (CTAG) AND A GORE® DRYSEAL FLEX INTRODUCER SHEATH (DSF). PRIOR TO THE CTAG BEING ADVANCED, A THROMBUS WAS DISLODGED FROM THE PROXIMAL THORACIC AORTA. THE THROMBUS TRAVELLED INTO THE SMA. AFTER DEPLOYMENT OF THE CTAG, THE PHYSICIAN ATTEMPTED TO PERFORM A THROMBECTOMY, AND ENDED UP DISSECTING THE SMA. THE PHYSICIAN CONVERTED TO OPEN REPAIR, AND REPAIRED THE SMA. AFTER CLOSING UP THE PATIENT AND TAKING THEM OUT OF THE OPERATING ROOM, THE PATIENT CODED. CHEST COMPRESSIONS WERE STARTED. THE PATIENT THEN DIED OF A LACERATED LIVER. IT IS UNCLEAR WHETHER THE THROMBUS WAS DISLODGED BY A GUIDEWIRE OR THE DSF. | H.6. INVESTIGATION FINDINGS CODE C20: THE DEVICE WAS DISCARDED AT THE TREATING FACILITY AND WAS THEREFORE NOT AVAILABLE FOR ANALYSIS. H.6. TYPE OF INVESTIGATION CODE B22: A REVIEW OF THE MANUFACTURING RECORDS FOR THE DEVICE COULD NOT BE CONDUCTED BECAUSE THE SERIAL/LOT NUMBER REMAINS UNKNOWN. H.6. INVESTIGATION CONCLUSIONS CODE D12: IT SHOULD BE NOTED THAT, PER THE GORE® DRYSEAL FLEX INTRODUCER SHEATH INSTRUCTIONS FOR USE, COMPLICATIONS ASSOCIATED WITH THE USE OF THE GORE® DRYSEAL FLEX INTRODUCER SHEATH THAT MAY OCCUR AND/OR REQUIRE INTERVENTION INCLUDE, BUT ARE NOT LIMITED TO, EMBOLIZATION. W. L. GORE & ASSOCIATES, INC. (GORE) IS SUBMITTING THIS REPORT TO COMPLY WITH 21 C.F.R. PART 803, THE MEDICAL DEVICE REPORTING REGULATION. THIS REPORT IS BASED UPON INFORMATION OBTAINED BY GORE, WHICH THE COMPANY MAY NOT HAVE BEEN ABLE TO FULLY INVESTIGATE OR VERIFY PRIOR TO THE DATE THE REPORT WAS REQUIRED BY THE FDA. BLANK FIELDS PRESENT ON THIS REPORT INCLUDE REQUIRED FIELDS AND FIELDS DETERMINED TO BE NOT APPLICABLE. BLANK REQUIRED FIELDS INDICATE THAT THE INFORMATION WAS NOT PROVIDED, WAS DEEMED UNAVAILABLE OR WAS NOT APPLICABLE. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, GORE, OR ITS ASSOCIATES THAT THE DEVICE, GORE OR ITS ASSOCIATES CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. IN PARTICULAR, THIS REPORT DOES NOT CONSTITUTE A LEGAL ADMISSION BY ANYONE THAT THE PRODUCT DESCRIBED IN THIS REPORT HAS ANY DEFECTS OR HAS MALFUNCTIONED, AS DEFINED FROM A LEGAL STANDPOINT. THESE WORDS ARE INCLUDED IN THE REPORT AND ARE FIXED ITEMS FOR SELECTION CREATED BY THE FDA, TO CATEGORIZE THE TYPE OF EVENT SOLELY FOR THE PURPOSE OF REPORTING PURSUANT TO PART 803. THIS STATEMENT SHOULD BE INCLUDED WITH ANY INFORMATION OR REPORT DISCLOSED TO THE PUBLIC UNDER THE FREEDOM OF INFORMATION ACT.
HEART LUNG MACHINE | K991864 | THE EVENT OCCURRED IN JAPAN. IT WAS REPORTED THAT THE ERROR MESSAGE ¿REFERENC ERROR¿ OCCURRED ON THE ROTAFLOW CONSOLE AND THE PUMP STOPPED DUE TO THE "REFERENC ERROR" DURING PATIENT USE. THE CUSTOMER USED THE EMERGENCY DRIVE AND REPLACED THE AFFECTED ROTAFLOW CONSOLE WITH ANOTHER ROTAFLOW CONSOLE. THE PUMP HAD BEEN USED ON A PATIENT WHO WAS IN A BAD CONDITION, AND THE PATIENT LATER PASSED AWAY, BUT THE PUMP HAD ONLY BEEN STOPPED FOR A SHORT TIME AND HIS BLOOD PRESSURE WAS STABLE, SO THERE IS NO CAUSAL RELATIONSHIP TO THIS ERROR ACCORDING TO THE CUSTOMER. DUE TO THE REPORTED DEATH OF THE PATIENT A REPORT IS REQUIRED. COMPLAINT ID: (B)(4). | THE INVESTIGATION IS ONGOING. FURTHER INFORMATION HAS BEEN REQUESTED BUT HAS NOT YET BEEN RECEIVED. A FOLLOW UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE JAPANESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "BASE UNIT, ROTAFLOW CONSOLE¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K991864. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR ROTAFLOW CONSOLE WITH CATALOG NUMBER 701051711.
POWERPORT M.R.I. IMPLANTABLE PORT, CHRONOFLEX SINGLE-LUMEN, 8F | K063377 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT APPROXIMATELY ONE YEAR AND TEN MONTHS POST PORT PLACEMENT PROCEDURE, THE PORT WAS REMOVED FROM THE PATIENT. LATER THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS 01-JAN-1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
VACCESS CT LOW-PROFILE POWER-INJECTABLE IMPLANTABLE PORT, CHRONOFLEX, 8F | K063377 | IT WAS REPORTED THROUGH LITIGATION PROCESS THAT APPROXIMATELY NINE MONTHS AND FOURTEEN DAYS LATER PORT PLACEMENT PROCEDURE, THE PATIENT DEVELOPED SEVERE SEPSIS. LATER THE PATIENT PASSED AWAY AND THE CAUSE OF DEATH WAS NOT PROVIDED. | H11: DATE OF DEATH FOR THIS PATIENT WAS NOT PROVIDED, DATE OF DEATH UPDATED AS (B)(6) 1900 FOR THE MDR SUBMISSION REQUIREMENT. AS THE LOT NUMBER FOR THE DEVICE WAS PROVIDED, A REVIEW OF THE DEVICE HISTORY RECORDS WILL BE PERFORMED. THE SAMPLE WAS NOT RETURNED TO THE MANUFACTURER FOR INSPECTION/EVALUATION. THEREFORE, THE INVESTIGATION OF THE REPORTED EVENT IS INCONCLUSIVE. BASED UPON THE AVAILABLE INFORMATION, THE DEFINITIVE ROOT CAUSE FOR THIS EVENT IS UNKNOWN. THE INSTRUCTIONS FOR USE (IFU) IS ADEQUATE FOR THE REPORTED DEVICE/PATIENT CODE(S) AND PROVIDES GENERAL INSTRUCTIONS FOR USE, AS WELL AS WARNINGS, PRECAUTIONS AND POTENTIAL COMPLICATIONS ASSOCIATED WITH THE DEVICE. UPON RECEIPT OF NEW OR ADDITIONAL INFORMATION, A FOLLOW-UP REPORT WILL BE SUBMITTED AS APPLICABLE. SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD.
ECHELON 3000 60MM STANDARD | K213633 | IT WAS REPORTED THAT DURING AN OPEN LEFT NEPHRECTOMY, THEY FIRED ACROSS THE RENAL AND SOMEHOW, THEY GOT INTO THE AORTA. THE PATIENT STARTS BLEEDING AND WAS STABLE OFF THE TABLE THEN WAS TRANSFERRED TO ICU BUT THE PATIENT PASSED AWAY. THERE'S NO ISSUE WITH THE DEVICE BUT THIS IS AN ADVERSE EVENT. | (B)(4). DATE SENT: 8/13/2025. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: I DO NOT KNOW THE INTERACTION BETWEEN THE AORTA AND THE STAPLER. LEFT RADICAL NEPHRECTOMY HE WILL USE THE 45 BUT IF HE CANT GET ACROSS THE RENAL VESSELS WITH THE 45 HE WILL GO TO THE 60 I DON'T KNOW IF IT WAS SKELETONIZED DON'T KNOW THE SIZE OF THE RENAL ARTERY HE SAID WHEN HE GOT THE KIDNEY REMOVED FROM THE BODY HE STARTED HAVING BLEEDING FROM THE AORTA. HE DOES NOT KNOW HOW IT REALLY HAPPENED. HE SAID HE FELT GOOD THAT HE GOT EVERYTHING FREED UP, AND ONCE THE KIDNEY WAS REMOVED BLOOD STARTED COMING FROM THE AORTA. DO NOT KNOW BLOOD LOSS IN ML THE PATIENT DID HAVE A TRANSFUSION THERE WERE CLAMPS NEAR THE AREA WERE STAPLER WAS FIRED HE HASN'T SAID YES OR NO IF HE THINKS THE STAPLER CONTRIBUTED TO THE BLEEDING THE DEVICE IS NOT AVAILABLE FOR ANALYSIS BECAUSE IT STAPLED AND CUT LIKE IT WAS SUPPOSED TO, TO FREE UP THE KIDNEY FOR REMOVAL. THEY DID NOT KEEP IT BECAUSE IT PERFORMED LIKE IT SUPPOSED TO. UNKNOWN ON THE AUTOPSY REPORT SURGEON HAS NOT REQUESTED TO SPEAK TO ETHICON MEDICAL OR ENGINEER.
ECHELON FLEX POWERED VASCULAR STAPLER WITH ADVANCED PLACEMENT TIP | K141952 | IT WAS REPORTED THAT POST-OP OF A LIVER RESECTION THAT THREE TO FOUR WEEKS AFTER THE PROCEDURE THE PATIENT PASSED. WHILE TRYING TO TAKE OUT THE PEDICLE BILE WAS COMING OUT, IN ORDER TO REINFORCE THE STAPLE LINE SUTURE WAS USED TO STOP THE BILE LEAKAGE. TWO DAYS AFTER THE CASE THEY HAD TO EMBOLIZE THE PATIENT. FIVE DAYS LATER THERE WAS SOME SORT OF VENOUS BLEEDING. | (B)(4). DATE SENT: 7/24/2025. B3: EXACT EVENT DATE UNK, ENTERED 6/1/2025 AS ONLY THE MONTH WAS PROVIDED. D4: BATCH # UNK. D4: UDI: AS THE LOT NUMBER FOR THE DEVICE INVOLVED IN THE EVENT WAS NOT PROVIDED, THE FULL UDI IS CURRENTLY NOT AVAILABLE. D4: UDI: THE EXPIRATION DATE IS CURRENTLY NOT AVAILABLE. THEREFORE, THE FULL UDI IS CURRENTLY NOT AVAILABLE. AN ANALYSIS OF THE PRODUCT COULD NOT BE PERFORMED SINCE A PHYSICAL SAMPLE WAS NOT RECEIVED FOR EVALUATION. AN EVALUATION OF THE MANUFACTURING RECORD COULD NOT BE PERFORMED AS THE REQUIRED PRODUCT IDENTIFICATION NUMBER WAS NOT PROVIDED TO COMPLETE THE EVALUATION. ADDITIONAL INFORMATION WAS REQUESTED AND THE FOLLOWING WAS OBTAINED: THE SURGEON HAD A PATIENT EXPIRE AFTER USING A PVS DEVICE. THE PATIENT PASSED ON (B)(6) 2025. THE INITIAL LIVER RESECTION PROCEDURE OCCURRED 3-4 WEEKS PRIOR. DURING THE INITIAL PROCEDURE THE SURGEON EXPERIENCED A BILE LEAK FROM THE STAPLE LINE AND REINFORCED WITH SUTURE. DAYS LATER THE SURGEON HAD TO EMBOLIZE THE PATIENT AND THE PATIENT HAD FURTHER COMPLICATIONS THAT LEAD TO THE EXPIRATION. HOW LONG WAS THE PROCEDURE? UNK. LIVER RESECTION TYPE? UNK. ANATOMICAL STRUCTURE ION OR DIVIDED? UNK. HAS THE REP BEEN IN LIVER CASES PRIOR? DOES THE SURGEON ISOLATE THE BILE DUCTS OR OFTEN INCORPORATE LIVER TISSUE? NO, THEY TRY NOT TO GET ANY LIVER TISSUE IN THE PVS DEVICE AND ISOLATE THE VESSEL. WERE ANY OTHER INSTRUMENTS USED IN THE DISSECTION? HARMONIC 1100 AND CLIPS. HAS AN AUTOPSY BEEN COMPLETED? UNK. IS THE SURGEON INTERESTED IN AN EXTERNAL RRT? UNK. THIS REPORT IS BEING SUBMITTED PURSUANT TO THE PROVISIONS OF 21 CFR, PART 803. THIS REPORT MAY BE BASED ON INFORMATION WHICH HAS NOT BEEN INVESTIGATED OR VERIFIED PRIOR TO THE REQUIRED REPORTING DATE. THIS REPORT DOES NOT REFLECT A CONCLUSION BY ETHICON, OR ITS EMPLOYEES THAT THE REPORT CONSTITUTES AN ADMISSION THAT THE PRODUCT, ETHICON, OR ITS EMPLOYEES CAUSED OR CONTRIBUTED TO THE POTENTIAL EVENT DESCRIBED IN THIS REPORT. IF INFORMATION IS OBTAINED THAT WAS NOT AVAILABLE FOR THE INITIAL REPORT, A FOLLOW-UP REPORT WILL BE FILED AS APPROPRIATE.
INFINITY ACUTE CARE SYSTEM (M540) | K113798 | A DIABETIC PATIENT WAS ADMITTED WITH HYPERGLYCEMIA AND TAKEN TO THE OPERATING ROOM FOR A RIGHT FOOT WOUND DEBRIDEMENT. THE PROCEDURE WAS COMPLETED, BUT UPON REVERSAL OF ANESTHESIA, THE PATIENT DID NOT RESPOND. RESUSCITATION EFFORTS WERE PERFORMED BUT WERE UNSUCCESSFUL, AND THE PATIENT PASSED AWAY. THE CUSTOMER REPORTED THAT THE HR (HEART RATE) WAVEFORM WAS VISIBLE ON THE M540 PATIENT MONITOR, BUT NO HR VALUES WERE DISPLAYED. ADDITIONALLY, HR VALUES WERE DASHED OUT ON THE PRINTOUTS. FOLLOW-UP INSPECTION (JULY 11, 2025): A SITE VISIT WAS CONDUCTED FOR INSPECTION AND ERROR ANALYSIS. DURING THE VISIT, THE MONITOR FILTER SETTING WAS CORRECTED FOR USE IN AN ELECTRICAL SURGICAL ENVIRONMENT FROM 'MONITOR' TO 'ESU'(ELECTRICAL SURGICAL UNIT), RESOLVING THE ISSUE WITH MISSING HR VALUES. HR VALUES APPEARED ON THE MONITOR AND WERE CORRECTLY PRINTED. THE PATIENT MONITOR WAS TESTED ON-SITE WITH THE CUSTOMER AND SUCCESSFULLY READ NIBP AND SPO2 DATA THROUGHOUT THE TEST. THE DATA WAS PRINTED CORRECTLY, CONFIRMING NORMAL DEVICE FUNCTIONALITY. | THIS INVESTIGATION IS ONGOING AS OF JULY 31ST, 2025. A FOLLOW UP REPORT WILL BE SUBMITTED UPON COMPLETION.
MICRA | K132030 | IT WAS FURTHER NOTED THAT WHILE BLOOD PRESSURE STABILIZED AFTER DRAINAGE, THE PATIENT LATER PASSED AWAY DUE TO ACUTE HEART FAILURE. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
TELESCOPE GUIDE EXTENSION CATHETER | K183353 | IT WAS REPORTED THAT DURING A PERCUTANEOUS CORONARY INTERVENTION (PCI) PROCEDURE INVOLVING ONE TELESCOPE GUIDE EXTENSION CATHETER (GEC) A DETACHMENT OF THE TIP OF THE DEVICE OCCURRED. THE GEC WAS INSPECTED PRIOR TO USE AND PREPPED PER IFU. THE LESION BEING TREATED HAD MODERATE TO SEVERE TORTUOSITY AND WAS SEVERELY CALCIFIED LOCATED IN THE MID CIRCUMFLEX (CX) ARTERY WITH 70% STENOSIS. THERE WAS POOR GUIDE SUPPORT WAS WITH ANGULATED CX. THE GEC WAS BEING USED TO FACILITATE NC BALLOON POSITIONING. IT WAS ATTEMPTED TO DELIVER THE STENT WITHOUT THE GEC HOWEVER THIS WAS UNSUCCESSFUL AND THE GEC WAS RE-DEPLOYED USING NC BALLOON ANCHOR TECHNIQUE TO ADVANCE. THE GEC WAS DELIVERED ACROSS THE STENOSIS WITHOUT DIFFICULTY FOLLOWING EXPANSION OF NC BALLOON AT NOMINAL PRESSURE IN THE HEAVILY CALCIFIED SEGMENT. MILD RESISTANCE WELL WITHIN NORMAL PRACTICE WAS EXPERIENCED ADVANCING THE GEC TO DELIVER THE GEC FOR ANCHOR SUPPORT WITH NO EXCESSIVE FORCE USED. THE STENT WAS THEN DELIVERED THROUGH THE GEC TO THE INTENDED DEPLOYMENT SITE WITHOUT RESISTANCE. IT WAS THEN ATTEMPTED TO WITHDRAW THE GEC OVER THE STENT IN ORDER TO DEPLOY THE STENT HOWEVER THE GEC WAS UNABLE TO BE WITHDRAWN. ATTEMPTS TO REMOVE THE STENT AND GEC EN BLOC FAILED. THERE WAS NO INTERACTION BETWEEN THE STENT AND THE GEC. A DETACHMENT OF THE SOFT POL YMER TIP OF THE GEC PROXIMAL TO THE RADIO-OPAQUE MARKER OCCURRED. THE STENT GUIDE AND WIRE WERE REMOVED EN BLOC. A NEW ACCESS WAS MADE AND THE GEC TIP WAS REWIRED WITH DIFFICULTY. MULTIPLE ATTEMPTS WERE MADE TO REMOVE THE GEC USING DISTAL NC BALLOON WHICH WAS INFLATED HOWEVER THIS WAS UNSUCCESSFUL. IT WAS UNABLE TO WIRE AROUND THE TIGHT STENOSIS PRESENT. THE PROCEDURE WAS ABANDONED. THE PATIENT WAS HAEMODYNAMICALLY STABLE, WITH NO ST ELEVATION, TIMI 3 FLOW IN VESSEL AT THE END OF THE PROCEDURE. THE PATIENT WAS TRANSFERRED TO THE SURGICAL THEATRE FOR CORONARY ARTERY BYPASS GRAFT (CABG) PROCEDURE TO REMOVE THE DETACHED TIP. CARDIAC SURGERY WAS WAS ALSO INDICATED FOR THE PATIENT DUE TO SEVERE LEFT ANTERIOR DESCENDING (LAD) ARTERY, CX AND RIGHT CORONARY ARTERY (RCA). THE PATIENT SUBSEQUENTLY DIED AS A RESULT OF POST-OPERATIVE COMPLICATIONS OF THE CABG SURGERY. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
BIPAP S/T C SERIES | K131982 | THE MANUFACTURER RECEIVED INFORMATION ALLEGING A PATIENT WITH A BIPAP S/T C SERIES CORE PACKAGE DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. | THE MANUFACTURER PREVIOUSLY RECEIVED INFORMATION ALLEGING A PATIENT WITH A BIPAP S/T C SERIES CORE PACKAGE DEVICE HAS PASSED AWAY. THERE WAS NO ALLEGATION THAT THE DEVICE CONTRIBUTED TO THE DEATH. ADDITIONAL INFORMATION HAS BEEN REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. PREVIOUSLY SUBMITTED REPORT WAS INCORRECTLY FILED WITH WRONG HEALTH IMPACT GRID. IN THIS REPORT, BOX H: HEALTH IMPACT GRID (1) HAS BEEN UPDATED AND CORRECTED.
MAGNETOM ESPREE | K041112 | SIEMENS HEALTHINEERS BECAME AWARE OF A FATAL INCIDENT INVOLVING THE MAGNETOM ESPREE SYSTEM. THE INCIDENT REPORTEDLY OCCURRED ON (B)(6) 2025, WHEN A 61-YEAR-OLD MALE ENTERED AN MRI ROOM WITH A METALLIC OBJECT AROUND HIS NECK. THE OBJECT WAS ATTRACTED TO THE MRI¿S MAGNETIC FIELD, RESULTING IN ALLEGED INJURY/MEDICAL EVENT(S) THAT RESULTED IN HIS DEATH. | D4: DUE TO THE AGE OF THE REPORTED DEVICE, A UDI IS NOT AVAILABLE. E1: A REPORTING FACILITY CONTACT NAME WAS NOT PROVIDED FOR REPORTING. H3, H6: THE INVESTIGATION IS ONGOING; HOWEVER, IT IS CURRENTLY UNKNOWN IF SIEMENS WILL BE ABLE TO EVALUATE THE REPORTED SYSTEM. THE SYSTEM WAS BOUGHT USED AND WAS INSTALLED AT THE FACILITY BY A THIRD-PARTY AND SIEMENS DOES NOT HAVE A SERVICE CONTRACT WITH THE FACILITY. SIEMENS HAS REACHED OUT TO THE FACILITY BY TELEPHONE FOR ADDITIONAL INFORMATION AND LEFT A MESSAGE WITH CONTACT INFORMATION WITH THE RECEPTIONIST. A RETURN CALL FROM THE FACILITY MANAGER IS PENDING RECEIPT. A SUPPLEMENTAL REPORT WILL BE SUBMITTED IF ADDITIONAL INFORMATION IS PROVIDED UPON COMPLETION OF THE INVESTIGATION.
TUBING SET | K112360 | THE EVENT OCCURRED IN CHINA DURING TREATMENT. IT WAS REPORTED THAT A PATIENT WAS CONNECTED TO THE ECMO DUE TO CARDIOGENIC DISEASE. A LEAKAGE ON THE QUICK VENT LUER CONNECTOR WAS REPORTED. THE CONNECTOR CAP WAS REPLACED BUT NOT SOLVED THE LEAKAGE. THE PATIENT PASSED WHILE NOT ON THE DEVICE. DUE TO THE LEAKAGE DURING USE AND THE HLS SET REPLACEMENT, A REPORT IS REQUIRED. COMPLAINT ID# (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
TUBING SET | K112360 | THE EVENT OCCURRED IN CHINA DURING TREATMENT. IT WAS REPORTED THAT A PATIENT WAS CONNECTED TO THE ECMO DUE TO CARDIOGENIC DISEASE. A LEAKAGE ON THE QUICK VENT LUER CONNECTOR WAS REPORTED. THE CONNECTOR CAP WAS REPLACED BUT NOT SOLVED THE LEAKAGE. THE PATIENT PASSED WHILE NOT ON THE DEVICE. DUE TO THE LEAKAGE DURING USE AND THE HLS SET REPLACEMENT, A REPORT IS REQUIRED. COMPLAINT ID#: (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. NOTE: THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069073.
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K150709 | IT WAS REPORTED THAT A PATIENT DEATH OCCURRED.DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. THE TRANSSEPTAL PUNCTURE WAS PERFORMED WITHOUT FLUOROSCOPY GUIDANCE. THE PUNCTURE WAS NOT DONE AT ITS MOST IDEAL POSITION. A 35MM WM FLX PRO DEVICE WAS OPENED AFTER CROSS VALIDATING APPENDAGE MEASUREMENTS FROM COMPUTED TOMOGRAPHY SCAN AND TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE). THE PHYSICIANS STILL TRIED TO SEE IF USING TRUSTEER SHEATH WOULD COMPENSATE FOR THE INADEQUATE TRANSSEPTAL POSITION. IT WAS TRIED MULTIPLE TIMES TO POSITION THE WM FLX PRO DEVICE CORRECTLY IN THE APPENDAGE, BUT THE POSITION OF THE DEVICE WAS NOT ACCEPTABLE. HENCE, IT WAS DECIDED TO FULLY RE-CAPTURE THE WM FLX PRO 35MM DEVICE WITHIN THE TRUSTEER SHEATH AND REMOVE BOTH. IN REMOVING THE WATCHMAN SHEATH AND DEVICE, A PERICARDIAL EFFUSION FLAGGED. THE PHYSICIANS RIGHT AWAY CALLED FOR STERILE TRAYS AND ALL THE EQUIPMENT NEEDED FOR DRAINAGE OF PERICARDIAL EFFUSION. AT THIS POINT, ABOUT TWO TO THREE 60CC SYRINGES OF BLOOD HAD BEEN EXTRACTED. THE PHYSICIAN THEN DECIDED TO GO FOR A SECOND TRANSSEPTAL. THE SECOND TRANSEPTAL PUNCTURE SUCCESSFULLY. THE TRUSTEER SHEATH AND THE 35MM WM FLX PRO DEVICE WERE RE-INTRODUCED AND DEPLOYED IN THE PATIENT. AT THIS TIME, THE PHYSICIANS WERE PLEASED THAT THE IMPLANT OF THE WM FLX PRO HELPED 'BLOCK' THE LEAK. THE PATIENT EFFUSION WAS STARTING TO POOL AGAIN. THE PHYSICIAN DECIDED TO RELEASE THE 35MM CLOSURE DEVICE WHICH WAS WELL POSITIONED IN THE APPENDAGE. ADDITIONAL DRAINAGE WAS THEN PERFORMED, BUT THE PATIENT WENT INTO CARDIAC ARREST. THE NURSES PERFORMED HEART MASSAGE. THE PATIENT EXPERIENCED SEVERE BRADYCARDIA, AND HYPOTENSION (ARTERIAL TENSION LOW 16 HIGH 34). A DEFIBRILLATOR (200 JOULESX1) WAS USED AND LIFE CELL WAS USED TO GIVE ADDITIONAL BLOOD UNITS. A CODE WAS CALLED IN THE HOSPITAL TO ADVISE THE SURGERY TEAM. PROTAMINE WAS GIVEN TO THE PATIENT. THE PATIENT SEEMED TO STABILIZE SLIGHTLY, BUT THE EFFUSION PERSISTED. THE PHYSICIAN KEPT REMOVING 60CC SYRINGES OF BLOOD FROM THE PATIENT. THE PATIENT REMAINED IN THE INTERVENTIONAL CARDIOLOGY LAB UNDER ACTIVE RESUSCITATION AND WAS REPORTED TO HAVE DIED. THE PHYSICIAN SUSPECTED THAT THE INADEQUATE TRANSEPTAL LOCATION CAUSED THE WATCHMAN TRUSTEER ACCESS SYSTEM TO RUB AGAINST THE RIDGE AND POSSIBLY CAUSED A SMALL TEAR IN THE APPENDAGE. THERE WAS NO PERFORATION NOTED ON ANY IMAGING. THE CAUSE OF DEATH WAS DETERMINED TO BE TAMPONADE. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED.
| K963898 | IT WAS REPORTED THAT WHILE THE EXTERNAL PULSE GENERATOR (EPG) WAS IN USE, THE PATIENT EXPERIENCED AN R WAVE ON T WAVE EVENT, WHICH PRECIPITATED A VENTRICULAR FIBRILLATION (VF) ARREST. THE PATIENT WAS RESUSCITATED BUT ASPIRATED DURING RE-INTUBATION. IT WAS NOTED THAT HE SUBSEQUENTLY DEVELOPED ADDITIONAL COMPLICATIONS AND, UNFORTUNATELY, PASSED AWAY ON POSTOPERATIVE DAY SIX. IT WAS ALSO NOTED THAT THE PATIENT HAD BEEN ALERT, EXTUBATED, AND AMBULATORY ON POSTOPERATIVE DAY ONE FOLLOWING OPEN-HEART SURGERY, WHICH INCLUDED AN AORTIC VALVE REPLACEMENT AND ANEURYSM REPAIR. NO FURTHER INFORMATION WAS AVAILABLE THE EPG IS EXPECTED TO BE RETURNED FOR EVALUATION. | MEDTRONIC SUBMITS THIS REPORT TO COMPLY WITH FDA REGULATIONS 21 CFR PARTS 4 AND 803. MEDTRONIC HAS MADE REASONABLE EFFORTS TO PROVIDE AS MUCH RELEVANT INFORMATION AS IS AVAILABLE TO THE COMPANY AS OF THE SUBMISSION DATE OF THIS REPORT. THIS REPORT DOES NOT CONSTITUTE AN ADMISSION OR A CONCLUSION BY FDA, MEDTRONIC, OR ITS EMPLOYEES THAT THE DEVICE, MEDTRONIC, OR ITS EMPLOYEE CAUSED OR CONTRIBUTED TO THE EVENT DESCRIBED IN THE REPORT. ANY REQUIRED FIELDS THAT ARE UNPOPULATED ARE BLANK BECAUSE THE INFORMATION IS CURRENTLY UNKNOWN OR UNAVAILABLE. MEDTRONIC WILL SUBMIT A SUPPLEMENTAL REPORT IF ADDITIONAL RELEVANT INFORMATION BECOMES KNOWN.
HERO® | K213845 | A LITERATURE ARTICLE DESCRIBES A 40-YEAR-OLD FEMALE WITH END-STAGE RENAL DISEASE AND A COMPLEX VASCULAR HISTORY UNDERWENT SUCCESSFUL IMPLANTATION OF A SUPER HERO GRAFT IN 2021. DESPITE CORRECT IMPLANTATION AND MULTIPLE SUCCESSFUL THROMBECTOMIES OVER THE FOLLOWING MONTHS, THE PATIENT EXPERIENCED REPEATED THROMBOTIC OCCLUSIONS OF THE GRAFT. AFTER RECEIVING A KIDNEY TRANSPLANT AND DISCONTINUING WARFARIN, DEVICE REMOVAL WAS PLANNED BUT NOT IMMEDIATELY EXECUTED. IN MARCH 2022, THE PATIENT PRESENTED WITH ACUTE ABDOMINAL SYMPTOMS AND UNDERWENT AN URGENT CHOLECYSTECTOMY. POSTOPERATIVELY, SHE DEVELOPED UNEXPLAINED SHOCK, AND IMAGING REVEALED EXTENSIVE THROMBOSIS OF THE INFERIOR VENA CAVA AND SUPRAHEPATIC VEINS CONSISTENT WITH BUDD-CHIARI SYNDROME. DESPITE URGENT TRANSFEMORAL THROMBECTOMY, PARTIAL THROMBUS REMAINED, AND THE PATIENT DEVELOPED MULTI-ORGAN FAILURE, INCLUDING DIALYSIS-DEPENDENT RENAL FAILURE, HEPATIC DYSFUNCTION, AND RESPIRATORY COMPROMISE. SHE DIED 10 DAYS AFTER ICU ADMISSION. | THE SUSPECT DEVICE WAS NOT RETURNED FOR EVALUATION. THE COMPLAINT COULD NOT BE CONFIRMED. THE ROOT CAUSE COULD NOT BE DETERMINED. A REVIEW OF THE DEVICE HISTORY AND COMPLAINT DATABASE COULD NOT BE PERFORMED SINCE THE LOT NUMBER WAS NOT PROVIDED.
TUBING SET | K112360 | THE EVENT OCCURRED IN CHINA. THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR TREATMENT DUE TO A CAR ACCIDENT. IT WAS REPORTED THAT THE HLS SET COULD NOT BE PRIMED DUE TO AIR BUBBLES IN THE OXYGENATOR. THIS HLS SET WAS NEVER CONNECTED TO THE PATIENT. AS THE FIRST SET COULD NOT BE USED IMMEDIATELY, THERE WAS A DELAY IN TREATMENT. A SECOND HLS SET WAS USED AND HAS BEEN CONNECTED THE PATIENT TO PROCESS FURTHER. THE PATIENT PASSED AWAY WHILE BEING CONNECTED TO THE SECOND HLS SET AFTER SEVERAL HOURS. AS THE PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID# (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069076.
AVALON FETAL MONITOR FM30 | K140535 | IT WAS REPORTED THE MONITOR MEASURED THE MATERNAL HEART RATE INSTEAD OF THE FETAL HEART RATE ON TWO SEPARATE CTG READINGS AND THE FETUS WAS STILLBORN AT TERM. NO OTHER CLINICAL INFORMATION OR MEDICAL INTERVENTION WAS REPORTED; THEREFORE, GOOD FAITH EFFORTS ARE BEING PERFORMED. | THE MANUFACTURING DATE FOR THE REPORTED DEVICE IS PRIOR TO 24SEPT2016 SO NO UDI REQUIRED. THE REPORTER'S NAME AND EMAIL WERE PROVIDED; HOWEVER, DUE TO PRIVACY LAWS WERE NOT INCLUDED IN THE REPORT. E1 REPORTER PHONE#: (B)(6). PHILIPS IS IN THE PROCESS OF OBTAINING ADDITIONAL INFORMATION CONCERNING THIS EVENT AND THE COMPLAINT IS STILL UNDER INVESTIGATION. A FINAL REPORT WILL BE SUBMITTED ONCE THE INVESTIGATION IS COMPLETE.
DREAMSTATION CPAP | K131982 | THE MANUFACTURER WAS CONTACTED IN REFERENCE TO THE VOLUNTARY FIELD SAFETY NOTICE/RECALL NOTIFICATION RELATED TO THE SOUND ABATEMENT FOAM IN CERTAIN CPAP, BIPAP AND MECHANICAL VENTILATOR DEVICES. THE MANUFACTURER RECEIVED INFORMATION THAT THE PATIENT ALLEGING KIDNEY CANCER DEATH. THERE WAS NO REPORT OF MEDICAL INTERVENTION. THE MANUFACTURER WAS MADE AWARE OF THIS COMPLAINT THROUGH A REPRESENTATIVE OF THE CUSTOMER. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER. AT THIS TIME, NO FURTHER INVESTIGATION CAN BE REQUESTED AT THIS TIME. IF ANY ADDITIONAL INFORMATION IS RECEIVED, A SUPPLEMENTAL REPORT WILL BE FILED. | SINCE NO DEVICE INFORMATION WAS PROVIDED, THE EXACT RECALL NUMBER IS UNKNOWN. POSSIBLE RECALL NUMBERS INCLUDE Z-1972-2021, Z-1973-2021, AND Z-1974-2021.
TUBING SET | K112360 | THE EVENT OCCURRED IN CHINA. THE PATIENT WAS ADMITTED TO THE EMERGENCY DEPARTMENT FOR TREATMENT DUE TO A CAR ACCIDENT. IT WAS REPORTED THAT THE HLS SET COULD NOT BE PRIMED DUE TO AIR BUBBLES IN THE OXYGENATOR. THIS HLS SET WAS NEVER CONNECTED TO THE PATIENT. AS THE FIRST SET COULD NOT BE USED IMMEDIATELY, THERE WAS A DELAY IN TREATMENT. A SECOND HLS SET WAS USED AND HAS BEEN CONNECTED THE PATIENT TO PROCESS FURTHER. THE PATIENT PASSED AWAY WHILE BEING CONNECTED TO THE SECOND HLS SET AFTER SEVERAL HOURS. AS THE PATIENT PASSED AWAY A REPORT IS REQUIRED. COMPLAINT ID # (B)(4). | A FOLLOW-UP MEDWATCH WILL BE SUBMITTED WHEN ADDITIONAL INFORMATION BECOMES AVAILABLE. THIS EVENT OCCURRED ON THE CHINESE MARKET. IT IS A SIGNIFICANTLY SIMILAR DEVICE TO "HLS SET¿ WHICH IS SOLD IN THE USA UNDER PREMARKET SUBMISSION NUMBER K112360. FOR SECTION D4 ALL AVAILABLE IDENTIFYING INFORMATION IS FOR THE OUT OF THE US DEVICE, SUBJECTED TO THIS REPORT. THEREFORE, NO GUDID INFORMATION EXISTS. FURTHERMORE, UDI INFORMATION (PRIMARY DI NUMBER) PROVIDED IN D4 IS FOR HLS SET WITH CATALOG NUMBER 701069076.
BENEVISION N1 PATIENT MONITOR | K213799 | IT WAS REPORTED THAT ON (B)(6) 2025, AT 10:00 P.M., A PATIENT WAS CONNECTED TO A BENEVISION N12 PATIENT MONITOR (SERIAL NUMBER: (B)(6) IN USE WITH A BENEVISION N1 (SERIAL NUMBER: (B)(6). THE PATIENT EXPERIENCED ASYSTOLE; HOWEVER, NO ALARM WAS GENERATED. | NARRATIVE MINDRAY REPRESENTATIVES TESTED THE OPERATION OF ALL BENEVISION N12 AND N1 MONITORS IN THE DEPARTMENT. THE DEVICES ARE FUNCTIONING NORMALLY, AND NO ALARM FAILURE PHENOMENON WAS FOUND. MORE MONITOR LOGS ARE BEING COLLECTED FOR INVESTIGATION, AND THIS ISSUE IS UNDER INVESTIGATION.
LIBERTY SELECT CYCLER ASSY(NON-VALUATED) | K181108 | A NURSE REPORTED THAT A PERITONEAL DIALYSIS (PD) PATIENT PASSED AWAY WHILE CONNECTED TO THE LIBERTY SELECT CYCLER. ADDITIONAL INFORMATION WAS OBTAINED THROUGH FOLLOW-UP WITH THE PD CLINIC. THE PATIENT¿S SPOUSE CONTACTED THE PERITONEAL DIALYSIS REGISTERED NURSE (PDRN) ON (B)(6) 2025 AROUND 10:30AM TO REPORT THAT THE PATIENT HAD PASSED AWAY. THE PATIENT HAD BEEN IN TREATMENT CONNECTED TO THE LIBERTY SELECT CYCLER WHEN SHE BEGAN VOMITING. THE VOMIT WAS BLACK IN COLOR. THE SPOUSE DID AN EMERGENCY DISCONNECT FROM THE TREATMENT AND CALLED 911. WHEN EMERGENCY MEDICAL SERVICES (EMS) ARRIVED, THE PATIENT WAS NOT RESPONSIVE. EMS ATTEMPTED TO RESUSCITATE THE PATIENT BUT WERE UNSUCCESSFUL. THE CAUSE OF DEATH IS UNKNOWN. THERE WERE NO REPORTED ISSUES WITH THE LIBERTY SELECT CYCLER OR THE PATIENT¿S TREATMENT PRIOR TO THE EVENT. NO ADDITIONAL INFORMATION WAS AVAILABLE PERTAINING TO THE DEATH. | CLINICAL REVIEW: THERE IS A TEMPORAL RELATIONSHIP BETWEEN PERITONEAL DIALYSIS UTILIZING THE LIBERTY SELECT CYCLER AND THE PATIENT EVENT OF DEATH. HOWEVER, THERE IS NO DOCUMENTATION IN THE COMPLAINT FILE TO SHOW A CAUSAL RELATIONSHIP BETWEEN THE ADVERSE EVENT AND USE OF THE LIBERTY SELECT CYCLER. ADDITIONALLY, THERE IS NO ALLEGATION OF A DEVICE MALFUNCTION OR DEFICIENCY REPORTED FOR THIS EVENT. THE PATIENT REPORTEDLY BEGAN VOMITING BLACK MATTER PRIOR TO BECOMING UNRESPONSIVE. ALTHOUGH THE CAUSE OF DEATH WAS NOT CONFIRMED, THIS IS AN INDICATION OF A GASTROINTESTINAL (GI) BLEED. IF THE PATIENT DID HAVE A SIGNIFICANT BLEED, SHE COULD HAVE GONE INTO HYPOVOLEMIC SHOCK CAUSING CARDIAC ARREST AND DEATH WHICH WOULD HAVE BEEN WHY THE PATIENT WAS UNRESPONSIVE WHEN PARAMEDICS ARRIVED. HOWEVER, WITHOUT ADDITIONAL INFORMATION, THAT CANNOT BE CONFIRMED. BASED ON THE AVAILABLE INFORMATION AND NO ALLEGATION OR EVIDENCE OF A MALFUNCTION OR DEFICIENCY, THE LIBERTY SELECT CYCLER CAN BE EXCLUDED AS THE CAUSE OF THE PATIENT¿S DEATH. THE PLANT INVESTIGATION IS IN PROCESS. A SUPPLEMENTAL MDR WILL BE SUBMITTED UPON COMPLETION OF THIS ACTIVITY.
DREAMSTATION AUTO BIPAP | K131982 | THE MANUFACTURER RECEIVED INFORMATION A USER OF UNKNOWN DEVICE HAD PASSED AWAY. THE REPORTED EVENT MAKES NO ALLEGATION OF ANY SPECIFIC DEVICE MALFUNCTION, USE ERROR, FAILURE, LABELING, OR OTHER DEFICIENCY THAT IS RELEVANT TO THE HARM REPORTED; THEREFORE, BASED ON INFORMATION AVAILABLE AT THIS TIME, THE DEVICE DID NOT CAUSE OR CONTRIBUTE TO A SERIOUS INJURY OR DEATH. ADDITIONAL INFORMATION IS BEING REQUESTED REGARDING THE DETAILS OF THE REPORTED DEATH. THE CAUSE OF DEATH IS NOT KNOWN. THE DEVICE HAS NOT BEEN RETURNED FOR EVALUATION. THE MANUFACTURER'S INVESTIGATION IS ONGOING. A FOLLOW-UP REPORT WILL BE SUBMITTED WHEN THE MANUFACTURER'S INVESTIGATION IS COMPLETE.
INTRAOSSEOUS NEEDLE KITS | K191976 | IT WAS REPORTED, AFTER INSERTION OF THE BD IO NEEDLE, THE PARAMEDIC COULD NOT SEPARATE THE BD IO NEEDLE HUB FROM THE BD IO DRILL. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. EACH EVENT REPORTED TO BD IS EVALUATED AND INVESTIGATED IN ACCORDANCE WITH OUR COMPLAINT INVESTIGATION PROCEDURES. THE INVESTIGATION PROCESS INCLUDES, BUT IS NOT LIMITED TO, EVALUATION OF THE EVENT DETAILS PROVIDED BY THE COMPLAINT FACILITY, A REVIEW OF COMPLAINT HISTORY, MANUFACTURING RECORDS AND RISK DOCUMENT WHERE APPLICABLE, AND AN EVALUATION OF THE SUBJECT DEVICE WHEN AVAILABLE TO IDENTIFY POTENTIAL CONTRIBUTING FACTORS. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION.
INTRAOSSEOUS NEEDLE KITS | K191976 | IT WAS REPORTED, AFTER INSERTION OF THE BD IO NEEDLE, THE PARAMEDIC COULD NOT SEPARATE THE BD IO NEEDLE HUB FROM THE BD IO DRILL. | H11: SECTION A THROUGH F - THE INFORMATION PROVIDED BY BD REPRESENTS ALL OF THE KNOWN INFORMATION AT THIS TIME. DESPITE GOOD FAITH EFFORTS TO OBTAIN ADDITIONAL INFORMATION, THE COMPLAINANT / REPORTER WAS UNABLE OR UNWILLING TO PROVIDE ANY FURTHER PATIENT, PRODUCT, OR PROCEDURAL DETAILS TO BD. EACH EVENT REPORTED TO BD IS EVALUATED AND INVESTIGATED IN ACCORDANCE WITH OUR COMPLAINT INVESTIGATION PROCEDURES. THE INVESTIGATION PROCESS INCLUDES, BUT IS NOT LIMITED TO, EVALUATION OF THE EVENT DETAILS PROVIDED BY THE COMPLAINT FACILITY, A REVIEW OF COMPLAINT HISTORY, MANUFACTURING RECORDS AND RISK DOCUMENT WHERE APPLICABLE, AND AN EVALUATION OF THE SUBJECT DEVICE WHEN AVAILABLE TO IDENTIFY POTENTIAL CONTRIBUTING FACTORS. THE DEVICE HAS NOT BEEN RETURNED TO THE MANUFACTURER FOR EVALUATION.
STEALTH | K220568 | IT WAS REPORTED THE STEALTH 360 PERIPHERAL ORBITAL ATHERECTOMY DEVICE (OAD) WAS USED TO TREAT A CALCIFIED, TORTUOUS COMMON ILIAC ARTERY. DURING DISTAL TO PROXIMAL TREATMENT, THE OAD CROWN SEEMED TO BE JUMPING SLIGHTLY AND NOT SPINNING AS PER USUAL. UPON REMOVAL, IT WAS OBSERVED THE OAD CROWN HAD DETACHED AND REMAINED ON THE VIPERWIRE GUIDE WIRE. A DISSECTION WAS OBSERVED. A STENT WAS PLACED AT THE DISSECTION SITE. TWO DAYS POST PROCEDURE, THE PATIENT EXPIRED FROM COMPLICATIONS. | MANUFACTURER'S INVESTIGATION IS STILL PENDING AT THIS TIME. RESULTS AND CONCLUSIONS WILL BE PROVIDED IN THE FINAL REPORT. D4: THE UDI NUMBER IS NOT KNOWN AS THE PART AND LOT NUMBER WERE NOT PROVIDED.
WATCHMAN FXD CURVE? ACCESS SYSTEM | K212228 | IT WAS REPORTED A PATIENT DEATH OCCURRED. A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE WAS BEING PERFORMED ON A PATIENT WITH MARFAN SYNDROME. A WATCHMAN FXD DOUBLE CURVE ACCESS SYSTEM (WAS) AND A WATCHMAN FLX PRO LAA CLOSURE DEVICE & DELIVERY SYSTEM (WDS) WAS SELECTED FOR USE. A SMALL EFFUSION WAS NOTED ON THE RIGHT VENTRICLE PRIOR TO THE START OF THE PROCEDURE. AFTER THE TRANSEPTAL PUNCTURE (TSP) USING THE VERSACROSS CONNECT SYSTEM AND THE WAS DOUBLE CURVE SHEATH, THE PATIENT DEVELOPED LARGER EFFUSION AROUND THE WHOLE PERICARDIUM ALONG WITH CARDIAC TAMPONADE. CARDIOPULMONARY RESUSCITATION (CPR) WAS STARTED. A PERICARDIAL TAP WAS ATTEMPTED TO TREAT THE EFFUSION BUT WAS UNSUCCESSFUL. RESUSCITATION WAS DONE FOR THIRTY (30) MINUTES, HOWEVER THE PATIENT DIED.
TIGHTRAIL ROTATING DILATOR SHEATH | K150360 | A LEAD EXTRACTION PROCEDURE COMMENCED TO REMOVE A RIGHT VENTRICULAR (RV) LEAD DUE TO FRACTURE. A RIGHT ATRIAL (RA), LEFT VENTRICULAR (LV), AND A CAPPED RV LEAD WERE PRESENT IN THE PATIENT, BUT NOT TARGETED FOR EXTRACTION. A SPECTRANETICS LLD LEAD LOCKING DEVICE (LLD) WAS INSERTED INTO THE LEAD TO PROVIDE TRACTION. BEGINNING WITH A SPECTRANETICS 11F TIGHTRAIL ROTATING DILATOR SHEATH, SOME RESISTANCE WAS ENCOUNTERED; THEREFORE, A COOK MEDICAL INTRODUCER SHEATH WAS USED. THE LEAD CAME FREE FROM THE WALL OF THE RV AND A MERIT MEDICAL EN SNARE DEVICE WAS USED TO REPOSITION THE LEAD TO IMPROVE ALIGNMENT AND TRACTION. FURTHER RESISTANCE WAS MET AT THE COIL OF THE LEAD, AND UPSIZED TO A SPECTRANETICS 13F TIGHTRAIL ROTATING DILATOR SHEATH. AFTER NO FURTHER ADVANCEMENT COULD BE MADE, THE COOK MEDICAL INTRODUCER SHEATH WAS LOADED ONTO THE TIGHTRAIL. WITH SOME DIFFICULTY GOING AROUND THE INNOMINATE/SUPERIOR VENA CAVA (SVC) JUNCTION, THE TIGHTRAIL MOVED FURTHER ALONG THE LEAD PAST THE SVC, WHEN PROGRESS STALLED. UPON REMOVAL OF THE TIGHTRAIL, A PERICARDIAL EFFUSION WAS DETECTED VIA TRANSESOPHAGEAL ECHOCARDIOGRAM (TEE), AND RESCUE EFFORTS BEGAN, INCLUDING BYPASS AND STERNOTOMY. A SVC PERFORATION THAT EXTENDED INTO THE RA WAS DISCOVERED AND ATTEMPTS WERE MADE TO REPAIR; HOWEVER, BLEEDING CONTINUED. DESPITE LIFESAVING EFFORTS, THE PATIENT DID NOT SURVIVE THE PROCEDURE. THIS REPORT CAPTURES THE 13F TIGHTRAIL IN USE IN THE AREA WHEN THE PERFORATION OCCURRED, REQUIRING INTERVENTION, BUT RESULTING IN DEATH. THERE WAS NO ALLEGED MALFUNCTION OF ANY SPECTRANETICS DEVICES IN USE DURING THE PROCEDURE. | A4) PATIENT WEIGHT - NOT AVAILABLE FROM FACILITY. A5) PATIENT ETHNICITY - NOT AVAILABLE FROM FACILITY. B7) OTHER RELEVANT HISTORY - NOT AVAILABLE FROM FACILITY. H3) THE TIGHTRAIL WAS DISCARDED, THUS NO PRODUCT INVESTIGATION WAS PERFORMED. H6) GREAT VESSEL PERFORATION, CARDIAC PERFORATION, AND DEATH ARE KNOWN RISKS OF COMPLICATION WITH USE OF THE TIGHTRAIL DEVICE. SUBMISSION OF THIS REPORT DOES NOT, IN ITSELF, REPRESENT A CONCLUSION BY THE MANUFACTURER AND/OR AUTHORIZED REPRESENTATIVE OR THE NATIONAL COMPETENT AUTHORITY THAT THE CONTENT OF THIS REPORT IS COMPLETE OR ACCURATE, THAT THE MEDICAL DEVICE(S) LISTED FAILED IN ANY MANNER AND/OR THAT THE MEDICAL DEVICE(S) CAUSED OR CONTRIBUTED TO AN ALLEGED DEATH OR DETERIORATION IN THE STATE OF THE HEALTH OF ANY PERSON.
DEXCOM G7 CONTINUOUS GLUCOSE MONITORING SYSTEM | K213919 | IT WAS REPORTED VIA MAUDE (MW5173056) AND FOLLOW UP CONTACT WITH THE PATIENT'S PARENT THAT THE PATIENT PASSED AWAY. THE PATIENT'S PARENT NOTED THAT SHE REPORTED THIS VIA MAUDE AFTER RECEIVING THE RECEIVER FIELD ACTION LETTER IN THE MAIL AND NOTING THAT THE SERIAL NUMBER ON THE LETTER MATCHED THE PATIENT'S RECEIVER. THE PATIENT'S PARENT REPORTED THAT THE PATIENT USED AN INSULIN PUMP BUT DID NOT KNOW WHICH ONE AND BELIEVED IT TO HAVE BEEN IN MODIFIED CLOSED LOOP. USE OF AN INSULIN PUMP DURING THE EVENT IS NOT CONFIRMED AND NO FURTHER DETAILS WERE AVAILABLE. THE PATIENT HAD REPORTEDLY BEEN SICK AND VOMITING. HE WAS NOTED TO HAVE BEEN ALONE IN HIS APARTMENT WHEN HE WAS FOUND AFTER BEING DECEASED 3 DAYS. THE PARENT REPORTED FINDING THE DEXCOM RECEIVER NEXT TO HIM, BUT COULD NOT RECALL ANY CGM READINGS, ALERTS, OR ALARMS FROM THE DEVICE AT THAT TIME. AN AUTOPSY NOTED THAT HE HAD DIED FROM DIABETIC KETOACIDOSIS (DKA). THE PATIENT HAD BEEN EVALUATED AND TREATED IN THE EMERGENCY DEPARTMENT FOR SYMPTOMATIC HYPERGLYCEMIA DUE TO INACCURACIES ON THE CGM APPROXIMATELY 8 MONTHS PRIOR TO HIS DEATH. THIS PREVIOUS ISSUE COUPLED WITH THE LETTER SHE RECEIVED MADE THE PARENT SUSPICIOUS THAT THE PATIENT'S DEATH WAS THE RESULT OF A DEXCOM MALFUNCTION, ALTHOUGH SHE COULD NOT SPECIFY WHAT SHE BELIEVED THE PROBLEM TO BE. THE PATIENT'S PARENT NOTED THAT SHE HAD THE PATIENT'S RECEIVER IN HER POSSESSION AND WOULD CONSIDER RETURNING IT FOR INVESTIGATION. DEXCOM HAS SENT THE PATIENT'S MOTHER A RETURN GOODS AUTHORIZATION TO SEND PRODUCT IN FOR INVESTIGATION. IF PRODUCT IS RETURNED, A FOLLOW UP REPORT WILL BE SUBMITTED UPON INVESTIGATION COMPLETION. NO PRODUCT OR DATA WAS PROVIDED FOR INVESTIGATION. PROBLEM COULD NOT BE CONFIRMED AND PROBABLE CAUSE CANNOT BE DETERMINED. NO ADDITIONAL PATIENT OR EVENT INFORMATION IS AVAILABLE. | (B)(4).
VERSACROSS CONNECT LAAC ACCESS SOLUTION | K150709 | IT WAS REPORTED THAT PERICARDIAL EFFUSION (PE), CARDIAC TAMPONADE, AND DEATH OCCURRED. THE PROCEDURE WAS CANCELLED. DURING A LEFT ATRIAL APPENDAGE CLOSURE (LAAC) PROCEDURE, A VERSACROSS CONNECT LAAC WAS SELECTED FOR USE. INTRACARDIAC ECHOCARDIOGRAM (ICE) IMAGING NOTED THERE WAS NO BASELINE PE. GROIN ACCESS OBTAINED AND NON-BOSTON SCIENTIFIC VASCULAR PRE-CLOSURE DEVICE WAS USED. A VERSACROSS RF WIRE WAS INSERTED UP TO THE SUPERIOR VENA CAVA (SVC), A WATCHMAN TRUSTEER ACCESS SYSTEM (WAS) AND VERSACROSS DILATOR TO FOLLOW. THE TRANSSEPTAL PUNCTURE (TSP) WAS PERFORMED WITHOUT RF DELIVERY REQUIRED. TSP LOCATION APPEARED TO BE MID/MID-POSTERIOR BASED ON ICE IMAGING. NO PE WAS NOTED AT THIS POINT IN THE PROCEDURE AND PATIENT VITALS WERE STABLE. TSP LOCATION WAS DILATED WITH THE WAS AND THEN THE SHEATH AND DILATOR WERE RETRACTED FOR LEFT ATRIUM (LA) ACCESS OF THE ICE CATHETER. TSP ACCESS OF ICE CATHETER WAS STILL UNSUCCESSFUL. A PE WAS THEN DISCOVERED FOUR (4) MINUTES LATER. WAS REMOVED FROM THE GROIN AND ICE VISUALIZED THE PE LOCATION ON THE RIGHT SIDE OF THE HEART. THE PROCEDURE WAS ABORTED. CARDIAC TAMPONADE OCCURRED, WHICH CAUSED THE NEED FOR CARDIOPULMONARY RESUSCITATION (CPR), BLOOD TRANSFUSION, PERICARDIOCENTESIS REMOVING BRIGHT RED BLOOD OF OVER THREE (3) LITERS, FRESH FROZEN PLASMA (FFP), AND CRYOPRECIPITATE. THE FAMILY WAS CONTACTED A LITTLE OVER TWO (2) HOURS LATER AND AS THE PATIENT WAS A DO NOT RESUSCITATE (DNR) STATUS AND THEY REQUESTED NO ADDITIONAL LIFE-SAVING MEASURES, SO TIME OF DEATH WAS CALLED BY THE PHYSICIAN. IT WAS NOTED THE PATIENT WAS ON ELIQUIS AND STOPPED TAKING ELIQUIS TWO (2) DAYS PRIOR TO THE PROCEDURE. LAST RECORDED INR LAB RESULT WAS 1.4. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION.
VERSACROSS ACCESS SOLUTION | K150709 | IT WAS REPORTED THAT THE PATIENT EXPERIENCED STROKE WHILE IN RECOVERY. THE PATIENT PASSED AWAY. DURING A PULSE FIELD ABLATION PROCEDURE, A VERSACROSS ACCESS SOLUTION WAS SELECTED FOR USE. A CLOT WAS RULED OUT PRIOR TO THE PROCEDURE, BY TEE IMAGE PERFORMED. TRANSEPTAL PUNCTURE (TSP) WAS DONE USING VERSACROSS RF WIRE TO GAIN ACCESS TO THE LEFT ATRIUM, WITH NO ISSUES REPORTED. SUCCESSFUL PULSE FIELD ABLATION DONE TO ALL FOUR PULMONARY VEINS. THE PHYSICIAN WAS COMPLETED WITHOUT ANY ISSUE FROM THE PHYSICIAN. DURING POST-PROCEDURE, WHILE IN RECOVERY, THE PATIENT HAD A STROKE, CONFIRMED BY A CT SCAN. ON THE DAY OF THE PROCEDURE, ACTIVATED CLOTTING TIME (ACT) WAS MAINTAINED ABOVE 300 SECONDS WITH HEPARIN ADMINISTERED PRIOR TO THE TRANSSEPTAL PUNCTURE. LEFT ATRIAL DWELL TIME WAS KEPT UNDER 15 MINUTES. PULMONARY VEIN ISOLATION (PVI) WAS PERFORMED EXCLUSIVELY, REQUIRING A TOTAL OF 48 APPLICATIONS. FOLLOWING PVI, A CAVOTRICUSPID ISTHMUS (CTI) LINE WAS CREATED USING A RADIOFREQUENCY (RF) CATHETER. THE ONLY ADDITIONAL CATHETER PLACED WAS IN THE CORONARY SINUS. POST-PROCEDURE, THE PATIENT EXPERIENCED DIFFICULTY AWAKENING FROM ANESTHESIA AND WAS CHALLENGING TO EXTUBATE. A CT SCAN FOLLOWED BY AN MRI REVEALED MULTIPLE CEREBRAL INFARCTS CONSISTENT WITH AN EMBOLIC SHOWER. THE PATIENT DID NOT REGAIN CONSCIOUSNESS FOLLOWING GENERAL ANESTHESIA. NO TREATMENT WAS ADMINISTERED BEYOND INTRACRANIAL PRESSURE (ICP) RELIEF, AS THE STROKE INVOLVED MULTIPLE AREAS WITH ASSOCIATED MASS EFFECTS. THE PATIENT HAS SINCE BEEN TRANSITIONED TO COMFORT CARE, WITH NO EXPECTATION OF RECOVERY. THE PHYSICIAN WAS UNABLE TO DETERMINE THE CAUSE OF THE STROKE. THE TEE WAS REVIEWED AND THOROUGHLY EXAMINED, SHOWING NO EVIDENCE OF A LEFT ATRIAL APPENDAGE (LAA) THROMBUS. IT IS HIGHLY UNLIKELY THAT THE CEREBROVASCULAR ACCIDENT (CVA) WAS DUE TO A MISSED THROMBUS, AS THE INFARCTS WERE DISTRIBUTED THROUGHOUT THE BRAIN. THE PATIENT PASSED AWAY. THE DEVICE IS NOT EXPECTED TO BE RETURNED FOR ANALYSIS (DISPOSED). | IT WAS INDICATED THAT THE DEVICE WILL NOT BE RETURNED FOR EVALUATION. IF THERE IS ANY FURTHER RELEVANT INFORMATION OBTAINED, A SUPPLEMENTAL MEDWATCH WILL BE FILED. DETAILED PRODUCT INFORMATION WAS NOT PROVIDED TO BSC. WE COMPLETED A GOOD FAITH EFFORT TO OBTAIN THE INFORMATION, BUT IT WAS NOT AVAILABLE. BECAUSE THE PRODUCT IS UNKNOWN, WE ARE UNABLE TO PROVIDE THE UNIQUE IDENTIFIER (UDI) AND OTHER SPECIFIC PRODUCT INFORMATION. GOOD FAITH EFFORT ATTEMPTS WERE MADE TO TRY AND RETRIEVE ADDITIONAL DETAILS REGARDING THE REPORTED EVENT, BUT FURTHER INFORMATION WAS UNABLE TO BE OBTAINED.
DAVINCI XI | K131861 | IT WAS REPORTED THAT A DA VINCI-ASSISTED RADICAL CYSTECTOMY WITH INTRACORPOREAL ILEAL CONDUIT CREATION, BILATERAL URETERAL STENT PLACEMENT, AND BILATERAL EXTENDED PELVIC LYMPH NODE DISSECTION WAS PERFORMED. THERE WERE NO INTRAOPERATIVE COMPLICATIONS. THE INITIAL POST-OPERATIVE COURSE WAS UNCOMPLICATED AND THE PATIENT WAS DISCHARGED ON POST-OPERATIVE DAY (POD) 5. THE PATIENT ULTIMATELY EXPIRED ON POD 50 DUE TO SUBSEQUENT COMPLICATIONS. THE PATIENT PRESENTED TO THE EMERGENCY DEPARTMENT ON POD 8 WITH INTRA-ABDOMINAL FLUID COLLECTION AND SUBSEQUENTLY DEVELOPED SIGNS OF SEPTIC SHOCK AND ACIDOSIS AND WAS STABILIZED WITH VASOPRESSORS. DURING EMERGENT EXPLORATORY SURGERY 2 LITERS OF FRANK STOOL WAS FOUND IN THE ABDOMEN, AND A LARGE ABSCESS IN THE PELVIS WAS IRRIGATED AND SUCTIONED OUT. A 1 CM MID-ILEAL ANTI-MESENTERIC SMALL BOWEL PERFORATION WAS FOUND AND RESECTED. ADDITIONALLY, THE COMMON ENTEROTOMY AND DISTAL ILEAL-ILEAL ANASTOMOSIS STAPLE LINE WAS COMPLETELY DISRUPTED WITH FRANK LEAKAGE OF STOOL; THE AREA WAS RESECTED (20CM TOTAL) AND REPAIRED. THE PATIENT REMAINED CRITICALLY ILL AFTER THIS PROCEDURE. A SUBSEQUENT EXPLORATORY LAPAROTOMY ON POD 10 OCCURRED AND PER THE OPERATIVE NOTES, A RIGHT COLECTOMY, ABDOMINAL WASHOUT, ILEOCOLIC ANASTOMOSIS, SMALL-BOWEL TO SMALL-BOWEL ANASTOMOSIS, PLACEMENT OF A WOUND VAC AND FASCIAL CLOSURE WERE COMPLETED. EXAMINATION OF THE ILEAL CONDUIT FROM THE ORIGINAL ROBOTIC PROCEDURE BY UROLOGY FOUND NO SIGNS OF LEAKAGE; THE ANASTOMOSES AT THE CONDUIT ITSELF AS WELL AS THE URETERS APPEARED VIABLE WITHOUT SIGNS OF NECROSIS OR ISCHEMIA. NOTES AT POD 33 DOCUMENT A POST-PROCEDURE COURSE THAT INCLUDED ACUTE RENAL FAILURE REQUIRING DIALYSIS, ACUTE HYPOXIC RESPIRATORY FAILURE REQUIRING VENTILATION, URINE LEAK REQUIRING BILATERAL NEPHROSTOMY TUBE/DRAIN PLACEMENT, SEPTIC SHOCK REQUIRING ONGOING VASOPRESSORS, MULTIPLE ANTIBIOTICS RELATED TO ANASTOMOTIC LEAK, FUNGEMIA, AND LOWER ABDOMINAL MIDLINE WOUND DEHISCENCE WITH PURULENT DRAINAGE. NO ADDITIONAL INFORMATION WAS PROVIDED UP TO THE PROVIDED DEATH CERTIFICATE AT POD 50 WHICH LISTED CARDIAC ARREST, SEPSIS AND ABDOMINAL INFECTION AS CAUSE OF DEATH. | A REVIEW OF THE SYSTEM LOGS FOUND THAT NO RELEVANT ERRORS OCCURRED DURING THE PROCEDURE. LOG REVIEW OF THE FIVE SUBSEQUENT PROCEDURES PERFORMED ON THE SYSTEM FOUND NO ERRORS OCCURRED THAT WOULD HAVE LIKELY CAUSED OR CONTRIBUTED TO THE EVENT. REVIEW FOUND THAT THE SYSTEM FUNCTIONED NORMALLY, NO INTRAOPERATIVE EVENTS OR ISSUES WERE OBSERVED. THE FOLLOWING CONCOMITANT PRODUCTS WERE USED DURING THIS PROCEDURE: 30 DEGREE ENDOSCOPE PLUS, VESSEL SEALER EXTEND, MONOPOLAR CURVED SCISSORS, CADIERE FORCEPS, MARYLAND BIPOLAR FORCEPS, SUREFORM 60 STAPLER, LARGE NEEDLE DRIVER, LARGE SUTURECUT NEEDLE DRIVER. A SITE HISTORY REVIEW FOUND NO RELATED COMPLAINTS WERE RECEIVED FOR THE INSTRUMENTS USED DURING THIS PROCEDURE. A DEVICE HISTORY RECORD (DHR) REVIEW FOR THE DEVICE(S) USED DURING THE PROCEDURE FOUND NO NON-CONFORMANCE'S WERE IDENTIFIED TO BE RELATED TO THIS EVENT. THE STAPLER INSTRUMENT LOGS FOR THIS PROCEDURE FOUND A SUREFORM 60 STAPLER INSTRUMENT WAS USED TO FIRE 12 RELOADS WITHOUT ANY ISSUES OBSERVED IN THE LOGS. THE VESSEL SEALER INSTRUMENT LOGS FOUND ONE CUT FAILURE ON THE INSTRUMENT'S 15TH CUT ATTEMPT. THE LOGS INDICATE THERE WERE TWO TIMES WHERE THE VESSEL SEALER EXTEND JAWS WERE OPEN TOO MUCH TO INITIATE A CUT OF THE TISSUE WITHIN THE JAWS. A REVIEW OF THE EVENT PERFORMED BY AN INTUITIVE SURGICAL MEDICAL SAFETY OFFICER (MSO) CONCLUDED THAT THE PATIENT IN THIS REPORT HAD A PRIOR HISTORY OF PROSTATE CANCER AND PRIOR PROSTATECTOMY WHO WAS UNDERGOING A CYSTECTOMY FOR BLADDER CANCER WITH CREATION OF AN ILEAL CONDUIT. THE PROCEDURE REPORTEDLY WENT WELL AND NO INTRAOPERATIVE ISSUES WERE NOTED. POST-OPERATIVELY, THE PATIENT DEVELOPED A SMALL BOWEL LEAK SEVERAL DAYS AFTER THE ROBOTIC PROCEDURE FROM THE ANASTOMOSIS WHERE THE ILEAL CONDUIT WAS HARVESTED. THE REASON FOR THE LEAK AND POTENTIAL FACTORS WHICH MAY HAVE LED TO THE LEAK IS NOT PROVIDED. THE PATIENT UNDERWENT MULTIPLE ADDITIONAL PROCEDURES AS NOTED ABOVE FOLLOWING THE INITIAL COMPLICATION AND EVENTUALLY DIED SEVERAL WEEKS LATER. NO ALLEGATION WAS MADE AGAINST ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT FROM THE INITIAL PROCEDURE. BASED ON THE INFORMATION PROVIDED IN THE SUMMARY OF EVENTS, INSUFFICIENT INFORMATION IS PROVIDED TO DETERMINE IF ANY INTUITIVE SURGICAL PRODUCT OR INSTRUMENT CONTRIBUTED TO THIS EVENT.